## Methodology
BIOCOMPATIBILITY SCREENING METHODOLOGY — SDS ANALYSIS
Regulatory framework: MDR 2017/745 (Annex I GSPR) ¦ ISO 10993-1:2025 ¦ REACH Reg. 1907/2006 ¦ CLP Reg. 1272/2008
SECTION 1 — CELL FORMAT RULES
Marker | Meaning | When to use | Example | Formula detection
🔴 RED | CRITICAL hazard signal | CMR 1A/1B confirmed ¦ IARC Group 1 or 2A ¦ Confirmed restricted aromatic amine (REACH Annex XVII Entry 43) ¦ ED on EU Official list | 🔴 CMR Cat.1B (Carc.) — H350 stated [SDS §2] | COUNTIF cell range "*🔴*" > 0 → CRITICAL
🟠 ORANGE | MAJOR hazard signal | CMR Cat.2 ¦ Bioaccumulation HIGH ¦ Heavy metal declared §3 ¦ STOT RE Cat.1 H372 ¦ H334 respiratory sensitiser ¦ ED TEDX only | 🟠 CMR Cat.2 (Repr.) — H361 stated [SDS §2] | COUNTIF cell range "*🟠*" > 0 → MAJOR (if no 🔴 found)
🟡 YELLOW | MINOR hazard signal | H317 skin sensitiser ¦ Eye/skin irritation H314/H315/H318/H319 ¦ Formaldehyde at any concentration ¦ Heavy metal mentioned outside §3 only ¦ STOT SE Cat.3 | 🟡 H317 — Skin sensitisation Cat.1B [SDS §2] | COUNTIF cell range "*🟡*" > 0 → MINOR (if no 🔴 or 🟠 found)
– | Negative / absent finding | Data found, no hazard signal OR data searched and not present | – Not classified (H317, H334 absent) [SDS §2] – Negative for cytotoxicity [SDS §11] – No sensitisation study data [SDS §11] | Not detected by formula — no alert
ℹ️ | Structural limitation (appears ONCE only in Composition cell) | Composition undisclosed ¦ No CAS numbers available — independent research not possible | ℹ️ Composition undisclosed — no CAS numbers provided... | Not detected by formula — no alert
ALERT LEVEL formula (col AC) | Auto-calculated from markers in cols G:AB | Reads markers across all endpoint and compliance cells (Cytotoxicity col G to REACH/SVHC compliance col AB) — returns worst-case level | Returns CRITICAL ¦ MAJOR ¦ MINOR ¦ NONE | Worst-case rule applied automatically: 1 🔴 anywhere = CRITICAL regardless of other markers
[SDS §X] | Data source: SDS section | Every data item extracted from the SDS document | – Not classified (H317 absent) [SDS §2] | Section number always required
[Research, TABLE v.X] | Data source: external reference table (version specified) | Every item from database tables loaded in this file | 🔴 Carc. 1B H350 [Research, CLP ATP22] | ALWAYS specify table name and version so data age is traceable
Copy this formula down for each new row — references are relative
SECTION 2 — BIOLOGICAL ENDPOINTS & COMPLIANCE COLUMNS (ISO 10993-1:2025)
Biological Endpoint | SDS Sections | What to look for | Alert triggers | SDS = primary source ¦ Listed tabs = supplementary (CAS numbers required)
Cytotoxicity | §11 | In vitro cytotoxicity data (neutral red uptake, MTT, colony-forming) | Any positive in vitro cytotoxicity result stated | SDS §§11 only — no supplementary lookup tab
Sensitisation | §2 ¦ §8 ¦ §11 | H317 (skin sens.) ¦ H334 (resp. sens.) ¦ EUH208 ¦ DNEL sensitisation route ¦ LLNA / GPMT / Buehler / human patch test data | H317 or H334 declared ¦ EUH208 declared ¦ Any positive sensitisation study ¦ No CAS = research impossible | SDS §§2/§8/§11 (primary) ¦ CLP_ATP22 (supplementary)
Skin irritation / Intracutaneous reactivity | §2 ¦ §8 ¦ §11 | H314 (severe burns) ¦ H315 (skin irritation) ¦ DNEL dermal ¦ Primary skin irritation study (rabbit/human) ¦ pH value (§9) | H314 or H315 declared ¦ Any positive skin irritation data ¦ pH <2 or >11.5 noted in §9 | SDS §§2/§8/§11 (primary) — no supplementary lookup tab
Eye irritation | §2 ¦ §11 | H318 (serious eye damage) ¦ H319 (eye irritation) ¦ Eye irritation study data | H318 or H319 declared ¦ Any positive eye irritation data | SDS §§2/§11 (primary) — no supplementary lookup tab
Acute systemic toxicity | §2 ¦ §8 ¦ §11 | H300 / H301 / H302 / H310 / H311 / H312 / H330 / H331 / H332 ¦ LD50 oral/dermal/inhalation (rat) ¦ LC50 inhalation ¦ DNEL acute systemic | Any H300–H332 declared ¦ LD50 < guidance threshold ¦ DNEL acute stated | SDS §§2/§8/§11 (primary) — no supplementary lookup tab
Sub-acute / Sub-chronic toxicity | §2 ¦ §8 ¦ §11 | H371 (STOT SE Cat.2) ¦ H373 (STOT RE Cat.2) ¦ 28-day / 90-day study data (NOAEL, target organs) ¦ DNEL sub-chronic | H371 or H373 declared ¦ Any repeated dose study data stated ¦ DNEL sub-chronic stated | SDS §§2/§8/§11 (primary) — no supplementary lookup tab
Chronic toxicity / Repeated dose | §2 ¦ §8 ¦ §11 | H372 (STOT RE Cat.1) ¦ Chronic study data (LOAEL, NOAEL, target organs) ¦ DNEL chronic systemic/local | H372 declared ¦ Any chronic toxicity data stated ¦ DNEL chronic stated | SDS §§2/§8/§11 (primary) ¦ CLP_ATP22 ¦ IARC_Groups (supplementary)
Genotoxicity / Mutagenicity | §2 ¦ §11 | H340 (Muta. Cat.1A/1B) ¦ H341 (Muta. Cat.2) ¦ Ames test (bacterial reverse mutation) ¦ Chromosomal aberration (CA) ¦ Micronucleus test (MNT) ¦ UDS ¦ In vivo genotoxicity data | H340 or H341 declared ¦ Any positive in vitro or in vivo genotoxicity result | SDS §§2/§11 (primary) ¦ CLP_ATP22 ¦ IARC_Groups (supplementary)
Carcinogenicity | §2 ¦ §11 ¦ §15 | H350 (Carc. Cat.1A/1B) ¦ H351 (Carc. Cat.2) ¦ Carcinogenicity study data ¦ IARC group mentioned in §15 | H350 or H351 declared ¦ IARC Group 1 or 2A ¦ Any positive carcinogenicity data | SDS §§2/§11/§15 (primary) ¦ CLP_ATP22 ¦ IARC_Groups (supplementary)
Reproductive / Developmental toxicity | §2 ¦ §8 ¦ §11 | H360 (Repr. Cat.1A/1B) ¦ H361 (Repr. Cat.2) ¦ H362 (lactation) ¦ Reproductive/developmental toxicity study data ¦ DNEL reproductive | H360, H361, or H362 declared ¦ Any reproductive or developmental toxicity data | SDS §§2/§8/§11 (primary) ¦ CLP_ATP22 (supplementary)
Endocrine disruption | §2 ¦ §11 ¦ §15 | Any mention of endocrine disruption ¦ SVHC designation for ED reason in §15 | Any ED mention ¦ Substance on EU Official ED list ¦ Substance listed in TEDX | SDS §§2/§11/§15 (primary) ¦ EU_ED_Official ¦ SVHC_Candidates ¦ TEDX (supplementary)
Bioaccumulation | §9 ¦ §12 ¦ §15 | Log Kow (§9) ¦ BCF (§12) ¦ Biodegradability: ready / inherent / not readily (§12) ¦ Water solubility (§9) ¦ PBT/vPvB statement (§12.5 / §15) | Log Kow >5 (HIGH) ¦ BCF >5000 (HIGH) ¦ Log Kow 4.5–5 (MODERATE) ¦ Substance designated PBT/vPvB | SDS §§9/§12/§15 (primary) — no supplementary lookup tab
Haemocompatibility | §11 | Any haemolysis / coagulation / platelet / complement data stated in §11 | Any positive haemolysis or blood interaction data | SDS §§11 only — no supplementary lookup tab
Pyrogenicity (material-mediated) | §11 | Any pyrogenicity or endotoxin data stated in §11 | Any positive pyrogenicity data | SDS §§11 only — no supplementary lookup tab
Implantation | §11 | Any implantation / tissue reaction data stated in §11 | Any positive implantation data | SDS §§11 only — no supplementary lookup tab
Immune responses / Immunotoxicity | §11 | Any immunotoxicity / immunosuppression / hypersensitivity data in §11 | Any positive immunotoxicity data | SDS §§11 only — no supplementary lookup tab
Other biological effects | §2 ¦ §11 ¦ §15 | Any hazard signal not covered by above endpoints ¦ Environmental classification (WGK, aquatic H400–H413) for context only ¦ Prop.65 ¦ Medical device use restrictions | Any new endpoint not listed in ISO 10993-1:2025 ¦ Medical device / skin contact restriction stated | SDS §§2/§11/§15 (primary) — no supplementary lookup tab
SECTION 3 — COMPLIANCE REGULATORY COLUMNS: WHAT IS SEARCHED
Compliance column | SDS sections consulted | Specific data searched in SDS | ⚠️ Trigger conditions | External research sources
CMR regulatory status | §2 ¦ §15 | H340/H341 (muta.) ¦ H350/H351 (carc.) ¦ H360/H361/H362 (repr.) ¦ SVHC designation for CMR reason in §15 | Any CMR H-statement declared ¦ Discrepancy between SDS and ATP22 harmonised classification | CLP Annex VI ATP22 (search by CAS) ¦ ECHA SVHC Candidate List
Azo dyes / Aromatic amines | §3 ¦ §11 ¦ §15 | Azo dye structure mentioned ¦ Aromatic amine mentioned ¦ REACH Annex XVII Entry 43 reference ¦ Cleavable aromatic amines listed in §3 | Azo dye confirmed ¦ Aromatic amine listed ¦ REACH Entry 43 applicable ¦ No CAS = Entry 43 not verifiable | REACH Annex XVII Entry 43 ¦ ECHA restriction dossier
Formaldehyde | §3 ¦ §11 ¦ §15 | Formaldehyde (CAS 50-00-0) ¦ Formaldehyde-releasing agents (methylol compounds, DMDM hydantoin, etc.) ¦ Any formaldehyde mention at any concentration | Formaldehyde or any releasing agent found at any concentration | ECHA dossier ¦ Textile formaldehyde guidance (EN ISO 14184)
Heavy metals | §3 ¦ §11 ¦ §15 | Ni ¦ Co ¦ Pb ¦ Cd ¦ Hg ¦ As ¦ Cr (total or CrVI) ¦ Sb ¦ Sn — any mention anywhere in SDS | Metal declared in §3 = MAJOR (M6) ¦ Metal mentioned outside §3 only = MINOR (Mi7) ¦ No CAS = not verifiable | REACH Annex XVII (Ni Entry 27 ¦ CrVI Entry 47) ¦ ECHA SVHC ¦ EN 1811 (Ni release)
REACH / SVHC compliance | §15 | SVHC statement (presence >0.1% w/w) ¦ REACH Annex XVII restrictions ¦ Seveso Directive category ¦ WGK class | SVHC identified >0.1% ¦ REACH restriction applicable ¦ No CAS = independent verification not possible | ECHA Candidate List (echa.europa.eu/candidate-list-table) ¦ REACH Annex XVII
SECTION 4 — SPECIAL CASE: COMPOSITION UNDISCLOSED (no CAS numbers in §3)
Rule | Description | Cell content | Regulatory basis
Negative finding | No signal in §2 and no SVHC in §15 | – No data identified (§2, §8, §11) | CLP Reg. 1272/2008 mandatory disclosure: CMR 1A/1B ≥0.1% and CMR Cat.2 ≥1% must be labelled if present
Critical limitation | No CAS = no independent verification possible | ⚠️ Independent research not possible — no CAS numbers available [Research] | Legally accepted under REACH Art. 15 (confidentiality of trade names) — supplier NOT required to disclose individual CAS
Research endpoints | IARC, ECHA dossier, ATP22, ED lists — all CAS-based | ⚠️ [Endpoint]: not assessable — no CAS numbers available [Research] | Applies to: genotoxicity, carcinogenicity, ED (EU Official Lists), TEDX, SVHC Candidate List, REACH Annex XVII Entry 43
SVHC compliance | Supplier declares no SVHC >0.1% in §15 | – No SVHC >0.1% (supplier statement) [SDS §15] | REACH Art. 59: substances of very high concern must be communicated if >0.1% w/w in an article or mixture
NO COMPONENT blocks | No CAS = no individual substance can be identified | Only one MIXTURE: block — no COMPONENT X blocks written | Composition non-disclosure is legally valid under REACH. Do NOT request CAS numbers as a corrective action.
Composition undisclosed rule does NOT escalate alert level | Alert level based on actual SDS signals ONLY (§2, §9, §12, §15) | If §2 not classified AND §15 no SVHC: alert level can be NONE | ISO 10993-1:2025 §5.4: precautionary principle applies but absence of data ≠ presence of hazard when CLP disclosure obligations are met
SECTION 5 — MEMO: WHY EACH PARAMETER IS SEARCHED (utility for biocompatibility evaluation)
Parameter | What it measures | Why it matters for biocompatibility | Threshold / Reference value | Important limitation for patient context
Log Kow (octanol-water partition coefficient) | Lipophilicity: how much a substance prefers fatty tissues over water | High Log Kow = substance accumulates in fat tissues / cell membranes rather than being eliminated in urine. Relevant for chronic patient exposure via skin contact. | Log Kow >5 = HIGH ¦ 4.5–5 = MODERATE ¦ <4.5 = LOW | Measures partitioning potential, not actual accumulation. Must be combined with BCF and toxicokinetic data.
BCF (Bioconcentration Factor) | Actual accumulation of a substance in a living organism (fish) relative to water concentration | High BCF confirms bioaccumulation in living tissues. Used as surrogate for human tissue retention when human data is absent. | BCF >5000 = HIGH (REACH PBT criterion) ¦ BCF 500–5000 = MODERATE ¦ <500 = LOW | Measured in aquatic organisms — extrapolation to human tissue is indirect. Conservative indicator only.
Water solubility | How easily a substance dissolves in water (mg/L) | Low solubility + high Log Kow = substance likely to accumulate rather than be excreted. High solubility = faster elimination. Also affects dermal absorption (aqueous skin layer). | No standalone threshold — always interpret with Log Kow and BCF | High water solubility does not automatically mean low risk — some water-soluble substances are highly toxic at low concentrations.
PBT / vPvB (Persistent, Bioaccumulative and Toxic / very Persistent and very Bioaccumulative) | REACH Annex XIII regulatory classification for substances of very high environmental and human concern | PBT/vPvB substances persist in the environment, accumulate in organisms AND are toxic. SVHC designation. Chronic patient exposure to a PBT substance is inherently high concern. | P: half-life >60d (marine) or >40d (freshwater) ¦ B: BCF >2000 ¦ T: CMR or STOT RE ¦ vP: >60d freshwater ¦ vB: BCF >5000 | PBT assessment is primarily environmental. Designation does not quantify risk to patient — requires further toxicokinetic evaluation.
DNEL (Derived No-Effect Level) | Maximum human exposure level below which no adverse effects are expected, derived from toxicological studies | Confirms that substance-specific toxicological data exists and has been used to set a reference level. Provides a quantitative anchor for hazard characterisation. | Substance-specific. Always note: route (dermal/inhalation/oral) ¦ population (worker/consumer) ¦ duration (acute/chronic) ¦ effect type (systemic/local) | CRITICAL LIMITATION: Worker DNEL = 8h/day, 5d/week. Consumer/General population DNEL = 24h/day, 7d/week and is more relevant for patients. Worker DNEL is NOT directly applicable to patient exposure. Correct extrapolation requires ISO 10993-17 risk assessment.
NOAEL (No Observed Adverse Effect Level) | Highest dose in an animal study at which no adverse effects were observed | When reported in SDS §11, confirms that formal toxicity studies exist. Basis for DNEL calculation. | Species-specific — requires uncertainty factors to extrapolate to humans (typically 10× interspecies × 10× intraspecies minimum) | Animal NOAEL − NOT a human safe level. Requires application of uncertainty/assessment factors before any patient risk conclusion.
LD50 / LC50 (Lethal Dose / Lethal Concentration, 50%) | Acute lethality: dose/concentration causing death in 50% of test animals | Indicates severity of acute systemic toxicity. Underpins H300–H302 (oral/dermal) and H330–H332 (inhalation) CLP classification. | H300 (oral ≤300 mg/kg) ¦ H301 (≤300 mg/kg Cat.3) ¦ H302 (≤2000 mg/kg) — CLP Reg. Annex I | Measures acute lethality only. High LD50 (low acute toxicity) does NOT exclude chronic, sensitising, CMR or endocrine disruptor effects.
H-statements (CLP Hazard Statements) | Standardised phrases describing the nature and severity of hazard — legally mandatory under CLP Reg. 1272/2008 when classification thresholds are exceeded | Primary hazard signal in SDS §2. Absence of H-statements is a regulated negative signal (not just an opinion): CLP requires labelling when thresholds are exceeded. Key signal for alert level assignment. | See CLP Annex I. Key biocompat. H-statements: H317 (skin sens.) ¦ H334 (resp. sens.) ¦ H340/350/360 (CMR 1A/1B) ¦ H341/351/361 (CMR Cat.2) ¦ H372/373 (STOT RE) | H-statements cover occupational exposure routes. They do not address patient scenarios, migration from finished products, or residual concentrations after manufacturing.
EUH208 (Contains [sensitising substance]) | Supplementary statement: the mixture contains a substance known to cause sensitisation, even below the H317 classification threshold | Sensitisation can occur at trace concentrations (below classification thresholds). EUH208 indicates a known sensitiser is present at any level in the mixture. | No concentration threshold — applies when any skin sensitiser is present, even <0.1% | EUH208 does not identify the sensitiser by name unless voluntarily disclosed. Limits traceability without further supplier information.
IARC Group (Carcinogenicity) | Evidence-based assessment of carcinogenicity: Group 1 (proven human carcinogen) ¦ 2A (probable) ¦ 2B (possible) ¦ 3 (not classifiable) | Independent of regulatory classification. Group 1 and 2A are critical signals. Provides scientific basis for carcinogenicity concern even when CLP classification is not yet harmonised. | Group 1 = CRITICAL ¦ Group 2A = CRITICAL ¦ Group 2B = MAJOR ¦ Group 3 = monitor | IARC classifies hazard, not risk. Group 1 at trace residual levels may pose negligible actual risk — requires dose-response and exposure assessment (ISO 10993-17).
CLP Annex VI ATP22 (Harmonised Classification) | Official EU mandatory CMR and respiratory sensitisation classifications — override supplier self-classifications | If a substance has a harmonised Annex VI classification (CMR, resp. sens.), it must be treated accordingly regardless of what the SDS declares. Closes gaps from incomplete supplier SDS. | Latest: ATP22 (2023). Search by CAS: echa.europa.eu. CMR 1A/1B = CRITICAL ¦ CMR Cat.2 = MAJOR ¦ Resp. Sens. = MAJOR | Covers specific endpoints only — not listed ≠ confirmed safe. Absence from Annex VI means no harmonised classification, not no hazard.
SVHC — REACH Candidate List | Substances of Very High Concern: CMR, PBT/vPvB, ED or equivalent concern — triggers mandatory supply chain communication >0.1% | SVHC listing is a strong regulatory hazard signal. Supplier must declare if >0.1% in mixture/article (REACH Art. 59). Its presence above threshold is an automatic MAJOR/CRITICAL trigger. | Updated regularly at echa.europa.eu/candidate-list-table. Always verify against current list. | Dynamic list — new substances added several times per year. Always check current version. Training knowledge cutoff may miss recent additions.
REACH Annex XVII Entry 43 (Azo dyes) | Restriction on azo colorants that can release carcinogenic aromatic amines by reductive cleavage on skin or saliva contact | Directly relevant for textile/compression hosiery. Certain azo dyes break down on skin to release amines classified as CMR. Presence of an azo dye structure in a product triggers this verification. | 22 restricted aromatic amines. Limit: 30 mg/kg in textile articles in contact with skin. Test method: EN ISO 14362. | Verification requires specific chemical testing (EN ISO 14362) — SDS data alone cannot confirm compliance or non-compliance. CAS-based research needed.
SECTION 6 — ALERT LEVEL CRITERIA: REGULATORY AND SCIENTIFIC BASIS
Alert level | Criterion | Regulatory / scientific basis | ISO 10993-1:2025 endpoint | Marker
🔴 CRITICAL | CMR Cat.1A/1B (Carc., Muta., or Repr.) | CLP Reg. 1272/2008 Cat.1A = known human hazard (human evidence) ¦ Cat.1B = presumed human hazard (strong animal evidence). MDR 2017/745 Annex I §10.4.3 explicitly requires CMR assessment. | Genotoxicity / Carcinogenicity / Reproductive toxicity | 🔴
🔴 CRITICAL | IARC Group 1 (proven human carcinogen) | IARC Group 1 = sufficient evidence of carcinogenicity in humans. Evidence-based, independent of CLP regulation. IARC Monographs are widely accepted by regulatory agencies globally. | Carcinogenicity | 🔴
🔴 CRITICAL | IARC Group 2A (probable human carcinogen) | IARC Group 2A = limited evidence in humans + sufficient evidence in animals. Precautionary principle applied in biocompatibility context: no threshold assumed. | Carcinogenicity | 🔴
🔴 CRITICAL | Endocrine disruptor — EU Official List (Reg. 2021/2045 or 2023/2055) | EU official designation = confirmed ED mechanism in biocide or plant protection context. EU regulatory approach: no safe threshold for ED effects. MDR 2017/745 Annex I §10.4 covers chemical safety. | Endocrine disruption | 🔴
🔴 CRITICAL | Restricted aromatic amine — REACH Annex XVII Entry 43 | 22 aromatic amines restricted in textile articles. Most are IARC Group 1/2A or CLP CMR 1A/1B. Release possible by reductive cleavage of azo bonds on skin surface. Limit: 30 mg/kg (EN ISO 14362). | Genotoxicity / Carcinogenicity (azo dye context) | 🔴
🟠 MAJOR | CMR Cat.2 (Carc., Muta., or Repr.) | CLP Cat.2 = suspected human hazard. Lower confidence than 1A/1B but still requires investigation. Cannot be excluded without further data. MDR Annex I §10.4.3. | Genotoxicity / Carcinogenicity / Reproductive toxicity | 🟠
🟠 MAJOR | REACH Annex XIII PBT/vPvB criteria. High bioaccumulation potential = risk of tissue accumulation during chronic patient exposure (relevant for all long-term and/or direct tissue contact devices). Threshold: log Kow >5 OR BCF >5000. | REACH Annex XIII PBT/vPvB criteria. High bioaccumulation potential = risk of tissue accumulation during chronic patient exposure (compression hosiery = long-term use, compromised skin). | Bioaccumulation | 🟠
🟠 MAJOR | Heavy metal declared in SDS §3 (Ni, Co, Pb, Cd, Hg, As, Cr, Sb, Sn) | Confirmed presence in composition. Known systemic toxicity and/or carcinogenicity (e.g., Cr(VI) = IARC Gr.1 ¦ Ni = IARC Gr.1 ¦ Pb = CMR 1B). EN 1811 (Ni release from articles). REACH Annex XVII. | Relevant endpoint depends on metal identity | 🟠
🟠 MAJOR | STOT RE Cat.1 (H372) — repeated dose target organ toxicity | H372 = causes damage to organs through prolonged or repeated exposure — Cat.1 based on human evidence or significant animal evidence at low exposure. ISO 10993-1:2025 endpoint: chronic / repeated dose toxicity. | Chronic toxicity / Repeated dose | 🟠
🟠 MAJOR | SVHC designated for PBT/vPvB reason | SVHC = Substance of Very High Concern. PBT/vPvB designation triggers REACH supply chain communication >0.1%. Long-term chronic patient exposure scenario amplifies persistence + bioaccumulation concern. | Bioaccumulation / REACH compliance | 🟠
🟡 MINOR | H317 — Skin sensitisation (Skin Sens. 1, 1A, or 1B) | Direct relevance: compression hosiery = prolonged skin contact. Sensitisation is irreversible once established. ISO 10993-1:2025 endpoint: sensitisation. | Sensitisation | 🟡
🟡 MINOR | H314/H315 (skin corrosion/irritation) or H318/H319 (eye damage/irritation) | Direct relevance: damaged/compromised skin contact scenario. Local effects, generally reversible. ISO 10993-1:2025 endpoints: skin/eye irritation. | Skin irritation / Eye irritation | 🟡
🟡 MINOR | H317 skin sensitiser (can induce sensitisation even at trace levels). IARC Group 1 for carcinogenicity via inhalation route (occupational concern — not a primary patient risk via skin contact). Any presence at any concentration requires toxicologist review. Quantitative risk threshold depends on exposure scenario and device contact nature — see ISO 10993-17. | H317 skin sensitiser. IARC Group 1 for carcinogenicity (inhalation route, not textile). EN ISO 14184 textile limit: 75 ppm (direct contact). Below this limit = MINOR; above = reassess. | Sensitisation + Carcinogenicity (at high concentration) | 🟡
🟡 MINOR | Heavy metal mentioned outside SDS §3 (information section only) | Not confirmed in declared composition. Lower confidence in actual presence. Precautionary flag pending further verification. | Relevant endpoint depends on metal | 🟡
SECTION 7 — ROW STATUS DEFINITIONS AND DUPLICATE MANAGEMENT
Status value | Definition | When to use | Action required
Draft | Row filled by AI screening agent — not yet reviewed by a human | Set automatically by AI on every new row | Toxicologist review required before any regulatory use
Validated | Review complete, signed off by a qualified toxicologist or biocompatibility engineer | Set manually after human review of the row | None — approved for use in biological evaluation
Superseded | A newer version of the SDS has been received and a new row created for the same product | Set manually when a newer SDS row replaces this one | Keep row for traceability — do not delete
Archived | Product out of scope or discontinued — no longer screened | Set manually when product is removed from scope | Keep row for traceability — do not delete
DUPLICATE MANAGEMENT RULES
Rule | Description | Detection method | Resolution
Rule 1 — True duplicate | Same Supplier + Product Name + SDS Date = duplicate. Should never occur. | Conditional formatting on col A: =COUNTIF($A:$A,A3)>1 → highlight red | Delete the duplicate row. Keep only one.
Rule 2 — SDS update (same product, new version) | New SDS published. New date = new analysis row. | New row has more recent SDS Date for same product | Add new row ¦ Set old row to 'Superseded' ¦ Note new SDS ID in old row col H
Rule 3 — Same name, different supplier | Same commercial name but different supplier = different product. Not a duplicate. | Check col F (Supplier): different = both valid | Keep both rows. Each row is a distinct product with its own screening record.
Rule 4 — Archived product re-introduced | Product previously archived, now re-introduced. | Check Archived rows before creating new row | Change status from Archived to Draft. Re-submit SDS if date has changed.
SECTION 8 — H-PHRASE REFERENCE LIST (biocompatibility-relevant phrases)
H-phrase | Full text (EN) | CLP Classification category | Biocompatibility endpoint (ISO 10993-1:2025) | Alert level
SENSITISATION
H317 | May cause an allergic skin reaction | Skin Sens. 1/1A/1B | Sensitisation | 🟡 MINOR
H334 | May cause allergy or asthma symptoms or breathing difficulties if inhaled | Resp. Sens. 1/1A/1B | Sensitisation (respiratory) | 🟠 MAJOR
EUH208 | Contains [name of sensitising substance]. May produce an allergic reaction. | Supplementary | Sensitisation | 🟡 MINOR (trace level)
SKIN / EYE IRRITATION
H314 | Causes severe skin burns and eye damage | Skin Corr. 1/1A/1B/1C | Skin irritation / Intracutaneous reactivity | 🟡 MINOR
H315 | Causes skin irritation | Skin Irrit. 2 | Skin irritation | 🟡 MINOR
H318 | Causes serious eye damage | Eye Dam. 1 | Eye irritation | 🟡 MINOR
H319 | Causes serious eye irritation | Eye Irrit. 2 | Eye irritation | 🟡 MINOR
ACUTE SYSTEMIC TOXICITY
H300 | Fatal if swallowed | Acute Tox. 1/2 (oral) | Acute systemic toxicity | 🟠 MAJOR
H301 | Toxic if swallowed | Acute Tox. 3 (oral) | Acute systemic toxicity | 🟠 MAJOR
H302 | Harmful if swallowed | Acute Tox. 4 (oral) | Acute systemic toxicity | 🟠 MAJOR
H310 | Fatal in contact with skin | Acute Tox. 1/2 (dermal) | Acute systemic toxicity | 🟠 MAJOR
H311 | Toxic in contact with skin | Acute Tox. 3 (dermal) | Acute systemic toxicity | 🟠 MAJOR
H312 | Harmful in contact with skin | Acute Tox. 4 (dermal) | Acute systemic toxicity | 🟠 MAJOR
H330 | Fatal if inhaled | Acute Tox. 1/2 (inhalation) | Acute systemic toxicity | 🟠 MAJOR
H331 | Toxic if inhaled | Acute Tox. 3 (inhalation) | Acute systemic toxicity | 🟠 MAJOR
H332 | Harmful if inhaled | Acute Tox. 4 (inhalation) | Acute systemic toxicity | Note only — inhalation route not relevant for patient
STOT — SPECIFIC TARGET ORGAN TOXICITY
H370 | Causes damage to organs (STOT SE Cat.1) | STOT SE 1 | Sub-acute / Sub-chronic toxicity | 🟠 MAJOR
H371 | May cause damage to organs (STOT SE Cat.2) | STOT SE 2 | Sub-acute / Sub-chronic toxicity | 🟠 MAJOR
H335 | May cause respiratory irritation (STOT SE Cat.3) | STOT SE 3 | Inhalation route only | Note only — irrelevant for patient
H336 | May cause drowsiness or dizziness (STOT SE Cat.3) | STOT SE 3 | Inhalation route only | Note only — irrelevant for patient
H372 | Causes damage to organs through prolonged or repeated exposure (STOT RE Cat.1) | STOT RE 1 | Chronic toxicity / Repeated dose | 🟠 MAJOR
H373 | May cause damage to organs through prolonged or repeated exposure (STOT RE Cat.2) | STOT RE 2 | Chronic toxicity / Repeated dose | 🟡 MINOR
CMR — MUTAGENICITY
H340 | May cause genetic defects | Muta. 1A/1B | Genotoxicity / Mutagenicity | 🔴 CRITICAL
H341 | Suspected of causing genetic defects | Muta. 2 | Genotoxicity / Mutagenicity | 🟠 MAJOR
CMR — CARCINOGENICITY
H350 | May cause cancer | Carc. 1A/1B | Carcinogenicity | 🔴 CRITICAL
H350i | May cause cancer by inhalation | Carc. 1A/1B (inhalation) | Carcinogenicity | 🔴 CRITICAL — note inhalation route
H351 | Suspected of causing cancer | Carc. 2 | Carcinogenicity | 🟠 MAJOR
CMR — REPRODUCTIVE TOXICITY
H360 | May damage fertility or the unborn child | Repr. 1A/1B | Reproductive / Developmental toxicity | 🔴 CRITICAL
H360F | May damage fertility | Repr. 1A/1B (fertility) | Reproductive / Developmental toxicity | 🔴 CRITICAL
H360D | May damage the unborn child | Repr. 1A/1B (developmental) | Reproductive / Developmental toxicity | 🔴 CRITICAL
H360FD | May damage fertility. May damage the unborn child. | Repr. 1A/1B (both) | Reproductive / Developmental toxicity | 🔴 CRITICAL
H361 | Suspected of damaging fertility or the unborn child | Repr. 2 | Reproductive / Developmental toxicity | 🟠 MAJOR
H361f | Suspected of damaging fertility | Repr. 2 (fertility) | Reproductive / Developmental toxicity | 🟠 MAJOR
H361d | Suspected of damaging the unborn child | Repr. 2 (developmental) | Reproductive / Developmental toxicity | 🟠 MAJOR
H362 | May cause harm to breast-fed children | Lact. (supplementary) | Reproductive toxicity — lactation | 🟠 MAJOR
ENVIRONMENTAL / BIOACCUMULATION PROXY
H410 | Very toxic to aquatic life with long lasting effects | Aquatic Chronic 1 | Bioaccumulation proxy | 🟠 MAJOR — log Kow typically >5 — verify
H411 | Toxic to aquatic life with long lasting effects | Aquatic Chronic 2 | Bioaccumulation proxy | 🟡 MINOR signal — check log Kow
H412 | Harmful to aquatic life with long lasting effects | Aquatic Chronic 3 | Bioaccumulation proxy | Note only — check log Kow
H413 | May cause long lasting harmful effects to aquatic life | Aquatic Chronic 4 | Bioaccumulation proxy | Note only — check log Kow
EUH070 | Toxic by eye contact | Supplementary EU-only | Eye irritation | 🟡 MINOR
NOTE | H-phrases for PHYSICAL HAZARDS ONLY (H220–H290: flammable, explosive, oxidizing, gases...) are OUT OF SCOPE for biocompatibility evaluation.

## Agent rules (generic)
METHODOLOGY (AI) — EVERGREEN RULES & VOCABULARY
Regulatory basis: MDR 2017/745 Annex I ¦ ISO 10993-1:2025 ¦ REACH Reg.1907/2006 ¦ CLP Reg.1272/2008 ¦ IARC Monographs
SECTION 1 - OVERALL ALERT LEVEL RULES (col AC — ALERT LEVEL)
Rule: WORST CASE across ALL declared substances ¦ Alert level based on ACTUAL signals only ¦ See default rules below for unverified data
LEVEL | CODE | TRIGGER CONDITION | REGULATORY BASIS
CRITICAL | C1 | CMR Cat.1A or 1B — any endpoint (Carc./Muta./Repr.) — confirmed in CLP Annex VI ATP22 or stated in SDS §2 | CLP Reg.1272/2008 Annex VI ¦ MDR Annex I §10.4.3
CRITICAL | C2 | Endocrine Disruptor CONFIRMED — EU Official List (Reg.2021/2045 biocides, Reg.2023/2055 PPP, or SVHC for ED reason in REACH Candidate List) | Reg.2021/2045 ¦ Reg.2023/2055 ¦ REACH Art.57f
CRITICAL | C3 | Carcinogenic aromatic amine — REACH Annex XVII Entry 43 (azo dye cleavage to listed amine) | REACH Annex XVII Entry 43 ¦ EN ISO 14362-1
CRITICAL | C4 | Chromium VI (CrVI) — any concentration — declared or identified in SDS or research | IARC Vol.100C ¦ REACH Annex XVII Entry 47 ¦ EN ISO 17075
CRITICAL | C5 | IARC Group 1 — proven human carcinogen | IARC Monographs Program
CRITICAL | C6 | IARC Group 2A — probable human carcinogen (sufficient animal evidence, limited human evidence) | IARC Monographs Program
MAJOR | M1 | CMR Cat.2 — any endpoint (Carc./Muta./Repr.) — suspected CMR | CLP Reg.1272/2008 ¦ MDR Annex I §10
MAJOR | M2 | Endocrine Disruptor SUSPECTED — TEDX database only (not EU Official list) — non-regulatory source | TEDX (The Endocrine Disruption Exchange) — non-regulatory
MAJOR | M3 | SVHC for PBT or vPvB reason — ECHA REACH Candidate List | REACH Art.57d/57e ¦ REACH Annex XIII
MAJOR | M4 | Bioaccumulation HIGH — log Kow >5 OR BCF >5000 L/kg OR confirmed tissue accumulation (metals/inorganics) | REACH Annex XIII ¦ ISO 10993-1:2025 §6.4.3
MAJOR | M5 | IARC Group 2B — possibly carcinogenic to humans | IARC Monographs Program
MAJOR | M6 | Heavy metal declared in SDS Section 3 — Ni / Co / Pb / Cd / Hg / As / Cr (total) / Sb / Sn | EN ISO 1811 ¦ ISO 10993-17 ¦ REACH Annex XVII
MAJOR | M7 | Respiratory sensitiser H334 confirmed — inhalation route (conservative flag regardless of device exposure route) | CLP Reg.1272/2008 ¦ GHS
MAJOR | M8 | STOT RE Cat.1 H372 — repeated exposure, specific organ damage confirmed | CLP Reg.1272/2008 ¦ GHS
MAJOR | M9 | Manufacturer restriction for medical device / body contact / food contact use stated in SDS §7/§15/§16 | Supplier SDS — flagged for toxicologist review
MAJOR | M10 | Product identified as colorant/dye AND REACH Annex XVII Entry 43 aromatic amine status = unverified or not stated in SDS | REACH Annex XVII Entry 43 ¦ EN ISO 14362-1
MAJOR | M11 | Product identified as silicone-based AND D4/D5 cyclic siloxanes (REACH Annex XVII Entries 70/71) absence not attested | REACH Annex XVII Entries 70/71
MINOR | Mi1 | H317 — skin sensitiser — Cat.1/1A/1B confirmed | CLP Reg.1272/2008 ¦ GHS ¦ ISO 10993-10
MINOR | Mi2 | STOT SE Cat.3 H336 (narcotic effects) OR STOT RE Cat.2 H373 (possible repeated exposure damage) | CLP Reg.1272/2008 ¦ GHS
MINOR | Mi3 | Skin/eye irritant — H314 (corrosive) / H315 (skin irritation) / H318 (serious eye damage) / H319 (eye irritation) | CLP Reg.1272/2008 ¦ GHS ¦ ISO 10993-10
MINOR | Mi4 | Bioaccumulation MODERATE — log Kow 4.5 to 5 (borderline accumulation potential) | REACH Annex XIII ¦ ISO 10993-1:2025 §6.4.3
MINOR | Mi5 | EUH208 — contains [named sensitising substance] — below classification threshold but sensitisation potential present | CLP Reg.1272/2008 ¦ GHS
MINOR | Mi6 | Formaldehyde or formaldehyde-releasing agent declared at any concentration — H317 sensitiser ¦ IARC Group 1 (inhalation route) | CLP Reg.1272/2008 ¦ ISO 10993-17 ¦ IARC Monographs
MINOR | Mi7 | Heavy metal mentioned in SDS for information only (not declared in §3) — monitoring signal only | SDS data ¦ ISO 10993-17
NONE | — | No criteria C1-C6 / M1-M11 / Mi1-Mi7 triggered after complete screening of cols L to AH | All screening columns
NOT AN SDS | — | Document submitted is not an SDS compliant with REACH Reg.1907/2006 / GHS (e.g. TDS, product spec, certificate) | Step 0 validation
DEFAULT RULES — unverified data
CMR 1A/1B or ED confirmed: potentially present but not yet verified | → CRITICAL (provisional) | Add 'To verify — confirm in CLP ATP22 / ECHA Candidate List' in col AK
CMR Cat.2 or ED suspected: potentially present but not yet verified | → MAJOR (provisional) | Add 'To verify — confirm in ATP22 / TEDX / SVHC' in col AK
SECTION 2 — COMPOSITION NOT DISCLOSED
Applies when: SDS §3 contains no CAS numbers (composition confidential or mixture with no substance breakdown)
CONDITION | No CAS numbers in §3 AND no individual substance names provided
ALERT RULE | Based on ACTUAL SDS signals in cols O-Z ONLY ¦ Do NOT escalate to CRITICAL/MAJOR solely because CAS numbers are absent
REGULATORY BASIS | CLP Reg.1272/2008: CMR 1A/1B ≥0.1% must be declared ¦ CMR Cat.2 ≥1% ¦ REACH Art.59: SVHC ≥0.1% must be declared
LEGAL CONSEQUENCE | Absence of declaration = legal obligation met by supplier = basis for limited confidence. Not proof of absence of hazard.
RESEARCH COLS AA-AH | Write ONE consolidated note in col AA only
Cols AB to AH: write exactly 'See col. AA — composition not disclosed'
Do NOT write 'To verify' in each research column individually
FORBIDDEN ACTION (col AK) | Do NOT write: 'Request full substance disclosure from supplier'
CAS disclosure is not required by REACH and is legally protected as trade secret
Do NOT write: 'Request CAS numbers from supplier' or any equivalent
CONCLUSION FORMAT | No individual substances identifiable — research not possible ¦ No CMR signal (§2) ¦ No SVHC >0.1% (§15 — supplier statement) ¦ CLP mandatory disclosure applied — limited confidence ¦ ACCEPTABLE (limited data)'
SECTION 3 - CONTROLLED VOCABULARY
Values marked PROJECT-SPECIFIC must be defined in the client tab Methodology (AI, [ClientName]). Universal values apply to all projects.
COLUMN | RULE | DEFINITION
Product Type | PROJECT-SPECIFIC - see Methodology (AI, [ClientName]) tab | Describes the role of the chemical in the manufacturing process from a patient exposure perspective. Values must be defined per project and industry context.
| Universal fallback: To verify | Use when product type cannot be determined from SDS data alone - requires manual classification
ALERT LEVEL | Universal - CRITICAL | One or more C1-C6 criteria triggered
| Universal - MAJOR | One or more M1-M11 triggered, no CRITICAL
| Universal - MINOR | One or more Mi1-Mi7 triggered, no CRITICAL or MAJOR
| Universal - NONE | No criteria triggered after complete screening
| Universal - NOT AN SDS | Document is not an SDS compliant with REACH Reg.1907/2006 / GHS
Row Status | Universal - Draft | Set by AI agent - not yet reviewed by a human
| Universal - Validated | Review complete - signed off by qualified toxicologist or biocompatibility engineer
| Universal - Superseded | Newer SDS version received and new row created for same product
| Universal - Archived | Product out of scope or discontinued - row kept for traceability
| Validated | Review complete - signed off by qualified toxicologist or biocompatibility engineer
| Superseded | Newer SDS version received and new row created for same product
| Archived | Product out of scope or discontinued - row kept for traceability
Source tags | [SDS sX] | Data extracted from SDS document, section X - mandatory for all SDS-sourced data
| [Research - source] | Data from external database or training knowledge - specify source name
| [Research, CLP ATP22] | CLP Annex VI ATP22 harmonized classification - lookup tab in this workbook
| [Research, IARC_Groups] | IARC Monographs carcinogenicity groups - lookup tab in this workbook
| [Research, REACH_AXVII_Ent43] | REACH Annex XVII Entry 43 (azo dyes/aromatic amines) - lookup tab
| [Research, EU_ED_Official] | EU Official Endocrine Disruptor lists (Reg.2021/2045 + Reg.2023/2055) - lookup tab
| [Research, TEDX] | The Endocrine Disruption Exchange - non-regulatory suspected ED list - lookup tab
| [Research, SVHC_Candidates] | ECHA REACH Candidate List - requires live verification at echa.europa.eu - training knowledge
SECTION 4 - ALERT LEVEL FORMULA (column: ALERT LEVEL)
LOGIC | Read markers (RED / ORANGE / YELLOW / dash) across ALL columns from Cytotoxicity to REACH/SVHC compliance
Apply worst case: RED present -> CRITICAL ¦ ORANGE (no RED) -> MAJOR ¦ YELLOW (no RED/ORANGE) -> MINOR ¦ else -> NONE
FORMULA PATTERN | COUNTIF across columns [Cytotoxicity] to [REACH/SVHC compliance] searching for RED marker -> CRITICAL ¦ then ORANGE -> MAJOR ¦ then YELLOW -> MINOR ¦ else NONE | Use column names in structured table references
SPECIAL | If document is not an SDS: write NOT AN SDS directly in column ALERT LEVEL - no formula applied
CONCLUSION TEMPLATES (column: Alert Justification)
NOT ACCEPTABLE | C1 - CMR Cat.1B (Carc.) - [substance] - NOT ACCEPTABLE ¦ Further quantitative risk assessment required (ISO 10993-17)
PENDING | M1 - CMR Cat.2 (Repr.) - [substance] + M6 - Heavy metal declared (Ni) ¦ PENDING ¦ Toxicologist review required before regulatory use
MINOR | Mi1 - H317 skin sensitiser + Mi3 - Eye irritant H319 ¦ MINOR ¦ Acceptable pending toxicologist review of patient exposure scenario
NONE | No hazard criteria triggered - complete screening performed columns Cytotoxicity to REACH/SVHC compliance ¦ ACCEPTABLE
NOT AN SDS | Document type: [specify] - NOT AN SDS ¦ Obtain compliant EU CLP SDS from supplier before screening
UNDISCLOSED | No individual substances identifiable - research not possible ¦ No CMR signal (s2) ¦ No SVHC >0.1% (s15) ¦ CLP mandatory disclosure applied - limited confidence ¦ ACCEPTABLE (limited data)
SPECIAL NOTE - col Residual on fiber | If column Residual on fiber = No - Sigvaris confirmed: append to Alert Justification -> Patient exposure unlikely - rinsed during process

## CLP Annex VI ATP22
ECHA - ATP 22
Harmonised classification and labelling, if applicable, as per the table in CLP Annex VI Part 3.
DISCLAIMER: This is an unofficial excel table containing the substances with harmonised classification and labelling up until the 22nd Adaptation to Technical Progress, i.e. Commission Delegated Regulation (EU) No 2024/2564 of 19 June 2024 amending Regulation (EC) No 1272/2008 of the European Parliament and of the Council as regards the harmonised classification and labelling of certain substances. ECHA assumes no responsibility of any kind for your use of the data inter alia in the context of fulfilling the requirements of the relevant legal obligations. Please note that this document should be treated only for informative purposes not including commercial activities or reproduction
ATP 22
|  |  |  | Classification |  | Labelling
Index No | International Chemical Identification | EC No | CAS No | Hazard Class and Category Code(s) | Hazard Statement Code(s) | Pictogram, Signal Word Code(s) | Hazard statement Code(s)2 | Suppl. Hazard statement Code(s) | Specific Conc. Limits, M-factors | Notes | ATP inserted/ATP Updated
001-001-00-9 | hydrogen | 215-605-7 | 1333-74-0 | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | U | CLP00
001-002-00-4 | aluminium lithium hydride | 240-877-9 | 16853-85-3 | Water-react. 1 Skin Corr. 1A | H260 H314 | GHS02 GHS05 Dgr | H260 H314 |  |  |  | CLP00/ATP01
001-003-00-X | sodium hydride | 231-587-3 | 7646-69-7 | Water-react. 1 | H260 | GHS02 Dgr | H260 |  |  |  | CLP00
001-004-00-5 | calcium hydride | 232-189-2 | 7789-78-8 | Water-react. 1 | H260 | GHS02 Dgr | H260 |  |  |  | CLP00
003-001-00-4 | lithium | 231-102-5 | 7439-93-2 | Water-react. 1 Skin Corr. 1B | H260 H314 | GHS02 GHS05 Dgr | H260 H314 | EUH014 |  |  | CLP00
003-002-00-X | n-hexyllithium | 404-950-0 | 21369-64-2 | Pyr. Sol. 1 Water-react. 1 Skin Corr. 1A | H250 H260 H314 | GHS02 GHS05 Dgr | H260 H250 H314 | EUH014 |  |  | CLP00
003-003-00-5 | (2-methylpropyl)lithium; isobutyllithium | 440-620-2 | 920-36-5 | Pyr. Liq. 1 Water-react. 1 STOT SE 3 Skin Corr. 1A Aquatic Acute 1 Aquatic Chronic 1 | H250 H260 H336 H314 H400 H410 | GHS02 GHS05 GHS07 GHS09 Dgr | H260 H250 H314 H336 H410 | EUH014 |  |  | ATP01
004-001-00-7 | beryllium | 231-150-7 | 7440-41-7 | Carc. 1B Acute Tox. 2 * Acute Tox. 3 * STOT SE 3 STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H350i H330 H301 H335 H372 ** H315 H319 H317 | GHS06 GHS08 Dgr | H350i H330 H301 H372 ** H319 H335 H315 H317 |  |  |  | CLP00
004-002-00-2 | beryllium compounds with the exception of aluminium beryllium silicates, and with those specified elsewhere in this Annex |  |  | Carc. 1B Acute Tox. 2 * Acute Tox. 3 * STOT SE 3 STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H350i H330 H301 H335 H372 ** H315 H319 H317 H411 | GHS06 GHS08 GHS09 Dgr | H350i H330 H301 H372 ** H319 H335 H315 H317 H411 |  |  | A | CLP00
004-003-00-8 | beryllium oxide | 215-133-1 | 1304-56-9 | Carc. 1B Acute Tox. 2 * Acute Tox. 3 * STOT SE 3 STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H350i H330 H301 H335 H372 ** H315 H319 H317 | GHS06 GHS08 Dgr | H350i H330 H301 H372 ** H319 H335 H315 H317 |  |  |  | CLP00
005-001-00-X | boron trifluoride | 231-569-5 | 7637-07-2 | Press. Gas Acute Tox. 2 * Skin Corr. 1A | H330 H314 | GHS04 GHS06 GHS05 Dgr | H330 H314 | EUH014 |  | U | CLP00
005-002-00-5 | boron trichloride | 233-658-4 | 10294-34-5 | Press. Gas Acute Tox. 2 * Acute Tox. 2 * Skin Corr. 1B | H330 H300 H314 | GHS04 GHS06 GHS05 Dgr | H330 H300 H314 | EUH014 |  | U | CLP00
005-003-00-0 | boron tribromide | 233-657-9 | 10294-33-4 | Acute Tox. 2 * Acute Tox. 2 * Skin Corr. 1A | H330 H300 H314 | GHS06 GHS05 Dgr | H330 H300 H314 | EUH014 |  |  | CLP00
005-004-00-6 | trialkylboranes |  |  | Pyr. Sol. 1 Skin Corr. 1B | H250 H314 | GHS02 GHS05 Dgr | H250 H314 |  |  | A | CLP00
005-004-01-3 | trialkylboranes, liquid |  |  | Pyr. Liq. 1 Skin Corr. 1B | H250 H314 | GHS02 GHS05 Dgr | H250 H314 |  |  | A | CLP00
005-005-00-1 | trimethyl borate | 204-468-9 | 121-43-7 | Flam. Liq. 3 Repr. 1B Acute Tox. 4* | H226 H360FD H312 | GHS02 GHS08 GHS07 Dgr | H226 H360FD H312 |  |  | 11 | CLP00/ATP22
005-006-00-7 | dibutyltin hydrogen borate | 401-040-5 | 75113-37-0 | Muta. 2 Repr. 1B Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H341 H360FD H312 H302 H372 ** H318 H317 H400 H410 | GHS05 GHS08 GHS07 GHS09 Dgr | H302 H312 H318 H317 H341 H360FD H372 ** H410 |  |  |  | CLP00/ATP01corr
005-007-00-2 | boric acid [1] boric acid [2] | 233-139-2 [1] 234-343-4 [2] | 10043-35-3 [1] 11113-50-1 [2] | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  | 11 | ATP01/ATP20
005-008-00-8 | diboron trioxide | 215-125-8 | 1303-86-2 | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  | 11 | ATP01/ATP20
005-009-00-3 | tetrabutylammonium butyltriphenylborate | 418-080-4 | 120307-06-4 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
005-010-00-9 | N,N-dimethylanilinium tetrakis(pentafluorophenyl)borate | 422-050-6 | 118612-00-3 | Carc. 2 Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 | H351 H302 H315 H318 | GHS08 GHS05 GHS07 Dgr | H351 H302 H315 H318 |  |  |  | CLP00
005-011-00-4 | tetraboron disodium heptaoxide, hydrate [1] disodium tetraborate, anhydrous [2] orthoboric acid, sodium salt [3] disodium tetraborate decahydrate [4] disodium tetraborate pentahydrate [5] | 235-541-3 [1] 215-540-4 [2] 237-560-2 [3] 215-540-4 [4] 215-540-4 [5] | 12267-73-1 [1] 1330-43-4 [2] 13840-56-7 [3] 1303-96-4 [4] 12179-04-3 [5] | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  | 11 | ATP01/ATP20
005-012-00-X | diethyl{}{4-[1,5,5-tris(4-diethylaminophenyl)penta-2,4-dienylidene]cyclohexa-2,5-dienylidene}}ammonium butyltriphenylborate | 418-070-1 | 141714-54-7 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
005-013-00-5 | diethylmethoxyborane | 425-380-9 | 7397-46-8 | Pyr. Liq. 1 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 4 | H250 H332 H312 H302 H373 ** H314 H317 H413 | GHS02 GHS05 GHS08 GHS07 Dgr | H250 H332 H312 H302 H373 ** H314 H317 H413 |  |  |  | ATP01
005-014-00-0 | 4-formylphenylboronic acid | 438-670-5 | 87199-17-5 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
005-015-00-6 | 1-chloromethyl-4-fluoro-1,4-diazoniabicyclo[2.2.2]octane bis(tetrafluoroborate) | 414-380-4 | 140681-55-6 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H302 H318 H317 H412 | GHS05 GHS07 Dgr | H302 H318 H317 H412 |  |  |  | ATP01
005-016-00-1 | tetrabutylammonium butyl tris-(4-tert-butylphenyl)borate | 431-370-5 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
005-020-00-3 | disodium octaborate anhydrous [1] disodium octaborate tetrahydrate [2] | 234-541-0 [1] 234-541-0 [2] | 12008-41-2 [1] 12280-03-4 [2] | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  |  | ATP09
005-022-00-4 | perboric acid, sodium salt [1] perboric acid, sodium salt, monohydrate [2] perboric acid (HBO(O 2 )), sodium salt, monohydrate [3] sodium peroxoborate [4] sodium perborate [5] | 234-390-0 [1] 234-390-0 [2] - [3] - [4] 239-172-9 [5] | 11138-47-9 [1] 12040-72-1 [2] 10332-33-9 [3] - [4] 15120-21-5 [5] | Ox. Sol. 3 Repr. 1B Acute Tox. 3 Acute Tox. 4 STOT SE 3 Eye Dam. 1 | H272 H360FD H331 H302 H335 H318 | GHS03 GHS08 GHS06 GHS05 Dgr | H272 H360FD H331 H302 H335 H318 |  | inhalation: ATE = 0,75 mg/L (dusts or mists) oral: ATE = 890 mg/kg bw Eye Dam. 1; H318: C ≥ 22 % Eye Irrit. 2; H319: 14 % ≤ C < 22 % | 11 | ATP22
005-023-00-X | perboric acid (H 3 BO 2(O2 )), monosodium salt trihydrate [1] perboric acid, sodium salt, tetrahydrate [2] perboric acid (HBO(O 2 )), sodium salt, tetrahydrate [3] sodium peroxoborate, hexahydrate [4] | 239-172-9 [1] 234-390-0 [2] - [3] - [4] | 13517-20-9 [1] 37244-98-7 [2] 10486-00-7 [3] - [4] | Repr. 1B Acute Tox. 4 STOT SE 3 Eye Dam. 1 | H360FD H332 H335 H318 | GHS08 GHS05 GHS07 Dgr | H360FD H332 H335 H318 |  | inhalation: ATE = 1,2 mg/L (dusts or mists) Eye Dam. 1; H318: C ≥ 36 % Eye Irrit. 2; H319: 22 % ≤ C < 36 % | 11 | ATP22
005-024-00-5 | sodium peroxometaborate | 231-556-4 | 7632-04-4 | Ox. Sol. 2 Repr. 1B Acute Tox. 3 Acute Tox. 4 STOT SE 3 Eye Dam. 1 | H272 H360FD H331 H302 H335 H318 | GHS03 GHS08 GHS06 GHS05 Dgr | H272 H360FD H331 H302 H335 H318 |  | inhalation: ATE = 0,62 mg/L (dusts or mists) oral: ATE = 730 mg/kg bw Eye Dam. 1; H318: C ≥ 22 % Eye Irrit. 2; H319: 14 % ≤ C < 22 % | 11 | ATP22
006-001-00-2 | carbon monoxide | 211-128-3 | 630-08-0 | Flam. Gas 1 Press. Gas Repr. 1A Acute Tox. 3 * STOT RE 1 | H220 H360D *** H331 H372 ** | GHS02 GHS04 GHS06 GHS08 Dgr | H220 H360D *** H331 H372 ** |  |  | U | CLP00
006-002-00-8 | phosgene; carbonyl chloride | 200-870-3 | 75-44-5 | Press. Gas Acute Tox. 2 * Skin Corr. 1B | H330 H314 | GHS04 GHS06 GHS05 Dgr | H330 H314 |  |  | U | CLP00
006-003-00-3 | carbon disulphide | 200-843-6 | 75-15-0 | Flam. Liq. 2 Repr. 2 STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 | H225 H361fd H372 ** H315 H319 | GHS02 GHS08 GHS07 Dgr | H225 H361fd H372 ** H319 H315 |  | Repr. 2; H361fd: C ≥ 1 % STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,2 % ≤C < 1 % |  | CLP00
006-004-00-9 | calcium carbide | 200-848-3 | 75-20-7 | Water-react. 1 | H260 | GHS02 Dgr | H260 |  |  | T | CLP00
006-005-00-4 | thiram (ISO); tetramethylthiuram disulphide | 205-286-2 | 137-26-8 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H373 ** H315 H319 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H332 H302 H373 ** H319 H315 H317 H410 |  | M=10 |  | CLP00
006-006-00-X | hydrogen cyanide; hydrocyanic acid | 200-821-6 | 74-90-8 | Flam. Liq. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H224 H330 H400 H410 | GHS02 GHS06 GHS09 Dgr | H224 H330 H410 |  |  |  | CLP00
006-006-01-7 | hydrogen cyanide ...%; hydrocyanic acid ...% | 200-821-6 | 74-90-8 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H400 H410 | GHS06 GHS09 Dgr | H330 H310 H300 H410 |  |  | B | CLP00
006-007-00-5 | salts of hydrogen cyanide with the exception of complex cyanides such as ferrocyanides, ferricyanides and mercuric oxycyanide and those specified elsewhere in this Annex | - | - | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H400 H410 | GHS06 GHS09 Dgr | H330 H310 H300 H410 | EUH032 |  | A | CLP00/ATP01
006-008-00-0 | antu (ISO); 1-(1-naphthyl)-2-thiourea | 201-706-3 | 86-88-4 | Carc. 2 Acute Tox. 2 * | H351 H300 | GHS06 GHS08 Dgr | H300 H351 |  |  |  | CLP00
006-009-00-6 | 1-isopropyl-3-methylpyrazol-5-yl dimethylcarbamate; Isolan | 204-318-2 | 119-38-0 | Acute Tox. 1 Acute Tox. 2 * | H310 H300 | GHS06 Dgr | H310 H300 |  |  |  | CLP00
006-010-00-1 | 5,5-dimethyl-3-oxocyclohex-1-enyl dimethylcarbamate 5,5-dimethyldihydroresorcinol dimethylcarbamate; Dimetan | 204-525-8 | 122-15-6 | Acute Tox. 3 * | H301 | GHS06 Dgr | H301 |  |  |  | CLP00
006-011-00-7 | carbaryl (ISO); 1-naphthyl methylcarbamate | 200-555-0 | 63-25-2 | Carc. 2 Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 | H351 H332 H302 H400 | GHS08 GHS07 GHS09 Wng | H351 H332 H302 H400 |  | M=100 |  | CLP00/ATP01
006-012-00-2 | ziram (ISO); zinc bis dimethyldithiocarbamate | 205-288-3 | 137-30-4 | Acute Tox. 2 * Acute Tox. 4 * STOT SE 3 STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H302 H335 H373 ** H318 H317 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H330 H302 H373 ** H335 H318 H317 H410 |  | M=100 |  | CLP00
006-013-00-8 | metam-sodium (ISO); sodium methyldithiocarbamate | 205-293-0 | 137-42-8 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H314 H317 H410 | EUH031 |  |  | CLP00
006-014-00-3 | nabam (ISO); disodium ethylenebis(N,N'-dithiocarbamate) | 205-547-0 | 142-59-6 | Acute Tox. 4 * STOT SE 3 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H335 H317 H400 H410 | GHS07 GHS09 Wng | H302 H335 H317 H410 |  |  |  | CLP00
006-015-00-9 | diuron (ISO); 3-(3,4-dichlorophenyl)-1,1-dimethylurea | 206-354-4 | 330-54-1 | Carc. 1B STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H350 H373 (blood system) H400 H410 | GHS08 GHS09 Dgr | H350 H373 (blood system) H410 |  | M = 100 M = 100 |  | CLP00/ATP21
006-016-00-4 | propoxur (ISO); 2-isopropyloxyphenyl N-methylcarbamate; 2-isopropoxyphenyl methylcarbamate | 204-043-8 | 114-26-1 | Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H400 H410 | GHS06 GHS09 Dgr | H301 H410 |  |  |  | CLP00
006-017-00-X | aldicarb (ISO); 2-methyl-2-(methylthio)propanal-O-(N-methylcarbamoyl)oxime | 204-123-2 | 116-06-3 | Acute Tox. 2 * Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H330 H300 H311 H400 H410 | GHS06 GHS09 Dgr | H330 H300 H311 H410 |  |  |  | CLP00
006-018-00-5 | aminocarb (ISO); 4-dimethylamino-3-tolyl methylcarbamate | 217-990-7 | 2032-59-9 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H311 H301 H400 H410 | GHS06 GHS09 Dgr | H311 H301 H410 |  |  |  | CLP00
006-019-00-0 | di-allate (ISO); S-(2,3-dichloroallyl)-N,N-diisopropylthiocarbamate | 218-961-1 | 2303-16-4 | Carc. 2 Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H302 H410 |  |  |  | CLP00
006-020-00-6 | barban (ISO); 4-chlorbut-2-ynyl N-(3-chlorophenyl)carbamate | 202-930-4 | 101-27-9 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
006-021-00-1 | linuron (ISO); 3-(3,4-dichlorophenyl)-1-methoxy-1-methylurea | 206-356-5 | 330-55-2 | Carc. 2 Repr. 1B Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H360Df H302 H373 ** H400 H410 | GHS08 GHS07 GHS09 Dgr | H360Df H351 H302 H373 ** H410 |  |  |  | CLP00
006-022-00-7 | decarbofuran (ISO); 2,3-dihydro-2-methylbenzofuran-7-yl methylcarbamate |  | 1563-67-3 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H331 H311 H301 | GHS06 Dgr | H331 H311 H301 |  |  |  | CLP00
006-023-00-2 | mercaptodimethur (ISO); methiocarb (ISO); 3,5-dimethyl-4-methylthiophenyl N-methylcarbamate | 217-991-2 | 2032-65-7 | Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H400 H410 | GHS06 GHS09 Dgr | H301 H410 |  |  |  | CLP00
006-024-00-8 | proxan-sodium (ISO); sodium O-isopropyldithiocarbonate | 205-443-5 | 140-93-2 | Acute Tox. 4 * Skin Irrit. 2 Aquatic Chronic 2 | H302 H315 H411 | GHS07 GHS09 Wng | H302 H315 H411 |  |  |  | CLP00
006-025-00-3 | allethrin; (RS)-3-allyl-2-methyl-4-oxocyclopent-2-enyl (1RS,3RS;1RS,3SR)-2,2-dimethyl-3-(2-methylprop-1-enyl)cyclopropanecarboxylate; bioallethrin; (RS)-3-allyl-2-methyl-4-oxocyclopent-2-enyl (1R,3R)-2,2-dimethyl-3-(2-methylprop-1-enyl)cyclopropanecarboxylate [1] S-bioallethrin; (S)-3-allyl-2-methyl-4-oxocyclopent-2-enyl (1R,3R)-2,2-dimethyl-3-(2-methylprop-1-enyl)cyclopropanecarboxylate [2] esbiothrin; (RS)-3-allyl-2-methyl-4-oxocyclopent-2-enyl (1R,3R)-2,2-dimethyl-3-(2-methylprop-1-enyl)cyclopropanecarboxylate [3] | 209-542-4 [1] 249-013-5 [2] 249-013-5 [3] | 584-79-2 [1] 28434-00-6 [2] 84030-86-4 [3] | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H400 H410 | GHS07 GHS09 Wng | H332 H302 H410 |  |  | C | CLP00
006-026-00-9 | carbofuran (ISO); 2,3-dihydro-2,2-dimethylbenzofuran-7-yl N-methylcarbamate | 216-353-0 | 1563-66-2 | Acute Tox. 2 * Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H330 H300 H400 H410 | GHS06 GHS09 Dgr | H330 H300 H410 |  |  |  | CLP00
006-028-00-X | dinobuton (ISO); 2-(1-methylpropyl)-4,6-dinitrophenyl isopropyl carbonate | 213-546-1 | 973-21-7 | Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H400 H410 | GHS06 GHS09 Dgr | H301 H410 |  |  |  | CLP00
006-029-00-5 | dioxacarb (ISO); 2-(1,3-dioxolan-2-yl)phenyl N-methylcarbamate | 230-253-4 | 6988-21-2 | Acute Tox. 3 * Aquatic Chronic 2 | H301 H411 | GHS06 GHS09 Dgr | H301 H411 |  |  |  | CLP00
006-030-00-0 | EPTC (ISO); S-ethyl dipropylthiocarbamate | 212-073-8 | 759-94-4 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
006-031-00-6 | formetanate (ISO); 3-[(EZ)-dimethylaminomethyleneamino]phenyl methylcarbamate | 244-879-0 | 22259-30-9 | Acute Tox. 2 * Acute Tox. 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H300 H317 H400 H410 | GHS06 GHS09 Dgr | H330 H300 H317 H410 |  |  |  | CLP00
006-032-00-1 | monolinuron (ISO); 3-(4-chlorophenyl)-1-methoxy-1-methylurea | 217-129-5 | 1746-81-2 | Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H410 |  |  |  | CLP00
006-033-00-7 | metoxuron (ISO); 3-(3-chloro-4-methoxyphenyl)-1,1-dimethylurea | 243-433-2 | 19937-59-8 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
006-034-00-2 | pebulate (ISO); N-butyl-N-ethyl-S-propylthiocarbamate | 214-215-4 | 1114-71-2 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
006-035-00-8 | pirimicarb (ISO); 2-(dimethylamino)-5,6-dimethylpyrimidin-4-yl dimethylcarbamate | 245-430-1 | 23103-98-2 | Carc. 2 Acute Tox. 3 Acute Tox. 3 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H331 H301 H317 H400 H410 | GHS08 GHS06 GHS09 Dgr | H331 H301 H317 H351 H410 |  | M=10 M=100 |  | CLP00/ATP09
006-036-00-3 | benzthiazuron (ISO); 1-benzothiazol-2-yl-3-methylurea | 217-685-9 | 1929-88-0 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
006-037-00-9 | promecarb (ISO); 3-isopropyl-5-methylphenyl N-methylcarbamate | 220-113-0 | 2631-37-0 | Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H400 H410 | GHS06 GHS09 Dgr | H301 H410 |  |  |  | CLP00
006-038-00-4 | sulfallate (ISO); 2-chloroallyl N,N-dimethyldithiocarbamate | 202-388-9 | 95-06-7 | Carc. 1B Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H302 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H302 H410 |  |  |  | CLP00
006-039-00-X | tri-allate (ISO); S-2,3,3-trichloroallyl diisopropylthiocarbamate | 218-962-7 | 2303-17-5 | Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H317 H410 |  |  |  | CLP00
006-040-00-5 | 3-methylpyrazol-5-yl-dimethylcarbamate; monometilan |  | 2532-43-6 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H331 H311 H301 | GHS06 Dgr | H331 H311 H301 |  |  |  | CLP00
006-041-00-0 | dimethylcarbamoyl chloride | 201-208-6 | 79-44-7 | Carc. 1B Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H350 H331 H302 H335 H315 H319 | GHS06 GHS08 Dgr | H350 H331 H302 H319 H335 H315 |  | Carc. 1B; H350: C ≥ 0,001 % |  | CLP00
006-042-00-6 | monuron (ISO); 3-(4-chlorophenyl)-1,1-dimethylurea | 205-766-1 | 150-68-5 | Carc. 2 Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H302 H410 |  |  |  | CLP00
006-043-00-1 | 3-(4-chlorophenyl)-1,1-dimethyluronium trichloroacetate; monuron-TCA |  | 140-41-0 | Carc. 2 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H315 H319 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H319 H315 H410 |  |  |  | CLP00
006-044-00-7 | isoproturon (ISO); 3-(4-isopropylphenyl)-1,1-dimethylurea | 251-835-4 | 34123-59-6 | Carc. 2 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H373 (blood) H400 H410 | GHS08 GHS09 Wng | H351 H373 (blood) H410 |  | M=10 M=10 |  | CLP00/ATP13
006-045-00-2 | methomyl (ISO); 1-(methylthio)ethylideneamino N-methylcarbamate | 240-815-0 | 16752-77-5 | Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H400 H410 | GHS06 GHS09 Dgr | H300 H410 |  | M=100 |  | CLP00/ATP01
006-046-00-8 | bendiocarb (ISO); 2,2-dimethyl-1,3-benzodioxol-4-yl N-methylcarbamate; 2,2-dimethyl-1,3-benzodioxol-4-yl methylcarbamate | 245-216-8 | 22781-23-3 | Acute Tox. 2 Acute Tox. 3 Acute Tox. 3 Aquatic Acute 1 Aquatic Chronic 1 | H300 H331 H311 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H300 H410 |  | M=10 M=100 |  | CLP00/ATP10
006-047-00-3 | bufencarb (ISO); reaction mass of 3-(1-methylbutyl)phenyl N-methylcarbamate and 3-(1-ethylpropyl)phenyl N-methylcarbamate |  | 8065-36-9 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H311 H301 H400 H410 | GHS06 GHS09 Dgr | H311 H301 H410 |  |  |  | CLP00
006-048-00-9 | ethiofencarb (ISO); 2-(ethylthiomethyl)phenyl N-methylcarbamate | 249-981-9 | 29973-13-5 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
006-049-00-4 | dixanthogen; O,O-diethyl dithiobis(thioformate) | 207-944-4 | 502-55-6 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
006-050-00-X | 1,1-dimethyl-3-phenyluronium trichloroacetate; fenuron-TCA |  | 4482-55-7 | Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H400 H410 | GHS07 GHS09 Wng | H315 H410 |  |  |  | CLP00
006-051-00-5 | ferbam (ISO); iron tris(dimethyldithiocarbamate) | 238-484-2 | 14484-64-1 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H335 H315 H319 H400 H410 | GHS07 GHS09 Wng | H319 H335 H315 H410 |  |  |  | CLP00
006-052-00-0 | formetanate hydrochloride; 3-(N,N-dimethylaminomethyleneamino)phenyl N-methylcarbamate | 245-656-0 | 23422-53-9 | Acute Tox. 2 * Acute Tox. 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H300 H317 H400 H410 | GHS06 GHS09 Dgr | H330 H300 H317 H410 |  |  |  | CLP00
006-053-00-6 | isoprocarb (ISO); 2-isopropylphenyl N-methylcarbamate | 220-114-6 | 2631-40-5 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
006-054-00-1 | mexacarbate (ISO); 3,5-dimethyl-4-dimethylaminophenyl N-methylcarbamate | 206-249-3 | 315-18-4 | Acute Tox. 2 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H312 H400 H410 | GHS06 GHS09 Dgr | H300 H312 H410 |  |  |  | CLP00
006-055-00-7 | xylylcarb (ISO); 3,4-dimethylphenyl N-methylcarbamate; 3,4-xylyl methylcarbamate; MPMC | 219-364-9 | 2425-10-7 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
006-056-00-2 | metolcarb (ISO); m-tolyl methylcarbamate; MTMC | 214-446-0 | 1129-41-5 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
006-057-00-8 | nitrapyrin (ISO); 2-chloro-6-trichloromethylpyridine | 217-682-2 | 1929-82-4 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
006-058-00-3 | noruron (ISO); 1,1-dimethyl-3-(perhydro-4,7-methanoinden-5-yl)urea |  | 2163-79-3 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
006-059-00-9 | oxamyl (ISO); N',N'-dimethylcarbamoyl(methylthio)methylenamine N-methylcarbamate | 245-445-3 | 23135-22-0 | Acute Tox. 2 * Acute Tox. 2 * Acute Tox. 4 * Aquatic Chronic 2 | H330 H300 H312 H411 | GHS06 GHS09 Dgr | H330 H300 H312 H411 |  |  |  | CLP00
006-060-00-4 | oxycarboxin (ISO); 2,3-dihydro-6-methyl-5-(N-phenylcarbamoyl)-1,4-oxothiine 4,4-dioxide | 226-066-2 | 5259-88-1 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
006-061-00-X | S-ethyl N-(dimethylaminopropyl)thiocarbamatehydrochloride; prothiocarb hydrochloride | 243-193-9 | 19622-19-6 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
006-062-00-5 | methyl 3,4-dichlorophenylcarbanilate; SWEP. |  | 1918-18-9 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
006-063-00-0 | thiobencarb (ISO); S-4-chlorobenzyl diethylthiocarbamate | 248-924-5 | 28249-77-6 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
006-064-00-6 | thiofanox (ISO); 3,3-dimethyl-1-(methylthio)butanone-O-(N-methylcarbamoyl)oxime | 254-346-4 | 39196-18-4 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H310 H300 H410 |  |  |  | CLP00
006-065-00-1 | 3-chloro-6-cyano-bicyclo(2,2,1)heptan-2-one-O-(N-methylcarbamoyl)oxime; triamid |  | 15271-41-7 | Acute Tox. 2 * Acute Tox. 3 * Aquatic Chronic 2 | H300 H311 H411 | GHS06 GHS09 Dgr | H300 H311 H411 |  |  |  | CLP00
006-066-00-7 | vernolate (ISO); S-propyl dipropylthiocarbamate | 217-681-7 | 1929-77-7 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
006-067-00-2 | XMC; 3,5-xylyl methylcarbamate |  | 2655-14-3 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
006-068-00-8 | diazomethane | 206-382-7 | 334-88-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
006-069-00-3 | thiophanate-methyl (ISO); dimethyl (1,2-phenylenedicarbamo-thioyl)biscarbamate; dimethyl 4,4′-(o-phenylene) bis(3-thioallophanate) | 245-740-7 | 23564-05-8 | Carc. 2 Muta. 2 Acute Tox. 4 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H341 H332 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H332 H317 H341 H351 H410 |  | Inhalation: ATE = 1.7 mg/L (dusts/mists) M=10 M=10 |  | CLP00/ATP17
006-070-00-9 | furmecyclox (ISO); N-cyclohexyl-N-methoxy-2,5-dimethyl-3-furamide | 262-302-0 | 60568-05-0 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  |  |  | CLP00
006-071-00-4 | cyclooct-4-en-1-yl methyl carbonate | 401-620-8 | 87731-18-8 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
006-072-00-X | prosulfocarb(ISO); S-benzyl N,N-dipropylthiocarbamate | 401-730-6 | 52888-80-9 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | CLP00
006-073-00-5 | 3-(dimethylamino)propylurea | 401-950-2 | 31506-43-1 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
006-074-00-0 | 2-(3-(prop-1-en-2-yl)phenyl)prop-2-yl isocyanate | 402-440-2 | 2094-99-7 | Acute Tox. 2 * STOT RE 2 * Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H373 ** H314 H334 H317 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H330 H314 H373 ** H334 H317 H410 |  |  |  | CLP00
006-076-00-1 | mancozeb (ISO); manganese ethylenebis(dithiocarbamate) (polymeric) complex with zinc salt |  | 8018-01-7 | Carc. 2 Repr. 1B STOT RE 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H360D H373 (thyroid, nervous system) H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H317 H351 H360D H373 (thyroid, nervous system) H410 |  | M=10 M=10 |  | CLP00/ATP17
006-077-00-7 | maneb (ISO); manganese ethylenebis(dithiocarbamate) (polymeric) | 235-654-8 | 12427-38-2 | Repr. 2 Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H332 H319 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H361d *** H332 H319 H317 H410 |  | M=10 |  | CLP00/ATP01
006-078-00-2 | zineb (ISO); zinc ethylenebis(dithiocarbamate) (polymeric) | 235-180-1 | 12122-67-7 | STOT SE 3 Skin Sens. 1 | H335 H317 | GHS07 Wng | H335 H317 |  |  |  | CLP00
006-079-00-8 | disulfiram; tetraethylthiuramdisulfide | 202-607-8 | 97-77-8 | Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H317 H410 |  |  |  | CLP00
006-080-00-3 | tetramethylthiuram monosulphide | 202-605-7 | 97-74-5 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | CLP00
006-081-00-9 | zinc bis(dibutyldithiocarbamate) | 205-232-8 | 136-23-2 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H335 H315 H319 H317 H400 H410 | GHS07 GHS09 Wng | H319 H335 H315 H317 H410 |  |  |  | CLP00
006-082-00-4 | zinc bis(diethyldithiocarbamate) | 238-270-9 | 14324-55-1 | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H335 H315 H319 H317 H400 H410 | GHS07 GHS09 Wng | H302 H319 H335 H315 H317 H410 |  |  |  | CLP00
006-083-00-X | butocarboxim (ISO); 3-(methylthio)-2-butanone O-[(methylamino)carbonyl]oxime | 252-139-3 | 34681-10-2 | Flam. Liq. 3 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H226 H331 H311 H301 H319 H400 H410 | GHS02 GHS06 GHS09 Dgr | H226 H331 H311 H301 H319 H410 |  |  |  | CLP00
006-084-00-5 | carbosulfan (ISO); 2,3-dihydro-2,2-dimethyl-7-benzofuryl [(dibutylamino)thio]methylcarbamate | 259-565-9 | 55285-14-8 | Acute Tox. 2 * Acute Tox. 3 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H301 H317 H400 H410 | GHS06 GHS09 Dgr | H330 H301 H317 H410 |  |  |  | CLP00/ATP01
006-085-00-0 | fenobucarb (ISO); 2-butylphenyl methylcarbamate | 223-188-8 | 3766-81-2 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
006-086-00-6 | fenoxycarb (ISO); ethyl [2-(4-phenoxyphenoxy)ethyl]carbamate | 276-696-7 | 72490-01-8 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  | M=1 M=10000 |  | CLP00/ATP06
006-087-00-1 | furathiocarb (ISO); 2,3-dihydro-2,2-dimethyl-7-benzofuryl 2,4-dimethyl-6-oxa-5-oxo-3-thia-2,4-diazadecanoate | 265-974-3 | 65907-30-4 | Acute Tox. 2 * Acute Tox. 3 * STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H301 H373 ** H315 H319 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H301 H330 H315 H319 H317 H373 ** H410 |  | M=100 |  | CLP00/ATP01corr
006-088-00-7 | benfuracarb (ISO); ethyl N-[2,3-dihydro-2,2-dimethylbenzofuran-7-yloxycarbonyl(methyl)aminothio]-N-isopropyl- β-alaninate | - | 82560-54-1 | Repr. 2 Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H361f *** H331 H302 H400 H410 | GHS06 GHS08 GHS09 Dgr | H361f *** H331 H302 H410 |  |  |  | CLP00/ATP01
006-090-00-8 | 2-(3-iodoprop-2-yn-1-yloxy)ethyl phenylcarbamate | 408-010-0 | 88558-41-2 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H332 H318 H412 | GHS05 GHS07 Dgr | H332 H318 H412 |  |  |  | CLP00
006-091-00-3 | propineb (ISO); polymeric zinc propylenebis(dithiocarbamate) | - | 9016-72-2 | Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 | H332 H373 ** H317 H400 | GHS08 GHS07 GHS09 Wng | H332 H373 ** H317 H400 |  |  |  | ATP01
006-092-00-9 | tert-butyl (1S)-N-[1-((2S)-2-oxiranyl)-2-phenylethyl]carbamate | 425-420-5 | 98737-29-2 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
006-093-00-4 | 2,2'-dithio di(ethylammonium)-bis(dibenzyldithiocarbamate) | 427-180-7 | - | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | ATP01
006-094-00-X | O-isobutyl-N-ethoxy carbonylthiocarbamate | 434-350-4 | 103122-66-3 | Flam. Liq. 3 Carc. 1B Muta. 1B Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Chronic 2 | H226 H350 H340 H302 H373 ** H317 H411 | GHS02 GHS08 GHS07 GHS09 Dgr | H226 H350 H340 H302 H373 ** H317 H411 |  |  |  | ATP01
006-095-00-5 | fosetyl-aluminium (ISO); aluminium triethyl triphosphonate | 254-320-2 | 39148-24-8 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
006-096-00-0 | chlorpropham (ISO); isopropyl 3-chlorocarbanilate | 202-925-7 | 101-21-3 | Carc. 2 STOT RE 2 * Aquatic Chronic 2 | H351 H373 ** H411 | GHS08 GHS09 Wng | H351 H373 ** H411 |  |  |  | ATP01
006-097-00-6 | 1-phenyl-3-(p-toluenesulfonyl)urea | 424-620-1 | 13909-63-2 | Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 3 | H302 H373 ** H412 | GHS08 GHS07 Wng | H302 H373 ** H412 |  |  |  | ATP01
006-098-00-1 | tert-butyl (1R,5S)-3-azabicyclo[3.1.0]hex-6-ylcarbamate | 429-170-8 | 134575-17-0 | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 | H302 H373 ** H318 H317 | GHS05 GHS08 GHS07 Dgr | H302 H318 H317 H373 ** |  |  |  | ATP01/ATP01corr
006-099-00-7 | N-(p-toluenesulfonyl)-N'-(3-(p-toluenesulfonyloxy)phenyl)urea; 3-({[(4-methylphenyl)sulfonyl]carbamoyl}amino)phenyl 4-methylbenzenesulfonate | 432-520-2 | 232938-43-1 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
006-101-00-6 | reaction mass of: N,N''-(methylenedi-4,1-phenylene)bis[N'-phenylurea]; N-(4-[[4-[[(phenylamino)carbonyl]amino]phenylmethyl]phenyl]-N'-cyclohexylurea; N,N''-(methylenedi-4,1-phenylene)bis[N'-cyclohexylurea] | 423-070-8 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
006-102-00-1 | O-hexyl-N-ethoxycarbonylthiocarbamate | 432-750-3 | - | Carc. 1B Muta. 1B Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Chronic 2 | H350 H340 H302 H373 ** H317 H411 | GHS08 GHS07 GHS09 Dgr | H350 H340 H302 H373 ** H317 H411 |  |  |  | ATP01
006-103-00-7 | N,N''-(methylenedi-4,1-phenylene)bis[N'-octyl]urea | 445-760-8 | - | Eye Dam. 1 Resp. Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H334 H400 H410 | GHS05 GHS08 GHS09 Dgr | H318 H334 H410 |  | M=100 |  | ATP01
006-104-00-2 | multi-walled carbon tubes (synthetic graphite in tubular shape) with a geometric tube diameter range ≥ 30 nm to < 3 μm and a length ≥ 5 μm and aspect ratio > 3:1, including multi-walled carbon nanotubes, MWC(N)T | - | - | Carc. 1B STOT RE 1 | H350i H372 (lung, inhalation) | GHS08 Dgr | H350i H372 (lung, inhalation) |  | STOT RE 1; H372: C ≥ 1 %; STOT RE 2; H373: 0,1 % ≤ C < 1 % |  | ATP22
007-001-00-5 | ammonia, anhydrous | 231-635-3 | 7664-41-7 | Flam. Gas 2 Press. Gas Acute Tox. 3 * Skin Corr. 1B Aquatic Acute 1 | H221 H331 H314 H400 | GHS04 GHS06 GHS05 GHS09 Dgr | H221 H331 H314 H400 |  |  | U | CLP00
007-001-01-2 | ammonia ....% | 215-647-6 | 1336-21-6 | Skin Corr. 1B Aquatic Acute 1 | H314 H400 | GHS05 GHS09 Dgr | H314 H400 |  | STOT SE 3; H335: C ≥ 5 % | B | CLP00
007-002-00-0 | nitrogen dioxide [1] dinitrogen tetraoxide [2] | 233-272-6 [1] 234-126-4 [2] | 10102-44-0 [1] 10544-72-6 [2] | Ox. Gas 1 Press. Gas Acute Tox. 2 * Skin Corr. 1B | H270 H330 H314 | GHS04 GHS03 GHS06 GHS05 Dgr | H270 H330 H314 |  | * STOT SE 3; H335: C ≥ 0,5 % | 5 | CLP00/ATP01
007-003-00-6 | chlormequat chloride (ISO); 2-chloroethyltrimethylammonium chloride | 213-666-4 | 999-81-5 | Acute Tox. 4 * Acute Tox. 4 * | H312 H302 | GHS07 Wng | H312 H302 |  |  |  | CLP00
007-004-00-1 | nitric acid ...% [C > 70 %] | 231-714-2 | 7697-37-2 | Ox. Liq. 2 Acute Tox. 1 Skin Corr. 1A | H272 H330 H314 | GHS03 GHS06 GHS05 Dgr | H272 H330 H314 | EUH071 | Ox. Liq. 2; H272: C ≥ 99 % Ox. Liq. 3; H272: 70 % ≤ C < 99 % | B | CLP00/ATP15
007-006-00-2 | ethyl nitrite | 203-722-6 | 109-95-5 | Flam. Gas 1 Press. Gas Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H220 H332 H312 H302 | GHS02 GHS04 GHS07 Dgr | H220 H332 H312 H302 |  |  | U | CLP00
007-007-00-8 | ethyl nitrate | 210-903-3 | 625-58-1 | Unst. Expl. | H200 | GHS01 Dgr | H200 |  |  |  | CLP00/ATP01corr
007-008-00-3 | hydrazine | 206-114-9 | 302-01-2 | Flam. Liq. 3 Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H226 H350 H331 H311 H301 H314 H317 H400 H410 | GHS02 GHS06 GHS08 GHS05 GHS09 Dgr | H226 H350 H331 H311 H301 H314 H317 H410 |  | Skin Corr. 1B; H314: C ≥ 10 % Skin Irrit. 2; H315: 3 % ≤ C < 10 % Eye Irrit. 2; H319: 3 % ≤ C < 10 % |  | CLP00
007-009-00-9 | dicyclohexylammonium nitrite | 221-515-9 | 3129-91-7 | Acute Tox. 4 * Acute Tox. 4 * | H332 H302 | GHS07 Wng | H332 H302 |  | * |  | CLP00
007-010-00-4 | sodium nitrite | 231-555-9 | 7632-00-0 | Ox. Sol. 3 Acute Tox. 3 * Aquatic Acute 1 | H272 H301 H400 | GHS03 GHS06 GHS09 Dgr | H272 H301 H400 |  | * |  | CLP00
007-011-00-X | potassium nitrite | 231-832-4 | 7758-09-0 | Ox. Sol. 2 Acute Tox. 3 * Aquatic Acute 1 | H272 H301 H400 | GHS03 GHS06 GHS09 Dgr | H272 H301 H400 |  | * |  | CLP00
007-012-00-5 | N,N-dimethylhydrazine | 200-316-0 | 57-14-7 | Flam. Liq. 2 Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Aquatic Chronic 2 | H225 H350 H331 H301 H314 H411 | GHS02 GHS06 GHS08 GHS05 GHS09 Dgr | H225 H350 H331 H301 H314 H411 |  |  |  | CLP00
007-013-00-0 | 1,2-dimethylhydrazine |  | 540-73-8 | Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Chronic 2 | H350 H331 H311 H301 H411 | GHS06 GHS08 GHS09 Dgr | H350 H331 H311 H301 H411 |  | Carc. 1B; H350: C ≥ 0,01 % |  | CLP00
007-014-00-6 | salts of hydrazine |  |  | Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H331 H311 H301 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H331 H311 H301 H317 H410 |  |  | A | CLP00
007-015-00-1 | O-ethylhydroxylamine | 402-030-3 | 624-86-2 | Flam. Liq. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 | H225 H331 H311 H301 H372 ** H319 H317 H400 | GHS02 GHS06 GHS08 GHS09 Dgr | H225 H331 H311 H301 H372 ** H319 H317 H400 |  |  |  | CLP00
007-016-00-7 | butyl nitrite | 208-862-1 | 544-16-1 | Flam. Liq. 2 Acute Tox. 3 * Acute Tox. 3 * | H225 H331 H301 | GHS02 GHS06 Dgr | H225 H331 H301 |  |  |  | CLP00
007-017-00-2 | isobutyl nitrite | 208-819-7 | 542-56-3 | Flam. Liq. 2 Carc. 1B Muta. 2 Acute Tox. 4 * Acute Tox. 4 * | H225 H350 H341 H332 H302 | GHS02 GHS08 GHS07 Dgr | H225 H350 H341 H332 H302 |  |  |  | CLP00
007-018-00-8 | sec-butyl nitrite | 213-104-8 | 924-43-6 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * | H225 H332 H302 | GHS02 GHS07 Dgr | H225 H332 H302 |  |  |  | CLP00
007-019-00-3 | tert-butyl nitrite | 208-757-0 | 540-80-7 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * | H225 H332 H302 | GHS02 GHS07 Dgr | H225 H332 H302 |  |  |  | CLP00
007-020-00-9 | pentyl nitrite [1] ‘amyl nitrite’, mixed isomers [2] | 207-332-7 [1] 203-770-8 [2] | 463-04-7 [1] 110-46-3 [2] | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * | H225 H332 H302 | GHS02 GHS07 Dgr | H225 H332 H302 |  |  |  | CLP00
007-021-00-4 | hydrazobenzene; 1,2-diphenylhydrazine | 204-563-5 | 122-66-7 | Carc. 1B Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H302 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H302 H410 |  |  |  | CLP00
007-022-00-X | hydrazine bis(3-carboxy-4-hydroxybenzensulfonate) | 405-030-1 |  | Carc. 1B Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 3 | H350 H302 H314 H317 H412 | GHS08 GHS05 GHS07 Dgr | H350 H302 H314 H317 H412 |  |  |  | CLP00
007-023-00-5 | sodium 3,5-bis(3-(2,4-di-tert-pentylphenoxy)propylcarbamoyl)benzenesulfonate | 405-510-0 |  | Skin Irrit. 2 Skin Sens. 1 | H315 H317 | GHS07 Wng | H315 H317 |  |  |  | CLP00
007-024-00-0 | 2-(decylthio)ethylammonium chloride | 405-640-8 | 36362-09-1 | STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H315 H318 H400 H410 | GHS08 GHS05 GHS09 Dgr | H373 ** H315 H318 H410 |  |  |  | CLP00
007-025-00-6 | (4-hydrazinophenyl)-N-methylmethanesulfonamide hydrochloride | 406-090-1 | 81880-96-8 | Muta. 2 Acute Tox. 3 * STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H341 H301 H372 ** H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H341 H301 H372 ** H317 H410 |  |  |  | CLP00
007-026-00-1 | oxo-((2,2,6,6-tetramethylpiperidin-4-yl)amino)carbonylacetohydrazide | 413-230-5 | 122035-71-6 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
007-027-00-7 | 1,6-bis(3,3-bis((1-methylpentylidenimino)propyl)ureido)hexane | 420-190-2 | 771478-66-1 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H373 ** H314 H317 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H312 H302 H373 ** H314 H317 H410 |  |  |  | CLP00
007-028-00-2 | hydroxylammonium nitrate | 236-691-2 | 13465-08-2 | Expl. 1.1 **** Carc. 2 Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 | H201 H351 H311 H302 H373 ** H315 H319 H317 H400 | GHS01 GHS06 GHS08 GHS09 Dgr | H201 H351 H311 H302 H373 ** H319 H315 H317 H400 |  |  |  | ATP01
007-029-00-8 | diethyldimethylammonium hydroxide | 419-400-5 | 95500-19-9 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A | H312 H302 H314 | GHS05 GHS07 Dgr | H312 H302 H314 |  |  |  | ATP01
007-030-00-3 | nitric acid …% [C ≤ 70 %] | 231-714-2 | 7697-37-2 | Ox. Liq. 3 Acute Tox. 3 Skin Corr. 1A | H272 H331 H314 | GHS03 GHS06 GHS05 Dgr | H272 H331 H314 | EUH071 | Inhalation: ATE = 2.65 mg/L (Vapours) Ox. Liq. 3; H272: C ≥ 65 % Skin Corr. 1A; H314: C ≥ 20 % Skin Corr. 1B; H314: 5 % ≤ C < 20 % | B | ATP15
008-001-00-8 | oxygen | 231-956-9 | 7782-44-7 | Ox. Gas 1 Press. Gas | H270 | GHS03 GHS04 Dgr | H270 |  |  | U | CLP00
008-003-00-9 | hydrogen peroxide solution ...% | 231-765-0 | 7722-84-1 | Ox. Liq. 1 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A | H271 H332 H302 H314 | GHS03 GHS05 GHS07 Dgr | H271 H332 H302 H314 |  | Ox. Liq. 1; H271: C ≥ 70 %**** Ox. Liq. 2; H272: 50 % ≤ C < 70 % **** * Skin Corr. 1A; H314: C ≥ 70 % Skin Corr. 1B; H314: 50 % ≤ C < 70 % Skin Irrit. 2; H315: 35 % ≤ C < 50 % Eye Dam. 1; H318: 8 % ≤ C < 50 % Eye Irrit. 2; H319: 5 % ≤ C < 8 % STOT SE 3; H335; C ≥ 35 % | B | CLP00
009-001-00-0 | fluorine | 231-954-8 | 7782-41-4 | Ox. Gas 1 Press. Gas Acute Tox. 2 * Skin Corr. 1A | H270 H330 H314 | GHS04 GHS03 GHS06 GHS05 Dgr | H270 H330 H314 |  |  |  | CLP00/ATP01
009-002-00-6 | hydrogen fluoride | 231-634-8 | 7664-39-3 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Skin Corr. 1A | H310 H330 H300 H314 | GHS06 GHS05 Dgr | H330 H310 H300 H314 |  |  |  | CLP00
009-003-00-1 | hydrofluoric acid ... % | 231-634-8 | 7664-39-3 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Skin Corr. 1A | H310 H330 H300 H314 | GHS06 GHS05 Dgr | H330 H310 H300 H314 |  | Skin Corr. 1A; H314: C ≥ 7 % Skin Corr. 1B; H314: 1 % ≤ C < 7 % Eye Irrit. 2; H319: 0,1 % ≤ C < 1 % | B | CLP00
009-004-00-7 | sodium fluoride | 231-667-8 | 7681-49-4 | Acute Tox. 3 * Skin Irrit. 2 Eye Irrit. 2 | H301 H315 H319 | GHS06 Dgr | H301 H319 H315 | EUH032 |  |  | CLP00
009-005-00-2 | potassium fluoride | 232-151-5 | 7789-23-3 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H331 H311 H301 | GHS06 Dgr | H331 H311 H301 |  |  |  | CLP00
009-006-00-8 | ammonium fluoride | 235-185-9 | 12125-01-8 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H331 H311 H301 | GHS06 Dgr | H331 H311 H301 |  |  |  | CLP00
009-007-00-3 | sodium bifluoride; sodium hydrogen difluoride | 215-608-3 | 1333-83-1 | Acute Tox. 3 * Skin Corr. 1B | H301 H314 | GHS06 GHS05 Dgr | H301 H314 |  | * Skin Corr. 1B; H314: C ≥ 1 % Skin Irrit. 2; H315: 0,1 % ≤ C < 1 % Eye Irrit. 2; H319: 0,1 % ≤ C < 1 % |  | CLP00
009-008-00-9 | potassium bifluoride; potassium hydrogen difluoride | 232-156-2 | 7789-29-9 | Acute Tox. 3 * Skin Corr. 1B | H301 H314 | GHS06 GHS05 Dgr | H301 H314 |  | * Skin Corr. 1B; H314: C ≥ 1 % Skin Irrit. 2; H315: 0,1 % ≤ C < 1 % Eye Irrit. 2; H319: 0,1 % ≤ C < 1 % |  | CLP00
009-009-00-4 | ammonium bifluoride; ammonium hydrogen difluoride | 215-676-4 | 1341-49-7 | Acute Tox. 3 * Skin Corr. 1B | H301 H314 | GHS06 GHS05 Dgr | H301 H314 |  | * Skin Corr. 1B; H314: C ≥ 1 % Skin Irrit. 2; H315: 0,1 % ≤ C < 1 % Eye Irrit. 2; H319: 0,1 % ≤ C < 1 % |  | CLP00
009-010-00-X | fluoroboric acid ... % | 240-898-3 | 16872-11-0 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  | Skin Corr. 1B; H314: C ≥ 25 % Skin Irrit. 2; H315: 10 % ≤ C < 25 % Eye Irrit. 2; H319: 10 % ≤ C < 25 % | B | CLP00
009-011-00-5 | fluorosilicic acid ... % | 241-034-8 | 16961-83-4 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  |  | B | CLP00
009-012-00-0 | alkali fluorosilicates(Na) [1] alkali fluorosilicates(K) [2] alkali fluorosilicates(NH4) [3] | 240-934-8 [1] 240-896-2 [2] 240-968-3 [3] | 16893-85-9 [1] 16871-90-2 [2] 16919-19-0 [3] | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H331 H311 H301 | GHS06 Dgr | H331 H311 H301 |  | * | A | CLP00
009-013-00-6 | fluorosilicates, with the exception of those specified elsewhere in this annex |  |  | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  | * | A | CLP00
009-014-00-1 | lead hexafluorosilicate | 247-278-1 | 25808-74-6 | Repr. 1A Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H360Df H332 H302 H373 ** H400 H410 | GHS08 GHS07 GHS09 Dgr | H360Df H332 H302 H373 ** H410 |  |  | 1 | CLP00
009-015-00-7 | sulphuryl difluoride | 220-281-5 | 2699-79-8 | Press. Gas Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 | H331 H373 ** H400 | GHS04 GHS06 GHS08 GHS09 Dgr | H331 H373 ** H400 |  |  | U | CLP00
009-016-00-2 | trisodium hexafluoroaluminate [1] trisodium hexafluoroaluminate (cryolite) [2] | 237-410-6 [1] 239-148-8 [2] | 13775-53-6 [1] 15096-52-3 [2] | Acute Tox. 4 STOT RE 1 Aquatic Chronic 2 | H332 H372 H411 | GHS07 GHS08 GHS09 Dgr | H372 H332 H411 |  |  |  | CLP00/ATP03
009-017-00-8 | potassium mu-fluoro-bis(triethylaluminium) | 400-040-2 | 12091-08-6 | Flam. Sol. 1 Water-react. 1 Acute Tox. 4 * Skin Corr. 1A | H228 H260 H332 H314 | GHS02 GHS05 GHS07 Dgr | H228 H260 H332 H314 | EUH014 |  | T | CLP00
009-018-00-3 | magnesium hexafluorosilicate | 241-022-2 | 16949-65-8 | Acute Tox. 3 * | H301 | GHS06 Dgr | H301 |  | * |  | CLP00
011-001-00-0 | sodium | 231-132-9 | 7440-23-5 | Water-react. 1 Skin Corr. 1B | H260 H314 | GHS02 GHS05 Dgr | H260 H314 | EUH014 |  |  | CLP00
011-002-00-6 | sodium hydroxide; caustic soda | 215-185-5 | 1310-73-2 | Skin Corr. 1A | H314 | GHS05 Dgr | H314 |  | Skin Corr. 1A; H314: C ≥ 5 % Skin Corr. 1B; H314: 2 % ≤ C < 5 % Skin Irrit. 2; H315: 0,5 % ≤ C < 2 % Eye Irrit. 2; H319: 0,5 % ≤ C < 2 % |  | CLP00
011-003-00-1 | sodium peroxide | 215-209-4 | 1313-60-6 | Ox. Sol. 1 Skin Corr. 1A | H271 H314 | GHS03 GHS05 Dgr | H271 H314 |  |  |  | CLP00
011-004-00-7 | sodium azide | 247-852-1 | 26628-22-8 | Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H400 H410 | GHS06 GHS09 Dgr | H300 H410 | EUH032 |  |  | CLP00
011-005-00-2 | sodium carbonate | 207-838-8 | 497-19-8 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
011-006-00-8 | sodium cyanate | 213-030-6 | 917-61-3 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
011-007-00-3 | propoxycarbazone-sodium |  | 181274-15-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=10 |  | CLP00
012-001-00-3 | magnesium powder (pyrophoric) | 231-104-6 | 7439-95-4 | Pyr. Sol. 1 Water-react. 1 | H250 H260 | GHS02 Dgr | H260 H250 |  |  | T | CLP00
012-002-00-9 | magnesium, powder or turnings | 231-104-6 |  | Flam. Sol. 1 Self-heat. 1 Water-react. 2 | H228 H252 H261 | GHS02 Dgr | H228 H261 H252 |  |  | T | CLP00
012-003-00-4 | magnesium alkyls |  |  | Pyr. Liq. 1 Water-react. 1 Skin Corr. 1B | H250 H260 H314 | GHS02 GHS05 Dgr | H250 H260 H314 | EUH014 |  | A | CLP00
012-004-00-X | aluminium-magnesium-carbonate-hydroxide-perchlorate-hydrate | 422-150-1 | - | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
013-001-00-6 | aluminium powder (pyrophoric) | 231-072-3 | 7429-90-5 | Pyr. Sol. 1 Water-react. 2 | H250 H261 | GHS02 Dgr | H261 H250 |  |  | T | CLP00
013-002-00-1 | aluminium powder (stabilised) | 231-072-3 | 7429-90-5 | Flam. Sol. 1 Water-react. 2 | H228 H261 | GHS02 Dgr | H261 H228 |  |  | T | CLP00/ATP01
013-003-00-7 | aluminium chloride, anhydrous | 231-208-1 | 7446-70-0 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
013-004-00-2 | aluminium alkyls |  |  | Pyr. Liq. 1 Water-react. 1 Skin Corr. 1B | H250 H260 H314 | GHS02 GHS05 Dgr | H250 H260 H314 | EUH014 |  | A | CLP00
013-005-00-8 | diethyl(ethyldimethylsilanolato)aluminium | 401-160-8 | 55426-95-4 | Pyr. Liq. 1 Water-react. 1 Skin Corr. 1A | H250 H260 H314 | GHS02 GHS05 Dgr | H260 H250 H314 | EUH014 |  |  | CLP00
013-006-00-3 | (ethyl-3-oxobutanoato-O'1,O'3)(2-dimethylaminoethanolato)(1-methoxypropan-2-olato)aluminium(III), dimerised | 402-370-2 |  | Flam. Liq. 3 Eye Dam. 1 | H226 H318 | GHS02 GHS05 Dgr | H226 H318 |  |  |  | CLP00
013-007-00-9 | poly(oxo(2-butoxyethyl-3-oxobutanoato-O'1,O'3)aluminium) | 403-430-0 |  | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
013-008-00-4 | di-n-octylaluminium iodide | 408-190-0 | 7585-14-0 | Pyr. Liq. 1 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H250 H314 H400 H410 | GHS02 GHS05 GHS09 Dgr | H250 H314 H410 | EUH014 |  |  | CLP00
013-009-00-X | sodium((n-butyl)x(ethyl)y-1,5-dihydro)aluminate) x = 0,5, y = 1,5 | 418-720-2 |  | Flam. Sol. 1 Pyr. Sol. 1 Water-react. 1 Acute Tox. 4 * Skin Corr. 1A | H228 H250 H260 H332 H314 | GHS02 GHS05 GHS07 Dgr | H228 H260 H250 H332 H314 | EUH014 |  | T | CLP00
013-010-00-5 | hydroxy aluminium bis(2,4,8,10-tetra-tert-butyl-6-hydroxy-12H-dibenzo[d,g][1.3.2]dioxaphosphocin-6-oxide) | 430-650-4 | 151841-65-5 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
014-001-00-9 | trichlorosilane | 233-042-5 | 10025-78-2 | Flam. Liq. 1 Water-react. 1 Acute Tox. 3 Acute Tox. 4 Skin Corr. 1A Eye Dam. 1 | H224 H260 H331 H302 H314 H318 | GHS02 GHS06 GHS05 Dgr | H224 H260 H331 H302 H314 | EUH014 EUH029 EUH071 | inhalation: ATE = 7,6 mg/L (vapour) oral: ATE=1 000 mg/kg bw |  | CLP00/ATP18
014-002-00-4 | silicon tetrachloride | 233-054-0 | 10026-04-7 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H335 H315 H319 | GHS07 Wng | H319 H335 H315 | EUH014 |  |  | CLP00
014-003-00-X | dimethyldichlorosilane | 200-901-0 | 75-78-5 | Flam. Liq. 2 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H225 H335 H315 H319 | GHS02 GHS07 Dgr | H225 H319 H335 H315 |  |  |  | CLP00
014-004-00-5 | trichloro(methyl)silane; methyltrichlorosilane | 200-902-6 | 75-79-6 | Flam. Liq. 2 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H225 H335 H315 H319 | GHS02 GHS07 Dgr | H225 H319 H335 H315 | EUH014 | Skin Irrit. 2; H315: C ≥ 1 % Eye Irrit. 2; H319: C ≥ 1 % STOT SE 3; H335: C ≥ 1 % |  | CLP00
014-005-00-0 | tetraethyl silicate; ethyl silicate | 201-083-8 | 78-10-4 | Flam. Liq. 3 Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 | H226 H332 H335 H319 | GHS02 GHS07 Wng | H226 H332 H319 H335 |  |  |  | CLP00
014-006-00-6 | bis(4-fluorophenyl)-methyl-(1,2,4-triazol-4-ylmethyl)silane hydrochloride | 401-380-4 |  | Eye Irrit. 2 Aquatic Chronic 2 | H319 H411 | GHS07 GHS09 Wng | H319 H411 |  |  |  | CLP00
014-007-00-1 | triethoxyisobutylsilane | 402-810-3 | 17980-47-1 | Skin Irrit. 2 | H315 | GHS07 Wng | H315 |  |  |  | CLP00
014-008-00-7 | (chloromethyl)bis(4-fluorophenyl)methylsilane | 401-200-4 | 85491-26-5 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
014-009-00-2 | isobutylisopropyldimethoxysilane | 402-580-4 | 111439-76-0 | Flam. Liq. 3 Acute Tox. 4 * Skin Irrit. 2 | H226 H332 H315 | GHS02 GHS07 Wng | H226 H332 H315 |  |  |  | CLP00
014-010-00-8 | disodium metasilicate | 229-912-9 | 6834-92-0 | STOT SE 3 Skin Corr. 1B | H335 H314 | GHS05 GHS07 Dgr | H314 H335 |  |  |  | CLP00
014-011-00-3 | cyclohexyldimethoxymethylsilane | 402-140-1 | 17865-32-6 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
014-012-00-9 | bis(3-(trimethoxysilyl)propyl)amine | 403-480-3 |  | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
014-013-00-4 | α-hydroxypoly(methyl-(3-(2,2,6,6-tetramethylpiperidin-4-yloxy)propyl)siloxane) | 404-920-7 |  | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 2 | H312 H302 H314 H411 | GHS05 GHS07 GHS09 Dgr | H312 H302 H314 H411 |  |  |  | CLP00
014-014-00-X | etacelasil (ISO); 6-(2-chloroethyl)-6-(2-methoxyethoxy)-2,5,7,10-tetraoxa-6-silaundecane | 253-704-7 | 37894-46-5 | Repr. 1B Acute Tox. 4 * STOT RE 2 * | H360D *** H302 H373 ** | GHS08 GHS07 Dgr | H360D *** H302 H373 ** |  |  |  | CLP00
014-015-00-5 | α-trimethylsilanyl-ω-trimethylsiloxypoly[oxy(methyl-3-(2-(2-methoxypropoxy)propoxy)propylsilanediyl]-co-oxy(dimethylsilane)) | 406-420-4 | 69430-40-6 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
014-016-00-0 | reaction mass of: 1,3-dihex-5-en-1-yl-1,1,3,3-tetramethyldisiloxane; 1,3-dihex-n-en-1-yl-1,1,3,3-tetramethyldisiloxane | 406-490-6 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
014-017-00-6 | flusilazole (ISO); bis(4-fluorophenyl)(methyl)(1H-1,2,4-triazol-1-ylmethyl)silane |  | 85509-19-9 | Carc. 2 Repr. 1B Acute Tox. 4 * Aquatic Chronic 2 | H351 H360D *** H302 H411 | GHS08 GHS07 GHS09 Dgr | H351 H360D *** H302 H411 |  |  |  | CLP00
014-018-00-1 | octamethylcyclotetrasiloxane; [D4] | 209-136-7 | 556-67-2 | Repr. 2 Aquatic Chronic 1 | H361f *** H410 | GHS08 GHS09 Wng | H361f *** H410 |  | M=10 |  | CLP00/ATP15
014-019-00-7 | reaction mass of: 4-[[bis-(4-fluorophenyl)methylsilyl]methyl]-4H-1,2,4-triazole; 1-[[bis-(4-fluorophenyl)methylsilyl]methyl]-1H-1,2,4-triazole | 403-250-2 |  | Carc. 2 Repr. 1B Acute Tox. 4 * Aquatic Chronic 2 | H351 H360D *** H302 H411 | GHS08 GHS07 GHS09 Dgr | H351 H360D *** H302 H411 |  |  |  | CLP00
014-020-00-2 | bis(1,1-dimethyl-2-propynyloxy)dimethylsilane | 414-960-7 | 53863-99-3 | Acute Tox. 4 * | H332 | GHS07 Wng | H332 |  |  |  | CLP00
014-021-00-8 | tris(isopropenyloxy)phenyl silane | 411-340-8 | 52301-18-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
014-022-00-3 | reaction product of: (2-hydroxy-4-(3-propenoxy)benzophenone and triethoxysilane) with (hydrolysis product of silica and methyltrimethoxysilane) | 401-530-9 |  | Flam. Sol. 1 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 1 | H228 H332 H312 H302 H370 ** | GHS02 GHS08 GHS07 Dgr | H228 H370 ** H332 H312 H302 |  |  | T | CLP00
014-023-00-9 | α, ω-dihydroxypoly(hex-5-en-1-ylmethylsiloxane)hoxysilane with (hydrolysis product of silica and methyltrimethoxysilane)iazole | 408-160-7 | 125613-45-8 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
014-024-00-4 | 1-((3-(3-chloro-4-fluorophenyl)propyl)dimethylsilanyl)-4-ethoxybenzene | 412-620-2 | 121626-74-2 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
014-025-00-X | 4-[3-(diethoxymethylsilylpropoxy)-2,2,6,6-tetramethyl]piperidine | 411-400-3 | 102089-33-8 | Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H302 H373 ** H315 H318 H412 | GHS08 GHS05 GHS07 Dgr | H302 H373 ** H315 H318 H412 |  |  |  | CLP00
014-026-00-5 | dichloro-(3-(3-chloro-4-fluorophenyl)propyl)methylsilane | 407-180-3 | 770722-36-6 | Skin Corr. 1A | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
014-027-00-0 | chloro(3-(3-chloro-4-fluorophenyl)propyl)dimethylsilane | 410-270-5 | 770722-46-8 | Skin Corr. 1A | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
014-028-00-6 | α-[3-(1-oxoprop-2-eny)l-1-oxypropyl]dimethoxysilyloxy-ω-[3(1-oxoprop-2-enyl)-1-oxypropyl]dimethoxysilyl poly(dimethylsiloxane) | 415-290-8 | 193159-06-7 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
014-029-00-1 | O,O'-(ethenylmethylsilylene)di[(4-methylpentan-2-one)oxime] | 421-870-1 | 156145-66-3 | Repr. 2 Acute Tox. 4 * STOT RE 2 * | H361f *** H302 H373 ** | GHS08 GHS07 Wng | H361f *** H302 H373 ** |  |  |  | CLP00
014-030-00-7 | [(dimethylsilylene)bis((1,2,3,3a,7a-η)-1H-inden-1-ylidene)dimethyl]hafnium | 422-060-0 | 137390-08-0 | Acute Tox. 2 * | H300 | GHS06 Dgr | H300 |  |  |  | CLP00
014-031-00-2 | bis(1-methylethyl)-dimethoxysilane | 421-540-7 | 18230-61-0 | Flam. Liq. 3 Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H226 H315 H317 H412 | GHS02 GHS07 Wng | H226 H315 H317 H412 |  |  |  | CLP00
014-032-00-8 | dicyclopentyldimethoxysilane | 404-370-8 | 126990-35-0 | Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H400 H410 | GHS05 GHS09 Dgr | H315 H318 H410 |  |  |  | CLP00
014-033-00-3 | 2-methyl-3-(trimethoxysilyl)propyl-2-propenoate hydrolysis product with silica | 419-030-4 | 125804-20-8 | Flam. Liq. 2 STOT SE 3 Eye Irrit. 2 | H225 H336 H319 | GHS02 GHS07 Dgr | H225 H319 H336 |  |  |  | ATP01
014-034-00-9 | 3-hexylheptamethyltrisiloxane | 428-700-5 | 1873-90-1 | Acute Tox. 4 * Aquatic Chronic 4 | H332 H413 | GHS07 Wng | H332 H413 |  |  |  | ATP01
014-035-00-4 | 2-(3,4-epoxycyclohexyl)ethyltriethoxy silane | 425-050-4 | 10217-34-2 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
014-036-00-X | (4-ethoxyphenyl)(3-(4-fluoro-3-phenoxyphenyl)propyl)dimethylsilane | 405-020-7 | 105024-66-6 | Repr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H360F *** H400 H410 | GHS08 GHS09 Dgr | H360F *** H410 |  | M=1000 |  | ATP01
014-037-00-5 | 2-butanone-O,O',O''-(phenylsilylidyne)trioxime | 433-360-6 | 34036-80-1 | STOT RE 2 * Skin Sens. 1 Aquatic Chronic 3 | H373 ** H317 H412 | GHS08 GHS07 Wng | H373 ** H317 H412 |  |  |  | ATP01
014-038-00-0 | S-(3-(triethoxysilyl)propyl)octanethioate | 436-690-9 | 220727-26-4 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
014-039-00-6 | (2,3-dimethylbut-2-yl)-trimethoxysilane | 439-360-2 | 142877-45-0 | Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H315 H318 H412 | GHS05 Dgr | H315 H318 H412 |  |  |  | ATP01
014-041-00-7 | N,N-bis(trimethylsilyl)aminopropylmethyldiethoxysilane | 445-890-5 | 201290-01-9 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  |  |  | ATP01
014-042-00-2 | reaction mass of: O,O',O'',O'''-silanetetrayl tetrakis(4-methyl-2-pentanone oxime) (3 stereoisomers) | 423-010-0 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
014-043-00-8 | reaction product of amorphous silica (50-85 %), butyl (1-methylpropyl) magnesium (3-15 %), tetraethyl orthosilicate (5-15 %) and titanium tetrachloride (5-20 %) | 432-200-2 |  | STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H335 H315 H318 H412 | GHS05 GHS07 Dgr | H315 H318 H335 H412 |  |  |  | ATP01/ATP01corr
014-044-00-3 | 3-[(4'-acetoxy-3'-methoxyphenyl) propyl]trimethoxysilane | 433-050-0 | - | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
014-045-00-9 | magnesium sodium fluoride silicate | 442-650-1 | - | STOT RE 2 * | H373 ** | GHS08 Wng | H373 ** |  |  |  | ATP01
014-046-00-4 | e-glass microfibres of representative composition; [Calcium-aluminium-silicate fibres with random orientation with the following representative composition (% given by weight): SiO2 50,0-56,0 %, Al2O3 13,0-16,0 %, B2O3 5,8-10,0 %, Na2O < 0,6 %, K2O < 0,4 %, CaO 15,0-24,0 %, MgO < 5,5 %, Fe2O3 < 0,5 %, F2 < 1,0 %. Process: typically produced by flame attenuation and rotary process. (Additional individual elements may be present at low levels; the process list does not preclude innovation).] |  |  | Carc. 1B | H350i | GHS08 Dgr | H350i |  |  | A | ATP09
014-047-00-X | glass microfibres of representative composition; [Calcium-aluminium-silicate fibres with random orientation with the following composition (% given by weight): SiO2 55,0-60,0 %, Al2O3 4,0-7,0 %, B2O3 8,0-11,0 %, ZrO2 0,0-4,0 %, Na2O 9,5-13,5 %, K2O 0,0-4,0 %, CaO 1,0-5,0 %, MgO 0,0-2,0 %, Fe2O3 < 0,2 %, ZnO 2,0-5,0 %, BaO 3,0-6,0 %, F2 < 1,0 %. Process: typically produced by flame attenuation and rotary process. (Additional individual elements may be present at low levels; the process list does not preclude innovation).] |  |  | Carc. 2 | H351 (Inhalation) | GHS08 Wng | H351 (Inhalation) |  |  | A | ATP09
014-048-00-5 | silicon carbide fibres (with diameter < 3 μm, length > 5 μm and as­pect ratio ≥ 3:1) | 206-991-8 | 409-21-2 | Carc. 1B | H350i | GHS08 Dgr | H350i |  |  |  | ATP15
014-049-00-0 | trimethoxyvinylsilane; trimethoxy(vinyl)silane | 220-449-8 | 2768-02-7 | Skin Sens. 1B | H317 | GHS07 Wng | H317 |  |  |  | ATP15
014-050-00-6 | tris(2-methoxyethoxy) vinylsilane; 6-(2-methoxyethoxy)- 6-vinyl-2,5,7,10-tetra­oxa-6-silaundecane | 213-934-0 | 1067-53-4 | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  |  | ATP15
014-052-00-7 | silanamine, 1,1,1-trimethyl-N-(trimethylsilyl)-, hydrolysis products with silica; pyrogenic, synthetic amorphous, nano, surface treated silicon dioxide | 272-697-1 | 68909-20-6 | STOT RE 2 | H373 (lungs, inhalation) | GHS08 Wng | H373 (lungs, inhalation) | EUH066 |  |  | ATP18
015-001-00-1 | white phosphorus | 231-768-7 | 12185-10-3 | Pyr. Sol. 1 Acute Tox. 2 * Acute Tox. 2 * Skin Corr. 1A Aquatic Acute 1 | H250 H330 H300 H314 H400 | GHS02 GHS06 GHS05 GHS09 Dgr | H250 H330 H300 H314 H400 |  |  |  | CLP00
015-002-00-7 | red phosphorus | 231-768-7 | 7723-14-0 | Flam. Sol. 1 Aquatic Chronic 3 | H228 H412 | GHS02 Dgr | H228 H412 |  |  |  | CLP00
015-003-00-2 | calcium phosphide; tricalcium diphosphide | 215-142-0 | 1305-99-3 | Water-react. 1 Acute Tox. 1 Acute Tox. 2 Acute Tox. 3 Eye Dam. 1 Aquatic Acute 1 | H260 H330 H300 H311 H318 H400 | GHS02 GHS06 GHS05 GHS09 Dgr | H260 H300 H311 H330 H318 H400 | EUH029 EUH032 | M=100 |  | CLP00/ATP07
015-004-00-8 | aluminium phosphide | 244-088-0 | 20859-73-8 | Water-react. 1 Acute Tox. 1 Acute Tox. 2 Acute Tox. 3 Aquatic Acute 1 | H260 H330 H300 H311 H400 | GHS02 GHS06 GHS09 Dgr | H260 H300 H311 H330 H400 | EUH029 EUH032 | M=100 |  | CLP00/ATP05
015-005-00-3 | magnesium phosphide; trimagnesium diphosphide | 235-023-7 | 12057-74-8 | Water-react. 1 Acute Tox. 1 Acute Tox. 2 Acute Tox. 3 Aquatic Acute 1 | H260 H330 H300 H311 H400 | GHS02 GHS06 GHS09 Dgr | H260 H300 H311 H330 H400 | EUH029 EUH032 | M=100 |  | CLP00/ATP05
015-006-00-9 | trizinc diphosphide; zinc phosphide | 215-244-5 | 1314-84-7 | Water-react. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H260 H300 H400 H410 | GHS02 GHS06 GHS09 Dgr | H260 H300 H410 | EUH029 EUH032 | M=100 | T | CLP00/ATP01
015-007-00-4 | phosphorus trichloride | 231-749-3 | 7719-12-2 | Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Skin Corr. 1A | H330 H300 H373 ** H314 | GHS06 GHS08 GHS05 Dgr | H330 H300 H373 ** H314 | EUH014 EUH029 |  |  | CLP00
015-008-00-X | phosphorus pentachloride | 233-060-3 | 10026-13-8 | Acute Tox. 2 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B | H330 H302 H373 ** H314 | GHS06 GHS08 GHS05 Dgr | H330 H302 H373 ** H314 | EUH014 EUH029 |  |  | CLP00
015-009-00-5 | phosphoryl trichloride | 233-046-7 | 10025-87-3 | Acute Tox. 2 * Acute Tox. 4 * STOT RE 1 Skin Corr. 1A | H330 H302 H372 ** H314 | GHS06 GHS08 GHS05 Dgr | H330 H372 ** H302 H314 | EUH014 EUH029 |  |  | CLP00
015-010-00-0 | phosphorus pentoxide | 215-236-1 | 1314-56-3 | Skin Corr. 1A | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
015-011-00-6 | phosphoric acid ... %, orthophosphoric acid ... % | 231-633-2 | 7664-38-2 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  | Skin Corr. 1B; H314: C ≥ 25 % Skin Irrit. 2; H315: 10 % ≤ C < 25 % Eye Irrit. 2; H319: 10 % ≤ C < 25 % | B | CLP00
015-012-00-1 | tetraphosphorus trisulphide; phosphorus sesquisulphid | 215-245-0 | 1314-85-8 | Flam. Sol. 2 Water-react. 1 Acute Tox. 4 * Aquatic Acute 1 | H228 H260 H302 H400 | GHS02 GHS07 GHS09 Dgr | H228 H260 H302 H400 |  |  | T | CLP00
015-013-00-7 | triethyl phosphate | 201-114-5 | 78-40-0 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
015-014-00-2 | tributyl phosphate | 204-800-2 | 126-73-8 | Carc. 2 Acute Tox. 4 * Skin Irrit. 2 | H351 H302 H315 | GHS08 GHS07 Wng | H351 H302 H315 |  |  |  | CLP00
015-015-00-8 | tricresyl phosphate (o-o-o-, o-o-m-, o-o-p-, o-m-m-, o-m-p-, o-p-p-); tritolyl phosphate (o-o-o-, o-o-m-, o-o-p-, o-m-m-, o-m-p-, o-p-p-) | 201-103-5 | 78-30-8 | STOT SE 1 Aquatic Chronic 2 | H370 ** H411 | GHS08 GHS09 Dgr | H370 ** H411 |  | STOT SE 1; H370: C ≥ 1 % STOT SE 2; H371: 0,2 % ≤ C < 1 % | C | CLP00
015-016-00-3 | tricresyl phosphate (m-m-m-, m-m-p-, m-p-p-, p-p-p-); tritolyl phosphate (m-m-m-, m-m-p-, m-p-p-, p-p-p-) | 201-105-6 | 78-32-0 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 2 | H312 H302 H411 | GHS07 GHS09 Wng | H312 H302 H411 |  | * | C | CLP00
015-019-00-X | dichlorvos (ISO); 2,2-dichlorovinyl dimethyl phosphate | 200-547-7 | 62-73-7 | Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * Skin Sens. 1 Aquatic Acute 1 | H330 H311 H301 H317 H400 | GHS06 GHS09 Dgr | H330 H311 H301 H317 H400 |  | M=1000 |  | CLP00/ATP01
015-020-00-5 | mevinphos (ISO); 2-methoxycarbonyl-1-methylvinyl dimethyl phosphate | 232-095-1 | 7786-34-7 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H310 H300 H410 |  | M=10000 |  | CLP00
015-021-00-0 | trichlorfon (ISO); dimethyl 2,2,2-trichloro-1-hydroxyethylphosphonate | 200-149-3 | 52-68-6 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  | M=1000 |  | CLP00
015-022-00-6 | phosphamidon (ISO); 2-chloro-2-diethylcarbamoyl-1-methylvinyl dimethyl phosphate | 236-116-5 | 13171-21-6 | Muta. 2 Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H341 H300 H311 H400 H410 | GHS06 GHS08 GHS09 Dgr | H341 H300 H311 H410 |  |  |  | CLP00
015-023-00-1 | pyrazoxon; diethyl 3-methylpyrazol-5-yl phosphate |  | 108-34-9 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * | H310 H330 H300 | GHS06 Dgr | H330 H310 H300 |  |  |  | CLP00
015-024-00-7 | triamiphos (ISO); 5-amino-3-phenyl-1,2,4-triazol-1-yl-N,N,N',N'-tetramethylphosphonic diamide |  | 1031-47-6 | Acute Tox. 1 Acute Tox. 2 * | H310 H300 | GHS06 Dgr | H310 H300 |  |  |  | CLP00
015-025-00-2 | TEPP (ISO); tetraethyl pyrophosphate | 203-495-3 | 107-49-3 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 | H310 H300 H400 | GHS06 GHS09 Dgr | H310 H300 H400 |  |  |  | CLP00
015-026-00-8 | schradan (ISO); octamethylpyrophosphoramide | 205-801-0 | 152-16-9 | Acute Tox. 1 Acute Tox. 2 * | H310 H300 | GHS06 Dgr | H310 H300 |  |  |  | CLP00
015-027-00-3 | sulfotep (ISO); O,O,O,O-tetraethyl dithiopyrophosphate | 222-995-2 | 3689-24-5 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H310 H300 H410 |  | M=1000 |  | CLP00
015-028-00-9 | demeton-O (ISO); O,O-diethyl-O-2-ethylthioethyl phosphorothioate | 206-053-8 | 298-03-3 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 | H310 H300 H400 | GHS06 GHS09 Dgr | H310 H300 H400 |  |  |  | CLP00
015-029-00-4 | demeton-S (ISO); diethyl-S-2-ethylthioethyl phosphorothioate | 204-801-8 | 126-75-0 | Acute Tox. 1 Acute Tox. 2 * | H310 H300 | GHS06 Dgr | H310 H300 |  |  |  | CLP00
015-030-00-X | demeton-O-methyl (ISO); O-2-ethylthioethyl O,O-dimethyl phosphorothioate | 212-758-1 | 867-27-6 | Acute Tox. 3 * | H301 | GHS06 Dgr | H301 |  |  |  | CLP00
015-031-00-5 | demeton-S-methyl (ISO); S-2-ethylthioethyl dimethyl phosphorothioate | 213-052-6 | 919-86-8 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Chronic 2 | H311 H301 H411 | GHS06 GHS09 Dgr | H311 H301 H411 |  |  |  | CLP00
015-032-00-0 | prothoate (ISO); O,O-diethyl isopropylcarbamoylmethyl phosphorodithioate | 218-893-2 | 2275-18-5 | Acute Tox. 1 Acute Tox. 2 * Aquatic Chronic 3 | H310 H300 H412 | GHS06 Dgr | H310 H300 H412 |  |  |  | CLP00
015-033-00-6 | phorate (ISO); O,O-diethyl ethylthiomethyl phosphorodithioate | 206-052-2 | 298-02-2 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H310 H300 H410 |  | M=1000 |  | CLP00
015-034-00-1 | parathion (ISO); O,O-diethyl O-4-nitrophenyl phosphorothioate | 200-271-7 | 56-38-2 | Acute Tox. 2 * Acute Tox. 2 * Acute Tox. 3 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H300 H311 H372 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H300 H311 H372 ** H410 |  | M=100 |  | CLP00
015-035-00-7 | parathion - methyl (ISO); O,O-dimethyl O-4-nitrophenyl phosphorothioate | 206-050-1 | 298-00-0 | Flam. Liq. 3 Acute Tox. 2 * Acute Tox. 2 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H226 H330 H300 H311 H373 ** H400 H410 | GHS02 GHS06 GHS08 GHS09 Dgr | H226 H330 H300 H311 H373 ** H410 |  | M=100 |  | CLP00
015-036-00-2 | O-ethyl O-4-nitrophenyl phenylphosphonothioate; EPN | 218-276-8 | 2104-64-5 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H310 H300 H410 |  |  |  | CLP00
015-037-00-8 | phenkapton (ISO); S-(2,5-dichlorophenylthiomethyl) O,O-diethyl phosphorodithioate | 218-892-7 | 2275-14-1 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H301 H410 |  |  |  | CLP00
015-038-00-3 | coumaphos (ISO); O-3-chloro-4-methylcoumarin-7-yl O,O-diethyl phosphorothioate | 200-285-3 | 56-72-4 | Acute Tox. 2 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H312 H400 H410 | GHS06 GHS09 Dgr | H300 H312 H410 |  |  |  | CLP00
015-039-00-9 | azinphos-methyl (ISO); O,O-dimethyl-4-oxobenzotriazin-3-ylmethyl phosphorodithioate | 201-676-1 | 86-50-0 | Acute Tox. 2 * Acute Tox. 2 * Acute Tox. 3 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H300 H311 H317 H400 H410 | GHS06 GHS09 Dgr | H330 H300 H311 H317 H410 |  |  |  | CLP00
015-040-00-4 | diazinon (ISO); O,O-diethyl O-2-isopropyl-6-methylpyrimidin-4-yl phosphorothioate | 206-373-8 | 333-41-5 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
015-041-00-X | malathion (ISO); 1,2-bis(ethoxycarbonyl)ethyl O,O-dimethyl phosphorodithioate; [containing ≤ 0.03 % isomalathion] | 204-497-7 | 121-75-5 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  | M=1000 |  | CLP00/ATP01
015-042-00-5 | chlorthion; O-(3-chloro-4-nitrophenyl) O,O-dimethyl phosphorothioate | 207-902-5 | 500-28-7 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  | M=100 |  | CLP00
015-043-00-0 | phosnichlor (ISO); O-4-chloro-3-nitrophenyl O,O-dimethyl phosphorothioate |  | 5826-76-6 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H332 H312 H302 | GHS07 Wng | H332 H312 H302 |  |  |  | CLP00
015-044-00-6 | carbophenothion (ISO); 4-chlorophenylthiomethyl O,O-diethyl phosphorodithioate | 212-324-1 | 786-19-6 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H311 H301 H400 H410 | GHS06 GHS09 Dgr | H311 H301 H410 |  |  |  | CLP00
015-045-00-1 | mecarbam (ISO); N-ethoxycarbonyl-N-methylcarbamoylmethyl O,O-diethyl phosphorodithioate | 219-993-9 | 2595-54-2 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H311 H301 H400 H410 | GHS06 GHS09 Dgr | H311 H301 H410 |  |  |  | CLP00
015-046-00-7 | oxydemeton-methyl; S-2-(ethylsulphinyl)ethyl O,O-dimethyl phosphorothioate | 206-110-7 | 301-12-2 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 | H311 H301 H400 | GHS06 GHS09 Dgr | H311 H301 H400 |  |  |  | CLP00
015-047-00-2 | ethion (ISO); O,O,O',O'-tetraethyl S,S'-methylenedi (phosphorodithioate); diethion | 209-242-3 | 563-12-2 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H312 H400 H410 | GHS06 GHS09 Dgr | H301 H312 H410 |  | M=10000 |  | CLP00
015-048-00-8 | fenthion (ISO); O,O-dimethyl-O-(4-methylthion-m-tolyl) phosphorothioate | 200-231-9 | 55-38-9 | Muta. 2 Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H341 H331 H312 H302 H372 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H341 H331 H312 H302 H372 ** H410 |  | M=100 |  | CLP00/ATP01
015-049-00-3 | endothion (ISO); S-5-methoxy-4-oxopyran-2-ylmethyl dimethyl phosphorothioate | 220-472-3 | 2778-04-3 | Acute Tox. 3 * Acute Tox. 3 * | H311 H301 | GHS06 Dgr | H311 H301 |  |  |  | CLP00
015-050-00-9 | thiometon (ISO); S-2-ethylthioethyl O,O-dimethyl phosphorodithioate | 211-362-6 | 640-15-3 | Acute Tox. 3 * Acute Tox. 4 * | H301 H312 | GHS06 Dgr | H301 H312 |  |  |  | CLP00
015-051-00-4 | dimethoate (ISO); O,O-dimethyl methylcarbamoylmethyl phosphorodithioate | 200-480-3 | 60-51-5 | Acute Tox. 4 * Acute Tox. 4 * | H312 H302 | GHS07 Wng | H312 H302 |  |  |  | CLP00
015-052-00-X | fenchlorphos (ISO); O,O-dimethyl O-2,4,5-trichlorophenyl phosphorothioate | 206-082-6 | 299-84-3 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H400 H410 | GHS07 GHS09 Wng | H312 H302 H410 |  |  |  | CLP00
015-053-00-5 | menazon (ISO); S-[(4,6-diamino-1,3,5-triazin-2-yl)methyl] O,O-dimethyl phosphorodithioate | 201-123-4 | 78-57-9 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
015-054-00-0 | fenitrothion (ISO); O,O-dimethyl O-4-nitro-m-tolyl phosphorothioate | 204-524-2 | 122-14-5 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
015-055-00-6 | naled (ISO); 1,2-dibromo-2,2-dichloroethyl dimethyl phosphate | 206-098-3 | 300-76-5 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 | H312 H302 H315 H319 H400 | GHS07 GHS09 Wng | H312 H302 H319 H315 H400 |  | M=1000 |  | CLP00
015-056-00-1 | azinphos-ethyl (ISO); O,O-diethyl 4-oxobenzotriazin-3-ylmethyl phosphorodithioate | 220-147-6 | 2642-71-9 | Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H311 H400 H410 | GHS06 GHS09 Dgr | H300 H311 H410 |  | M=100 |  | CLP00/ATP01
015-057-00-7 | formothion (ISO); N-formyl-N-methylcarbamoylmethyl O,O-dimethyl phosphorodithioate | 219-818-6 | 2540-82-1 | Acute Tox. 4 * Acute Tox. 4 * | H312 H302 | GHS07 Wng | H312 H302 |  |  |  | CLP00
015-058-00-2 | morphothion (ISO); O,O-dimethyl-S-(morpholinocarbonylmethyl) phosphorodithioate | 205-628-0 | 144-41-2 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H301 H410 |  |  |  | CLP00
015-059-00-8 | vamidothion (ISO); O,O-dimethyl S-2-(1-methylcarbamoylethylthio) ethyl phosphorothioate | 218-894-8 | 2275-23-2 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 | H301 H312 H400 | GHS06 GHS09 Dgr | H301 H312 H400 |  |  |  | CLP00
015-060-00-3 | disulfoton (ISO); O,O-diethyl 2-ethylthioethyl phosphorodithioate | 206-054-3 | 298-04-4 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H310 H300 H410 |  |  |  | CLP00
015-061-00-9 | dimefox (ISO); tetramethylphosphorodiamidic fluoride | 204-076-8 | 115-26-4 | Acute Tox. 1 Acute Tox. 2 * | H310 H300 | GHS06 Dgr | H310 H300 |  |  |  | CLP00
015-062-00-4 | mipafox (ISO); N,N'- di-isopropylphosphorodiamidic fluoride | 206-742-3 | 371-86-8 | STOT SE 1 | H370 ** | GHS08 Dgr | H370 ** |  |  |  | CLP00
015-063-00-X | dioxathion (ISO); 1,4-dioxan-2,3-diyl-O,O,O',O'-tetraethyl di(phosphorodithioate) | 201-107-7 | 78-34-2 | Acute Tox. 2 * Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H330 H300 H311 H400 H410 | GHS06 GHS09 Dgr | H330 H300 H311 H410 |  | M=1000 |  | CLP00
015-064-00-5 | bromophos-ethyl (ISO); O-4-bromo-2,5-dichlorophenyl O,O-diethyl phosphorothioate | 225-399-0 | 4824-78-6 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H312 H400 H410 | GHS06 GHS09 Dgr | H301 H312 H410 |  |  |  | CLP00
015-065-00-0 | S-[2-(ethylsulphinyl)ethyl] O,O-dimethyl phosphorodithioate |  | 2703-37-9 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Aquatic Chronic 2 | H310 H330 H300 H411 | GHS06 GHS09 Dgr | H330 H310 H300 H411 |  |  |  | CLP00
015-066-00-6 | omethoate (ISO); O,O-dimethyl S-methylcarbamoylmethyl phosphorothioate | 214-197-8 | 1113-02-6 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 | H301 H312 H400 | GHS06 GHS09 Dgr | H301 H312 H400 |  |  |  | CLP00
015-067-00-1 | phosalone (ISO); S-(6-chloro-2-oxobenzoxazolin-3-ylmethyl) O,O-diethyl phosphorodithioate | 218-996-2 | 2310-17-0 | Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H332 H312 H317 H400 H410 | GHS06 GHS09 Dgr | H301 H332 H312 H317 H410 |  | M=1000 |  | CLP00/ATP01
015-068-00-7 | dichlofenthion (ISO); O-,4-dichlorophenyl O,O-diethyl phosphorothioate | 202-564-5 | 97-17-6 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
015-069-00-2 | methidathion (ISO); 2,3-dihydro-5-methoxy-2-oxo-1,3,4-thiadiazol-3-ylmethyl-O,O-dimethylphosphorodithioate | 213-449-4 | 950-37-8 | Acute Tox. 2 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H312 H400 H410 | GHS06 GHS09 Dgr | H300 H312 H410 |  |  |  | CLP00
015-070-00-8 | cyanthoate (ISO); S-(N-(1-cyano-1-methylethyl)carbamoylmethyl) O,O-diethyl phosphorothioate | 223-099-4 | 3734-95-0 | Acute Tox. 2 * Acute Tox. 3 * | H300 H311 | GHS06 Dgr | H300 H311 |  |  |  | CLP00
015-071-00-3 | chlorfenvinphos (ISO); 2-chloro-1-(2,4 dichlorophenyl) vinyl diethyl phosphate | 207-432-0 | 470-90-6 | Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H311 H400 H410 | GHS06 GHS09 Dgr | H300 H311 H410 |  |  |  | CLP00
015-072-00-9 | monocrotophos (ISO); dimethyl-1-methyl-2-(methylcarbamoyl)vinyl phosphate | 230-042-7 | 6923-22-4 | Muta. 2 Acute Tox. 2 * Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H341 H330 H300 H311 H400 H410 | GHS06 GHS08 GHS09 Dgr | H341 H330 H300 H311 H410 |  |  |  | CLP00
015-073-00-4 | dicrotophos (ISO); (Z)-2-dimethylcarbamoyl-1-methylvinyl dimethyl phosphate | 205-494-3 | 141-66-2 | Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H311 H400 H410 | GHS06 GHS09 Dgr | H300 H311 H410 |  |  |  | CLP00
015-074-00-X | crufomate (ISO); 4-tert-butyl-2-chlorophenyl methyl methylphosphoramidate | 206-083-1 | 299-86-5 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H400 H410 | GHS07 GHS09 Wng | H312 H302 H410 |  |  |  | CLP00
015-075-00-5 | S-[2-(isopropylsulphinyl)ethyl] O,O-dimethyl phosphorothioate |  | 2635-50-9 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H331 H311 H301 | GHS06 Dgr | H331 H311 H301 |  |  |  | CLP00
015-076-00-0 | potasan; O,O-diethyl O-(4-methylcoumarin-7-yl) phosphorothioate |  | 299-45-6 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H400 H410 | GHS06 GHS09 Dgr | H330 H310 H300 H410 |  | M=1000 |  | CLP00
015-077-00-6 | 2,2-dichlorovinyl 2-ethylsulphinylethyl methyl phosphate |  | 7076-53-1 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H331 H311 H301 | GHS06 Dgr | H331 H311 H301 |  |  |  | CLP00
015-078-00-1 | demeton-S-methylsulphon (ISO); S-2-ethylsulphonylethyl dimethyl phosphorothioate | 241-109-5 | 17040-19-6 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Chronic 2 | H301 H312 H411 | GHS06 GHS09 Dgr | H301 H312 H411 |  |  |  | CLP00
015-079-00-7 | acephate (ISO); O,S-dimethyl acetylphosphoramidothioate | 250-241-2 | 30560-19-1 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
015-080-00-2 | amidithion (ISO); 2-methoxyethylcarbamoylmethyl O,O-dimethyl phosphorodithioate |  | 919-76-6 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
015-081-00-8 | O,O,O',O'-tetrapropyl dithiopyrophosphate | 221-817-0 | 3244-90-4 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H400 H410 | GHS07 GHS09 Wng | H312 H302 H410 |  |  |  | CLP00
015-082-00-3 | azothoate (ISO); O-4-(4-chlorophenylazo)phenyl O,O-dimethyl phosphorothioate | 227-419-3 | 5834-96-8 | Acute Tox. 4 * Acute Tox. 4 * | H332 H302 | GHS07 Wng | H332 H302 |  |  |  | CLP00
015-083-00-9 | bensulide (ISO); O,O-diisopropyl 2-phenylsulphonylaminoethyl phosphorodithioate | 212-010-4 | 741-58-2 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
015-084-00-4 | chlorpyrifos (ISO); O,O-diethyl O-3,5,6-trichloro-2-pyridyl phosphorothioate | 220-864-4 | 2921-88-2 | Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H400 H410 | GHS06 GHS09 Dgr | H301 H410 |  | M=10000 |  | CLP00
015-085-00-X | chlorphonium chloride (ISO); tributyl (2,4-dichlorobenzyl) phosphonium chloride | 204-105-4 | 115-78-6 | Acute Tox. 3 * Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 | H301 H312 H315 H319 | GHS06 Dgr | H301 H312 H319 H315 |  |  |  | CLP00
015-086-00-5 | coumithoate (ISO); O,O-diethyl O-,8,9,10-tetrahydro-6-oxo-benzo(c)chromen-3-yl phosphorothioate |  | 572-48-5 | Acute Tox. 3 * | H301 | GHS06 Dgr | H301 |  |  |  | CLP00
015-087-00-0 | cyanophos (ISO); O-4-cyanophenyl O,O-dimethyl phosphorothioate | 220-130-3 | 2636-26-2 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H400 H410 | GHS07 GHS09 Wng | H312 H302 H410 |  |  |  | CLP00
015-088-00-6 | dialifos (ISO); 2-chloro-1-phthalimidoethyl O,O-diethyl phosphorodithioate | 233-689-3 | 10311-84-9 | Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H311 H400 H410 | GHS06 GHS09 Dgr | H311 H300 H410 |  |  |  | CLP00
015-089-00-1 | ethoate-methyl (ISO); ethylcarbamoylmethyl O,O-dimethyl phosphorodithioate | 204-121-1 | 116-01-8 | Acute Tox. 4 * Acute Tox. 4 * | H312 H302 | GHS07 Wng | H312 H302 |  |  |  | CLP00
015-090-00-7 | fensulfothion (ISO); O,O-diethyl O-4-methylsulfinylphenyl phosphorothioate | 204-114-3 | 115-90-2 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H310 H300 H410 |  |  |  | CLP00
015-091-00-2 | fonofos (ISO); O-ethyl phenyl ethylphosphonodithioate | 213-408-0 | 944-22-9 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H310 H300 H410 |  |  |  | CLP00
015-092-00-8 | phosacetim (ISO); O,O-bis(4-chlorophenyl) N-acetimidoylphosphoramidothioate | 223-874-7 | 4104-14-7 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H310 H300 H410 |  |  |  | CLP00
015-093-00-3 | leptophos (ISO); O-4-bromo-2,5-dichlorophenyl O-methyl phenylphosphorothioate | 244-472-8 | 21609-90-5 | Acute Tox. 3 * Acute Tox. 4 * STOT SE 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H312 H370 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H301 H370 ** H312 H410 |  |  |  | CLP00
015-094-00-9 | mephosfolan (ISO); diethyl 4-methyl-1,3-dithiolan-2-ylidenephosphoramidate | 213-447-3 | 950-10-7 | Acute Tox. 1 Acute Tox. 2 * Aquatic Chronic 2 | H310 H300 H411 | GHS06 GHS09 Dgr | H310 H300 H411 |  |  |  | CLP00
015-095-00-4 | methamidophos (ISO); O,S-dimethyl phosphoramidothioate | 233-606-0 | 10265-92-6 | Acute Tox. 2 * Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 | H330 H300 H311 H400 | GHS06 GHS09 Dgr | H330 H300 H311 H400 |  |  |  | CLP00
015-096-00-X | oxydisulfoton (ISO); O, O-diethyl S-2-ethylsulphinylethyl phosphorodithioate | 219-679-1 | 2497-07-6 | Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H311 H400 H410 | GHS06 GHS09 Dgr | H300 H311 H410 |  | M=10 |  | CLP00
015-097-00-5 | phenthoate (ISO); ethyl 2-(dimethoxyphosphinothioylthio)-2-phenylacetate | 219-997-0 | 2597-03-7 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H400 H410 | GHS07 GHS09 Wng | H312 H302 H410 |  | M=100 |  | CLP00
015-098-00-0 | trichloronate (ISO); O-ethyl O-2,4,5-trichlorophenyl ethylphosphonothioate | 206-326-1 | 327-98-0 | Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H311 H400 H410 | GHS06 GHS09 Dgr | H300 H311 H410 |  |  |  | CLP00
015-099-00-6 | pirimiphos-ethyl (ISO); O,O-diethyl O-2-diethylamino-6-methylpyrimidin-4-yl phosphorothioate | 245-704-0 | 23505-41-1 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H312 H400 H410 | GHS06 GHS09 Dgr | H301 H312 H410 |  |  |  | CLP00
015-100-00-X | phoxim (ISO); α-(diethoxyphosphinothioylimino) phenylacetonitrile | 238-887-3 | 14816-18-3 | Repr. 2 Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361f *** H302 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H361f *** H302 H317 H410 |  | M=1000 |  | CLP00/ATP01
015-101-00-5 | phosmet (ISO); S-[(1,3-dioxo-1,3-dihydro-2H-isoindol-2-yl)methyl] O,O-di-methyl phosphorodithioate; O,O-dimethyl-S-phthalimido-methyl phosphorodithioate | 211-987-4 | 732-11-6 | Repr. 2 Acute Tox. 3 Acute Tox. 4 STOT SE 1 Aquatic Acute 1 Aquatic Chronic 1 | H361f H301 H332 H370 (nervous system) H400 H410 | GHS08 GHS06 GHS09 Dgr | H332 H301 H361f H370 (nervous system) H410 |  | M=100 M=100 |  | CLP00/ATP13
015-102-00-0 | tris(2-chloroethyl)phosphate | 204-118-5 | 115-96-8 | Carc. 2 Repr. 1B Acute Tox. 4 * Aquatic Chronic 2 | H351 H360F *** H302 H411 | GHS08 GHS07 GHS09 Dgr | H351 H360F *** H302 H411 |  |  |  | CLP00/ATP01
015-103-00-6 | phosphorus tribromide | 232-178-2 | 7789-60-8 | STOT SE 3 Skin Corr. 1B | H335 H314 | GHS05 GHS07 Dgr | H314 H335 | EUH014 |  |  | CLP00
015-104-00-1 | diphosphorus pentasulphide; phosphorus pentasulphide | 215-242-4 | 1314-80-3 | Flam. Sol. 1 Water-react. 1 Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 | H228 H260 H332 H302 H400 | GHS02 GHS07 GHS09 Dgr | H228 H260 H332 H302 H400 | EUH029 |  | T | CLP00
015-105-00-7 | triphenyl phosphite | 202-908-4 | 101-02-0 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H319 H400 H410 | GHS07 GHS09 Wng | H319 H315 H410 |  | Skin Irrit. 2; H315: C ≥ 5 % Eye Irrit. 2; H319: C ≥ 5 % |  | CLP00
015-106-00-2 | hexamethylphosphoric triamide; hexamethylphosphoramide | 211-653-8 | 680-31-9 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  | Carc. 1B; H350: C ≥ 0,01 % |  | CLP00
015-107-00-8 | ethoprophos (ISO); ethyl-S,S-dipropyl phosphorodithioate | 236-152-1 | 13194-48-4 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 3 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H301 H317 H400 H410 | GHS06 GHS09 Dgr | H330 H310 H301 H317 H410 |  |  |  | CLP00
015-108-00-3 | bromophos (ISO); O-4-bromo-2,5-dichlorophenyl O,O-dimethyl phosphorothioate | 218-277-3 | 2104-96-3 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  | M=100 |  | CLP00
015-109-00-9 | crotoxyphos (ISO); 1-phenylethyl 3-(dimethoxyphosphinyloxy) isocrotonate | 231-720-5 | 7700-17-6 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H311 H301 H400 H410 | GHS06 GHS09 Dgr | H311 H301 H410 |  | M=10 |  | CLP00
015-110-00-4 | cyanofenphos (ISO); O-4-cyanophenyl O-ethyl phenylphosphonothioate |  | 13067-93-1 | Acute Tox. 3 * Acute Tox. 4 * STOT SE 1 Eye Irrit. 2 Aquatic Chronic 2 | H301 H312 H370 ** H319 H411 | GHS06 GHS08 GHS09 Dgr | H301 H370 ** H312 H319 H411 |  |  |  | CLP00
015-111-00-X | phosfolan (ISO); diethyl 1,3-dithiolan-2-ylidenephosphoramidate | 213-423-2 | 947-02-4 | Acute Tox. 1 Acute Tox. 2 * | H310 H300 | GHS06 Dgr | H310 H300 |  |  |  | CLP00
015-112-00-5 | thionazin (ISO); O,O-diethyl O-pyrazin-2-yl phosphorothioate | 206-049-6 | 297-97-2 | Acute Tox. 1 Acute Tox. 2 * | H310 H300 | GHS06 Dgr | H310 H300 |  |  |  | CLP00
015-113-00-0 | tolclofos-methyl (ISO); O-(2,6-dichloro-p-tolyl)-O,O-dimethyl thiophosphate | 260-515-3 | 57018-04-9 | Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  | M=1 M=1 |  | ATP01/ATP17
015-114-00-6 | chlormephos (ISO); S-chloromethyl O,O-diethyl phosphorodithioate | 246-538-1 | 24934-91-6 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H300 H310 H410 |  | M=10 |  | CLP00/ATP01corr
015-115-00-1 | chlorthiophos (ISO); [isomeric reaction mass in which O-2,5-dichlorophenyl-4-methylthiophenyl O,O-diethyl phosphorothioate predominates] | 244-663-6 | 21923-23-9 | Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H311 H400 H410 | GHS06 GHS09 Dgr | H300 H311 H410 |  | M=1000 |  | CLP00/ATP01corr
015-116-00-7 | demephion-O (ISO); O,O-dimethyl O-2-methylthioethyl phosphorothioate | 211-666-9 | 682-80-4 | Acute Tox. 2 * Acute Tox. 3 * | H300 H311 | GHS06 Dgr | H300 H311 |  |  |  | CLP00
015-117-00-2 | demephion-S (ISO); O,O-dimethyl S-2-methylthioethyl phosphorothioate | 219-971-9 | 2587-90-8 | Acute Tox. 2 * Acute Tox. 3 * | H300 H311 | GHS06 Dgr | H300 H311 |  |  |  | CLP00
015-118-00-8 | demeton |  | 8065-48-3 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 | H310 H300 H400 | GHS06 GHS09 Dgr | H310 H300 H400 |  |  |  | CLP00
015-119-00-3 | dimethyl 4-(methylthio)phenyl phosphate |  | 3254-63-5 | Acute Tox. 1 Acute Tox. 2 * | H310 H300 | GHS06 Dgr | H310 H300 |  |  |  | CLP00
015-120-00-9 | ditalimfos (ISO); O,O-diethyl phthalimidophosphonothioate | 225-875-8 | 5131-24-8 | Skin Irrit. 2 Skin Sens. 1 | H315 H317 | GHS07 Wng | H315 H317 |  |  |  | CLP00
015-121-00-4 | edifenphos (ISO); O-ethyl S,S-diphenyl phosphorodithioate | 241-178-1 | 17109-49-8 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H312 H317 H400 H410 | GHS06 GHS09 Dgr | H331 H301 H312 H317 H410 |  |  |  | CLP00
015-122-00-X | etrimfos (ISO); O-6-ethoxy-2-ethylpyrimidin-4-yl O,O-dimethylphosphorothioate | 253-855-9 | 38260-54-7 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  | M=10 |  | CLP00
015-123-00-5 | fenamiphos (ISO); ethyl-4-methylthio-m-tolyl isopropyl phosphoramidate | 244-848-1 | 22224-92-6 | Acute Tox. 2 Acute Tox. 2 Acute Tox. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H330 H310 H300 H319 H400 H410 | GHS06 GHS09 Dgr | H300 H310 H330 H319 H410 |  | M=100 M=100 |  | CLP00/ATP05
015-124-00-0 | fosthietan (ISO); diethyl 1,3-dithietan-2-ylidenephosphoramidate | 244-437-7 | 21548-32-3 | Acute Tox. 1 Acute Tox. 2 * | H310 H300 | GHS06 Dgr | H310 H300 |  |  |  | CLP00
015-125-00-6 | glyphosine (ISO); N,N-bis(phosphonomethyl)glycine | 219-468-4 | 2439-99-8 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
015-126-00-1 | heptenophos (ISO); 7-chlorobicyclo(3.2.0)hepta-2,6-dien-6-yl dimethyl phosphate | 245-737-0 | 23560-59-0 | Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H400 H410 | GHS06 GHS09 Dgr | H301 H410 |  | M=100 |  | CLP00
015-127-00-7 | iprobenfos (ISO); S-benzyl diisopropyl phosphorothioate | 247-449-0 | 26087-47-8 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
015-128-00-2 | IPSP; S-ethylsulphinylmethyl O,O-diisopropylphosphorodithioate |  | 5827-05-4 | Acute Tox. 1 Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H301 H400 H410 | GHS06 GHS09 Dgr | H310 H301 H410 |  | M=100 |  | CLP00
015-129-00-8 | isofenphos (ISO); O-ethyl O-2-isopropoxycarbonylphenyl-isopropylphosphoramidothioate | 246-814-1 | 25311-71-1 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H311 H301 H400 H410 | GHS06 GHS09 Dgr | H311 H301 H410 |  | M=100 |  | CLP00
015-130-00-3 | isothioate (ISO); S-2-isopropylthioethyl O,O-dimethyl phosphorodithioate |  | 36614-38-7 | Acute Tox. 3 * Acute Tox. 3 * | H311 H301 | GHS06 Dgr | H311 H301 |  |  |  | CLP00
015-131-00-9 | isoxathion (ISO); O,O-diethyl O-5-phenylisoxazol-3-ylphosphorothioate | 242-624-8 | 18854-01-8 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H311 H301 H400 H410 | GHS06 GHS09 Dgr | H311 H301 H410 |  |  |  | CLP00
015-132-00-4 | S-(chlorophenylthiomethyl) O,O-dimethylphosphorodithioate; methylcarbophenothione |  | 953-17-3 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H311 H301 H400 H410 | GHS06 GHS09 Dgr | H311 H301 H410 |  | M=1000 |  | CLP00
015-133-00-X | piperophos (ISO); S-2-methylpiperidinocarbonylmethyl-O,O-dipropyl phosphorodithioate |  | 24151-93-7 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  | M=10 |  | CLP00
015-134-00-5 | pirimiphos-methyl (ISO); O-[2-(diethylamino)-6-methylpyrimidin-4-yl] O,O-dimethyl phosphorothioate | 249-528-5 | 29232-93-7 | Acute Tox. 4 STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H372 (nervous system) H400 H410 | GHS07 GHS08 GHS09 Dgr | H302 H372 (nervous system) H410 |  | Oral: ATE = 1414 mg/kg bw M=1000 M=1000 |  | CLP00/ATP15
015-135-00-0 | profenofos (ISO); O-(4-bromo-2-chlorophenyl) O-ethyl S-propyl phosphorothioate | 255-255-2 | 41198-08-7 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  | M=1000 |  | CLP00
015-136-00-6 | trans-isopropyl-3-[[(ethylamino)methoxyfosfinothioyl]oxy]crotonate; isopropyl 3-[[(ethylamino)methoxyphosphinothioyl]oxy]isocrotonate; propetamphos (ISO) | 250-517-2 | 31218-83-4 | Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H400 H410 | GHS06 GHS09 Dgr | H301 H410 |  | M=100 |  | CLP00
015-137-00-1 | pyrazophos (ISO); O,O-diethyl O-(6-ethoxycarbonyl-5-methylpyrazolo[2,3-a]pyrimidin-2-yl) phosphorothioate | 236-656-1 | 13457-18-6 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H400 H410 | GHS07 GHS09 Wng | H332 H302 H410 |  |  |  | CLP00
015-138-00-7 | quinalphos (ISO); O,O-diethyl-O-quinoxalin-2-yl phosphorothioate | 237-031-6 | 13593-03-8 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H312 H400 H410 | GHS06 GHS09 Dgr | H301 H312 H410 |  | M=1000 |  | CLP00
015-139-00-2 | terbufos (ISO); S-tert-butylthiomethyl O,O-diethylphosphorodithioate | 235-963-8 | 13071-79-9 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H310 H300 H410 |  | M=1000 |  | CLP00
015-140-00-8 | triazophos (ISO); O,O-diethyl-O-1-phenyl-1H-1,2,4-triazol-3-yl phosphorothioate | 245-986-5 | 24017-47-8 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H312 H400 H410 | GHS06 GHS09 Dgr | H331 H301 H312 H410 |  | M=100 |  | CLP00/ATP01
015-141-00-3 | ethylenediammonium O,O-bis(octyl) phosphorodithioate, mixed isomers | 400-520-1 |  | Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H400 H410 | GHS05 GHS07 GHS09 Dgr | H314 H302 H410 |  |  |  | CLP00
015-142-00-9 | butyl (dialkyloxy(dibutoxyphosphoryloxy))titanium (trialkyloxy)titanium phosphate | 401-100-0 |  | Flam. Liq. 2 Eye Irrit. 2 Aquatic Chronic 2 | H225 H319 H411 | GHS02 GHS07 GHS09 Dgr | H225 H319 H411 |  |  | T | CLP00
015-143-00-4 | reaction mass of 2-chloroethyl chloropropyl 2-chloroethylphosphonate, mixture reaction mass of isomers and 2-chloroethyl chloropropyl 2-chloropropylphosphonate, reaction mass of isomers | 401-740-0 |  | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
015-144-00-X | reaction mass of pentyl methylphosphinate and 2-methylbutyl methylphosphinate | 402-090-0 | 87025-52-3 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
015-145-00-5 | reaction mass of copper(I) O,O-diisopropyl phosphorodithioate and copper(I) O-isopropyl O-(4-methylpent-2-yl) phosphorodithioate and copper(I) O,O-bis(4-methylpent-2-yl) phosphorodithioate | 401-520-4 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
015-146-00-0 | S-(tricyclo(5.2.1.02,6)deca-3-en-8(or 9)-yl O-(isopropyl or isobutyl or 2-ethylhexyl) O-(isopropyl or isobutyl or 2-ethylhexyl) phosphorodithioate | 401-850-9 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
015-147-00-6 | reaction mass of C12-14-tert-alkylammonium diphenyl phosphorothioate and dinonyl sulphide (or disulphide) | 400-930-0 |  | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H315 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H315 H318 H317 H411 |  |  |  | CLP00
015-148-00-1 | 2-(diphosphonomethyl)succinic acid | 403-070-4 | 51395-42-7 | Skin Corr. 1B Skin Sens. 1 | H314 H317 | GHS05 GHS07 Dgr | H314 H317 |  |  |  | CLP00
015-149-00-7 | reaction mass of: hexyldioctylphosphineoxide; dihexyloctylphosphineoxide; trioctylphosphineoxide | 403-470-9 |  | Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H314 H400 H410 | GHS05 GHS09 Dgr | H314 H410 |  |  |  | CLP00
015-150-00-2 | (2-(1,3-dioxolan-2-yl)ethyl)triphenylphosphonium bromide | 404-940-6 | 86608-70-0 | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Aquatic Chronic 3 | H302 H373 ** H318 H412 | GHS08 GHS05 GHS07 Dgr | H302 H318 H373 ** H412 |  |  |  | CLP00
015-151-00-8 | tris(isopropyl/tert-butylphenyl) phosphate | 405-010-2 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
015-152-00-3 | dioxabenzofos (ISO); 2-methoxy-4H-1,3,2-benzodioxaphosphorin 2-sulphide | 223-292-3 | 3811-49-2 | Acute Tox. 3 * Acute Tox. 3 * STOT SE 1 Aquatic Chronic 2 | H311 H301 H370 ** H411 | GHS06 GHS08 GHS09 Dgr | H311 H301 H370 ** H411 |  |  |  | CLP00
015-153-00-9 | isazofos (ISO); O-(5-chloro-1-isopropyl-1,2,4-triazol-3-yl) O,O-diethyl phosphorothioate | 255-863-8 | 42509-80-8 | Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H311 H301 H373 ** H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H311 H301 H373 ** H317 H410 |  |  |  | CLP00
015-154-00-4 | ethephon; 2-chloroethylphosphonic acid | 240-718-3 | 16672-87-0 | Acute Tox. 3 Acute Tox. 4 Acute Tox. 4 Skin Corr. 1C Aquatic Chronic 2 | H311 H332 H302 H314 H411 | GHS06 GHS05 GHS09 Dgr | H302 H311 H332 H314 H411 | EUH071 |  |  | CLP00/ATP06
015-155-00-X | glufosinate ammonium (ISO); ammonium 2-amino-4-(hydroxymethylphosphinyl)butyrate | 278-636-5 | 77182-82-2 | Repr. 1B Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * | H360Fd H332 H312 H302 H373 ** | GHS08 GHS07 Dgr | H360Fd H332 H312 H302 H373 ** |  |  |  | CLP00/ATP01
015-156-00-5 | methyl 3-[(dimethoxyphosphinothioyl)oxy]methacrylate [1] methacrifos (ISO); methyl (E)-3-[(dimethoxyphosphinothioyl)oxy]methacrylate [2] | 250-366-2 [1] 250-366-2 [2] | 30864-28-9 [1] 62610-77-9 [2] | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
015-157-00-0 | phosphonic acid [1] phosphorous acid [2] | 237-066-7 [1] 233-663-1 [2] | 13598-36-2 [1] 10294-56-1 [2] | Acute Tox. 4 * Skin Corr. 1A | H302 H314 | GHS05 GHS07 Dgr | H302 H314 |  |  |  | CLP00
015-158-00-6 | (η-cyclopentadienyl)(η-cumenyl)iron(1+)hexafluorophosphate(1-) | 402-340-9 | 32760-80-8 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
015-159-00-1 | hydroxyphosphonoacetic acid | 405-710-8 | 23783-26-8 | Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Skin Sens. 1 | H302 H373 ** H314 H317 | GHS08 GHS05 GHS07 Dgr | H302 H373 ** H314 H317 |  |  |  | CLP00
015-160-00-7 | vanadyl pyrophosphate | 406-260-5 | 58834-75-6 | Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H319 H317 H412 | GHS07 Wng | H319 H317 H412 |  |  |  | CLP00
015-161-00-2 | divanadyl pyrophosphate | 407-130-0 | 65232-89-5 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H317 H411 |  |  |  | CLP00
015-162-00-8 | vanadium(IV) oxide hydrogen phosphate hemihydrate, lithium, zinc, molybdenum, iron and chlorine-doped | 407-350-7 |  | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Aquatic Chronic 2 | H332 H373 ** H318 H411 | GHS08 GHS05 GHS07 GHS09 Dgr | H332 H373 ** H318 H411 |  |  |  | CLP00
015-163-00-3 | bis(2,6-dimethoxybenzoyl)-2,4,4-trimethylpentylphosphinoxide | 412-010-6 | 145052-34-2 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
015-164-00-9 | calcium P,P'-(1-hydroxyethylene)bis(hydrogen phosphonate)dihydrate | 400-480-5 | 36669-85-9 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
015-165-00-4 | reaction mass of: thiobis(4,1-phenylene)-S,S,S',S'-tetraphenyldisulfonium bishexafluorophosphate; diphenyl(4-phenylthiophenyl)sulfonium hexafluorophosphate | 404-986-7 |  | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | CLP00
015-166-00-X | 3,9-bis(2,6-di-tert-butyl-4-methylphenoxy)-2,4,8,10-tetraoxa-3,9-diphosphaspiro[5.5]undecane | 410-290-4 | 80693-00-1 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
015-167-00-5 | 3-(hydroxyphenylphosphinyl)propanoic acid | 411-200-6 | 14657-64-8 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
015-168-00-0 | fosthiazate (ISO); (RS)-S-sec-butyl-O-ethyl-2-oxo-1,3-thiazolidin-3-ylphosphonothioate |  | 98886-44-3 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H312 H317 H400 H410 | GHS06 GHS09 Dgr | H331 H301 H312 H317 H410 | EUH070 |  |  | CLP00
015-169-00-6 | tributyltetradecylphosphonium tetrafluoroborate | 413-520-1 |  | Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H314 H317 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H302 H373 ** H314 H317 H410 |  |  |  | CLP00
015-170-00-1 | reaction mass of: di-(1-octane-N,N,N-trimethylammonium) octylphosphate; 1-octane-N,N,N-trimethylammonium di-octylphosphate; 1-octane-N,N,N-trimethylammonium octylphosphate | 407-490-9 |  | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H312 H302 H314 | GHS05 GHS07 Dgr | H312 H302 H314 |  |  |  | CLP00
015-171-00-7 | O,O,O-tris(2(or 4)-C9-10-isoalkylphenyl) phosphorothioate | 406-940-1 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
015-172-00-2 | reaction mass of: bis(isotridecylammonium)mono(di-(4-methylpent-2-yloxy)thiophosphorothionylisopropyl)phosphate; isotridecylammonium bis(di-(4-methylpent-2-yloxy)thiophosphorothionylisopropyl)phosphate | 406-240-6 |  | Flam. Liq. 3 Skin Corr. 1B Aquatic Chronic 2 | H226 H314 H411 | GHS02 GHS05 GHS09 Dgr | H226 H314 H411 |  |  |  | CLP00
015-173-00-8 | methyl [2-(1,1-dimethylethyl)-6-methoxypyrimidin-4-yl]ethylphosphonothioate | 414-080-3 | 117291-73-3 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
015-174-00-3 | 1-chloro-N,N-diethyl-1,1-diphenyl-1-(phenylmethyl)phosphoramine | 411-370-1 | 82857-68-9 | Acute Tox. 3 * Eye Dam. 1 Aquatic Chronic 2 | H301 H318 H411 | GHS06 GHS05 GHS09 Dgr | H301 H318 H411 |  |  |  | CLP00
015-175-00-9 | tert-butyl (triphenylphosphoranylidene) acetate | 412-880-7 | 35000-38-5 | Acute Tox. 3 * STOT RE 2 * Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H301 H373 ** H319 H317 H411 | GHS06 GHS08 GHS09 Dgr | H301 H373 ** H319 H317 H411 |  |  |  | CLP00
015-176-00-4 | P,P,P',P'-tetrakis-(o-methoxyphenyl)propane-1,3-diphosphine | 413-430-2 | 116163-96-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
015-177-00-X | ((4-phenylbutyl)hydroxyphosphoryl)acetic acid | 412-170-7 | 83623-61-4 | STOT RE 2 * Eye Dam. 1 Skin Sens. 1 | H373 ** H318 H317 | GHS08 GHS05 Dgr | H373 ** H318 H317 |  |  |  | CLP00
015-178-00-5 | (R)-α-phenylethylammonium (-)-(1R, 2S)-(1,2-epoxypropyl)phosphonate monohydrate | 418-570-8 | 25383-07-7 | Repr. 2 Aquatic Chronic 2 | H361f *** H411 | GHS08 GHS09 Wng | H361f *** H411 |  |  |  | CLP00
015-179-00-0 | UVCB condensation product of: tetrakis-hydroxymethylphosphonium chloride, urea and distilled hydrogenated C16-18 tallow alkylamine | 422-720-8 | 166242-53-1 | Carc. 2 Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H373 ** H314 H317 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H351 H302 H373 ** H314 H317 H410 |  |  |  | CLP00
015-180-00-6 | [R-(R*,S*)]-[[2-methyl-1-(1-oxopropoxy)propoxy]-(4-phenylbutyl)phosphinyl] acetic acid, (-)-cinchonidine (1:1) salt | 415-820-8 | 137590-32-0 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | CLP00
015-181-00-1 | phosphine | 232-260-8 | 7803-51-2 | Flam. Gas 1 Press. Gas Acute Tox. 1 Skin Corr. 1B Aquatic Acute 1 | H220 H330 H314 H400 | GHS02 GHS04 GHS06 GHS05 GHS09 Dgr | H220 H330 H314 H400 |  | Inhalation: ATE = 10 ppmV (Gases) | U | CLP00/ATP15
015-182-00-7 | tetrapropan-2-yl (dichloromethanediyl)bis(phosphonate) | 430-630-5 | 10596-22-2 | Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 | H302 H319 H317 | GHS07 Wng | H302 H319 H317 |  |  |  | ATP01/ATP01corr
015-183-00-2 | (1-hydroxydodecylidene)diphosphonic acid | 425-230-2 | 16610-63-2 | Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H314 H400 H410 | GHS05 GHS09 Dgr | H314 H410 |  |  |  | ATP01
015-184-00-8 | Salts of glyphosate, with the exception of those specified elsewhere in this Annex |  |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  | A | CLP00
015-186-00-9 | chlorpyrifos-methyl (ISO),; O, O-dimethyl O-3,5,6-trichloro-2-pyridyl phosphorothioate | 227-011-5 | 5598-13-0 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  | M=10000 |  | CLP00
015-187-00-4 | reaction mass of: tetrasodium(((2-hydroxyethyl)imino)bis(methylene))bisphosphonate, N-oxide; trisodium ((tetrahydro-2-hydroxy-4H-1,4,2-oxazaphosphorin-4-yl)-methyl)phosphonate, N-oxide, P-oxide | 417-540-1 |  | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
015-189-00-5 | phenyl bis(2,4,6-trimethylbenzoyl)-phosphine oxide | 423-340-5 | 162881-26-7 | Skin Sens. 1A Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00/ATP14
015-190-00-0 | bis(2,4-dicumylphenyl) neopentyl diphosphite; 3,9-bis[2,4-bis(1-methyl-1-phenylethyl)phenoxy]-2,4,8,10-tetraoxa-3,9-diphosphaspiro[5.5]undecane | 421-920-2 | 154862-43-8 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
015-191-00-6 | dodecyldiphenyl phosphate | 431-760-5 | 27460-02-2 | Skin Irrit. 2 Aquatic Chronic 3 | H315 H412 | GHS07 Wng | H315 H412 |  |  |  | ATP01
015-193-00-7 | triphenyl(phenylmethyl)phosphonium 1,1,2,2,3,3,4,4,4-nonafluoro-N-methyl-1-butanesulfonamide (1:1) | 442-960-7 | 332350-93-3 | Acute Tox. 3 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H318 H400 H410 | GHS05 GHS06 GHS09 Dgr | H301 H318 H410 |  |  |  | ATP01
015-194-00-2 | tetrabutyl-phosphonium nonafluoro-butane-1-sulfonate | 444-440-5 | 220689-12-3 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | ATP01
015-195-00-8 | reaction mass of: potassium o-toluenephosphonate; potassium m-toluenephosphonate; potassium p-toluenephosphonate | 433-860-4 | - | Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H319 H317 H412 | GHS07 Wng | H319 H317 H412 |  |  |  | ATP01
015-196-00-3 | reaction mass of: dimethyl (2-(hydroxymethylcarbamoyl)ethyl)phosphonate; diethyl (2-(hydroxymethylcarbamoyl)ethyl)phosphonate; methyl ethyl (2-(hydroxymethylcarbamoyl)ethyl)phosphonate | 435-960-3 | - | Carc. 1B Muta. 1B Skin Sens. 1 | H350 H340 H317 | GHS08 GHS07 Dgr | H350 H340 H317 |  |  |  | ATP01
015-197-00-9 | bis(2,4,4-trimethylpentyl)dithiophosphonic acid | 420-160-9 | 107667-02-7 | Flam. Liq. 3 Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 2 | H226 H331 H302 H314 H411 | GHS02 GHS06 GHS05 GHS09 Dgr | H226 H331 H302 H314 H411 |  |  |  | ATP01
015-198-00-4 | (4-phenylbutyl)phosphinic acid | 420-450-5 | 86552-32-1 | Carc. 2 Eye Dam. 1 | H351 H318 | GHS05 GHS08 Dgr | H351 H318 |  |  |  | ATP01
015-199-00-X | tris[2-chloro-1-(chloromethyl)ethyl] phosphate | 237-159-2 | 13674-87-8 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | ATP03
015-200-00-3 | indium phosphide | 244-959-5 | 22398-80-7 | Carc. 1B Repr. 2 STOT RE 1 | H350 H361f H372 (lungs) | GHS08 Dgr | H350 H361f H372 (lungs) |  | Carc 1B; H350: C ≥ ,01 % STOT RE 1; : C ≥ ,1 % STOT RE 2; H373: ,01 % ≤ C < ,1 % |  | ATP03
015-201-00-9 | trixylyl phosphate | 246-677-8 | 25155-23-1 | Repr. 1B | H360F | GHS08 Dgr | H360F |  |  |  | ATP03
015-202-00-4 | tris(nonylphenyl) phosphite | 247-759-6 | 26523-78-4 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | ATP03
015-203-00-X | diphenyl(2,4,6-trimethylbenzoyl)phosphine oxide | 278-355-8 | 75980-60-8 | Repr. 1B Skin Sens. 1B | H360Fd H317 | GHS08 GHS07 Dgr | H360Fd H317 |  |  |  | ATP03/ATP21
015-204-00-5 | benzyl(diethylamino)diphenylphosphonium 4-[1,1,1,3,3,3-hexafluoro-2-(4-hydroxyphenyl)propan-2-yl]phenolate | 479-100-5 | 577705-90-9 | Repr. 1B | H360F | GHS08 Dgr | H360F |  |  |  | ATP21
015-205-00-0 | benzyltriphenylphosphonium, salt with 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]bis[phenol] (1:1) | 278-305-5 | 75768-65-9 | Repr. 1B | H360F | GHS08 Dgr | H360F |  |  |  | ATP21
015-206-00-6 | reaction mass of 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]diphenol and benzyl(diethylamino)diphenylphosphonium 4-[1,1,1,3,3,3-hexafluoro-2-(4-hydroxyphenyl)propan-2-yl]phenolate (1:1) |  |  | Repr. 1B | H360F | GHS08 Dgr | H360F |  |  |  | ATP21
015-207-00-1 | reaction mass of 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]diphenol and benzyltriphenylphosphonium, salt with 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]diphenol (1:1) |  |  | Repr. 1B | H360F | GHS08 Dgr | H360F |  |  |  | ATP21
015-208-00-7 | dimethyl propylphosphonate | 242-555-3 | 18755-43-6 | Muta. 1B Repr. 1B | H340 H360Df | GHS08 Dgr | H340 H360Df |  |  |  | ATP21
016-001-00-4 | hydrogen sulphide, hydrogen sulfide | 231-977-3 | 7783-06-4 | Flam. Gas 1A Press. Gas Acute Tox. 2 Aquatic Acute 1 | H220 H330 H400 | GHS02 GHS06 GHS09 Dgr | H220 H330 H400 |  | inhalation: ATE = 440 ppmV(gases) | U | CLP00/ATP21
016-002-00-X | barium sulphide | 244-214-4 | 21109-95-5 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 | H332 H302 H400 | GHS07 GHS09 Wng | H332 H302 H400 | EUH031 |  |  | CLP00
016-003-00-5 | barium polysulphides | 256-814-3 | 50864-67-0 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 | H335 H315 H319 H400 | GHS07 GHS09 Wng | H319 H335 H315 H400 | EUH031 |  |  | CLP00
016-004-00-0 | calcium sulphide | 243-873-5 | 20548-54-3 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 | H335 H315 H319 H400 | GHS07 GHS09 Wng | H319 H335 H315 H400 | EUH031 |  |  | CLP00
016-005-00-6 | calcium polysulphides | 215-709-2 | 1344-81-6 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 | H335 H315 H319 H400 | GHS07 GHS09 Wng | H319 H335 H315 H400 | EUH031 |  |  | CLP00
016-006-00-1 | dipotassium sulphide; potassium sulphide | 215-197-0 | 1312-73-8 | Skin Corr. 1B Aquatic Acute 1 | H314 H400 | GHS05 GHS09 Dgr | H314 H400 | EUH031 |  |  | CLP00
016-007-00-7 | potassium polysulphides | 253-390-1 | 37199-66-9 | Skin Corr. 1B Aquatic Acute 1 | H314 H400 | GHS05 GHS09 Dgr | H314 H400 | EUH031 |  |  | CLP00
016-008-00-2 | ammonium polysulphides | 232-989-1 | 9080-17-5 | Skin Corr. 1B Aquatic Acute 1 | H314 H400 | GHS05 GHS09 Dgr | H314 H400 | EUH031 | EUH031: C ≥ 1 % |  | CLP00
016-009-00-8 | disodium sulfide; sodium sulfide | 215-211-5 | 1313-82-2 | Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 | H311 H302 H314 H400 | GHS06 GHS05 GHS09 Dgr | H311 H302 H314 H400 |  |  |  | CLP00/ATP01
016-010-00-3 | sodium polysulphides | 215-686-9 | 1344-08-7 | Acute Tox. 3 * Skin Corr. 1B Aquatic Acute 1 | H301 H314 H400 | GHS06 GHS05 GHS09 Dgr | H301 H314 H400 | EUH031 |  |  | CLP00
016-011-00-9 | sulphur dioxide; sulfur dioxide | 231-195-2 | 7446-09-5 | Press. Gas Acute Tox. 3 STOT SE 1 Skin. Corr. 1B | H331 H370(respiratory system, inhalation) H314 | GHS04 GHS06 GHS08 GHS05 Dgr | H331 H370(respiratory system, inhalation) H314 |  | inhalation: ATE = 1000 ppmV(gases) | U 5 | CLP00/ATP21
016-012-00-4 | disulphur dichloride; sulfur monochloride | 233-036-2 | 10025-67-9 | Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1A Aquatic Acute 1 | H301 H332 H314 H400 | GHS06 GHS05 GHS09 Dgr | H301 H332 H314 H400 | EUH014 EUH029 | STOT SE 3; H335: C ≥ 1 % |  | CLP00
016-013-00-X | sulphur dichloride | 234-129-0 | 10545-99-0 | STOT SE 3 Skin Corr. 1B Aquatic Acute 1 | H335 H314 H400 | GHS05 GHS07 GHS09 Dgr | H314 H335 H400 | EUH014 | STOT SE 3; H335: C ≥ 5 % |  | CLP00
016-014-00-5 | sulphur tetrachloride |  | 13451-08-6 | Skin Corr. 1B Aquatic Acute 1 | H314 H400 | GHS05 GHS09 Dgr | H314 H400 | EUH014 | STOT SE 3; H335: C ≥ 5 % |  | CLP00
016-015-00-0 | thionyl dichloride; thionyl chloride | 231-748-8 | 7719-09-7 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A | H332 H302 H314 | GHS05 GHS07 Dgr | H332 H302 H314 | EUH014 EUH029 | STOT SE 3; H335: C ≥ 1 % |  | CLP00
016-016-00-6 | sulphuryl chloride | 232-245-6 | 7791-25-5 | STOT SE 3 Skin Corr. 1B | H335 H314 | GHS05 GHS07 Dgr | H314 H335 | EUH014 |  |  | CLP00
016-017-00-1 | chlorosulphonic acid | 232-234-6 | 7790-94-5 | STOT SE 3 Skin Corr. 1A | H335 H314 | GHS05 GHS07 Dgr | H314 H335 | EUH014 |  |  | CLP00
016-018-00-7 | fluorosulphonic acid | 232-149-4 | 7789-21-1 | Acute Tox. 4 * Skin Corr. 1A | H332 H314 | GHS05 GHS07 Dgr | H332 H314 |  |  |  | CLP00
016-019-00-2 | oleum ... % SO3 |  |  | STOT SE 3 Skin Corr. 1A | H335 H314 | GHS05 GHS07 Dgr | H314 H335 | EUH014 |  | B | CLP00
016-020-00-8 | sulphuric acid ... % | 231-639-5 | 7664-93-9 | Skin Corr. 1A | H314 | GHS05 Dgr | H314 |  | Skin Corr. 1A; H314: C ≥ 15 % Skin Irrit. 2; H315: 5 % ≤ C < 15 % Eye Irrit. 2; H319: 5 % ≤ C < 15 % | B | CLP00
016-021-00-3 | methanethiol; methyl mercaptan | 200-822-1 | 74-93-1 | Flam. Gas 1 Press. Gas Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H220 H331 H400 H410 | GHS02 GHS04 GHS06 GHS09 Dgr | H220 H331 H410 |  |  | U | CLP00
016-022-00-9 | ethanethiol; ethyl mercaptan | 200-837-3 | 75-08-1 | Flam. Liq. 1 Acute Tox. 3 Acute Tox. 4 Aquatic Acute 1 Aquatic Chronic 1 | H224 H331 H302 H400 H410 | GHS02 GHS06 GHS09 Dgr | H224 H331 H302 H410 |  | inhalation: ATE = 7,1 mg/L (vapours) oral: ATE = 680 mg/kg bw |  | CLP00/ATP22
016-023-00-4 | dimethyl sulphate | 201-058-1 | 77-78-1 | Carc. 1B Muta. 2 Acute Tox. 2 * Acute Tox. 3 * Skin Corr. 1B Skin Sens. 1 | H350 H341 H330 H301 H314 H317 | GHS06 GHS08 GHS05 Dgr | H350 H341 H330 H301 H314 H317 |  | Carc. 1B; H350: C ≥ 0,01 % Muta. 2; H341: C ≥ 0,01 % STOT SE 3; H335: C ≥ 5 % |  | CLP00
016-024-00-X | dimexano (ISO); bis(methoxythiocarbonyl) disulphide | 215-993-8 | 1468-37-7 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
016-025-00-5 | disul (ISO); 2-(2,4-dichlorophenoxy)ethyl hydrogensulphate; 2,4-DES | 205-259-5 | 149-26-8 | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 | H302 H315 H318 | GHS05 GHS07 Dgr | H302 H315 H318 |  |  |  | CLP00
016-026-00-0 | sulphamidic acid; sulphamic acid; sulfamic acid | 226-218-8 | 5329-14-6 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 3 | H315 H319 H412 | GHS07 Wng | H319 H315 H412 |  |  |  | CLP00
016-027-00-6 | diethyl sulphate | 200-589-6 | 64-67-5 | Carc. 1B Muta. 1B Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H350 H340 H332 H312 H302 H314 | GHS05 GHS08 GHS07 Dgr | H350 H340 H332 H312 H302 H314 |  |  |  | CLP00
016-028-00-1 | sodium dithionite; sodium hydrosulphite | 231-890-0 | 7775-14-6 | Self-heat. 1 Acute Tox. 4 * | H251 H302 | GHS02 GHS07 Dgr | H251 H302 | EUH031 |  |  | CLP00
016-029-00-7 | p-toluenesulphonic acid, containing more than 5 % H2SO4 |  |  | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  | Skin Corr. 1B; H314: C ≥ 25 % Skin Irrit. 2; H315: 10 % ≤ C < 25 % Eye Irrit. 2; H319: 10 % ≤ C < 25 % |  | CLP00
016-030-00-2 | p-toluenesulphonic acid (containing a maximum of 5 % H2SO4) | 203-180-0 | 104-15-4 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H335 H315 H319 | GHS07 Wng | H319 H335 H315 |  | STOT SE 3; H335: C ≥ 20 % |  | CLP00
016-031-00-8 | tetrahydrothiophene-1,1-dioxide; sulpholane | 204-783-1 | 126-33-0 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
016-032-00-3 | 1,3-propanesultone; 1,2-oxathiolane 2,2-dioxide | 214-317-9 | 1120-71-4 | Carc. 1B Acute Tox. 4 * Acute Tox. 4 * | H350 H312 H302 | GHS08 GHS07 Dgr | H350 H312 H302 |  | Carc. 1B; H350: C ≥ 0,01 % |  | CLP00
016-033-00-9 | dimethylsulfamoylchloride | 236-412-4 | 13360-57-1 | Carc. 1B Acute Tox. 2 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H350 H330 H312 H302 H314 | GHS06 GHS05 GHS08 Dgr | H350 H330 H312 H302 H314 |  |  |  | CLP00
016-034-00-4 | tetrasodium 3,3'-(piperazine-1,4-diylbis((6-chloro-1,3,5-triazine-2,4-diyl)imino(2-acetamido)-4,1-phenyleneazo))bis(naphthalene-1,5-disulphonate) | 400-010-9 | 81898-60-4 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
016-035-00-X | pentasodium 5-anilino-3-(4-(4-(6-chloro-4-(3-sulphonatoanilino)-1,3,5-triazin-2-ylamino)-2,5-dimethylphenylazo)-2,5-disulphonatophenylazo)-4-hydroxynaphthalene-2,7-disulphonate | 400-120-7 |  | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
016-036-00-5 | tetrasodium 5-(4,6-dichloro-5-cyanopyrimidin-2-ylamino)-4-hydroxy-2,3-azodinaphthalene-1,2,5,7-disulphonate | 400-130-1 |  | Resp. Sens. 1 Aquatic Chronic 2 | H334 H411 | GHS08 GHS09 Dgr | H334 H411 |  |  |  | CLP00
016-037-00-0 | disodium 1-amino-4-(4-benzenesulphonamido-3-sulphonatoanilino)anthraquinone-2-sulphonate | 400-350-8 | 85153-93-1 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
016-038-00-6 | disodium 6-((4-chloro-6-(N-methyl)-2-toluidino)-1,3,5-triazin-2-ylamino)-1-hydroxy-2-(4-methoxy-2-sulphonatophenylazo)naphthalene-3-sulphonate | 400-380-1 | 86393-35-3 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
016-039-00-1 | tetrasodium 2-(6-chloro-4-(4-(2,5-dimethyl-4-(2,5-disulphonatophenylazo)phenylazo)-3-ureidoanilino)-1,3,5-triazin-2-ylamino)benzene-1,4-disulphonate | 400-430-2 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
016-040-00-7 | reaction mass of disodium 6-(2,4-dihydroxyphenylazo)-3-(4-(4-(2,4-dihydroxyphenylazo)anilino)-3-sulphonatophenylazo)-4-hydroxynaphthalene-2-sulphonate and disodium 6-(2,4-diaminophenylazo)-3-(4-(4-(2,4-diaminophenylazo)anilino)-3-sulphonatophenylazo)-4-hydroxynaphthalene-2-sulphonate and trisodium 6-(2,4-dihydroxyphenylazo)-3-(4-(4-(7-(2,4-dihydroxyphenylazo)-1-hydroxy-3-sulphonato-2-naphthylazo)anilino)-3-sulphonatophenylazo)-4-hydroxynaphthalene-2-sulphonate | 400-570-4 |  | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
016-041-00-2 | calcium 2,5-dichloro-4-(4-((5-chloro-4-methyl-2-sulphonatophenyl)azo)-5-hydroxy-3-methylpyrazol-1-yl)benzenesulphonate | 400-710-4 |  | Acute Tox. 4 * | H332 | GHS07 Wng | H332 |  |  |  | CLP00
016-042-00-8 | tetrasodium 5-benzamido-3-(5-(4-fluoro-6-(1-sulphonato-2-naphthylamino)-1,3,5-triazin-2-ylamino)-2-sulphonatophenylazo)-4-hydroxynaphthalene-2,7- disulphonate | 400-790-0 | 85665-97-0 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H315 H319 H317 | GHS07 Wng | H319 H315 H317 |  |  |  | CLP00
016-043-00-3 | dilithium 6-acetamido-4-hydroxy-3-(4-((2-sulphonatooxy)ethylsulphonyl)phenylazo)naphthalene-2-sulphonate | 401-010-1 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
016-044-00-9 | disodium S,S-hexane-1,6-diyldi(thiosulphate) dihydrate | 401-320-7 |  | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
016-045-00-4 | lithium sodium hydrogen 4-amino-6-(5-(5-chloro-2,6-difluoropyrimidin-4-ylamino)-2-sulphonatophenylazo)-5-hydroxy-3-(4-(2-(sulphonatooxy)ethylsulphonyl)phenylazo)naphthalene-2,7-disulphonate | 401-560-2 | 108624-00-6 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
016-046-00-X | sodium hydrogensulphate | 231-665-7 | 7681-38-1 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
016-047-00-5 | hexasodium 7-(4-(4-(4-(2,5-disulphonatoanilino)-6-fluoro-1,3,5-triazin-2-ylamino)-2-methylphenylazo)-7-sulphonatonaphthylazo)naphthalene-1,3,5- trisulphonate | 401-650-1 | 85665-96-9 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
016-048-00-0 | sodium 3,5-dichloro-2-(5-cyano-2,6-bis(3-hydroxypropylamino)-4-methylpyridin-3-ylazo)benzenesulphonate | 401-870-8 |  | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
016-049-00-6 | calcium octadecylxylenesulphonate | 402-040-8 |  | Skin Corr. 1B Aquatic Chronic 2 | H314 H411 | GHS05 GHS09 Dgr | H314 H411 |  |  |  | CLP00
016-050-00-1 | potassium sodium 5-(4-chloro-6-(N-(4-(4-chloro-6-(5-hydroxy-2,7-disulphonato-6-(2-sulphonatophenylazo)-4-naphthylamino)-1,3,5-triazin-2-ylamino) phenyl-N-methyl)amino)-1,3,5-triazin-2-ylamino)-4-hydroxy-3-(2-sulphonatophenylazo)naphthalene-2,7-disulphonat | 402-150-6 |  | Eye Irrit. 2 Skin Sens. 1 | H319 H317 | GHS07 Wng | H319 H317 |  |  |  | CLP00
016-051-00-7 | trisodium 7-(4-(6-fluoro-4-(2-(2-vinylsulphonylethoxy)ethylamino)-1,3,5-triazin-2-ylamino)-2-ureidophenylazo)naphthalene-1,3,6- trisulphonate | 402-170-5 | 106359-91-5 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
016-052-00-2 | benzyltributylammonium 4-hydroxynaphthalene-1-sulphonate | 402-240-5 | 102561-46-6 | Acute Tox. 4 * Aquatic Chronic 2 | H332 H411 | GHS07 GHS09 Wng | H332 H411 |  |  |  | CLP00
016-053-00-8 | (C16 or C18-n-alkyl)(C16 or C18-n-alkyl)ammonium 2-((C16 or C18-n-alkyl)(C16 or C18-n-alkyl)carbamoyl)benzenesulphonate | 402-460-1 |  | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 4 | H315 H317 H413 | GHS07 Wng | H315 H317 H413 |  |  |  | CLP00
016-054-00-3 | sodium 4-(2,4,4-trimethylpentylcarbonyloxy)benzenesulfonate | 400-030-8 |  | Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 STOT RE 1 Eye Irrit. 2 Skin Sens. 1 | H331 H302 H335 H372 ** H319 H317 | GHS06 GHS08 Dgr | H331 H372 ** H302 H319 H335 H317 |  |  |  | CLP00
016-055-00-9 | tetrasodium 4-amino-3,6-bis(5-(6-chloro-4-(2-hydroxyethylamino)-1,3,5-triazin-2-ylamino)-2-sulfonatophenylazo)-5-hydroxynaphthalene-2,7-sulfonate (containing > 35 % sodium chloride and sodium acetate) | 400-510-7 |  | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
016-056-00-4 | potassium hydrogensulphate | 231-594-1 | 7646-93-7 | STOT SE 3 Skin Corr. 1B | H335 H314 | GHS05 GHS07 Dgr | H314 H335 |  |  |  | CLP00
016-057-00-X | styrene-4-sulfonyl chloride | 404-770-2 | 2633-67-2 | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H315 H318 H317 | GHS05 GHS07 Dgr | H315 H318 H317 |  |  |  | CLP00
016-058-00-5 | thionyl chloride, reaction products with 1,3,4-thiadiazol-2,5-dithiol, tert-nonanethiol and C12-14-tert-alkylamine | 404-820-3 |  | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H315 H317 H412 | GHS07 Wng | H315 H317 H412 |  |  |  | CLP00
016-059-00-0 | N,N,N',N'-tetramethyldithiobis(ethylene)diamine dihydrochloride | 405-300-9 | 17339-60-5 | Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H319 H317 H400 H410 | GHS07 GHS09 Wng | H302 H319 H317 H410 |  |  |  | CLP00
016-060-00-6 | diammonium peroxodisulphate; ammonium persulphate | 231-786-5 | 7727-54-0 | Ox. Sol. 3 Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 Skin Sens. 1 | H272 H302 H335 H315 H319 H334 H317 | GHS03 GHS08 GHS07 Dgr | H272 H302 H319 H335 H315 H334 H317 |  |  |  | CLP00
016-061-00-1 | dipotassium peroxodisulphate; potassium persulphate | 231-781-8 | 7727-21-1 | Ox. Sol. 3 Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 Skin Sens. 1 | H272 H302 H335 H315 H319 H334 H317 | GHS03 GHS08 GHS07 Dgr | H272 H302 H319 H335 H315 H334 H317 |  |  |  | CLP00
016-062-00-7 | bensultap (ISO); 1,3-bis(phenylsulfonylthio)-2-(N,N-dimethylamino)propane |  | 17606-31-4 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
016-063-00-2 | sodium metabisulphite | 231-673-0 | 7681-57-4 | Acute Tox. 4 * Eye Dam. 1 | H302 H318 | GHS05 GHS07 Dgr | H302 H318 | EUH031 |  |  | CLP00
016-064-00-8 | sodium hydrogensulphite … %; sodium bisulphite … % | 231-548-0 | 7631-90-5 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 | EUH031 |  | B | CLP00
016-065-00-3 | sodium 1-amino-4-[2-methyl-5-(4-methylphenylsulfonylamino)phenylamino]anthraquinone-2-sulfonate | 400-100-8 | 84057-97-6 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
016-066-00-9 | tetrasodium [5-((4-amino-6-chloro-1,3,5-triazin-2-yl)amino)-2-((2-hydroxy-3,5-disulfonatophenylazo)-2- sulfonatobenzylidenehydrazino)benzoate]copper(II) | 404-070-7 | 116912-62-0 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
016-067-00-4 | (4-methylphenyl)mesitylene sulfonate | 407-530-5 | 67811-06-7 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
016-068-00-X | sodium 3,5-bis(tetradecyloxycarbonyl)benzenesulfinate | 407-720-8 | 155160-86-4 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
016-069-00-5 | 3,5-bis-(tetradecyloxycarbonyl)benzenesulfinic acid | 407-990-7 | 141915-64-2 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
016-070-00-0 | 4-benzyloxy-4'-(2,3-epoxy-2-methylprop-1-yloxy)diphenylsulfone | 408-220-2 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
016-071-00-6 | trisodium 3-amino-6,13-dichloro-10-((3-((4-chloro-6-(2-sulfophenylamino)-1,3,5-triazin-2-yl)amino)propyl) amino)-4,11-triphenoxydioxazinedisulfonate | 410-130-3 | 136248-03-8 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
016-072-00-1 | 3-amino-4-hydroxy-N-(2-methoxyethyl)-benzenesulfonamide | 411-520-6 | 112195-27-4 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H318 H317 H411 |  |  |  | CLP00
016-073-00-7 | tetrakis(phenylmethyl)thioperoxydi(carbothioamide) | 404-310-0 | 10591-85-2 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
016-074-00-2 | 6-fluoro-2-methyl-3-(4-methylthiobenzyl)indene | 405-410-7 |  | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H315 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H315 H318 H317 H411 |  |  |  | CLP00
016-075-00-8 | 2,2'-diallyl-4,4'-sulfonyldiphenol | 411-570-9 | 41481-66-7 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
016-076-00-3 | 2,3-bis((2-mercaptoethyl)thio)-1-propanethiol | 411-290-7 | 131538-00-6 | Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H410 |  |  |  | CLP00
016-077-00-9 | 2-chloro-p-toluenesulfochloride | 412-890-1 | 42413-03-6 | Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 3 | H314 H317 H412 | GHS05 GHS07 Dgr | H314 H317 H412 |  |  |  | CLP00
016-078-00-4 | 4-methyl-N,N-bis(2-(((4-methylphenyl)sulfonyl)amino)ethyl)benzenesulfonamide | 413-300-5 | 56187-04-3 | Aquatic Chronic 4 | H413 | [5 empty] | CLP00
016-079-00-X | N,N-bis(2-(p-toluenesulfonyloxy)ethyl)-p-toluenesulfonamide | 412-920-3 | 16695-22-0 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
016-080-00-5 | sodium 2-anilino-5-(2-nitro-4-(N-phenylsulfamoyl))anilinobenzenesulfonate | 412-320-1 | 31361-99-6 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
016-081-00-0 | hexahydrocyclopenta[c]pyrrole-1-(1H)-ammonium N-ethoxycarbonyl-N-(p-tolylsulfonyl)azanide | 418-350-1 |  | Muta. 2 Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H341 H302 H319 H317 H411 | GHS08 GHS07 GHS09 Wng | H341 H302 H319 H317 H411 |  |  |  | CLP00
016-082-00-6 | ethoxysulfuron (ISO); 1-(4,6-dimethoxypyrimidin-2-yl)-3-(2-ethoxyphenoxysulfonyl)urea |  | 126801-58-9 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
016-083-00-1 | acibenzolar-S-methyl; benzo[1,2,3]thiadiazole-7-carbothioic acid S-methyl ester | 420-050-0 | 135158-54-2 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H335 H315 H319 H317 H400 H410 | GHS07 GHS09 Wng | H319 H335 H315 H317 H410 |  |  |  | CLP00
016-084-00-7 | prosulfuron (ISO); 1-(4-methoxy-6-methyl-1,3,5-triazin-2-yl)-3-[2-(3,3,3-trifluoropropyl)phenylsulfonyl]urea | - | 94125-34-5 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  | M=100 |  | CLP00/ATP01
016-085-00-2 | flazasulfuron (ISO); 1-(4,6-dimethoxypyrimidin-2-yl)-3-(3-trifluoromethyl-2-pyridylsulfonyl)urea |  | 104040-78-0 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
016-086-00-8 | tetrasodium 10-amino-6,13-dichloro-3-(3-(4-(2,5-disulfonatoanilino)-6-fluoro-1,3,5-triazin-2-ylamino)prop-3-ylamino)-5,12-dioxa-7,14-diazapentacene-4,11-disulfonate | 402-590-9 | 109125-56-6 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
016-087-00-3 | reaction mass of: thiobis(4,1-phenylene)-S,S,S',S'-tetraphenyldisulfonium bishexafluorophosphate; diphenyl(4-phenylthiophenyl)sulfonium hexafluorophosphate; propylene carbonate | 403-490-8 | 104558-95-4 | Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H319 H317 H400 H410 | GHS07 GHS09 Wng | H319 H317 H410 |  |  |  | CLP00
016-088-00-9 | 4-(bis(4-(diethylamino)phenyl)methyl)benzene-1,2-dimethanesulfonic acid | 407-280-7 | 71297-11-5 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
016-089-00-4 | reaction mass of esters of 5,5',6,6',7,7'-hexahydroxy-3,3,3',3'-tetramethyl-1,1'-spirobiindan and 2-diazo-1,2-dihydro-1-oxo-5-sulfonaphthalene | 413-840-1 |  | Self-react. C **** Aquatic Chronic 4 | H242 H413 | GHS02 Dgr | H242 H413 |  |  |  | CLP00
016-090-00-X | 4-methyl-N-(methylsulfonyl)benzenesulfonamide | 415-040-8 | 14653-91-9 | Acute Tox. 4 * STOT SE 3 Eye Dam. 1 | H302 H335 H318 | GHS05 GHS07 Dgr | H302 H335 H318 |  |  |  | CLP00
016-091-00-5 | C12-14-tert-alkyl ammonium 1-amino-9,10-dihydro-9,10-dioxo-4-(2,4,6-trimethylanilino)-anthracen-2-sulfonate | 414-110-5 |  | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | CLP00
016-092-00-0 | reaction mass of: 4,7-bis(mercaptomethyl)-3,6,9-trithia-1,11-undecanedithiol; 4,8-bis(mercaptomethyl)-3,6,9-trithia-1,11-undecanedithiol; 5,7-bis(mercaptomethyl)-3,6,9-trithia-1,11-undecanedithiol | 427-050-1 |  | Repr. 2 Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361f H315 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H315 H317 H361f H410 |  |  |  | ATP01/ATP01corr
016-093-00-6 | reaction mass of: 4-(7-hydroxy-2,4,4-trimethyl-2-chromanyl)resorcinol-4-yl-tris(6-diazo-5,6-dihydro-5-oxonaphthalen-1-sulfonate); 4-(7-hydroxy-2,4,4-trimethyl-2-chromanyl)resorcinolbis(6-diazo-5,6-dihydro-5-oxonaphthalen-1-sulfonate) (2:1) | 414-770-4 | 140698-96-0 | Self-react. C **** Carc. 2 | H242 H351 | GHS02 GHS08 Dgr | H242 H351 |  |  |  | CLP00
016-094-00-1 | sulfur | 231-722-6 | 7704-34-9 | Skin Irrit. 2 | H315 | GHS07 Wng | H315 |  |  |  | ATP01/ATP22
016-095-00-7 | reaction mass of: reaction product of 4,4'-methylenebis[2-(4-hydroxybenzyl)-3,6-dimethylphenol] and 6-diazo-5,6-dihydro-5-oxo-naphthalenesulfonate (1:2); Reaction product of 4,4'-methylenebis[2-(4-hydroxybenzyl)-3,6-dimethylphenol] and 6-diazo-5,6-dihydro-5-oxo-naphthalenesulfonate (1:3) | 417-980-4 |  | Self-react. C **** Carc. 2 | H242 H351 | GHS02 GHS08 Dgr | H242 H351 |  |  |  | CLP00
016-096-00-2 | thifensulfuron-methyl (ISO); methyl 3-(4-methoxy-6-methyl-1,3,5-triazin-2-ylcarbamoylsulfamoyl)thiophene-2-carboxylate |  | 79277-27-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=100 M=100 |  | CLP00/ATP13
016-097-00-8 | 1-amino-2-methyl-2-propanethiol hydrochloride | 434-480-1 | 32047-53-3 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 3 | H302 H314 H317 H412 | GHS05 GHS07 Dgr | H302 H314 H317 H412 |  |  |  | ATP01
016-098-00-3 | dimethyl disulphide | 210-871-0 | 624-92-0 | Flam. Liq. 2 Acute Tox. 3 Acute Tox. 3 STOT SE 3 STOT SE 1 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H225 H331 H301 H336 H370 (upper respiratory tract) (Inhalation) H319 H317 H400 H410 | GHS02 GHS06 GHS08 GHS09 Dgr | H225 H331 H301 H319 H317 H370 (upper respiratory tract) (Inhalation) H336 H410 |  | Inhalation: ATE = 5 mg/L (Vapours) Oral: ATE = 190 mg/kg bw M=1 M=10 |  | ATP15
017-001-00-7 | chlorine | 231-959-5 | 7782-50-5 | Ox. Gas 1 Press. Gas Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 | H270 H331 H335 H315 H319 H400 | GHS03 GHS04 GHS06 GHS09 Dgr | H270 H331 H315 H319 H335 H400 |  | M=100 | U | CLP00/ATP01corr
017-002-00-2 | hydrogen chloride | 231-595-7 | 7647-01-0 | Press. Gas Acute Tox. 3 * Skin Corr. 1A | H331 H314 | GHS04 GHS06 GHS05 Dgr | H331 H314 |  |  | 5 U | CLP00
017-002-01-X | hydrochloric acid ... % | 231-595-7 |  | STOT SE 3 Skin Corr. 1B | H335 H314 | GHS05 GHS07 Dgr | H314 H335 |  | Skin Corr. 1B; H314: C ≥ 25 % Skin Irrit. 2; H315: 10 % ≤ C < 25 % Eye Irrit. 2; H319: 10 % ≤ C < 25 % STOT SE 3; H335: C ≥ 10 % | B | CLP00
017-003-00-8 | barium chlorate | 236-760-7 | 13477-00-4 | Ox. Sol. 1 Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 2 | H271 H332 H302 H411 | GHS03 GHS07 GHS09 Dgr | H271 H332 H302 H411 |  |  |  | CLP00
017-004-00-3 | potassium chlorate | 223-289-7 | 3811-04-9 | Ox. Sol. 1 Acute Tox. 3 | H271 H301 | GHS03 GHS06 Dgr | H271 H301 |  | oral: ATE = 100 mg/kg bw |  | CLP00/ATP21
017-005-00-9 | sodium chlorate | 231-887-4 | 7775-09-9 | Ox. Sol. 1 Acute Tox. 3 | H271 H301 | GHS03 GHS06 Dgr | H271 H301 |  | oral: ATE = 100 mg/kg bw |  | CLP00/ATP21
017-006-00-4 | perchloric acid ... % | 231-512-4 | 7601-90-3 | Ox. Liq. 1 Skin Corr. 1A | H271 H314 | GHS03 GHS05 Dgr | H271 H314 |  | Skin Corr. 1A; H314: C ≥ 50 % Skin Corr. 1B; H314: 10 % ≤ C < 50 % Skin Irrit. 2; H315: 1 % ≤ C < 10 % Eye Irrit. 2; H319: 1 % ≤ C < 10 % Ox. Liq. 1; H271: C > 50 % Ox. Liq. 2; H272: C ≤ 50 % | B | CLP00
017-007-00-X | barium perchlorate | 236-710-4 | 13465-95-7 | Ox. Sol. 1 Acute Tox. 4 * Acute Tox. 4 * | H271 H332 H302 | GHS03 GHS07 Dgr | H271 H332 H302 |  |  |  | CLP00
017-008-00-5 | potassium perchlorate | 231-912-9 | 7778-74-7 | Ox. Sol. 1 Acute Tox. 4 * | H271 H302 | GHS03 GHS07 Dgr | H271 H302 |  |  |  | CLP00
017-009-00-0 | ammonium perchlorate; [containing ≥ 80 % of 0-30 µm particles] | 232-235-1 | 7790-98-9 | Expl. 1.1 Ox. Sol. 1 | H201 H271 | GHS01 Dgr | H201 H271 |  |  | T | CLP00/ATP01
017-010-00-6 | sodium perchlorate | 231-511-9 | 7601-89-0 | Ox. Sol. 1 Acute Tox. 4 * | H271 H302 | GHS03 GHS07 Dgr | H271 H302 |  |  |  | CLP00
017-011-00-1 | sodium hypochlorite, solution ... % Cl active | 231-668-3 | 7681-52-9 | Skin Corr. 1B Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H314 H318 H400 H410 | GHS05 GHS09 Dgr | H314 H410 | EUH031 | EUH031: C ≥ 5 % M=10 M=1 | B | CLP00/ATP13
017-012-00-7 | calcium hypochlorite | 231-908-7 | 7778-54-3 | Ox. Sol. 2 Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 | H272 H302 H314 H400 | GHS03 GHS05 GHS07 GHS09 Dgr | H272 H302 H314 H400 | EUH031 | Skin Corr. 1B; H314: C ≥ 5 % Skin Irrit. 2; H315: 1 % ≤ C < 5 % Eye Dam. 1; H318: 3 % ≤ C < 5 % Eye Irrit. 2; H319: 0,5 % < C < 3 % M=10 | T | CLP00/ATP01corr
017-013-00-2 | calcium chloride | 233-140-8 | 10043-52-4 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
017-014-00-8 | ammonium chloride | 235-186-4 | 12125-02-9 | Acute Tox. 4 * Eye Irrit. 2 | H302 H319 | GHS07 Wng | H302 H319 |  |  |  | CLP00
017-015-00-3 | (2-(aminomethyl)phenyl)acetylchloride hydrochloride | 417-410-4 | 61807-67-8 | Acute Tox. 4 * Skin Corr. 1A Skin Sens. 1 | H302 H314 H317 | GHS05 GHS07 Dgr | H302 H314 H317 |  |  |  | CLP00
017-016-00-9 | methyltriphenylphosphonium chloride | 418-400-2 | 1031-15-8 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 2 | H312 H302 H315 H318 H411 | GHS05 GHS07 GHS09 Dgr | H312 H302 H315 H318 H411 |  |  |  | CLP00
017-017-00-4 | (Z)-13-docosenyl-N,N-bis(2-hydroxyethyl)-N-methyl-ammonium-chloride | 426-210-6 | 120086-58-0 | Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H314 H400 H410 | GHS05 GHS09 Dgr | H314 H410 |  |  |  | CLP00
017-018-00-X | N,N,N-trimethyl-2,3-bis(stearoyloxy)propylammonium chloride | 405-660-7 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
017-019-00-5 | (R)-1,2,3,4-tetrahydro-6,7-dimethoxy-1-veratrylisoquinoline hydrochloride | 415-110-8 | 54417-53-7 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
017-020-00-0 | ethyl propoxy aluminium chloride | 421-790-7 | 13014-29-4 | Water-react. 1 Skin Corr. 1A | H260 H314 | GHS02 GHS05 Dgr | H260 H314 | EUH014 |  |  | CLP00
017-021-00-6 | behenamidopropyl-dimethyl-(dihydroxypropyl) ammonium chloride | 423-420-1 | 136920-10-0 | Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H318 H317 H410 |  |  |  | CLP00
017-023-00-7 | [phosphinyldynetris(oxy)] tris[3-aminopropyl-2-hydroxy-N,N-dimethyl-N-(C6-18)-alkyl] trichlorides | 425-520-9 | 197179-61-6 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | ATP01
017-026-00-3 | chlorine dioxide | 233-162-8 | 10049-04-4 | Ox. Gas 1 Press. Gas Acute Tox. 2 * Skin Corr. 1B Aquatic Acute 1 | H270 H330 H314 H400 | GHS04 GHS03 GHS06 GHS05 GHS09 Dgr | H270 H330 H314 H400 |  | M=10 | 5 | CLP00/ATP01corr
017-026-01-0 | chlorine dioxide … % | 233-162-8 | 10049-04-4 | Acute Tox. 3 * Skin Corr. 1B Aquatic Acute 1 | H301 H314 H400 | GHS06 GHS05 GHS09 Dgr | H301 H314 H400 |  | STOT SE 3; H335: C ≥ 3 % Skin Corr. 1B; H314: C ≥ 5 % Skin Irrit. 2; H315: 1 % ≤ C < 5 % Eye Dam. 1; H318: 3 % ≤ C < 5 % Eye Irrit. 2; H319: 0,3 % ≤ C < 3 % M=10 | B | CLP00/ATP01corr
019-001-00-2 | potassium | 231-119-8 | 7440-09-7 | Water-react. 1 Skin Corr. 1B | H260 H314 | GHS02 GHS05 Dgr | H260 H314 | EUH014 |  |  | CLP00
019-002-00-8 | potassium hydroxide; caustic potash | 215-181-3 | 1310-58-3 | Acute Tox. 4 * Skin Corr. 1A | H302 H314 | GHS05 GHS07 Dgr | H302 H314 |  | Skin Corr. 1A; H314: C ≥ 5 % Skin Corr. 1B; H314: 2 % ≤ C < 5 % Skin Irrit. 2; H315: 0,5 % ≤ C < 2 % Eye Irrit. 2; H319: 0,5 % ≤ C < 2 % |  | CLP00
019-003-00-3 | potassium (E,E)-hexa-2,4-dienoate | 246-376-1 | 24634-61-5 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | ATP07
020-001-00-X | calcium | 231-179-5 | 7440-70-2 | Water-react. 2 | H261 | GHS02 Dgr | H261 |  |  |  | CLP00
020-002-00-5 | calcium cyanide | 209-740-0 | 592-01-8 | Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H400 H410 | GHS06 GHS09 Dgr | H300 H410 | EUH032 |  |  | CLP00
020-003-00-0 | reaction mass of: dicalcium (bis(2-hydroxy-5-tetra-propenylphenylmethyl)methylamine)dihydroxide; tri-calcium (tris(2-hydroxy-5-tetra-propenylphenylmethyl)methylamine)tri-hydroxide; poly[calcium ((2-hydroxy-5-tetra-propenyl-phenylmethyl)methylamine)hydroxide] | 420-470-4 |  | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H315 H319 H317 | GHS07 Wng | H319 H315 H317 |  |  |  | CLP00
022-001-00-5 | titanium tetrachloride | 231-441-9 | 7550-45-0 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 | EUH014 |  |  | CLP00
022-002-00-0 | titanium(4+) oxalate | 403-260-7 |  | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
022-003-00-6 | bis(η5-cyclopentadienyl)-bis(2,6-difluoro-3-[pyrrol-1-yl]-phenyl)titanium | 412-000-1 | 125051-32-3 | Flam. Sol. 1 Repr. 2 STOT RE 2 * Aquatic Chronic 2 | H228 H361f *** H373 ** H411 | GHS02 GHS08 GHS09 Dgr | H228 H361f *** H373 ** H411 |  |  | T | CLP00
022-004-00-1 | potassium titanium oxide (K2Ti6O13) | 432-240-0 | 12056-51-8 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | ATP01/ATP01corr
022-005-00-7 | [N-(1,1-dimethylethyl)-1,1-dimethyl-1-[(1,2,3,4,5-η)-2,3,4,5-tetramethyl-2,4-cyclopentadien-1-yl]silanaminato(2-)-κN][(1,2,3,4-η)-1,3-pentadiene]-titanium | 419-840-8 | 169104-71-6 | Flam. Sol. 1 **** Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 4 | H228 H314 H317 H413 | GHS02 GHS05 GHS07 Dgr | H228 H314 H317 H413 |  |  |  | ATP01
022-006-00-2 | titanium dioxide; [in powder form containing 1 % or more of particles with aerodynamic diameter ≤ 10 μm] | 236-675-5 | 13463-67-7 | Carc. 2 | H351 (Inhalation) | GHS08 Wng | H351 (Inhalation) |  |  |  | ATP14
023-001-00-8 | divanadium pentaoxide; vanadium pentoxide | 215-239-8 | 1314-62-1 | Muta. 2 Carc. 1B Repr. 2 Lact. Acute Tox. 3 Acute Tox. 2 STOT SE 3 STOT RE 1 Aquatic Chronic 2 | H341 H350 H361fd H362 H301 H330 H335 H372 (respiratory tract, inhalation) H411 | GHS06 GHS08 GHS09 Dgr | H341 H350 H361fd H362 H301 H330 H335 H372 (respiratory tract, inhalation) H411 |  | inhalation: ATE = 0,05 mg/L (dusts or mists) oral: ATE = 220 mg/kg bw |  | CLP00/ATP18
024-001-00-0 | chromium (VI) trioxide | 215-607-8 | 1333-82-0 | Ox. Sol. 1 Carc. 1A Muta. 1B Repr. 2 Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 Skin Corr. 1A Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H271 H350 H340 H361f *** H330 H311 H301 H372 ** H314 H334 H317 H400 H410 | GHS03 GHS06 GHS08 GHS05 GHS09 Dgr | H271 H350 H340 H361f *** H330 H311 H301 H372 ** H314 H334 H317 H410 |  | STOT SE 3; H335: C ≥ 1 % |  | CLP00
024-002-00-6 | potassium dichromate | 231-906-6 | 7778-50-9 | Ox. Sol. 2 Carc. 1B Muta. 1B Repr. 1B Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 4 * STOT RE 1 Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H272 H350 H340 H360FD H330 H301 H312 H372 ** H314 H334 H317 H400 H410 | GHS03 GHS06 GHS08 GHS05 GHS09 Dgr | H272 H350 H340 H360FD H330 H301 H372 ** H312 H314 H334 H317 H410 |  | STOT SE 3; H335: C ≥ 5 % | 3 | CLP00
024-003-00-1 | ammonium dichromate | 232-143-1 | 7789-09-5 | Ox. Sol. 2 **** Carc. 1B Muta. 1B Repr. 1B Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 4 * STOT RE 1 Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H272 H350 H340 H360FD H330 H301 H312 H372 ** H314 H334 H317 H400 H410 | GHS03 GHS06 GHS08 GHS05 GHS09 Dgr | H272 H350 H340 H360FD H330 H301 H372 ** H312 H314 H334 H317 H410 |  | STOT SE 3; H335: C ≥ 5 % Resp. Sens.; H334: C ≥ 0,2 % Skin Sens.; H317: C ≥ 0,2 % | 3 G | CLP00
024-004-00-7 | sodium dichromate | 234-190-3 | 10588-01-9 | Ox. Sol. 2 Carc. 1B Muta. 1B Repr. 1B Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 4 * STOT RE 1 Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H272 H350 H340 H360FD H330 H301 H312 H372 ** H314 H334 H317 H400 H410 | GHS03 GHS06 GHS05 GHS08 GHS09 Dgr | H272 H301 H312 H330 H314 H334 H317 H340 H350 H360FD H372 ** H410 |  | STOT SE 3; H335: C ≥ 5 % Resp. Sens. 1; H334: C ≥ 0,2 % Skin Sens. 1; H317: C ≥ 0,2 % | 3 | CLP00/ATP01corr
024-005-00-2 | chromyl dichloride; chromic oxychloride | 239-056-8 | 14977-61-8 | Ox. Liq. 1 Carc. 1B Muta. 1B Skin Corr. 1A Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H271 H350i H340 H314 H317 H400 H410 | GHS03 GHS08 GHS05 GHS07 GHS09 Dgr | H271 H350i H340 H314 H317 H410 |  | Skin Corr. 1A; H314: C ≥ 10 % Skin Corr. 1B; H314: 5 % ≤ C < 10 % Skin Irrit. 2; H315: 0,5 % ≤ C < 5 % Eye Irrit. 2; H319: 0,5 % ≤ C < 5 % STOT SE 3; H335: 0,5 % ≤ C < 5 % Skin Sens. 1; H317: C ≥ 0,5 % | 3 T | CLP00
024-006-00-8 | potassium chromate | 232-140-5 | 7789-00-6 | Carc. 1B Muta. 1B STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H340 H335 H315 H319 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H340 H319 H335 H315 H317 H410 |  | Skin Sens. 1; H317: C ≥ 0,5 % | 3 | CLP00
024-007-00-3 | zinc chromates including zinc potassium chromate |  |  | Carc. 1A Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H302 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H302 H317 H410 |  |  | A | CLP00
024-008-00-9 | calcium chromate | 237-366-8 | 13765-19-0 | Carc. 1B Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H302 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H302 H410 |  |  |  | CLP00
024-009-00-4 | strontium chromate | 232-142-6 | 7789-06-2 | Carc. 1B Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H302 H400 H410 | GHS08 GHS07 GHS09 Dgr | H302 H350 H410 |  |  |  | CLP00
024-010-00-X | dichromium tris(chromate); chromium III chromate; chromic chromate | 246-356-2 | 24613-89-6 | Ox. Sol. 1 Carc. 1B Skin Corr. 1A Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H271 H350 H314 H317 H400 H410 | GHS03 GHS08 GHS05 GHS07 GHS09 Dgr | H271 H350 H314 H317 H410 |  |  | T | CLP00
024-011-00-5 | ammonium bis(1-(3,5-dinitro-2-oxidophenylazo)-3-(N-phenylcarbamoyl)-2-naphtholato)chromate(1-) | 400-110-2 | 109125-51-1 | Self-react. C **** Aquatic Acute 1 Aquatic Chronic 1 | H242 H400 H410 | GHS02 GHS09 Dgr | H242 H410 |  |  |  | CLP00
024-012-00-0 | trisodium bis(7-acetamido-2-(4-nitro-2-oxidophenylazo)-3-sulphonato-1-naphtholato)chromate(1-) | 400-810-8 |  | Muta. 2 | H341 | GHS08 Wng | H341 |  |  |  | CLP00
024-013-00-6 | trisodium (6-anilino-2-(5-nitro-2-oxidophenylazo)-3-sulphonato-1-naphtholato)(4-sulphonato-1,1'-azodi-2,2'naphtholato)chromate(1-) | 402-500-8 |  | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
024-014-00-1 | trisodium bis(2-(5-chloro-4-nitro-2-oxidophenylazo)-5-sulphonato-1-naphtholato)chromate(1-) | 402-870-0 | 93952-24-0 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
024-015-00-7 | disodium (3-methyl-4-(5-nitro-2-oxidophenylazo)-1-phenylpyrazololato)(1-(3-nitro-2-oxido-5-sulfonatophenylazo)-2-naphtholato)chromate(1-) | 404-930-1 |  | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H332 H318 H411 | GHS05 GHS07 GHS09 Dgr | H332 H318 H411 |  |  |  | CLP00
024-016-00-2 | tetradecylammonium bis(1-(5-chloro-2-oxidophenylazo)-2-naphtholato)chromate(1-) | 405-110-6 | 88377-66-6 | STOT RE 2 * Aquatic Chronic 4 | H373 ** H413 | GHS08 Wng | H373 ** H413 |  |  |  | CLP00
024-017-00-8 | Chromium (VI) compounds, with the exception of barium chromate and of compounds specified elsewhere in this Annex |  |  | Carc. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H317 H410 |  |  | A | CLP00
024-018-00-3 | sodium chromate | 231-889-5 | 7775-11-3 | Carc. 1B Muta. 1B Repr. 1B Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 4 * STOT RE 1 Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H340 H360FD H330 H301 H312 H372 ** H314 H334 H317 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H350 H340 H360FD H330 H301 H372 ** H312 H314 H334 H317 H410 |  | Resp. Sens.; H334: C ≥ 0,2 % Skin Sens.; H317: C ≥ 0,2 % | 3 | CLP00
024-019-00-9 | Main component: acetoacetic acid anilide/3-amino-1-hydroxybenzene (ATAN-MAP): trisodium {}{6-[(2 or 3 or 4)-amino-(4 or 5 or 6)-hydroxyphenylazo]-5'-(phenylsulfamoyl)-3-sulfonatonaphthalene-2-azobenzene-1,2'-diolato}}-{}{6''-[1-(phenylcarbamoyl)ethylazo]-5'''-(phenylsulfamoyl)-3''-sulfonatonaphthalene-2''-azobenzene-1'',2'''-diolato}}chromate (III); by-product 1: acetoacetic acid anilide/acetoacetic acid anilide (ATAN-ATAN): trisodium bis{}{6-[1-(phenylcarbamoyl)ethylazo]-5'-(phenylsulfonyl)-3-sulfonatonaphthalene-2-azobenzene-1,2'-diolato}}chromate (III); by-product 2: 3-amino-1-hydroxybenzene/3-amino-1-hydroxybenzene (MAP-MAP): trisodium bis{}{6-[(2 or 3 or 4)-amino-(4 or 5 or 6)-hydroxyphenylazo]-5'-(phenylsulfamoyl)-3-sulfonatonaphthalene-2-azobenzene-1,2'-diolato}} chromate (III) | 419-230-1 |  | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
024-020-00-4 | trisodium bis[(3'-nitro-5'-sulfonato(6-amino-2-[4-(2-hydroxy-1-naphtylazo)phenylsulfonylamino]pyrimidin-5-azo)benzene-2',4-diolato)]chromate (III) | 418-220-4 |  | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
024-021-00-X | potassium tetrasodium bis[(N,N'-n)-1'-(phenylcarbamoyl)-3,5-disulfonatobenzeneazo-1'-prop-1'-ene-2,2'-diolato]chromate(III) | 425-830-4 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
025-001-00-3 | manganese dioxide | 215-202-6 | 1313-13-9 | Acute Tox. 4 * Acute Tox. 4 * | H332 H302 | GHS07 Wng | H332 H302 |  |  |  | CLP00
025-002-00-9 | potassium permanganate | 231-760-3 | 7722-64-7 | Ox. Sol. 2 Repr. 2 Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H272 H361d H302 H400 H410 | GHS03 GHS08 GHS07 GHS09 Dgr | H272 H302 H361d H410 |  |  |  | CLP00/ATP13
025-003-00-4 | manganese sulphate | 232-089-9 | 7785-87-7 | STOT RE 2 * Aquatic Chronic 2 | H373 ** H411 | GHS08 GHS09 Wng | H373 ** H411 |  |  |  | CLP00
025-004-00-X | bis(N,N',N''-trimethyl-1,4,7-triazacyclononane)-trioxo-dimanganese (IV) di(hexafluorophosphate) monohydrate | 411-760-1 | 116633-53-5 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
025-005-00-5 | reaction mass of: tri-sodium [29H, 31H-phthalocyanine-C,C,C-trisulfonato (6-)-N29,N30,N31,N32] manganate (3-); tetrasodium [29H,31H-phthalocyanine-C,C,C,C-tetrasulfonato (6-)-N29,N30,N31,N32], manganate (3-); pentasodium [29H,31H-phthalocyanine-C,C,C,C,C-pentasulfonato (6-)-N29,N30,N31,N32] manganate (3-) | 417-660-4 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
026-001-00-6 | (η-cumene)-(η-cyclopentadienyl)iron(II) hexafluoroantimonate | 407-840-0 | 100011-37-8 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | CLP00
026-002-00-1 | (η-cumene)-(η-cyclopentadienyl)iron(II) trifluoromethane-sulfonate | 407-880-9 | 117549-13-0 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
026-003-00-7 | iron (II) sulfate | 231-753-5 | 7720-78-7 | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 | H302 H315 H319 | GHS07 Wng | H302 H319 H315 |  |  |  | ATP01
026-003-01-4 | iron (II) sulfate (1:1) heptahydrate; sulfuric acid, iron(II) salt (1:1), heptahydrate; ferrous sulfate heptahydrate | 231-753-5 | 7782-63-0 | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 | H302 H315 H319 | GHS07 Wng | H302 H319 H315 |  | Skin Irrit. 2; H315: C ≥ 25 % |  | ATP01
026-004-00-2 | potassium ferrite | 430-010-4 | 12160-44-0 | Skin Corr. 1B Skin Sens. 1 | H314 H317 | GHS05 GHS07 Dgr | H314 H317 |  |  |  | ATP01
027-001-00-9 | cobalt | 231-158-0 | 7440-48-4 | Carc. 1B Muta. 2 Repr. 1B Resp. Sens. 1 Skin Sens. 1 Aquatic Chronic 4 | H350 H341 H360F H334 H317 H413 | GHS08 Dgr | H334 H317 H341 H350 H360F H413 |  |  |  | CLP00/ATP14
027-002-00-4 | cobalt oxide | 215-154-6 | 1307-96-6 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  | M=10 |  | CLP00/ATP01
027-003-00-X | cobalt sulfide | 215-273-3 | 1317-42-6 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  | M=10 |  | CLP00/ATP01
027-004-00-5 | cobalt dichloride | 231-589-4 | 7646-79-9 | Carc. 1B Muta. 2 Repr. 1B Acute Tox. 4 * Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360F *** H302 H334 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H341 H360F *** H302 H334 H317 H410 |  | Carc. 1B; H350i: C ≥ 0,01 % M=10 | 1 | CLP00/ATP01
027-005-00-0 | cobalt sulfate | 233-334-2 | 10124-43-3 | Carc. 1B Muta. 2 Repr. 1B Acute Tox. 4 * Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360F *** H302 H334 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H341 H360F *** H302 H334 H317 H410 |  | Carc. 1B; H350i: C ≥ ,01 % M=10 | 1 | CLP00/ATP01
027-006-00-6 | cobalt di(acetate) | 200-755-8 | 71-48-7 | Carc. 1B Muta. 2 Repr. 1B Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360F *** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H334 H317 H341 H350i H360F *** H410 |  | Carc. 1B; H350i: C ≥ 0,01 % M=10 | 1 | ATP01/ATP01corr
027-007-00-1 | zinc hexacyanocobaltate(III), tertiary butyl alcohol/polypropylene glycol complex | 425-240-7 | - | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | ATP01
027-008-00-7 | complex of cobalt(III)-bis(N-phenyl-4-(5-ethylsulfonyl-2-hydroxyphenylazo)-3-hydroxynaphthylamide), hydrated (n H2O, 2 | 427-390-9 | - | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
027-009-00-2 | cobalt dinitrate | 233-402-1 | 10141-05-6 | Carc. 1B Muta. 2 Repr. 1B Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360F *** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H334 H317 H341 H350i H360F *** H410 |  | Carc. 1B; H350i: C ≥ 0,01 % M=10 | 1 | ATP01/ATP01corr
027-010-00-8 | cobalt carbonate | 208-169-4 | 513-79-1 | Carc. 1B Muta. 2 Repr. 1B Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360F *** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360F *** H334 H317 H410 |  | Carc. 1B; H350i: C ≥ 0,01 % M=10 | 1 | ATP01
028-001-00-1 | tetracarbonylnickel; nickel tetracarbonyl | 236-669-2 | 13463-39-3 | Flam. Liq. 2 Carc. 2 Repr. 1B Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H225 H351 H360D *** H330 H400 H410 | GHS02 GHS06 GHS08 GHS09 Dgr | H225 H351 H360D *** H330 H410 |  |  |  | CLP00
028-002-00-7 | nickel | 231-111-4 | 7440-02-0 | Carc. 2 STOT RE 1 Skin Sens. 1 | H351 H372 ** H317 | GHS08 GHS07 Dgr | H351 H372 ** H317 |  |  | 7 S | CLP00/ATP01
028-002-01-4 | nickel powder; [particle diameter < 1 mm] | 231-111-4 | 7440-02-0 | Carc. 2 STOT RE 1 Skin Sens. 1 Aquatic Chronic 3 | H351 H372 ** H317 H412 | GHS08 GHS07 Dgr | H351 H372 ** H317 H412 |  |  |  | ATP01
028-003-00-2 | nickel monoxide [1] nickel oxide [2] bunsenite [3] | 215-215-7 [1] 234-323-5 [2] | 1313-99-1 [1] 11099-02-8 [2] 34492-97-2 [3] | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Chronic 4 | H350i H372 ** H317 H413 | GHS08 GHS07 Dgr | H350i H372 ** H317 H413 |  |  |  | CLP00/ATP01
028-004-00-8 | nickel dioxide | 234-823-3 | 12035-36-8 | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Chronic 4 | H350i H372 ** H317 H413 | GHS08 GHS07 Dgr | H350i H372 ** H317 H413 |  |  |  | CLP00/ATP01
028-005-00-3 | dinickel trioxide | 215-217-8 | 1314-06-3 | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Chronic 4 | H350i H372 ** H317 H413 | GHS08 GHS07 Dgr | H350i H372 ** H317 H413 |  |  |  | CLP00/ATP01
028-006-00-9 | nickel (II) sulfide [1] nickel sulfide [2] millerite [3] | 240-841-2 [1] 234-349-7 [2] | 16812-54-7 [1] 11113-75-0 [2] 1314-04-1 [3] | Carc. 1A Muta. 2 STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H341 H372 ** H317 H410 |  |  |  | CLP00/ATP01
028-007-00-4 | trinickel disulphide; nickel subsulfide [1] heazlewoodite [2] | 234-829-6 [1] | 12035-72-2 [1] 12035-71-1 [2] | Carc. 1A Muta. 2 Acute Tox. 3 STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H331 H372 ** H317 H400 H410 | GHS08 GHS06 GHS09 Dgr | H331 H317 H341 H350i H372 ** H410 |  | Inhalation: ATE = 0.92 mg/L (dusts/mists) |  | CLP00/ATP17
028-008-00-X | nickel dihydroxide [1] nickel hydroxide [2] | 235-008-5 [1] 234-348-1 [2] | 12054-48-7 [1] 11113-74-9 [2] | Carc. 1A Muta. 2 Repr. 1B Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Skin Irrit. 2 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H332 H302 H372 ** H315 H334 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H360D *** H341 H372 ** H332 H302 H315 H334 H317 H410 |  |  |  | CLP00/ATP01
028-009-00-5 | nickel sulfate | 232-104-9 | 7786-81-4 | Carc. 1A Muta. 2 Repr. 1B Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Skin Irrit. 2 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H332 H302 H372 ** H315 H334 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H302 H332 H315 H334 H317 H341 H350i H360D *** H372 ** H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Irrit. 2; H315: C ≥ 20 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | CLP00/ATP01corr
028-010-00-0 | nickel carbonate; basic nickel carbonate; carbonic acid, nickel (2+) salt [1] carbonic acid, nickel salt [2] [µ-[carbonato(2-)-O:O’]] dihydroxy trinickel [3] [carbonato(2-)] tetrahydroxytrinickel [4] | 222-068-2 [1] 240-408-8 [2] 265-748-4 [3] 235-715-9 [4] | 3333-67-3 [1] 16337-84-1 [2] 65405-96-1 [3] 12607-70-4 [4] | Carc. 1A Muta. 2 Repr. 1B Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Skin Irrit. 2 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H332 H302 H372 ** H315 H334 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H341 H360D *** H372 ** H332 H302 H315 H334 H317 H410 |  |  |  | CLP00/ATP01
028-011-00-6 | nickel dichloride | 231-743-0 | 7718-54-9 | Carc. 1A Muta. 2 Repr. 1B Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 Skin Irrit. 2 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H331 H301 H372 ** H315 H334 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H301 H331 H315 H334 H317 H341 H350i H360D *** H372 ** H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % < C < 1 % Skin Irrit. 2; H315: C ≥ 20 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01/ATP01corr
028-012-00-1 | nickel dinitrate [1] nitric acid, nickel salt [2] | 236-068-5 [1] 238-076-4 [2] | 13138-45-9 [1] 14216-75-2 [2] | Ox. Sol. 2 Carc. 1A Muta. 2 Repr. 1B Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Skin Irrit. 2 Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H272 H350i H341 H360D *** H332 H302 H372 ** H315 H318 H334 H317 H400 H410 | GHS03 GHS05 GHS08 GHS07 GHS09 Dgr | H272 H302 H332 H315 H318 H334 H317 H341 H350i H360D *** H372 ** H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: ,1 % < C < 1 % Skin Irrit. 2; H315: C ≥ 20 % Skin Sens. 1; H317: C > ,01 % M=1 |  | ATP01/ATP01corr
028-013-00-7 | nickel matte | 273-749-6 | 69012-50-6 | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H372 ** H317 H410 |  |  |  | ATP01
028-014-00-2 | slimes and sludges, copper electrolytic refining, decopperised, nickel sulfate | 295-859-3 | 92129-57-2 | Carc. 1A Muta. 2 Repr. 1B Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Skin Irrit. 2 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H332 H302 H372 ** H315 H334 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H341 H360D *** H372 ** H332 H302 H315 H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-015-00-8 | slimes and sludges, copper electrolyte refining, decopperised | 305-433-1 | 94551-87-8 | Carc. 1A Muta. 2 Repr. 1A STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  |  |  | ATP01
028-016-00-3 | nickel diperchlorate; perchloric acid, nickel(II) salt | 237-124-1 | 13637-71-3 | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H314 H334 H317 H400 H410 | GHS05 GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H314 H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-017-00-9 | nickel dipotassium bis(sulfate) [1] diammonium nickel bis(sulfate) [2] | 237-563-9 [1] 239-793-5 [2] | 13842-46-1 [1] 15699-18-0 [2] | Carc. 1A Muta. 2 Repr. 1B Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H332 H302 H372 ** H334 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H341 H360D *** H372 ** H332 H302 H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-018-00-4 | nickel bis(sulfamidate); nickel sulfamate | 237-396-1 | 13770-89-3 | Carc. 1A Muta. 2 Repr. 1B Acute Tox. 4 STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H302 H372 ** H334 H317 H400 H410 | GHS07 GHS08 GHS09 Dgr | H334 H317 H341 H350i H360D *** H372 ** H410 |  | Oral: ATE = 853 mg/kg bw STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: ,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ ,01 % M=1 |  | ATP01/ATP14
028-019-00-X | nickel bis(tetrafluoroborate) | 238-753-4 | 14708-14-6 | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-021-00-0 | nickel diformate [1] formic acid, nickel salt [2] formic acid, copper nickel salt [3] | 222-101-0 [1] 239-946-6 [2] 268-755-0 [3] | 3349-06-2 [1] 15843-02-4 [2] 68134-59-8 [3] | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-022-00-6 | nickel di(acetate) [1] nickel acetate [2] | 206-761-7 [1] 239-086-1 [2] | 373-02-4 [1] 14998-37-9 [2] | Carc. 1A Muta. 2 Repr. 1B Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H332 H302 H372 ** H334 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H302 H332 H334 H317 H341 H350i H360D *** H372 ** H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01/ATP01corr
028-024-00-7 | nickel dibenzoate | 209-046-8 | 553-71-9 | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-025-00-2 | nickel bis(4-cyclohexylbutyrate) | 223-463-2 | 3906-55-6 | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-026-00-8 | nickel(II) stearate; nickel(II) octadecanoate | 218-744-1 | 2223-95-2 | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-027-00-3 | nickel dilactate | - | 16039-61-5 | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-028-00-9 | nickel(II) octanoate | 225-656-7 | 4995-91-9 | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Skin Corr. 1A Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H314 H334 H317 H400 H410 | GHS05 GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H314 H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-029-00-4 | nickel difluoride [1] nickel dibromide [2] nickel diiodide [3] nickel potassium fluoride [4] | 233-071-3 [1] 236-665-0 [2] 236-666-6 [3] | 10028-18-9 [1] 13462-88-9 [2] 13462-90-3 [3] 11132-10-8 [4] | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-030-00-X | nickel hexafluorosilicate | 247-430-7 | 26043-11-8 | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-031-00-5 | nickel selenate | 239-125-2 | 15060-62-5 | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-032-00-0 | nickel hydrogen phosphate [1] nickel bis(dihydrogen phosphate) [2] trinickel bis(orthophosphate) [3] dinickel diphosphate [4] nickel bis(phosphinate) [5] nickel phosphinate [6] phosphoric acid, calcium nickel salt [7] diphosphoric acid, nickel(II) salt [8] | 238-278-2 [1] 242-522-3 [2] 233-844-5 [3] 238-426-6 [4] 238-511-8 [5] 252-840-4 [6] | 14332-34-4 [1] 18718-11-1 [2] 10381-36-9 [3] 14448-18-1 [4] 14507-36-9 [5] 36026-88-7 [6] 17169-61-8 [7] 19372-20-4 [8] | Carc. 1A STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H372 ** H334 H317 H410 |  |  |  | ATP01
028-033-00-6 | diammonium nickel hexacyanoferrate | - | 74195-78-1 | Carc. 1A STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H372 ** H334 H317 H410 |  |  |  | ATP01
028-034-00-1 | nickel dicyanide | 209-160-8 | 557-19-7 | Carc. 1A STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H372 ** H334 H317 H410 | EUH032 |  |  | ATP01
028-035-00-7 | nickel chromate | 238-766-5 | 14721-18-7 | Carc. 1A STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H372 ** H334 H317 H410 |  |  |  | ATP01
028-036-00-2 | nickel(II) silicate [1] dinickel orthosilicate [2] nickel silicate (3:4) [3] silicic acid, nickel salt [4] trihydrogen hydroxybis[orthosilicato(4-)]trinickelate(3-) [5] | 244-578-4 [1] 237-411-1 [2] 250-788-7 [3] 253-461-7 [4] 235-688-3 [5] | 21784-78-1 [1] 13775-54-7 [2] 31748-25-1 [3] 37321-15-6 [4] 12519-85-6 [5] | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H372 ** H317 H410 |  |  |  | ATP01
028-037-00-8 | dinickel hexacyanoferrate | 238-946-3 | 14874-78-3 | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H372 ** H317 H410 |  |  |  | ATP01
028-038-00-3 | trinickel bis(arsenate); nickel(II) arsenate | 236-771-7 | 13477-70-8 | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H372 ** H317 H410 |  |  |  | ATP01
028-039-00-9 | nickel oxalate [1] oxalic acid, nickel salt [2] | 208-933-7 [1] 243-867-2 [2] | 547-67-1 [1] 20543-06-0 [2] | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H372 ** H317 H410 |  |  |  | ATP01
028-040-00-4 | nickel telluride | 235-260-6 | 12142-88-0 | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H372 ** H317 H410 |  |  |  | ATP01
028-041-00-X | trinickel tetrasulfide | - | 12137-12-1 | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H372 ** H317 H410 |  |  |  | ATP01
028-042-00-5 | trinickel bis(arsenite) | - | 74646-29-0 | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H372 ** H317 H410 |  |  |  | ATP01
028-043-00-0 | cobalt nickel gray periclase; C.I. Pigment Black 25; C.I. 77332 [1] cobalt nickel dioxide [2] cobalt nickel oxide [3] | 269-051-6 [1] 261-346-8 [2] | 68186-89-0 [1] 58591-45-0 [2] 12737-30-3 [3] | Carc. 1A STOT RE 1 Skin Sens. 1 | H350i H372 ** H317 | GHS08 GHS07 Dgr | H350i H372 ** H317 |  |  |  | ATP01
028-044-00-6 | nickel tin trioxide; nickel stannate | 234-824-9 | 12035-38-0 | Carc. 1A STOT RE 1 Skin Sens. 1 | H350i H372 ** H317 | GHS08 GHS07 Dgr | H350i H372 ** H317 |  |  |  | ATP01
028-045-00-1 | nickel triuranium decaoxide | 239-876-6 | 15780-33-3 | Carc. 1A STOT RE 1 Skin Sens. 1 | H350i H372 ** H317 | GHS08 GHS07 Dgr | H350i H372 ** H317 |  |  |  | ATP01
028-046-00-7 | nickel dithiocyanate | 237-205-1 | 13689-92-4 | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 | EUH032 | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-047-00-2 | nickel dichromate | 239-646-5 | 15586-38-6 | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-048-00-8 | nickel(II) selenite | 233-263-7 | 10101-96-9 | Carc. 1A STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H372 ** H334 H317 H410 |  |  |  | ATP01
028-049-00-3 | nickel selenide | 215-216-2 | 1314-05-2 | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H372 ** H317 H410 |  |  |  | ATP01
028-050-00-9 | silicic acid, lead nickel salt | - | 68130-19-8 | Carc. 1A Repr. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H360Df H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H360Df H372 ** H317 H410 |  |  |  | ATP01
028-051-00-4 | nickel diarsenide [1] nickel arsenide [2] | 235-103-1 [1] 248-169-1 [2] | 12068-61-0 [1] 27016-75-7 [2] | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H372 ** H317 H410 |  |  |  | ATP01
028-052-00-X | nickel barium titanium primrose priderite; C.I. Pigment Yellow 157; C.I. 77900 | 271-853-6 | 68610-24-2 | Carc. 1A STOT RE 1 Skin Sens. 1 | H350i H372 ** H317 | GHS08 GHS07 Dgr | H317 H350i H372 ** |  |  |  | ATP01/ATP01corr
028-053-00-5 | nickel dichlorate [1] nickel dibromate [2] ethyl hydrogen sulfate, nickel(II) salt [3] | 267-897-0 [1] 238-596-1 [2] 275-897-7 [3] | 67952-43-6 [1] 14550-87-9 [2] 71720-48-4 [3] | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 %1 M=1 |  | ATP01
028-054-00-0 | nickel(II) trifluoroacetate [1] nickel(II) propionate [2] nickel bis(benzenesulfonate) [3] nickel(II) hydrogen citrate [4] citric acid, ammonium nickel salt [5] citric acid, nickel salt [6] nickel bis(2-ethylhexanoate) [7] 2-ethylhexanoic acid, nickel salt [8] dimethylhexanoic acid nickel salt [9] nickel(II) isooctanoate [10] nickel isooctanoate [11] nickel bis(isononanoate) [12] nickel(II) neononanoate [13] nickel(II) isodecanoate [14] nickel(II) neodecanoate [15] neodecanoic acid, nickel salt [16] nickel(II) neoundecanoate [17] bis(.sc.d.sc.-gluconato-O1,O2)nickel [18] nickel 3,5-bis(tert-butyl)-4-hydroxybenzoate (1:2) [19] nickel(II) palmitate [20] (2-ethylhexanoato-O)(isononanoato-O)nickel [21] (isononanoato-O)(isooctanoato-O)nickel [22] (isooctanoato-O)(neodecanoato-O)nickel [23] (2-ethylhexanoato-O)(isodecanoato-O)nickel [24] (2-ethylhexanoato-O)(neodecanoato-O)nickel [25] (isodecanoato-O)(isooctanoato-O)nickel [26] (isodecanoato-O)(isononanoato-O)nickel [27] (isononanoato-O)(neodecanoato-O)nickel [28] fatty acids, C6-19-branched, nickel salts [29] fatty acids, C8-18 and C18-unsaturated, nickel salts [30] 2,7-naphthalenedisulfonic acid, nickel(II) salt [31] | 240-235-8 [1] 222-102-6 [2] 254-642-3 [3] 242-533-3 [4] 242-161-1 [5] 245-119-0 [6] 224-699-9 [7] 231-480-1 [8] 301-323-2 [9] 249-555-2 [10] 248-585-3 [11] 284-349-6 [12] 300-094-6 [13] 287-468-1 [14] 287-469-7 [15] 257-447-1 [16] 300-093-0 [17] 276-205-6 [18] 258-051-1 [19] 237-138-8 [20] 287-470-2 [21] 287-471-8 [22] 284-347-5 [23] 284-351-7 [24] 285-698-7 [25] 285-909-2 [26] 284-348-0 [27] 287-592-6 [28] 294-302-1 [29] 283-972-0 [30] | 16083-14-0 [1] 3349-08-4 [2] 39819-65-3 [3] 18721-51-2 [4] 18283-82-4 [5] 22605-92-1 [6] 4454-16-4 [7] 7580-31-6 [8] 93983-68-7 [9] 29317-63-3 [10] 27637-46-3 [11] 84852-37-9 [12] 93920-10-6 [13] 85508-43-6 [14] 85508-44-7 [15] 51818-56-5 [16] 93920-09-3 [17] 71957-07-8 [18] 52625-25-9 [19] 13654-40-5 [20] 85508-45-8 [21] 85508-46-9 [22] 84852-35-7 [23] 84852-39-1 [24] 85135-77-9 [25] 85166-19-4 [26] 84852-36-8 [27] 85551-28-6 [28] 91697-41-5 [29] 84776-45-4 [30] 72319-19-8 [31] | Carc. 1A Muta. 2 Repr. 1B STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H341 H360D *** H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H341 H360D *** H372 ** H334 H317 H410 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % Skin Sens. 1; H317: C ≥ 0,01 % M=1 |  | ATP01
028-055-00-6 | nickel(II) sulfite [1] nickel tellurium trioxide [2] nickel tellurium tetraoxide [3] molybdenum nickel hydroxide oxide phosphate [4] | 231-827-7 [1] 239-967-0 [2] 239-974-9 [3] 268-585-7 [4] | 7757-95-1 [1] 15851-52-2 [2] 15852-21-8 [3] 68130-36-9 [4] | Carc. 1A STOT RE 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H334 H317 H400 H410 | GHS08 GHS09 Dgr | H350i H372 ** H334 H317 H410 |  |  |  | ATP01
028-056-00-1 | nickel boride (NiB) [1] dinickel boride [2] trinickel boride [3] nickel boride [4] dinickel silicide [5] nickel disilicide [6] dinickel phosphide [7] nickel boron phosphide [8] | 234-493-0 [1] 234-494-6 [2] 234-495-1 [3] 235-723-2 [4] 235-033-1 [5] 235-379-3 [6] 234-828-0 [7] | 12007-00-0 [1] 12007-01-1 [2] 12007-02-2 [3] 12619-90-8 [4] 12059-14-2 [5] 12201-89-7 [6] 12035-64-2 [7] 65229-23-4 [8] | Carc. 1A STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350i H372 ** H317 H410 |  |  |  | ATP01
028-057-00-7 | dialuminium nickel tetraoxide [1] nickel titanium trioxide [2] nickel titanium oxide [3] nickel divanadium hexaoxide [4] cobalt dimolybdenum nickel octaoxide [5] nickel zirkonium trioxide [6] molybdenum nickel tetraoxide [7] nickel tungsten tetraoxide [8] olivine, nickel green [9] lithium nickel dioxide [10] molybdenum nickel oxide [11] | 234-454-8 [1] 234-825-4 [2] 235-752-0 [3] 257-970-5 [4] 268-169-5 [5] 274-755-1 [6] 238-034-5 [7] 238-032-4 [8] 271-112-7 [9] | 12004-35-2 [1] 12035-39-1 [2] 12653-76-8 [3] 52502-12-2 [4] 68016-03-5 [5] 70692-93-2 [6] 14177-55-0 [7] 14177-51-6 [8] 68515-84-4 [9] 12031-65-1 [10] 12673-58-4 [11] | Carc. 1A STOT RE 1 Skin Sens. 1 | H350i H372 ** H317 | GHS08 GHS07 Dgr | H350i H372 ** H317 |  |  |  | ATP01
028-058-00-2 | cobalt lithium nickel oxide | 442-750-5 | - | Carc. 1A Acute Tox. 2 * STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350i H330 H372 ** H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H350i H330 H372 ** H317 H410 |  |  |  | ATP01
029-001-00-4 | copper chloride; copper (I) chloride; cuprous chloride | 231-842-9 | 7758-89-6 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
029-002-00-X | dicopper oxide; copper (I) oxide | 215-270-7 | 1317-39-1 | Acute Tox. 4 Acute Tox. 4 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H318 H400 H410 | GHS07 GHS05 GHS09 Dgr | H332 H302 H318 H410 |  | Inhalation: ATE = 3.34 mg/L (dusts/mists) Oral: ATE = 500 mg/kg bw M=100 M=10 |  | CLP00/ATP17
029-003-00-5 | Naphthenic acids, copper salts; copper naphthenate | 215-657-0 | 1338-02-9 | Flam. Liq. 3 Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H226 H302 H400 H410 | GHS02 GHS07 GHS09 Wng | H226 H302 H410 |  |  |  | CLP00
029-004-00-0 | copper sulphate | 231-847-6 | 7758-98-7 | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H315 H410 |  |  |  | CLP00
029-005-00-6 | (tris(chloromethyl)phthalocyaninato)copper(II), reaction products with N-methylpiperazine and methoxyacetic acid | 401-260-1 |  | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
029-006-00-1 | tris(octadec-9-enylammonium) (trisulfonatophthalocyaninato)copper(II) | 403-210-4 |  | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
029-007-00-7 | (trisodium (2-((3-(6-(2-chloro-5-sulfonato)anilino)-4-(3-carboxypyridinio)-1,3,5-triazin-2-ylamino)-2-oxido-5-sulfonatophenylazo)phenylmethylazo)-4-sulfonatobenzoato)copper(3-)) hydroxide | 404-670-9 | 89797-01-3 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  | G | CLP00
029-008-00-2 | copper(II) methanesulfonate | 405-400-2 | 54253-62-2 | Acute Tox. 4 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H318 H410 |  |  |  | CLP00
029-009-00-8 | phthalocyanine-N-[3-(diethylamino)propyl]sulfonamide copper complex | 413-650-9 | 93971-95-0 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
029-010-00-3 | reaction mass of compounds from (dodecakis(p-tolylthio)phthalocyaninato)copper(II) to (hexadecakis(p-tolylthio)phthalocyaninato)copper(II) | 407-700-9 | 101408-30-4 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
029-011-00-9 | sodium [29H,31H-phthalocyaninato-(2-)-N29,N30,N31,N32]-((3-(N-methyl-N-(2-hydroxyethyl)amino)propyl)amino)sulfonyl-sulfonato, copper complex | 412-730-0 | 150522-10-4 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
029-012-00-4 | sodium ((N-(3-trimethylammoniopropyl)sulfamoyl)methylsulfonatophthalocyaninato)copper(II) | 407-340-2 | 124719-24-0 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
029-013-00-X | trisodium(2-(α-(3-(4-chloro-6-(2-(2-(vinylsulfonyl)ethoxy)ethylamino)-1,3,5-triazin-2-ylamino)-2-oxido-5-sulfonatophenylazo)benzylidenehydrazino)-4-sulfonatobenzoato)copper(II) | 407-580-8 | 130201-51-3 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00/ATP01
029-014-00-5 | reaction mass of: 2,2'-[[cis-1,2-cyclohexanediylbis(nitrilomethylidene)]bis[phenolate]](2-)N,N',O,O'-copper complex; 2,2'-[[trans-1,2-cyclohexanediylbis(nitrilomethylidyne)]bis[phenolate]](2-)N,N',O,O'-copper complex | 419-610-7 | 171866-24-3 | STOT RE 2 * Aquatic Chronic 2 | H373 ** H411 | GHS08 GHS09 Wng | H373 ** H411 |  |  |  | ATP01
029-015-00-0 | copper thiocyanate | 214-183-1 | 1111-67-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 | EUH032 | M=10 M=10 |  | ATP09/ATP17
029-016-00-6 | copper(II) oxide | 215-269-1 | 1317-38-0 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=100 M=10 |  | ATP09/ATP17
029-017-00-1 | dicopper chloride trihydroxide | 215-572-9 | 1332-65-6 | Acute Tox. 3 Acute Tox. 4 Aquatic Acute 1 Aquatic Chronic 1 | H301 H332 H400 H410 | GHS06 GHS09 Dgr | H332 H301 H410 |  | Oral: ATE = 299 mg/kg bw Inhalation: ATE = 2.83 mg/L (dusts/mists) M=10 M=10 |  | ATP09/ATP17
029-018-00-7 | tetracopper hexahydroxide sulphate [1] tetracopper hexahydroxide sulphate hydrate [2] | 215-582-3 [1] 215-582-3 [2] | 1333-22-8 [1] 12527-76-3 [2] | Acute Tox. 4 Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  | Oral: ATE = 500 mg/kg bw M=10 M=10 |  | ATP09/ATP17
029-019-01-X | copper flakes (coated with aliphatic acid) |  |  | Acute Tox. 3 Acute Tox. 4 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H331 H302 H319 H400 H410 | GHS06 GHS09 Dgr | H331 H302 H319 H410 |  | Inhalation: ATE = 0.733 mg/L (dusts/mists) Oral: ATE = 500 mg/kg bw M=10 M=1 |  | ATP09/ATP22
029-020-00-8 | copper(II) carbonate--copper(II) hydroxide (1:1) | 235-113-6 | 12069-69-1 | Acute Tox. 4 Acute Tox. 4 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H319 H400 H410 | GHS07 GHS09 Wng | H332 H302 H319 H410 |  | Inhalation: ATE = 1.2 mg/L (dusts/mists) Oral: ATE = 500 mg/kg bw M=10 M=10 |  | ATP09/ATP17
029-021-00-3 | copper dihydroxide; copper(II) hydroxide | 243-815-9 | 20427-59-2 | Acute Tox. 2 Acute Tox. 4 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H302 H318 H400 H410 | GHS06 GHS05 GHS09 Dgr | H330 H302 H318 H410 |  | Inhalation: ATE = 0.47 mg/L (dusts/mists) Oral: ATE = 500 mg/kg bw M=10 M=10 |  | ATP09/ATP17
029-022-00-9 | bordeaux mixture; reaction products of copper sulphate with calcium dihydroxide |  | 8011-63-0 | Acute Tox. 4 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H332 H318 H400 H410 | GHS07 GHS05 GHS09 Dgr | H332 H318 H410 |  | Inhalation: ATE = 1.97 mg/L (dusts/mists) M=10 M=1 |  | ATP09/ATP17
029-023-00-4 | copper sulphate pentahydrate | 231-847-6 | 7758-99-8 | Acute Tox. 4 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H318 H400 H410 | GHS07 GHS05 GHS09 Dgr | H302 H318 H410 |  | Oral: ATE = 481 mg/kg bw M=10 M=1 |  | ATP09/ATP17
029-025-00-5 | bis(N-hydroxy-N-ni­trosocyclohexylamina­to-O,O’)copper; bis(N-cyclohexyl-diazenium-dioxy)-copper; [Cu-HDO] | 239-703-4 | 312600-89-8 | Flam. Sol. 1 Acute Tox. 4 STOT RE 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H228 H302 H373 (liver) H318 H400 H410 | GHS02 GHS07 GHS08 GHS05 GHS09 Dgr | H228 H302 H318 H373 (liver) H410 |  | Oral: ATE = 360 mg/kg bw M=1 M=1 |  | ATP15
029-026-00-0 | copper [specific surface area > 0,67 mm2/mg] | 231-159-6 | 7440-50-8 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=10 M=1 |  | ATP22
030-001-00-1 | zinc powder - zinc dust (pyrophoric) | 231-175-3 | 7440-66-6 | Pyr. Sol. 1 Water-react. 1 Aquatic Acute 1 Aquatic Chronic 1 | H250 H260 H400 H410 | GHS02 GHS09 Dgr | H260 H250 H410 |  |  | T | CLP00
030-001-01-9 | zinc powder - zinc dust (stabilised) | 231-175-3 | 7440-66-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
030-003-00-2 | zinc chloride | 231-592-0 | 7646-85-7 | Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H314 H410 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
030-004-00-8 | dimethylzinc [1] diethylzinc [2] | 208-884-1 [1] 209-161-3 [2] | 544-97-8 [1] 557-20-0 [2] | Pyr. Liq. 1 Water-react. 1 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H250 H260 H314 H400 H410 | GHS02 GHS05 GHS09 Dgr | H250 H260 H314 H410 | EUH014 |  |  | CLP00
030-005-00-3 | diamminediisocyanatozinc | 401-610-3 |  | Acute Tox. 4 * Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 | H302 H318 H334 H317 H400 | GHS05 GHS08 GHS07 GHS09 Dgr | H302 H318 H334 H317 H400 |  |  |  | CLP00
030-006-00-9 | zinc sulphate (hydrous) (mono-, hexa- and hepta hydrate) [1] zinc sulphate (anhydrous) [2] | 231-793-3 [1] 231-793-3 [2] | 7446-19-7 [1] 7733-02-0 [2] | Acute Tox. 4 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H318 H410 |  |  |  | CLP00
030-007-00-4 | bis(3,5-di-tert-butylsalicylato-O1,O2)zinc | 403-360-0 | 42405-40-3 | Flam. Sol. 1 Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H228 H302 H400 H410 | GHS02 GHS07 GHS09 Dgr | H228 H302 H410 |  |  | T | CLP00
030-008-00-X | hydroxo(2-(benzenesulfonamido)benzoato)zinc(II) | 403-750-0 | 113036-91-2 | Acute Tox. 4 * Aquatic Chronic 2 | H332 H411 | GHS07 GHS09 Wng | H332 H411 |  |  |  | CLP00
030-009-00-5 | zinc-bis(4-(n-octyloxycarbonylamino)salicylate) dihydrate | 417-130-2 | - | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | ATP01
030-010-00-0 | 2-dodec-1-enylbutanedioic acid, 4-methyl ester zinc salt | 430-740-3 | - | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
030-011-00-6 | trizinc bis(orthophosphate) | 231-944-3 | 7779-90-0 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
030-012-00-1 | aluminium-magnesium-zinc-carbonate-hydroxide | 423-570-6 | 169314-88-9 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01/ATP05
030-013-00-7 | zinc oxide | 215-222-5 | 1314-13-2 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
030-015-00-8 | tetrazinc(2+)bis(hexacyanocobalt(3+))diacetate | 440-060-9 | - | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
031-001-00-4 | gallium arsenide | 215-114-8 | 1303-00-0 | Carc. 1B Repr. 1B STOT RE 1 | H350 H360F H372 (respiratory and haematopoietic systems) | GHS08 Dgr | H350 H360F H372 (respiratory and haematopoietic systems) |  |  |  | ATP05/ATP07
033-001-00-X | arsenic | 231-148-6 | 7440-38-2 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H301 H410 |  |  |  | CLP00
033-002-00-5 | arsenic compounds, with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H301 H410 |  | * | 1 A | CLP00
033-003-00-0 | diarsenic trioxide; arsenic trioxide | 215-481-4 | 1327-53-3 | Carc. 1A Acute Tox. 2 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H300 H314 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H350 H300 H314 H410 |  |  |  | CLP00
033-004-00-6 | diarsenic pentaoxide; arsenic pentoxide; arsenic oxide | 215-116-9 | 1303-28-2 | Carc. 1A Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H331 H301 H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H331 H301 H410 |  |  |  | CLP00
033-005-00-1 | arsenic acid and its salts with the exception of those specified elsewhere in this Annex | - | - | Carc. 1A Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H331 H301 H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H331 H301 H410 |  |  | A | CLP00/ATP01
033-006-00-7 | arsine | 232-066-3 | 7784-42-1 | Flam. Gas 1 Press. Gas Acute Tox. 2 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H220 H330 H373 ** H400 H410 | GHS02 GHS04 GHS06 GHS08 GHS09 Dgr | H220 H330 H373 ** H410 |  |  | U | CLP00
033-007-00-2 | tert-butylarsine | 423-320-6 | 4262-43-5 | Pyr. Liq. 1 Acute Tox. 2 * | H250 H330 | GHS02 GHS06 Dgr | H250 H330 |  |  |  | CLP00
034-001-00-2 | selenium | 231-957-4 | 7782-49-2 | Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 4 | H331 H301 H373 ** H413 | GHS06 GHS08 Dgr | H331 H301 H373 ** H413 |  |  |  | CLP00
034-002-00-8 | selenium compounds with the exception of cadmium sulphoselenide and those specified elsewhere in this Annex | - | - | Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H301 H373 ** H410 |  |  | A | CLP00/ATP01
034-003-00-3 | sodium selenite | 233-267-9 | 10102-18-8 | Acute Tox. 2 * Acute Tox. 3 * Skin Sens. 1 Aquatic Chronic 2 | H300 H331 H317 H411 | GHS06 GHS09 Dgr | H300 H331 H317 H411 | EUH031 |  |  | CLP00
035-001-00-5 | bromine | 231-778-1 | 7726-95-6 | Acute Tox. 2 * Skin Corr. 1A Aquatic Acute 1 | H330 H314 H400 | GHS06 GHS05 GHS09 Dgr | H330 H314 H400 |  |  |  | CLP00
035-002-00-0 | hydrogen bromide | 233-113-0 | 10035-10-6 | Press. Gas STOT SE 3 Skin Corr. 1A | H335 H314 | GHS04 GHS05 GHS07 Dgr | H314 H335 |  |  | U | CLP00
035-002-01-8 | hydrobromic acid ... % |  |  | STOT SE 3 Skin Corr. 1B | H335 H314 | GHS05 GHS07 Dgr | H314 H335 |  | Skin Corr. 1B; H314: C ≥ 40 % Skin Irrit. 2; H315: 10 % ≤ C < 40 % Eye Irrit. 2; H319: 10 % ≤ C < 40 % STOT SE 3; H335: C ≥ 10 % | B | CLP00
035-003-00-6 | potassium bromate | 231-829-8 | 7758-01-2 | Ox. Sol. 1 Carc. 1B Acute Tox. 3 * | H271 H350 H301 | GHS03 GHS06 GHS08 Dgr | H271 H350 H301 |  |  |  | CLP00
035-004-00-1 | 2-hydroxyethylammonium perbromide | 407-440-6 |  | Ox. Sol. 2 **** Acute Tox. 4 * Skin Corr. 1A Skin Sens. 1 Aquatic Acute 1 | H272 H302 H314 H317 H400 | GHS03 GHS05 GHS07 GHS09 Dgr | H272 H302 H314 H317 H400 |  |  |  | CLP00
035-005-00-7 | ammonium bromide | 235-183-8 | 12124-97-9 | Repr. 1B Lact. STOT SE 3 STOT RE 1 Eye Irrit. 2 | H360FD H362 H336 H372 (nervous system) H319 | GHS08 GHS07 Dgr | H360FD H362 H336 H372 (nervous system) H319 |  |  |  | ATP18
040-001-00-3 | zirconium powder (pyrophoric) | 231-176-9 | 7440-67-7 | Pyr. Sol. 1 Water-react. 1 | H250 H260 | GHS02 Dgr | H260 H250 |  |  | T | CLP00
040-002-00-9 | zirconium powder (non pyrophoric) |  |  | Self-heat. 1 | H251 | GHS02 Dgr | H251 |  |  | T | CLP00
040-003-00-4 | reaction product of 3,5-di-tert-butylsalicylic acid and zirconium oxychloride, dehydrated, basic Zr : DTBS = 1.0 : 1.0 to 1.0 : 1.5 | 430-610-6 | 226996-19-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
042-001-00-9 | molybdenum trioxide | 215-204-7 | 1313-27-5 | Carc. 2 STOT SE 3 Eye Irrit. 2 | H351 H335 H319 | GHS08 GHS07 Wng | H351 H319 H335 |  |  |  | CLP00/ATP01
042-002-00-4 | tetrakis(dimethylditetradecylammonium) hexa-μ-oxotetra-μ3-oxodi-μ5-oxotetradecaoxooctamolybdate(4-) | 404-760-8 | 117342-25-3 | Acute Tox. 3 * Eye Dam. 1 | H331 H318 | GHS06 GHS05 Dgr | H331 H318 |  |  |  | CLP00/ATP01
042-003-00-X | tetrakis(trimethylhexadecylammonium) hexa-mu-oxotetra-mu3-oxodi-mu5-oxotetradecaoxooctamolybdate(4-) | 404-860-1 | 116810-46-9 | Flam. Sol. 1 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H228 H318 H400 H410 | GHS02 GHS05 GHS09 Dgr | H228 H318 H410 |  |  | T | CLP00
042-004-00-5 | Reaction product of ammonium molybdate and C12-C24-diethoxylated alkylamine (1:5-1:3) | 412-780-3 |  | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H315 H317 H411 | GHS07 GHS09 Wng | H315 H317 H411 |  |  |  | CLP00
042-005-00-0 | reaction mass of: mono- and di-glycerols of canola oil; canola oil acid amide of branched 1,3-propanediamine,N-[3-(tridecyloxy)-propyl]; N,N-diorgano dithiocarbamate molybdenum complex | 434-240-6 | - | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
046-001-00-X | tetraammine palladium (II) hydrogen carbonate | 425-270-0 | 134620-00-1 | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H318 H317 H400 H410 | GHS05 GHS08 GHS07 GHS09 Dgr | H302 H373 ** H318 H317 H410 |  |  |  | ATP01
047-001-00-2 | silver nitrate | 231-853-9 | 7761-88-8 | Ox. Sol. 2 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H272 H314 H400 H410 | GHS03 GHS05 GHS09 Dgr | H272 H314 H410 |  |  |  | CLP00/ATP01
047-002-00-8 | polyphosphoric acid, copper, sodium, magnesium, calcium, silver and zinc salt | 416-850-4 | - | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
047-003-00-3 | silver zinc zeolite (Zeolite, LTA framework type, surface-modified with silver and zinc ions) [This entry covers LTA (Linde Type A) framework type zeolite which has been surface-modified with both silver and zinc ions at contents Ag+ 0,5 %-6 %, Zn2 + 5 %-16 %, and potentially with phosphorus, NH4+, Mg2+ and/or Ca2+ each at level < 3 %] |  | 130328-20-0 | Repr. 2 Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d H315 H318 H400 H410 | GHS08 GHS05 GHS09 Dgr | H315 H318 H361d H410 |  | M=100 M=100 |  | ATP10
047-004-00-9 | silver massive: [particle diameter ≥ 1 mm] | 231-131-3 | 7440-22-4 | Repr. 2 STOT RE 2 | H361f H373 (nervous system) | GHS08 Wng | H361f H373 (nervous system) |  |  |  | ATP22
047-005-00-4 | silver powder: [particle diameter > 100 nm < 1 mm] | 231-131-3 | 7440-22-4 | Repr. 2 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H361f H373 (nervous system) H400 H410 | GHS08 GHS09 Wng | H361f H373 (nervous system) H410 |  | M=10 M=10 |  | ATP22
047-006-00-X | silver nano: [particle diameter > 1 nm ≤ 100 nm] | 231-131-3 | 7440-22-4 | Repr. 2 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H361f H373 (nervous system) H400 H410 | GHS08 GHS09 Wng | H361f H373 (nervous system) H410 |  | M=1000 M=1000 |  | ATP22
048-001-00-5 | cadmium compounds, with the exception of cadmium sulphoselenide (xCdS.yCdSe), reaction mass of cadmium sulphide with zinc sulphide (xCdS.yZnS), reaction mass of cadmium sulphide with mercury sulphide (xCdS.yHgS), and those specified elsewhere in this Annex |  |  | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  | * | 1 A | CLP00
048-002-00-0 | cadmium (non-pyrophoric) [1] cadmium oxide (non-pyrophoric) [2] | 231-152-8 [1] 215-146-2 [2] | 7440-43-9 [1] 1306-19-0 [2] | Carc. 1B Muta. 2 Repr. 2 Acute Tox. 2 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H341 H361fd H330 H372 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H341 H361fd H330 H372 ** H410 |  |  |  | CLP00
048-003-00-6 | cadmium diformate; cadmiumformate | 224-729-0 | 4464-23-7 | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H331 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H301 H351 H373 ** H410 |  | * STOT RE 2; H373: C ≥ 0,25 % |  | CLP00
048-004-00-1 | cadmium cyanide | 208-829-1 | 542-83-6 | Carc. 2 Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H310 H330 H300 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H310 H300 H351 H373 ** H410 | EUH032 | STOT RE 2; H373: C ≥ 0,1 % EUH032: C ≥ 1 % |  | CLP00
048-005-00-7 | cadmiumhexafluorosilicate(2-); cadmium fluorosilica | 241-084-0 | 17010-21-8 | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H331 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H301 H351 H373 ** H410 |  | * STOT RE 2; H373: C ≥ 0,1 % |  | CLP00
048-006-00-2 | cadmium fluoride | 232-222-0 | 7790-79-6 | Carc. 1B Muta. 1B Repr. 1B Acute Tox. 2 * Acute Tox. 3 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H340 H360FD H330 H301 H372 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H340 H360FD H330 H301 H372 ** H410 |  | * Oral Carc. 1B; H350: C ≥ ,01 % STOT RE 1; H372: C ≥ 7 % STOT RE 2; H373: ,1 % ≤ C < 7 % |  | CLP00
048-007-00-8 | cadmium iodide | 232-223-6 | 7790-80-9 | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H331 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H301 H351 H373 ** H410 |  | * STOT RE 2; H373: C ≥ 0,1 % |  | CLP00
048-008-00-3 | cadmium chloride | 233-296-7 | 10108-64-2 | Carc. 1B Muta. 1B Repr. 1B Acute Tox. 2 * Acute Tox. 3 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H340 H360FD H330 H301 H372 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H340 H360FD H330 H301 H372 ** H410 |  | Carc. 1B; H350: C ≥ 0,01 % * oral STOT RE 1; H372: C ≥ 7 % STOT RE 2; H373: 0,1 % ≤ C < 7 % |  | CLP00
048-009-00-9 | cadmium sulphate | 233-331-6 | 10124-36-4 | Carc. 1B Muta. 1B Repr. 1B Acute Tox. 2 * Acute Tox. 3 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H340 H360FD H330 H301 H372 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H340 H360FD H330 H301 H372 ** H410 |  | Carc. 1B; H350: C ≥ 0,01 % * oral STOT RE 1; H372: C ≥ 7 % STOT RE 2; H373: 0,1 % ≤ C < 7 % |  | CLP00
048-010-00-4 | cadmium sulphide | 215-147-8 | 1306-23-6 | Carc. 1B Muta. 2 Repr. 2 Acute Tox. 4 * STOT RE 1 Aquatic Chronic 4 | H350 H341 H361fd H302 H372 ** H413 | GHS08 GHS07 Dgr | H350 H341 H361fd H372 ** H302 H413 |  | * STOT RE 1; H372: C ≥ 10 % STOT RE 2; H373: 0,1 % ≤ C < 10 % | 1 | CLP00
048-011-00-X | cadmium (pyrophoric) | 231-152-8 | 7440-43-9 | Pyr. Sol. 1 Carc. 1B Muta. 2 Repr. 2 Acute Tox. 2 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H250 H350 H341 H361fd H330 H372 ** H400 H410 | GHS02 GHS06 GHS08 GHS09 Dgr | H250 H350 H341 H361fd H330 H372 ** H410 |  |  |  | CLP00
048-012-00-5 | cadmium carbonate | 208-168-9 | 513-78-0 | Carc. 1B Muta. 1B Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H340 H332 H312 H302 H372 (kidney, bone) H400 H410 | GHS08 GHS07 GHS09 Dgr | H332 H312 H302 H340 H350 H372 (kidney, bone) H410 |  |  | 1 A | ATP10
048-013-00-0 | cadmium hydroxide; cadmium dihydroxide | 244-168-5 | 21041-95-2 | Carc. 1B Muta. 1B Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H340 H332 H312 H302 H372 (kidney, bone) H400 H410 | GHS08 GHS07 GHS09 Dgr | H332 H312 H302 H340 H350 H372 (kidney, bone) H410 |  |  | 1 A | ATP10
048-014-00-6 | cadmium nitrate; cadmium dinitrate | 233-710-6 | 10325-94-7 | Carc. 1B Muta. 1B Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H340 H332 H312 H302 H372 (kidney, bone) H400 H410 | GHS08 GHS07 GHS09 Dgr | H332 H312 H302 H340 H350 H372 (kidney, bone) H410 |  | Carc. 1B; H350: C ≥ 0,01 % | 1 A | ATP10
050-001-00-5 | tin tetrachloride; stannic chloride | 231-588-9 | 7646-78-8 | Skin Corr. 1B Aquatic Chronic 3 | H314 H412 | GHS05 Dgr | H314 H412 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
050-002-00-0 | cyhexatin (ISO); hydroxytricyclohexylstannane; tri(cyclohexyl)tin hydroxide | 236-049-1 | 13121-70-5 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  | M=1000 |  | CLP00/ATP01
050-003-00-6 | fentin acetate (ISO); triphenyltin acetate | 212-984-0 | 900-95-8 | Carc. 2 Repr. 2 Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 STOT RE 1 Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H361d *** H330 H311 H301 H335 H372 ** H315 H318 H400 H410 | GHS06 GHS05 GHS08 GHS09 Dgr | H351 H361d *** H330 H311 H301 H372 ** H335 H315 H318 H410 |  | M=10 |  | CLP00/ATP01
050-004-00-1 | fentin hydroxide (ISO); triphenyltin hydroxide | 200-990-6 | 76-87-9 | Carc. 2 Repr. 2 Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 STOT RE 1 Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H361d *** H330 H311 H301 H335 H372 ** H315 H318 H400 H410 | GHS06 GHS05 GHS08 GHS09 Dgr | H351 H361d *** H330 H311 H301 H372 ** H335 H315 H318 H410 |  | M=10 |  | CLP00/ATP01
050-005-00-7 | trimethyltin compounds, with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H400 H410 | GHS06 GHS09 Dgr | H330 H310 H300 H410 |  | * | 1 A | CLP00
050-006-00-2 | triethyltin compounds, with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H400 H410 | GHS06 GHS09 Dgr | H330 H310 H300 H410 |  | * | 1 A | CLP00
050-007-00-8 | tripropyltin compounds, with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H301 H410 |  | * | 1 A | CLP00
050-008-00-3 | tributyltin compounds, with the exception of those specified elsewhere in this annex |  |  | Repr. 1B Acute Tox. 3 Acute Tox. 4 * STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H360FD H301 H312 H372 ** H315 H319 H400 H410 | GHS06 GHS08 GHS09 Dgr | H312 H301 H315 H319 H360FD H372 ** H410 |  | * STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,25 % ≤ C < 1 % Skin Irrit. 2; H315: C ≥ 1 % Eye Irrit. 2; H319: C ≥ 1 % M=10 | 1 A | CLP00/ATP07
050-009-00-9 | fluorotripentylstannane [1] hexapentyldistannoxane [2] | 243-546-7 [1] 247-143-7 [2] | 20153-49-5 [1] 25637-27-8 [2] | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  | * | 1 | CLP00
050-010-00-4 | fluorotrihexylstannane | 243-547-2 | 20153-50-8 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  | * | 1 | CLP00
050-011-00-X | triphenyltin compounds, with the exception of those specified elsewhere in this Annex | - | - | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H301 H410 |  | * M=100 | 1 A | CLP00/ATP01
050-012-00-5 | tetracyclohexylstannane [1] chlorotricyclohexylstannane [2] butyltricyclohexylstannane [3] | 215-910-5 [1] 221-437-5 [2] 230-358-5 [3] | 1449-55-4 [1] 3091-32-5 [2] 7067-44-9 [3] | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  | * | 1 A | CLP00
050-013-00-0 | trioctyltin compounds, with the exception of those specified elsewhere in this Annex |  |  | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 4 | H335 H315 H319 H413 | GHS07 Wng | H319 H335 H315 H413 |  | Skin Irrit. 2; H315: C ≥ 1 % Eye Irrit. 2; H319: C ≥ 1 % STOT SE 3; H335: C ≥ 1 % | 1 A | CLP00
050-017-00-2 | fenbutatin oxide (ISO); bis(tris(2-methyl-2-phenylpropyl)tin)oxide | 236-407-7 | 13356-08-6 | Acute Tox. 2 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H330 H315 H319 H400 H410 | GHS06 GHS09 Dgr | H330 H319 H315 H410 |  |  |  | CLP00
050-018-00-8 | tin(II) methanesulphonate | 401-640-7 | 53408-94-9 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H302 H314 H317 H411 | GHS05 GHS07 GHS09 Dgr | H314 H302 H317 H411 |  |  |  | CLP00/ATP01
050-019-00-3 | azocyclotin (ISO); 1-(tricyclohexylstannyl)-1H-1,2,4-triazole | 255-209-1 | 41083-11-8 | Acute Tox. 2 * Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H301 H335 H315 H318 H400 H410 | GHS06 GHS05 GHS09 Dgr | H330 H301 H335 H315 H318 H410 |  |  |  | CLP00
050-020-00-9 | trioctylstannane | 413-320-4 | 869-59-0 | STOT RE 1 Skin Irrit. 2 Aquatic Chronic 4 | H372 ** H315 H413 | GHS08 GHS07 Dgr | H372 ** H315 H413 |  |  |  | CLP00
050-021-00-4 | dichlorodioctylstannane | 222-583-2 | 3542-36-7 | Repr. 1B Acute Tox. 2 STOT RE 1 Aquatic Chronic 3 | H360D H330 H372 ** H412 | GHS08 GHS06 Dgr | H330 H360D H372 ** H412 |  | Inhalation: ATE = 0.098 mg/L (dusts/mists) Repr. 1B; H360D: C ≥ 0,03 % |  | ATP01/ATP15
050-022-00-X | dibutyltin dichloride; (DBTC) | 211-670-0 | 683-18-1 | Muta. 2 Repr. 1B Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 4 * STOT RE 1 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H341 H360FD H330 H301 H312 H372 ** H314 H400 H410 | GHS06 GHS05 GHS08 GHS09 Dgr | H341 H360FD H330 H301 H312 H372 ** H314 H410 |  | Skin Corr. 1B; H314: C ≥ 5 % Skin Irrit. 2; H315: 0,01 % ≤ C < 5 % Eye Dam. 1; H318: 3 % ≤ C < 5 % Eye Irrit. 2; H319: 0,01 % ≤ C < 3 % M=10 |  | ATP01
050-023-00-5 | reaction mass of: bis[(2-ethyl-1-oxohexyl)oxy]dioctyl stannane; bis[((2-ethyl-1-oxohexyl)oxy)dioctylstannyl]oxide; bis(1-phenyl-1,3-decanedionyl)dioctyl stannane; ((2-ethyl-1-oxohexyl)oxy)-(1-phenyl-1,3-decanedionyl)dioctyl stannane | 422-920-5 | - | STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H400 H410 | GHS08 GHS09 Wng | H373 ** H410 |  | M=10 |  | ATP01
050-024-00-0 | reaction mass of: tri-p-tolyltin hydroxide; hexa-p-tolyl-distannoxane | 432-230-6 | - | Acute Tox. 4 * STOT RE 1 Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H372 ** H315 H318 H317 H400 H410 | GHS05 GHS08 GHS07 GHS09 Dgr | H372 ** H302 H315 H318 H317 H410 |  |  |  | ATP01
050-025-00-6 | trichloromethylstannane | 213-608-8 | 993-16-8 | Repr. 2 | H361d | GHS08 Wng | H361d |  |  |  | ATP05
050-026-00-1 | 2-ethylhexyl 10-ethyl-4-[[2-[(2-ethylhexyl)oxy]-2-oxoethyl]thio]-4-methyl-7-oxo-8-oxa-3,5-dithia-4-stannatetradecanoate | 260-828-5 | 57583-34-3 | Repr. 2 | H361d | GHS08 Wng | H361d |  |  |  | ATP05
050-027-00-7 | 2-ethylhexyl 10-ethyl-4,4-dioctyl-7-oxo-8-oxa-3,5-dithia-4-stannatetradecanoate; [DOTE] | 239-622-4 | 15571-58-1 | Repr. 1B STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D H372 (immune system) H400 H410 | GHS08 GHS09 Dgr | H360D H372 (immune system) H410 |  |  |  | ATP05/ATP15
050-028-00-2 | 2-ethylhexyl 10-ethyl-4,4-dimethyl-7-oxo-8-oxa-3,5-dithia-4-stannatetradecanoate | 260-829-0 | 57583-35-4 | Repr. 2 Acute Tox. 4 STOT RE 1 Skin Sens. 1A | H361d H302 H372 (nervous system, immune system) H317 | GHS08 GHS07 Dgr | H302 H317 H361d H372 (nervous system, immune system) |  |  |  | ATP06
050-029-00-8 | dimethyltin dichloride | 212-039-2 | 753-73-1 | Repr. 2 Acute Tox. 2 Acute Tox. 3 Acute Tox. 3 STOT RE 1 Skin Corr. 1B | H361d H330 H311 H301 H372 (nervous system, immune system) H314 | GHS08 GHS06 GHS05 Dgr | H301 H311 H330 H314 H361d H372 (nervous system, immune system) | EUH071 |  |  | ATP06
050-030-00-3 | dibutyltin dilaurate; dibutyl[bis(dodecanoyloxy)]stannane | 201-039-8 | 77-58-7 | Muta. 2 Repr. 1B STOT RE 1 | H341 H360FD H372 (immune system) | GHS08 Dgr | H341 H360FD H372 (immune system) |  |  |  | ATP10
050-031-00-9 | dioctyltin dilaurate [1] stannane, dioctyl-, bis (coco acyloxy) derivs. [2] | 222-883-3 [1] 293-901-5 [2] | 3648-18-8 [1] 91648-39-4 [2] | Repr. 1B STOT RE 1 | H360D H372 (immune system) | GHS08 Dgr | H360D H372 (immune system) |  |  |  | ATP15
050-032-00-4 | dibutyltin bis (2-ethylhexanoate) | 220-481-2 | 2781-10-4 | Muta. 2 Repr. 1B STOT RE 1 | H341 H360FD H372 (immune system) | GHS08 Dgr | H341 H360FD H372 (immune system) |  |  |  | ATP18
050-033-00-X | dibutyltin di(acetate) | 213-928-8 | 1067-33-0 | Muta. 2 Repr. 1B STOT RE 1 | H341 H360FD H372 (immune system) | GHS08 Dgr | H341 H360FD H372 (immune system) |  |  |  | ATP18
050-034-00-5 | dibutyltin maleate | 201-077-5 | 78-04-6 | Muta. 2 Repr. 1B Acute Tox. 2 Acute Tox. 4 STOT RE 1 Skin Corr. 1 Eye Dam. 1 | H341 H360FD H330 H302 H372 (immune system) H314 H318 | GHS06 GHS08 GHS05 Dgr | H341 H360FD H330 H302 H372 (immune system) H314 |  | inhalation: ATE = 0.317 mg/L(dusts or mists) oral: ATE = 510 mg/kg bw |  | ATP21
050-035-00-0 | dibutyltin oxide | 212-449-1 | 818-08-6 | Muta. 2 Repr. 1B Acute Tox. 3 STOT RE 1 Skin Irrit. 2 Eye Dam. 1 | H341 H360FD H301 H372 (immune system) H315 H318 | GHS06 GHS08 GHS05 Dgr | H341 H360FD H301 H372 (immune system) H315 H318 |  | oral: ATE = 170 mg/kg bw |  | ATP21
051-001-00-8 | antimony trichloride | 233-047-2 | 10025-91-9 | Skin Corr. 1B Aquatic Chronic 2 | H314 H411 | GHS05 GHS09 Dgr | H314 H411 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
051-002-00-3 | antimony pentachloride | 231-601-8 | 7647-18-9 | Skin Corr. 1B Aquatic Chronic 2 | H314 H411 | GHS05 GHS09 Dgr | H314 H411 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
051-003-00-9 | antimony compounds, with the exception of the tetroxide (Sb2O4), pentoxide (Sb2O5), trisulphide (Sb2S3), pentasulphide (Sb2S5) and those specified elsewhere in this Annex |  |  | Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 2 | H332 H302 H411 | GHS07 GHS09 Wng | H332 H302 H411 |  | * | 1 A | CLP00
051-004-00-4 | antimony trifluoride | 232-009-2 | 7783-56-4 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Chronic 2 | H331 H311 H301 H411 | GHS06 GHS09 Dgr | H331 H311 H301 H411 |  |  |  | CLP00
051-005-00-X | antimony trioxide | 215-175-0 | 1309-64-4 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | CLP00
051-006-00-5 | diphenyl(4-phenylthiophenyl)sulfonium hexafluoroantimonate | 403-500-0 |  | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
051-007-00-0 | bis(4-dodecylphenyl)iodonium hexafluoroantimonate | 404-420-9 | 71786-70-4 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
052-001-00-0 | tellurium | 236-813-4 | 13494-80-9 | Repr. 1B Lact. | H360Df H362 | GHS08 DGR | H360Df H362 |  |  |  | ATP18
052-002-00-6 | tellurium dioxide | 231-193-1 | 7446-07-3 | Repr. 1B Lact. | H360Df H362 | GHS08 DGR | H360Df H362 |  |  |  | ATP18
053-001-00-3 | iodine | 231-442-4 | 7553-56-2 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 | H332 H312 H400 | GHS07 GHS09 Wng | H332 H312 H400 |  |  |  | CLP00
053-002-00-9 | hydrogen iodide | 233-109-9 | 10034-85-2 | Press. Gas Skin Corr. 1A | H314 | GHS04 GHS05 Dgr | H314 |  | Skin Corr. 1A; H314: C ≥ 10 % Skin Corr. 1B; H314: 0,2 % ≤ C < 10 % Skin Irrit. 2; H315: 0,02 % ≤ C < 0,2 % Eye Irrit. 2; H319: 0,02 % ≤ C < 0,2 % STOT SE 3; H335: C ≥ 0,02 % | 5 U | CLP00
053-002-01-6 | hydriodic acid ... % |  |  | Skin Corr. 1B | H314 | GHS05 Dgr |  |  | Skin Corr. 1B; H314: C ≥ 25 % Skin Irrit. 2; H315: 10 % ≤ C < 25 % Eye Irrit. 2; H319: 10 % ≤ C < 25 % | B | CLP00
053-003-00-4 | iodoxybenzene |  | 696-33-3 | Expl. **** | **** | **** | **** |  |  |  | CLP00/ATP01corr
053-004-00-X | calcium iodoxybenzoate |  |  | Expl. **** | **** | **** | **** |  |  | C | CLP00/ATP01corr
053-005-00-5 | (4-(1-methylethyl)phenyl)-(4-methylphenyl)iodonium tetrakis(pentafluorophenyl)borate (1-) | 422-960-3 | 178233-72-2 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H373 ** H400 H410 | GHS08 GHS07 GHS09 Wng | H312 H302 H373 ** H410 |  |  |  | CLP00
056-001-00-1 | barium peroxide | 215-128-4 | 1304-29-6 | Ox. Sol. 2 Acute Tox. 4 * Acute Tox. 4 * | H272 H332 H302 | GHS03 GHS07 Dgr | H272 H332 H302 |  |  |  | CLP00
056-002-00-7 | barium salts, with the exception of barium sulphate, salts of 1-azo-2-hydroxynaphthalenyl aryl sulphonic acid, and of salts specified elsewhere in this Annex |  |  | Acute Tox. 4 * Acute Tox. 4 * | H332 H302 | GHS07 Wng | H332 H302 |  | * | 1 A | CLP00
056-003-00-2 | barium carbonate | 208-167-3 | 513-77-9 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
056-004-00-8 | barium chloride | 233-788-1 | 10361-37-2 | Acute Tox. 3 * Acute Tox. 4 * | H301 H332 | GHS06 Dgr | H301 H332 |  |  |  | CLP00
056-005-00-3 | barium diboron tetraoxide | 237-222-4 | 13701-59-2 | Repr. 1B Acute Tox. 4 Acute Tox. 3 | H360FD H332 H301 | GHS08 GHS06 Dgr | H360FD H332 H301 |  | inhalation: ATE = 1,5 mg/L (dusts or mists) oral: ATE = 100 mg/kg bw |  | ATP18
064-001-00-8 | gadolinium(III)sulfite trihydrate | 456-900-2 | 51285-81-5 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
072-001-00-4 | hafnium tetra-n-butoxide | 411-740-2 | 22411-22-9 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
074-001-00-X | hexasodium tungstate hydrate | 412-770-9 | 12141-67-2 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | CLP00
074-002-00-5 | Reaction products of tungsten hexachloride with 2-methylpropan-2-ol, nonylphenol and pentane-2,4-dione | 408-250-6 |  | Flam. Liq. 2 Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H225 H332 H314 H317 H400 H410 | GHS02 GHS05 GHS07 GHS09 Dgr | H225 H332 H314 H317 H410 |  |  |  | CLP00
076-001-00-5 | osmium tetraoxide; osmic acid | 244-058-7 | 20816-12-0 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Skin Corr. 1B | H310 H330 H300 H314 | GHS06 GHS05 Dgr | H330 H310 H300 H314 |  |  |  | CLP00
078-001-00-0 | tetrachloroplatinates with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 3 * Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H301 H318 H334 H317 | GHS06 GHS05 GHS08 Dgr | H301 H318 H334 H317 |  |  | A | CLP00
078-002-00-6 | diammonium tetrachloroplatinate | 237-499-1 | 13820-41-2 | Acute Tox. 3 * Skin Irrit. 2 Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H301 H315 H318 H334 H317 | GHS06 GHS05 GHS08 Dgr | H301 H315 H318 H334 H317 |  |  |  | CLP00
078-003-00-1 | disodium tetrachloroplatinate | 233-051-4 | 10026-00-3 | Acute Tox. 3 * Skin Irrit. 2 Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H301 H315 H318 H334 H317 | GHS06 GHS05 GHS08 Dgr | H301 H315 H318 H334 H317 |  |  |  | CLP00
078-004-00-7 | dipotassium tetrachloroplatinate | 233-050-9 | 10025-99-7 | Acute Tox. 3 * Skin Irrit. 2 Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H301 H315 H318 H334 H317 | GHS06 GHS05 GHS08 Dgr | H301 H315 H318 H334 H317 |  |  |  | CLP00
078-005-00-2 | hexachloroplatinates with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 3 * Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H301 H318 H334 H317 | GHS06 GHS05 GHS08 Dgr | H301 H318 H334 H317 |  |  | A | CLP00
078-006-00-8 | disodium hexachloroplatinate | 240-983-5 | 16923-58-3 | Acute Tox. 3 * Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H301 H318 H334 H317 | GHS06 GHS05 GHS08 Dgr | H301 H318 H334 H317 |  |  |  | CLP00
078-007-00-3 | dipotassium hexachloroplatinate | 240-979-3 | 16921-30-5 | Acute Tox. 3 * Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H301 H318 H334 H317 | GHS06 GHS05 GHS08 Dgr | H301 H318 H334 H317 |  |  |  | CLP00
078-008-00-9 | diammonium hexachloroplatinate | 240-973-0 | 16919-58-7 | Acute Tox. 3 * Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H301 H318 H334 H317 | GHS06 GHS05 GHS08 Dgr | H301 H318 H334 H317 |  |  |  | CLP00
078-009-00-4 | hexachloroplatinic acid | 241-010-7 | 16941-12-1 | Acute Tox. 3 * Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1 | H301 H314 H334 H317 | GHS06 GHS05 GHS08 Dgr | H301 H314 H334 H317 |  |  |  | CLP00
078-010-00-X | tetraammine platinum (II) hydrogen carbonate | 426-730-3 | 123439-82-7 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | ATP01
078-011-00-5 | hydroxydisulfito platinum(II) acid | 423-310-1 | 61420-92-6 | Acute Tox. 4 * STOT RE 2 * Skin Corr. 1A Resp. Sens. 1 Skin Sens. 1 Aquatic Chronic 3 | H302 H373 H314 H334 H317 H412 | GHS05 GHS08 GHS07 Dgr | H302 H373 H314 H334 H317 H412 |  |  |  | ATP01
078-012-00-0 | platinum(IV) nitrate/nitric acid solution | 432-400-1 | - | Skin Corr. 1A Aquatic Acute 1 Aquatic Chronic 1 | H314 H400 H410 | GHS05 GHS09 Dgr | H314 H410 |  |  |  | ATP01
080-001-00-0 | mercury | 231-106-7 | 7439-97-6 | Repr. 1B Acute Tox. 2 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D *** H330 H372 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H360D *** H330 H372 ** H410 |  |  |  | CLP00/ATP01
080-002-00-6 | inorganic compounds of mercury with the exception of mercuric sulphide and those specified elsewhere in this Annex |  |  | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H310 H300 H373 ** H410 |  | * STOT RE 2; H373: C ≥ 0,1 % | 1 A | CLP00
080-003-00-1 | dimercury dichloride; mercurous chloride; calomel | 233-307-5 | 10112-91-1 | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H335 H315 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H335 H315 H410 |  |  |  | CLP00
080-004-00-7 | organic compounds of mercury with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H310 H300 H373 ** H410 |  | * STOT RE 2; H373: C ≥ 0,1 % | 1 A | CLP00
080-005-00-2 | mercury difulminate; mercuric fulminate; fulminate of mercury | 211-057-8 | 628-86-4 | Unst. Expl. Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H200 H331 H311 H301 H373 ** H400 H410 | GHS01 GHS06 GHS08 GHS09 Dgr | H200 H331 H311 H301 H373 ** H410 |  |  |  | CLP00
080-005-01-X | mercury difulminate; mercuric fulminate; fulminate of mercury [≥ 20 % phlegmatiser] | 211-057-8 | 628-86-4 | Expl. 1.1 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H201 H331 H311 H301 H373 ** H400 H410 | GHS01 GHS06 GHS08 GHS09 Dgr | H201 H331 H311 H301 H373 ** H410 |  |  |  | CLP00
080-006-00-8 | dimercury dicyanide oxide; mercuric oxycyanide | 215-629-8 | 1335-31-5 | Expl. 1.1 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H201 H331 H311 H301 H373 ** H400 H410 | GHS01 GHS06 GHS08 GHS09 Dgr | H201 H331 H311 H301 H373 ** H410 |  |  |  | CLP00/ATP01
080-007-00-3 | dimethylmercury [1] diethylmercury [2] | 209-805-3 [1] 211-000-7 [2] | 593-74-8 [1] 627-44-1 [2] | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H310 H300 H373 ** H410 |  | * STOT RE 2; H373: C ≥ 0,05 % | 1 | CLP00
080-008-00-9 | phenylmercury nitrate [1] phenylmercury hydroxide [2] basic phenylmercury nitrate [3] | 200-242-9 [1] 202-866-7 [2] | 55-68-5 [1] 100-57-2 [2] 8003-05-2 [3] | Acute Tox. 3 * STOT RE 1 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H301 H372 ** H314 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H301 H372 ** H314 H410 |  |  |  | CLP00
080-009-00-4 | 2-methoxyethylmercury chloride | 204-659-7 | 123-88-6 | Acute Tox. 3 * STOT RE 1 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H301 H372 ** H314 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H301 H372 ** H314 H410 |  |  |  | CLP00
080-010-00-X | mercury dichloride; mercuric chloride | 231-299-8 | 7487-94-7 | Muta. 2 Repr. 2 Acute Tox. 2 * STOT RE 1 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H341 H361f *** H300 H372 ** H314 H400 H410 | GHS06 GHS05 GHS08 GHS09 Dgr | H341 H361f *** H300 H372 ** H314 H410 |  |  |  | CLP00/ATP01
080-011-00-5 | phenylmercury acetate | 200-532-5 | 62-38-4 | Acute Tox. 3 * STOT RE 1 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H301 H372 ** H314 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H301 H372 ** H314 H410 |  |  |  | CLP00
080-012-00-0 | methylmercuric chloride | 204-064-2 | 115-09-3 | Carc. 2 Repr. 1A Lact. Acute Tox. 2 Acute Tox. 2 Acute Tox. 2 STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H360Df H362 H330 H310 H300 H372 (nervous system, kidneys) H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H310 H300 H351 H360Df H362 H372 (nervous system, kidneys) H410 |  | Inhalation: ATE = 0.05 mg/L (dusts/mists) Dermal: ATE = 50 mg/kg bw Oral: ATE = 5 mg/kg bw | 1 | ATP14
081-001-00-3 | thallium | 231-138-1 | 7440-28-0 | Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Chronic 4 | H330 H300 H373 ** H413 | GHS06 GHS08 Dgr | H330 H300 H373 ** H413 |  |  |  | CLP00
081-002-00-9 | thallium compounds, with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Chronic 2 | H330 H300 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H330 H300 H373 ** H411 |  |  | A | CLP00
081-003-00-4 | dithallium sulphate; thallic sulphate | 231-201-3 | 7446-18-6 | Acute Tox. 2 * STOT RE 1 Skin Irrit. 2 Aquatic Chronic 2 | H300 H372 ** H315 H411 | GHS06 GHS08 GHS09 Dgr | H300 H372 ** H315 H411 |  |  |  | CLP00
082-001-00-6 | lead compounds with the exception of those specified elsewhere in this Annex |  |  | Repr. 1A Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H360Df H332 H302 H373 ** H400 H410 | GHS08 GHS07 GHS09 Dgr | H360Df H332 H302 H373 ** H410 |  | Repr. 2; H361f: C ≥ 2,5 % * STOT RE 2; H373: C ≥ 0,5 % | 1 A | CLP00
082-002-00-1 | lead alkyls |  |  | Repr. 1A Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H360Df H310 H330 H300 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H360Df H330 H310 H300 H373 ** H410 |  | Repr. 1A; H360D: C ≥ 0,1 % * STOT RE 2; H373: C ≥ 0,05 % | 1 A | CLP00
082-003-00-7 | lead diazide; lead azide | 236-542-1 | 13424-46-9 | Unst. Expl. Repr. 1A Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H200 H360Df H332 H302 H373 ** H400 H410 | GHS01 GHS08 GHS07 GHS09 Dgr | H200 H360Df H332 H302 H373 ** H410 |  |  | 1 | CLP00
082-003-01-4 | lead diazide; lead azide [≥ 20 % phlegmatiser] | 236-542-1 | 13424-46-9 | Expl. 1.1 Repr. 1A Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H201 H360Df H332 H302 H373 ** H400 H410 | GHS01 GHS08 GHS07 GHS09 Dgr | H201 H360Df H332 H302 H373 ** H410 |  |  | 1 | CLP00
082-004-00-2 | lead chromate | 231-846-0 | 7758-97-6 | Carc. 1B Repr. 1A STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H350 H360Df H373 ** H400 H410 | GHS08 GHS09 Dgr | H350 H360Df H373 ** H410 |  |  | 1 | CLP00/ATP01
082-005-00-8 | lead di(acetate) | 206-104-4 | 301-04-2 | Repr. 1A STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H360Df H373 ** H400 H410 | GHS08 GHS09 Dgr | H360Df H373 ** H410 |  |  | 1 | CLP00
082-006-00-3 | trilead bis(orthophosphate) | 231-205-5 | 7446-27-7 | Repr. 1A STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H360Df H373 ** H400 H410 | GHS08 GHS09 Dgr | H360Df H373 ** H410 |  |  | 1 | CLP00
082-007-00-9 | lead acetate, basic | 215-630-3 | 1335-32-6 | Carc. 2 Repr. 1A STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H360Df H373 ** H400 H410 | GHS08 GHS09 Dgr | H351 H360Df H373 ** H410 |  |  | 1 | CLP00
082-008-00-4 | lead(II) methanesulphonate | 401-750-5 | 17570-76-2 | Repr. 1A Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 | H360Df H332 H302 H373 ** H315 H318 | GHS08 GHS05 GHS07 Dgr | H360Df H332 H302 H373 ** H315 H318 |  |  | 1 | CLP00
082-009-00-X | lead sulfochromate yellow; C.I. Pigment Yellow 34; [This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77603.] | 215-693-7 | 1344-37-2 | Carc. 1B Repr. 1A STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H350 H360Df H373 ** H400 H410 | GHS08 GHS09 Dgr | H350 H360Df H373 ** H410 |  |  | 1 | CLP00/ATP01
082-010-00-5 | lead chromate molybdate sulfate red; C.I. Pigment Red 104; [This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77605.] | 235-759-9 | 12656-85-8 | Carc. 1B Repr. 1A STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H350 H360Df H373 ** H400 H410 | GHS08 GHS09 Dgr | H350 H360Df H373 ** H410 |  |  | 1 | CLP00/ATP01
082-011-00-0 | lead hydrogen arsenate | 232-064-2 | 7784-40-9 | Carc. 1A Repr. 1A Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H360Df H331 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H360Df H331 H301 H373 ** H410 |  |  | 1 | CLP00
082-012-00-6 | barium calcium cesium lead samarium strontium bromide chloride fluoride iodide europium doped | 431-780-4 | 199876-46-5 | Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 2 | H302 H373 ** H411 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H411 |  |  |  | ATP01
082-013-00-1 | lead powder; [particle diameter < 1 mm] | 231-100-4 | 7439-92-1 | Repr. 1A Lact. Aquatic Acute 1 Aquatic Chronic 1 | H360FD H362 H400 H410 | GHS08 GHS09 Dgr | H360FD H362 H410 |  | Repr. 1A; H360D: C >= 0.03 % M = 10 M = 100 |  | ATP09/ATP21
082-014-00-7 | lead massive: [particle diameter ≥ 1 mm] | 231-100-4 | 7439-92-1 | Repr. 1A Lact. Aquatic Chronic 1 | H360FD H362 H410 | GHS08 GHS09 Dgr | H360FD H362 H410 |  | M = 10 |  | ATP09/ATP21
092-001-00-8 | uranium | 231-170-6 | 7440-61-1 | Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Chronic 4 | H330 H300 H373 ** H413 | GHS06 GHS08 Dgr | H330 H300 H373 ** H413 |  |  |  | CLP00
092-002-00-3 | uranium compounds with the exception of those specified elsewhere in this Annex | - | - | Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 Aquatic Chronic 2 | H330 H300 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H330 H300 H373 ** H411 |  |  | A | CLP00/ATP01
601-001-00-4 | methane | 200-812-7 | 74-82-8 | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | U | CLP00
601-002-00-X | ethane | 200-814-8 | 74-84-0 | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | U | CLP00
601-003-00-5 | propane | 200-827-9 | 74-98-6 | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | U | CLP00
601-004-00-0 | butane [1] and isobutane [2] | 203-448-7 [1] 200-857-2 [2] | 106-97-8 [1] 75-28-5 [2] | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | C U | CLP00
601-004-01-8 | butane (containing ≥ 0,1 % butadiene (203-450-8)) [1] isobutane (containing ≥ 0,1 % butadiene (203-450-8)) [2] | 203-448-7 [1] 200-857-2 [2] | 106-97-8 [1] 75-28-5 [2] | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS02 GHS04 GHS08 Dgr | H220 H350 H340 |  |  | C S U | CLP00
601-005-00-6 | 2,2-dimethylpropane; neopentane | 207-343-7 | 463-82-1 | Flam. Gas 1 Press. Gas Aquatic Chronic 2 | H220 H411 | GHS02 GHS04 GHS09 Dgr | H220 H411 |  |  | U | CLP00
601-006-00-1 | pentane | 203-692-4 | 109-66-0 | Flam. Liq. 2 Asp. Tox. 1 STOT SE 3 Aquatic Chronic 2 | H225 H304 H336 H411 | GHS02 GHS08 GHS07 GHS09 Dgr | H225 H304 H336 H411 | EUH066 |  | C | CLP00
601-007-00-7 | hexane (containing < 5 % n-hexane (203-777-6)); 2-methylpentane [1] 3-methylpentane [2] 2,2-dimethylbutane [3] 2,3-dimethylbutane [4] | 203-523-4 [1] 202-481-4 [2] 200-906-8 [3] 201-193-6 [4] | 107-83-5 [1] 96-14-0 [2] 75-83-2 [3] 79-29-8 [4] | Flam. Liq. 2 Asp. Tox. 1 STOT SE 3 Skin Irrit. 2 Aquatic Chronic 2 | H225 H304 H336 H315 H411 | GHS02 GHS08 GHS07 GHS09 Dgr | H225 H304 H315 H336 H411 |  |  | C | CLP00/ATP01
601-008-00-2 | heptane; n-heptane [1] 2,4-dimethylpentane [2] 2,2,3-trimethylbutane [3] 3,3-dimethylpentane [4] 2,3-dimethylpentane [5] 3-methylhexane [6] 2,2-dimethylpentane [7] 2-methylhexane [8] 3-ethylpentane [9] isoheptane [10] | 205-563-8 [1] 203-548-0 [2] 207-346-3 [3] 209-230-8 [4] 209-280-0 [5] 209-643-3 [6] 209-680-5 [7] 209-730-6 [8] 210-529-0 [9] 250-610-8 [10] | 142-82-5 [1] 108-08-7 [2] 464-06-2 [3] 562-49-2 [4] 565-59-3 [5] 589-34-4 [6] 590-35-2 [7] 591-76-4 [8] 617-78-7 [9] 31394-54-4 [10] | Flam. Liq. 2 Asp. Tox. 1 STOT SE 3 Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H225 H304 H336 H315 H400 H410 | GHS02 GHS08 GHS07 GHS09 Dgr | H225 H304 H315 H336 H410 |  |  | C | CLP00/ATP01
601-009-00-8 | octane; n-octane [1] 2,2,4-trimethylpentane [2] 2,3,3-trimethylpentane [3] 3,3-dimethylhexane [4] 2,2,3-trimethylpentane [5] 2,3,4-trimethylpentane [6] 3,4-dimethylhexane [7] 2,3-dimethylhexane [8] 2,4-dimethylhexane [9] 4-methylheptane [10] 3-methylheptane [11] 2,2-dimethylhexane [12] 2,5-dimethylhexane [13] 2-methylheptane [14] 2,2,3,3-tetramethylbutane [15] 3-ethyl-2-methylpentane [16] 3-ethylhexane [17] 3-ethyl-3-methylpentane [18] isooctane [19] | 203-892-1 [1] 208-759-1 [2] 209-207-2 [3] 209-243-9 [4] 209-266-4 [5] 209-292-6 [6] 209-504-7 [7] 209-547-1 [8] 209-649-6 [9] 209-650-1 [10] 209-660-6 [11] 209-689-4 [12] 209-745-8 [13] 209-747-9 [14] 209-855-6 [15] 210-187-2 [16] 210-621-0 [17] 213-923-0 [18] 247-861-0 [19] | 111-65-9 [1] 540-84-1 [2] 560-21-4 [3] 563-16-6 [4] 564-02-3 [5] 565-75-3 [6] 583-48-2 [7] 584-94-1 [8] 589-43-5 [9] 589-53-7 [10] 589-81-1 [11] 590-73-8 [12] 592-13-2 [13] 592-27-8 [14] 594-82-1 [15] 609-26-7 [16] 619-99-8 [17] 1067-08-9 [18] 26635-64-3 [19] | Flam. Liq. 2 Asp. Tox. 1 STOT SE 3 Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H225 H304 H336 H315 H400 H410 | GHS02 GHS08 GHS07 GHS09 Dgr | H225 H304 H315 H336 H410 |  |  | C | CLP00/ATP01
601-010-00-3 | ethylene | 200-815-3 | 74-85-1 | Flam. Gas 1 Press. Gas STOT SE 3 | H220 H336 | GHS02 GHS04 GHS07 Dgr | H220 H336 |  |  | U | CLP00
601-011-00-9 | propene; propylene | 204-062-1 | 115-07-1 | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | U | CLP00
601-012-00-4 | but-1-ene [1] butene, mixed-1-and-2-isomers [2] 2-methylpropene [3] (Z)-but-2-ene [4] (E)-but-2-ene [5] | 203-449-2 [1] 203-452-9 [2] 204-066-3 [3] 209-673-7 [4] 210-855-3 [5] | 106-98-9 [1] 107-01-7 [2] 115-11-7 [3] 590-18-1 [4] 624-64-6 [5] | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | C U | CLP00
601-013-00-X | 1,3-butadiene; buta-1,3-diene | 203-450-8 | 106-99-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS02 GHS04 GHS08 Dgr | H220 H350 H340 |  |  | D U | CLP00
601-014-00-5 | isoprene (stabilised); 2-methyl-1,3-butadiene | 201-143-3 | 78-79-5 | Flam. Liq. 1 Carc. 1B Muta. 2 Aquatic Chronic 3 | H224 H350 H341 H412 | GHS02 GHS08 Dgr | H224 H350 H341 H412 |  |  | D | CLP00
601-015-00-0 | acetylene; ethyne | 200-816-9 | 74-86-2 | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | U | CLP00
601-016-00-6 | cyclopropane | 200-847-8 | 75-19-4 | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | U | CLP00
601-017-00-1 | cyclohexane | 203-806-2 | 110-82-7 | Flam. Liq. 2 Asp. Tox. 1 STOT SE 3 Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H225 H304 H336 H315 H400 H410 | GHS02 GHS08 GHS07 GHS09 Dgr | H225 H304 H315 H336 H410 |  |  |  | CLP00
601-018-00-7 | methylcyclohexane | 203-624-3 | 108-87-2 | Flam. Liq. 2 Asp. Tox. 1 STOT SE 3 Skin Irrit. 2 Aquatic Chronic 2 | H225 H304 H336 H315 H411 | GHS02 GHS08 GHS07 GHS09 Dgr | H225 H304 H315 H336 H411 |  |  |  | CLP00
601-019-00-2 | 1,4-dimethylcyclohexane | 209-663-2 | 589-90-2 | Flam. Liq. 2 Asp. Tox. 1 STOT SE 3 Skin Irrit. 2 Aquatic Chronic 2 | H225 H304 H336 H315 H411 | GHS02 GHS08 GHS07 GHS09 Dgr | H225 H304 H315 H336 H411 |  |  |  | CLP00
601-020-00-8 | benzene | 200-753-7 | 71-43-2 | Flam. Liq. 2 Carc. 1A Muta. 1B Asp. Tox. 1 STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 | H225 H350 H340 H304 H372 ** H315 H319 | GHS02 GHS08 GHS07 Dgr | H225 H315 H319 H340 H350 H372 ** H304 |  |  |  | CLP00
601-021-00-3 | toluene | 203-625-9 | 108-88-3 | Flam. Liq. 2 Repr. 2 Asp. Tox. 1 STOT SE 3 STOT RE 2 * Skin Irrit. 2 | H225 H361d *** H304 H336 H373 ** H315 | GHS02 GHS08 GHS07 Dgr | H225 H361d *** H304 H373 ** H315 H336 |  |  |  | CLP00
601-022-00-9 | o-xylene [1] p-xylene [2] m-xylene [3] xylene [4] | 202-422-2 [1] 203-396-5 [2] 203-576-3 [3] 215-535-7 [4] | 95-47-6 [1] 106-42-3 [2] 108-38-3 [3] 1330-20-7 [4] | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 | H226 H332 H312 H315 | GHS02 GHS07 Wng | H226 H332 H312 H315 |  | * | C | CLP00
601-023-00-4 | ethylbenzene | 202-849-4 | 100-41-4 | Flam. Liq. 2 Acute Tox. 4 * Asp. Tox. 1 STOT RE 2 | H225 H332 H304 H373 (hearing organs) | GHS02 GHS07 GHS08 Dgr | H225 H332 H373 (hearing organs) H304 |  |  |  | CLP00/ATP06
601-024-00-X | Cumene | 202-704-5 | 98-82-8 | Flam. Liq. 3 Carc. 1B Asp. Tox. 1 STOT SE 3 Aquatic Chronic 2 | H226 H350 H304 H335 H411 | GHS02 GHS08 GHS07 GHS09 Dgr | H226 H350 H304 H335 H411 |  |  |  | CLP00/ATP18
601-025-00-5 | mesitylene; 1,3,5-trimethylbenzene | 203-604-4 | 108-67-8 | Flam. Liq. 3 STOT SE 3 Aquatic Chronic 2 | H226 H335 H411 | GHS02 GHS07 GHS09 Wng | H226 H335 H411 |  | STOT SE 3; H335: C ≥ 25 % |  | CLP00
601-026-00-0 | styrene | 202-851-5 | 100-42-5 | Flam. Liq. 3 Repr. 2 Acute Tox. 4 * STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 | H226 H361d H332 H372 (hearing organs) H315 H319 | GHS02 GHS08 GHS07 Dgr | H226 H332 H315 H319 H361d H372 (hearing organs) |  | * | D | CLP00/ATP06
601-027-00-6 | 2-phenylpropene; α-methylstyrene | 202-705-0 | 98-83-9 | Flam. Liq. 3 STOT SE 3 Eye Irrit. 2 Aquatic Chronic 2 | H226 H335 H319 H411 | GHS02 GHS07 GHS09 Wng | H226 H319 H335 H411 |  | STOT SE 3; H335: C ≥ 25 % |  | CLP00
601-028-00-1 | 2-methylstyrene; 2-vinyltoluene | 210-256-7 | 611-15-4 | Acute Tox. 4 * Aquatic Chronic 2 | H332 H411 | GHS07 GHS09 Wng | H332 H411 |  |  |  | CLP00
601-029-00-7 | dipentene; limonene [1] (S)-p-mentha-1,8-diene; l-limonene [2] trans-1-methyl-4-(1-methylvinyl)cyclohexene [3] (±)-1-methyl-4-(1-methylvinyl)cyclohexene [4] | 205-341-0 [1] 227-815-6 [2] 229-977-3 [3] 231-732-0 [4] | 138-86-3 [1] 5989-54-8 [2] 6876-12-6 [3] 7705-14-8 [4] | Flam. Liq. 3 Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H226 H315 H317 H400 H410 | GHS02 GHS07 GHS09 Wng | H226 H315 H317 H410 |  |  | C | CLP00/ATP17
601-030-00-2 | cyclopentane | 206-016-6 | 287-92-3 | Flam. Liq. 2 Aquatic Chronic 3 | H225 H412 | GHS02 Dgr | H225 H412 |  |  |  | CLP00
601-031-00-8 | 2,4,4-trimethylpent-1-ene | 203-486-4 | 107-39-1 | Flam. Liq. 2 Aquatic Chronic 2 | H225 H411 | GHS02 GHS09 Dgr | H225 H411 |  |  |  | CLP00
601-032-00-3 | benzo[a]pyrene; benzo[def]chrysene | 200-028-5 | 50-32-8 | Carc. 1B Muta. 1B Repr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H340 H360FD H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H340 H360FD H317 H410 |  | Carc. 1B; H350: C ≥ 0,01 % |  | CLP00
601-033-00-9 | benz[a]anthracene | 200-280-6 | 56-55-3 | Carc. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H400 H410 | GHS08 GHS09 Dgr | H350 H410 |  | M=100 |  | CLP00/ATP01
601-034-00-4 | benz[e]acephenanthrylene | 205-911-9 | 205-99-2 | Carc. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H400 H410 | GHS08 GHS09 Dgr | H350 H410 |  |  |  | CLP00
601-035-00-X | benzo[j]fluoranthene | 205-910-3 | 205-82-3 | Carc. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H400 H410 | GHS08 GHS09 Dgr | H350 H410 |  |  |  | CLP00
601-036-00-5 | benzo[k]fluoranthene | 205-916-6 | 207-08-9 | Carc. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H400 H410 | GHS08 GHS09 Dgr | H350 H410 |  |  |  | CLP00
601-037-00-0 | n-hexane | 203-777-6 | 110-54-3 | Flam. Liq. 2 Repr. 2 Asp. Tox. 1 STOT SE 3 STOT RE 1 Skin Irrit. 2 Aquatic Chronic 2 | H225 H361f *** H304 H336 H372 (nervous system) H315 H411 | GHS02 GHS08 GHS07 GHS09 Dgr | H225 H361f *** H304 H336 H372 (nervous system) H315 H411 |  |  |  | CLP00/ATP22
601-041-00-2 | dibenz[a,h]anthracene | 200-181-8 | 53-70-3 | Carc. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H400 H410 | GHS08 GHS09 Dgr | H350 H410 |  | Carc. 1B; H350: C ≥ 0,01 % M=100 |  | CLP00/ATP01
601-042-00-8 | biphenyl; diphenyl | 202-163-5 | 92-52-4 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H335 H315 H319 H400 H410 | GHS07 GHS09 Wng | H319 H335 H315 H410 |  |  |  | CLP00
601-043-00-3 | 1,2,4-trimethylbenzene | 202-436-9 | 95-63-6 | Flam. Liq. 3 Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H226 H332 H335 H315 H319 H411 | GHS02 GHS07 GHS09 Wng | H226 H332 H319 H335 H315 H411 |  |  |  | CLP00
601-044-00-9 | 3a,4,7,7a-tetrahydro-4,7-methanoindene | 201-052-9 | 77-73-6 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H225 H332 H302 H335 H315 H319 H411 | GHS02 GHS07 GHS09 Dgr | H225 H332 H302 H319 H335 H315 H411 |  |  |  | CLP00
601-045-00-4 | 1,2,3,4-tetrahydronaphthalene | 204-340-2 | 119-64-2 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H315 H319 H411 | GHS07 GHS09 Wng | H319 H315 H411 | EUH019 |  |  | CLP00
601-046-00-X | 7-methylocta-1,6-diene | 404-210-7 | 42152-47-6 | Flam. Liq. 3 Aquatic Acute 1 Aquatic Chronic 1 | H226 H400 H410 | GHS02 GHS09 Wng | H226 H410 |  |  |  | CLP00
601-047-00-5 | m-mentha-1,3(8)-diene | 404-150-1 | 17092-80-7 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
601-048-00-0 | chrysene | 205-923-4 | 218-01-9 | Carc. 1B Muta. 2 Aquatic Acute 1 Aquatic Chronic 1 | H350 H341 H400 H410 | GHS08 GHS09 Dgr | H350 H341 H410 |  |  |  | CLP00
601-049-00-6 | benzo[e]pyrene | 205-892-7 | 192-97-2 | Carc. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H400 H410 | GHS08 GHS09 Dgr | H350 H410 |  |  |  | CLP00
601-051-00-7 | 4-phenylbut-1-ene | 405-980-7 | 768-56-9 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
601-052-00-2 | naphthalene | 202-049-5 | 91-20-3 | Carc. 2 Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H400 H410 | GHS07 GHS08 GHS09 Wng | H351 H302 H410 |  |  |  | CLP00
601-053-00-8 | nonylphenol [1] 4-nonylphenol, branched [2] | 246-672-0 [1] 284-325-5 [2] | 25154-52-3 [1] 84852-15-3 [2] | Repr. 2 Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H361fd H302 H314 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H361fd H302 H314 H410 |  |  |  | CLP00
601-054-00-3 | reaction mass of isomers of: dibenzylbenzene; dibenzyl(methyl)benzene; dibenzyl(dimethyl)benzene; dibenzyl(trimethyl)benzene | 405-570-8 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
601-055-00-9 | reaction mass of isomers of: mono-(2-tetradecyl)naphthalenes; di-(2-tetradecyl)naphthalenes; tri-(2-tetradecyl)naphthalenes | 410-190-0 | 132983-41-6 | Eye Irrit. 2 Aquatic Chronic 4 | H319 H413 | GHS07 Wng | H319 H413 |  |  |  | CLP00
601-056-00-4 | reaction mass of isomers of: methyldiphenylmethane; dimethyldiphenylmethane | 405-470-4 | 73807-39-3 | Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H400 H410 | GHS07 GHS09 Wng | H315 H410 |  |  |  | CLP00
601-057-00-X | N-dodecyl-[3-(4-(dimethylamino)benzamido)-propyl]dimethylammonium tosylate | 421-130-8 | 156679-41-3 | Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H318 H317 H410 |  |  |  | CLP00
601-058-00-5 | di-L-para-menthene | 417-870-6 | 83648-84-4 | Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H317 H400 H410 | GHS07 GHS09 Wng | H315 H317 H410 |  |  |  | CLP00
601-059-00-0 | methyl 2-benzylidene-3-oxobutyrate | 420-940-9 | 15768-07-7 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H315 H319 H411 | GHS07 GHS09 Wng | H319 H315 H411 |  |  |  | CLP00
601-060-00-6 | 1,2-bis[4-fluoro-6-{}{4-sulfo-5-(2-(4-sulfonaphtalene-3-ylazo)-1-hydroxy-3,6-disulfo-8-aminonaphthalene-7-ylazo)phenylamino}}-1,3,5-triazin-2ylamino]ethane; x-sodium, y-potassium salts x = 7,755 y = 0,245 | 417-610-1 | 155522-09-1 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
601-061-00-1 | (ethyl-1,2-ethanediyl)[-2-[[[(2-hydroxyethyl)methylamino]acetyl]-propyl]ω-(nonylphenoxy)poly]oxy-(methyl-1,2-ethanediyl) | 418-960-8 |  | Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H314 H317 H411 | GHS05 GHS07 GHS09 Dgr | H314 H317 H411 |  |  |  | CLP00
601-062-00-7 | reaction mass of: branched triacontane; branched dotriacontane; branched tetratriacontane; branched hexatriacontane | 417-030-9 | 151006-59-6 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
601-063-00-2 | reaction mass of isomers of branched tetracosane | 417-060-2 | 151006-61-0 | Acute Tox. 4 * Aquatic Chronic 4 | H332 H413 | GHS07 Wng | H332 H413 |  |  |  | CLP00
601-065-00-3 | reaction mass of: (1'α,3'α,6'α)-2,2,3',7',7'-pentamethylspiro(1,3-dioxane-5,2'-norcarane); (1'α,3'β,6'α)-2,2,3',7',7'-pentamethylspiro(1,3-dioxane-5,2'-norcarane) | 416-930-9 | - | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00/ATP01
601-066-00-9 | 1-(4-(trans-4-heptylcyclohexyl)phenyl)ethanone | 426-820-2 | 78531-60-9 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
601-067-00-4 | triethyl arsenate | 427-700-2 | 15606-95-8 | Carc. 1A Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H331 H301 H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H331 H301 H410 |  |  |  | CLP00
601-068-00-X | 1,2-diacetoxybut-3-ene | 421-720-5 | 18085-02-4 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
601-069-00-5 | 2-ethyl-1-(2-(1,3-dioxanyl)ethyl)-pyridinium bromide | 422-680-1 | 287933-44-2 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
601-070-00-0 | reaction mass of: branched icosane; branched docosane; branched tetracosane | 417-050-8 | 151006-58-5 | Acute Tox. 4 * Aquatic Chronic 4 | H332 H413 | GHS07 Wng | H332 H413 |  |  |  | ATP01
601-071-00-6 | 1-dimethoxymethyl-2-nitro-benzene | 423-830-9 | 20627-73-0 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
601-072-00-1 | reaction mass of: 1-(4-isopropylphenyl)-1-phenylethane; 1-(3-isopropylphenyl)-1-phenylethane; 1-(2-isopropylphenyl)-1-phenylethane | 430-690-2 | 52783-21-8 | Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H400 H410 | GHS07 GHS09 Wng | H315 H410 |  |  |  | ATP01
601-073-00-7 | 1-bromo-3,5-difluorobenzene | 416-710-2 | 461-96-1 | Flam. Liq. 3 Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H226 H302 H373 ** H315 H317 H400 H410 | GHS02 GHS08 GHS07 GHS09 Wng | H226 H302 H373 ** H315 H317 H410 |  |  |  | CLP00
601-074-00-2 | reaction mass of: 4-(2,2,3-trimethylcyclopent-3-en-1-yl)-1-methyl-2-oxabicyclo[2.2.2]octane; 1-(2,2,3-trimethylcyclopent-3-en-1-yl)-5-methyl-6-oxabicyclo[3.2.1]octane; spiro[cyclohex-3-en-1-yl-[(4,5,6,6a-tetrahydro-3,6',6',6'a-tetramethyl)-1,3'(3'aH)-[2H]cyclopenta[b]furan]; spiro[cyclohex-3-en-1-yl-[4,5,6,6a-tetrahydro-4,6',6',6'a-tetramethyl)-1,3'(3'aH)-[2H]cyclopenta[b]]furan] | 422-040-1 |  | Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H315 H319 H411 | GHS07 GHS09 Wng | H319 H315 H411 |  |  |  | CLP00
601-075-00-8 | 4,4'-bis(N-carbamoyl-4-methylbenzenesulfonamide)diphenylmethane | 418-770-5 | 151882-81-4 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | ATP01
601-076-00-3 | ethynyl cyclopropane | 425-430-1 | 6746-94-7 | Flam. Liq. 2 Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H225 H315 H318 H412 | GHS02 GHS05 Dgr | H225 H315 H318 H412 |  |  |  | ATP01
601-077-00-9 | reaction mass of: 1-heptyl-4-ethyl-2,6,7-trioxabicyclo[2.2.2]octane; 1-nonyl-4-ethyl-2,6,7-trioxabicyclo[2.2.2]octane | 426-510-7 | 196965-91-0 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
601-078-00-4 | reaction mass of: 1,7-dimethyl-2-[(3-methylbicyclo[2.2.1]hept-2-yl)methyl]bicyclo[2.2.1]heptane; 2,3-dimethyl-2-[(3-methylbicyclo[2.2.1]hept-2-yl)methyl]bicyclo[2.2.1]heptane | 427-040-5 | - | Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H314 H400 H410 | GHS05 GHS09 Dgr | H314 H410 |  |  |  | ATP01
601-079-00-X | reaction mass of: trans-trans-cyclohexadeca-1,9-diene; cis-trans-cyclohexadeca-1,9-diene | 429-620-3 | - | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 4 | H315 H317 H413 | GHS07 Wng | H315 H317 H413 |  |  |  | ATP01
601-080-00-5 | reaction mass of: sec-butylphenyl(phenyl)methane, mixed isomers; 1-(sec-butylphenyl(phenyl)-2-phenylethane, mixed isomers; 1-(sec-butylphenyl-1-phenylethane, mixed isomers | 431-100-6 | - | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
601-081-00-0 | cyclohexadeca-1,9-diene | 431-730-1 | 4277-06-9 | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 4 | H315 H317 H413 | GHS07 Wng | H315 H317 H413 |  |  |  | ATP01
601-082-00-6 | reaction mass of: endo-2-methyl-exo-3-methyl-exo-2-[(exo-3-methylbicyclo[2.2.1]hept-exo-2-yl)methyl]bicyclo[2.2.1]heptane; exo-2-methyl-exo-3-methyl-endo-2-[(endo-3-methylbicyclo[2.2.1]hept-exo-2-yl)methyl]bicyclo[2.2.1]heptane | 434-420-4 | - | Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H400 H410 | GHS05 GHS09 Dgr | H315 H318 H410 |  |  |  | ATP01
601-083-00-1 | 5-endo-hexyl-bicyclo[2.2.1]hept-2-ene | 435-000-3 | 22094-83-3 | Asp. Tox. 1 Skin Irrit. 2 Aquatic Chronic 4 | H304 H315 H413 | GHS08 GHS07 Dgr | H304 H315 H413 |  |  |  | ATP01
601-084-00-7 | reaction mass of: 5-endo-butyl-bicyclo[2.2.1]hept-2-ene; 5-exo-butyl-bicyclo[2.2.1]hept-2-ene (80:20) | 435-180-3 | - | Asp. Tox. 1 Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H304 H315 H400 H410 | GHS08 GHS07 GHS09 Dgr | H304 H315 H410 |  |  |  | ATP01
601-085-00-2 | isopentane; 2-methylbutane | 201-142-8 | 78-78-4 | Flam. Liq. 1 Asp. Tox. 1 STOT SE 3 Aquatic Chronic 2 | H224 H304 H336 H411 | GHS02 GHS08 GHS07 GHS09 Dgr | H224 H304 H336 H411 | EUH066 |  |  | CLP00
601-087-00-3 | 2,4,4-trimethylpentene | 246-690-9 | 25167-70-8 | Flam. Liq. 2 Asp. Tox. 1 STOT SE 3 | H225 H304 H336 | GHS02 GHS07 GHS08 Dgr | H225 H336 H304 |  |  | D | ATP05
601-088-00-9 | 4-vinylcyclohexene | 202-848-9 | 100-40-3 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | ATP06
601-089-00-4 | muscalure; cis-tricos-9-ene | 248-505-7 | 27519-02-4 | Skin Sens. 1B | H317 | GHS07 Wng | H317 |  |  |  | ATP06
601-090-00-X | benzo[rst]pentaphene | 205-877-5 | 189-55-9 | Carc. 1B Muta. 2 | H350 H341 | GHS08 Dgr | H341 H350 |  |  |  | ATP14
601-091-00-5 | dibenzo[b,def]chrysene; dibenzo[a,h]pyrene | 205-878-0 | 189-64-0 | Carc. 1B Muta. 2 | H350 H341 | GHS08 Dgr | H341 H350 |  |  |  | ATP14
601-092-00-0 | dibenzo[def,p]chrysene; dibenzo[a,l]pyrene | 205-886-4 | 191-30-0 | Carc. 1B Muta. 2 | H350 H341 | GHS08 Dgr | H341 H350 |  | Carc. 1B; H350: C ≥ 0,001 % |  | ATP15
601-093-00-6 | 1,4-dimethylnaphthalene | 209-335-9 | 571-58-4 | Acute Tox. 4 Asp. Tox. 1 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 3 | H302 H304 H319 H400 H412 | GHS07 GHS08 GHS09 Dgr | H302 H319 H304 H410 |  | Oral: ATE = 1300 mg/kg bw M=1 |  | ATP17
601-094-00-1 | 1-isopropyl-4-methylbenzene; p-cymene | 202-796-7 | 99-87-6 | Flam. Liq. 3 Acute Tox. 3 Asp. Tox. 1 Aquatic Chronic 2 | H226 H331 H304 H411 | GHS02 GHS06 GHS08 GHS09 Dgr | H226 H331 H304 H411 |  | Inhalation: ATE = 3 mg/L (Vapours) |  | ATP17
601-095-00-7 | p-mentha-1,3-diene; 1-isopropyl-4-methylcyclohexa-1,3-diene; alpha-terpinene | 202-795-1 | 99-86-5 | Flam. Liq. 3 Acute Tox. 4 Asp. Tox. 1 Skin Sens. 1 Aquatic Chronic 2 | H226 H302 H304 H317 H411 | GHS02 GHS07 GHS08 GHS09 Dgr | H226 H302 H317 H304 H411 |  | Oral: ATE = 1680 mg/kg bw |  | ATP17
601-096-00-2 | (R)-p-mentha-1,8-diene;d-limonene | 227-813-5 | 5989-27-5 | Flam. Liq. 3 Asp. Tox. 1 Skin Irrit. 2 Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 3 | H226 H304 H315 H317 H400 H412 | GHS02 GHS07 GHS08 GHS09 Dgr | H226 H315 H317 H304 H410 |  | M=1 |  | ATP17
601-097-00-8 | Propylbenzene | 203-132-9 | 103-65-1 | Flam. Liq. 3 Asp. Tox. 1 STOT SE 3 Aquatic Chronic 2 | H226 H304 H335 H411 | GHS02 GHS08 GHS07 GHS09 Dgr | H226 H304 H335 H411 |  |  |  | ATP18
602-001-00-7 | chloromethane; methyl chloride | 200-817-4 | 74-87-3 | Flam. Gas 1 Press. Gas Carc. 2 STOT RE 2 * | H220 H351 H373 ** | GHS02 GHS04 GHS08 Dgr | H220 H351 H373 ** |  |  | U | CLP00
602-002-00-2 | bromomethane; methylbromide | 200-813-2 | 74-83-9 | Press. Gas Muta. 2 Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Ozone 1 | H341 H331 H301 H335 H373 H315 H319 H400 H420 | GHS04 GHS06 GHS08 GHS09 Dgr | H331 H301 H315 H319 H341 H335 H373 ** H400 H420 |  |  | U | CLP00/ATP02
602-003-00-8 | dibromomethane | 200-824-2 | 74-95-3 | Acute Tox. 4 * Aquatic Chronic 3 | H332 H412 | GHS07 Wng | H332 H412 |  | * |  | CLP00
602-004-00-3 | dichloromethane; methylene chloride | 200-838-9 | 75-09-2 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | CLP00
602-005-00-9 | methyl iodide; iodomethane | 200-819-5 | 74-88-4 | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 | H351 H331 H301 H312 H335 H315 | GHS06 GHS08 Dgr | H351 H312 H331 H301 H335 H315 |  |  |  | CLP00
602-006-00-4 | chloroform; trichloromethane | 200-663-8 | 67-66-3 | Carc. 2 Repr. 2 Acute Tox. 3 Acute Tox. 4 STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 | H351 H361d H331 H302 H372 H315 H319 | GHS06 GHS08 Dgr | H302 H331 H315 H319 H351 H361d H372 |  |  |  | CLP00/ATP05
602-007-00-X | bromoform; tribromomethane | 200-854-6 | 75-25-2 | Acute Tox. 3 * Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H331 H302 H315 H319 H411 | GHS06 GHS09 Dgr | H331 H302 H319 H315 H411 |  |  |  | CLP00/ATP01
602-008-00-5 | carbon tetrachloride; tetrachloromethane | 200-262-8 | 56-23-5 | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 Aquatic Chronic 3 Ozone 1 | H351 H331 H311 H301 H372 ** H412 H420 | GHS06 GHS08 Dgr | H351 H331 H311 H301 H372 ** H412 H420 |  | * STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: ,2 % ≤ C < 1 % |  | CLP00/ATP02
602-009-00-0 | chloroethane | 200-830-5 | 75-00-3 | Flam. Gas 1 Press. Gas Carc. 2 Aquatic Chronic 3 | H220 H351 H412 | GHS02 GHS04 GHS08 Dgr | H220 H351 H412 |  |  | U | CLP00
602-010-00-6 | 1,2-dibromoethane | 203-444-5 | 106-93-4 | Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H350 H331 H311 H301 H335 H315 H319 H411 | GHS06 GHS08 GHS09 Dgr | H350 H331 H311 H301 H319 H335 H315 H411 |  | * |  | CLP00
602-011-00-1 | 1,1-dichloroethane | 200-863-5 | 75-34-3 | Flam. Liq. 2 Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 Aquatic Chronic 3 | H225 H302 H335 H319 H412 | GHS02 GHS07 Dgr | H225 H302 H319 H335 H412 |  | * |  | CLP00
602-012-00-7 | 1,2-dichloroethane; ethylene dichloride | 203-458-1 | 107-06-2 | Flam. Liq. 2 Carc. 1B Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H225 H350 H302 H335 H315 H319 | GHS02 GHS08 GHS07 Dgr | H225 H350 H302 H319 H335 H315 |  |  |  | CLP00
602-013-00-2 | 1,1,1-trichloroethane; methyl chloroform | 200-756-3 | 71-55-6 | Acute Tox. 4 * Ozone 1 | H332 H420 | GHS07 Wng | H332 H420 |  |  | F | CLP00/ATP02
602-014-00-8 | 1,1,2-trichloroethane | 201-166-9 | 79-00-5 | Carc. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H351 H332 H312 H302 | GHS08 GHS07 Wng | H351 H332 H312 H302 | EUH066 | * |  | CLP00
602-015-00-3 | 1,1,2,2-tetrachloroethane | 201-197-8 | 79-34-5 | Acute Tox. 1 Acute Tox. 2 * Aquatic Chronic 2 | H310 H330 H411 | GHS06 GHS09 Dgr | H330 H310 H411 |  |  |  | CLP00
602-016-00-9 | 1,1,2,2-tetrabromoethane | 201-191-5 | 79-27-6 | Acute Tox. 2 * Eye Irrit. 2 Aquatic Chronic 3 | H330 H319 H412 | GHS06 Dgr | H330 H319 H412 |  |  |  | CLP00
602-017-00-4 | pentachloroethane | 200-925-1 | 76-01-7 | Carc. 2 STOT RE 1 Aquatic Chronic 2 | H351 H372 ** H411 | GHS08 GHS09 Dgr | H351 H372 ** H411 |  | STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,2 % ≤ C < 1 % |  | CLP00
602-018-00-X | 1-chloropropane [1] 2-chloropropane [2] | 208-749-7 [1] 200-858-8 [2] | 540-54-5 [1] 75-29-6 [2] | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H225 H332 H312 H302 | GHS02 GHS07 Dgr | H225 H332 H312 H302 |  |  | C | CLP00
602-019-00-5 | 1-bromopropane; n-propyl bromide | 203-445-0 | 106-94-5 | Flam. Liq. 2 Repr. 1B STOT SE 3 STOT SE 3 STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 | H225 H360FD H335 H336 H373 ** H315 H319 | GHS02 GHS08 GHS07 Dgr | H225 H360FD H373 ** H319 H335 H315 H336 |  |  |  | CLP00
602-020-00-0 | 1,2-dichloropropane; propylene dichloride | 201-152-2 | 78-87-5 | Flam. Liq. 2 Carc. 1B Acute Tox. 4 * Acute Tox. 4 * | H225 H350 H332 H302 | GHS02 GHS07 GHS08 Dgr | H225 H332 H302 H350 |  |  |  | CLP00/ATP09
602-021-00-6 | 1,2-dibromo-3-chloropropane | 202-479-3 | 96-12-8 | Carc. 1B Muta. 1B Repr. 1A Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 3 | H350 H340 H360F *** H301 H373 ** H412 | GHS06 GHS08 Dgr | H350 H340 H360F *** H301 H373 ** H412 |  |  |  | CLP00
602-022-00-1 | 1-chloropentane [1] 2-chloropentane [2] 3-chloropentane [3] | 208-846-4 [1] 210-885-7 [2] 210-467-4 [3] | 543-59-9 [1] 625-29-6 [2] 616-20-6 [3] | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H225 H332 H312 H302 | GHS02 GHS07 Dgr | H225 H332 H312 H302 |  |  | C | CLP00
602-023-00-7 | vinyl chloride; chloroethylene | 200-831-0 | 75-01-4 | Flam. Gas 1 Press. Gas Carc. 1A | H220 H350 | GHS02 GHS08 Dgr | H220 H350 |  |  | D U | CLP00
602-024-00-2 | bromoethylene | 209-800-6 | 593-60-2 | Flam. Gas 1 Press. Gas Carc. 1B | H220 H350 | GHS02 GHS08 Dgr | H220 H350 |  |  | U | CLP00
602-025-00-8 | 1,1-dichloroethylene; vinylidene chloride | 200-864-0 | 75-35-4 | Flam. Liq. 1 Carc. 2 Acute Tox. 4 * | H224 H351 H332 | GHS02 GHS08 GHS07 Dgr | H224 H351 H332 |  | * | D | CLP00
602-026-00-3 | 1,2-dichloroethylene [1] cis-dichloroethylene [2] trans-dichloroethylene [3] | 208-750-2 [1] 205-859-7 [2] 205-860-2 [3] | 540-59-0 [1] 156-59-2 [2] 156-60-5 [3] | Flam. Liq. 2 Acute Tox. 4 * Aquatic Chronic 3 | H225 H332 H412 | GHS02 GHS07 Dgr | H225 H332 H412 |  | * | C | CLP00
602-027-00-9 | trichloroethylene; trichloroethene | 201-167-4 | 79-01-6 | Carc. 1B Muta. 2 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 3 | H350 H341 H336 H315 H319 H412 | GHS08 GHS07 Dgr | H350 H341 H319 H315 H336 H412 |  |  |  | CLP00
602-028-00-4 | tetrachloroethylene | 204-825-9 | 127-18-4 | Carc. 2 Aquatic Chronic 2 | H351 H411 | GHS08 GHS09 Wng | H351 H411 |  |  |  | CLP00
602-029-00-X | 3-chloropropene; allyl chloride | 203-457-6 | 107-05-1 | Flam. Liq. 2 Carc. 2 Muta. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 | H225 H351 H341 H332 H312 H302 H335 H373 ** H315 H319 H400 | GHS02 GHS08 GHS07 GHS09 Dgr | H225 H351 H341 H332 H312 H302 H373 ** H319 H335 H315 H400 |  |  | D | CLP00
602-030-00-5 | 1,3-dichloropropene [1] (Z)-1,3-dichloropropene [2] | 208-826-5 [1] 233-195-8 [2] | 542-75-6 [1] 10061-01-5 [2] | Flam. Liq. 3 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * Asp. Tox. 1 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H226 H311 H301 H332 H304 H335 H315 H319 H317 H400 H410 | GHS02 GHS06 GHS08 GHS09 Dgr | H226 H311 H301 H332 H304 H319 H335 H315 H317 H410 |  |  | C D | CLP00/ATP01
602-031-00-0 | 1,1-dichloropropene | 209-253-3 | 563-58-6 | Flam. Liq. 2 Acute Tox. 3 * Aquatic Chronic 3 | H225 H301 H412 | GHS02 GHS06 Dgr | H225 H301 H412 |  |  |  | CLP00
602-032-00-6 | 3-chloro-2-methylpropene | 209-251-2 | 563-47-3 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H225 H332 H302 H314 H317 H411 | GHS02 GHS05 GHS07 GHS09 Dgr | H225 H332 H302 H314 H317 H411 |  |  |  | CLP00
602-033-00-1 | chlorobenzene | 203-628-5 | 108-90-7 | Flam. Liq. 3 Acute Tox. 4 Skin Irrit. 2 Aquatic Chronic 2 | H226 H332 H315 H411 | GHS02 GHS07 GHS09 Wng | H226 H332 H315 H411 |  |  |  | CLP00/ATP09
602-034-00-7 | 1,2-dichlorobenzene; o-dichlorobenzene | 202-425-9 | 95-50-1 | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H335 H315 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H335 H315 H410 |  | * |  | CLP00
602-035-00-2 | 1,4-dichlorobenzene; p-dichlorobenzene | 203-400-5 | 106-46-7 | Carc. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H319 H400 H410 | GHS08 GHS09 Wng | H351 H319 H410 |  |  |  | CLP00
602-036-00-8 | chloroprene (stabilised); 2-chlorobuta-1,3-diene (stabilised) | 204-818-0 | 126-99-8 | Flam. Liq. 2 Carc. 1B Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 | H225 H350 H332 H302 H335 H373 ** H315 H319 | GHS02 GHS08 GHS07 Dgr | H225 H350 H332 H302 H373 ** H319 H335 H315 |  |  | D | CLP00
602-037-00-3 | α-chlorotoluene; benzyl chloride | 202-853-6 | 100-44-7 | Carc. 1B Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 | H350 H331 H302 H335 H373 ** H315 H318 | GHS06 GHS08 GHS05 Dgr | H350 H331 H302 H373 ** H335 H315 H318 |  |  |  | CLP00
602-038-00-9 | α,α,α-trichlorotoluene; benzotrichloride | 202-634-5 | 98-07-7 | Carc. 1B Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 | H350 H331 H302 H335 H315 H318 | GHS06 GHS08 GHS05 Dgr | H350 H331 H302 H335 H315 H318 |  |  |  | CLP00
602-039-00-4 | polychlorobiphenyls; PCB | 215-648-1 | 1336-36-3 | STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H400 H410 | GHS08 GHS09 Wng | H373 ** H410 |  | STOT RE 2; H373: C ≥ 0,005 % | C | CLP00
602-040-00-X | 2-chlorotoluene [1] 3-chlorotoluene [2] 4-chlorotoluene [3] chlorotoluene [4] | 202-424-3 [1] 203-580-5 [2] 203-397-0 [3] 246-698-2 [4] | 95-49-8 [1] 108-41-8 [2] 106-43-4 [3] 25168-05-2 [4] | Acute Tox. 4 * Aquatic Chronic 2 | H332 H411 | GHS07 GHS09 Wng | H332 H411 |  |  | C | CLP00
602-041-00-5 | penthachloronaphthalene | 215-320-8 | 1321-64-8 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H315 H319 H400 H410 | GHS07 GHS09 Wng | H312 H302 H319 H315 H410 |  |  | C | CLP00
602-042-00-0 | 1,2,3,4,5,6-hexachlorcyclohexanes with the exception of those specified elsewhere in this Annex |  |  | Carc. 2 Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H301 H312 H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H301 H312 H410 |  |  | A C | CLP00
602-043-00-6 | lindane (ISO); γ-HCH or γ-BHC; γ-1,2,3,4,5,6-hexachlorocyclohexane | 200-401-2 | 58-89-9 | Lact. Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H362 H301 H332 H312 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H301 H332 H312 H373 ** H362 H410 |  | M=10 |  | CLP00
602-044-00-1 | camphechlor (ISO); toxaphene | 232-283-3 | 8001-35-2 | Carc. 2 Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H301 H312 H335 H315 H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H301 H312 H335 H315 H410 |  |  |  | CLP00
602-045-00-7 | DDT (ISO); clofenotane (INN); dicophane; 1,1,1-trichloro-2,2-bis(4-chlorophenyl)ethane; dichlorodiphenyltrichloroethane | 200-024-3 | 50-29-3 | Carc. 2 Acute Tox. 3 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H301 H372 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H301 H372 ** H410 |  |  |  | CLP00
602-046-00-2 | heptachlor (ISO); 1,4,5,6,7,8,8-heptachloro-3a,4,7,7a-tetrahydro-4,7-methanoindene | 200-962-3 | 76-44-8 | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H311 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H311 H301 H373 ** H410 |  |  |  | CLP00
602-047-00-8 | chlordane (ISO); 1,2,4,5,6,7,8,8-octachloro-3a,4,7,7a-tetrahydro-4,7-methanoindan | 200-349-0 | 57-74-9 | Carc. 2 Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H312 H302 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H312 H302 H410 |  |  |  | CLP00
602-048-00-3 | aldrin (ISO) | 206-215-8 | 309-00-2 | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H311 H301 H372 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H311 H301 H372 ** H410 |  |  |  | CLP00
602-049-00-9 | dieldrin (ISO) | 200-484-5 | 60-57-1 | Carc. 2 Acute Tox. 1 Acute Tox. 3 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H310 H301 H372 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H310 H301 H372 ** H410 |  |  |  | CLP00
602-050-00-4 | isodrin; (1α,4α,4aβ,5β,8β,8aβ)-1,2,3,4,10,10-hexachloro-1,4,4a,5,8,8a-hexahydro-1,4:5,8-dimethanonaphthalene | 207-366-2 | 465-73-6 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H400 H410 | GHS06 GHS09 Dgr | H330 H310 H300 H410 |  | M=100 |  | CLP00/ATP01
602-051-00-X | endrin (ISO); 1,2,3,4,10,10-hexachloro-6,7-epoxy-1,4,4a,5,6,7,8,8a-octahydro-1,4:5,8-dimethanonaphthalene | 200-775-7 | 72-20-8 | Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H300 H311 H400 H410 | GHS06 GHS09 Dgr | H300 H311 H410 |  |  |  | CLP00
602-052-00-5 | endosulfan (ISO); 1,2,3,4,7,7-hexachloro-8,9,10-trinorborn-2-en-5,6-ylenedimethylene sulfite; 1,4,5,6,7,7-hexachloro-8,9,10-trinorborn-5-en-2,3-ylenedimethylene sulfite | 204-079-4 | 115-29-7 | Acute Tox. 2 * Acute Tox. 2 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H330 H300 H312 H400 H410 | GHS06 GHS09 Dgr | H330 H300 H312 H410 |  |  |  | CLP00/ATP01
602-053-00-0 | isobenzan (ISO); 1,3,4,5,6,7,8,8-octachloro-1,3,3a,4,7,7a-hexahydro-4,7-methanoisobenzofuran | 206-045-4 | 297-78-9 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 | H310 H300 H400 | GHS06 GHS09 Dgr | H310 H300 H400 |  |  |  | CLP00
602-054-00-6 | 3-iodpropene; allyl iodide | 209-130-4 | 556-56-9 | Flam. Liq. 2 Skin Corr. 1B | H225 H314 | GHS02 GHS05 Dgr | H225 H314 |  |  |  | CLP00/ATP01
602-055-00-1 | bromoethane; ethyl bromide | 200-825-8 | 74-96-4 | Flam. Liq. 2 Carc. 2 Acute Tox. 4 * Acute Tox. 4 * | H225 H351 H332 H302 | GHS02 GHS08 GHS07 Dgr | H225 H351 H332 H302 |  |  |  | CLP00
602-056-00-7 | α,α,α-trifluorotoluene; benzotrifluoride | 202-635-0 | 98-08-8 | Flam. Liq. 2 Aquatic Chronic 2 | H225 H411 | GHS02 GHS09 Dgr | H225 H411 |  |  |  | CLP00
602-057-00-2 | α-bromotoluene; benzyl bromide | 202-847-3 | 100-39-0 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H335 H315 H319 | GHS07 Wng | H319 H335 H315 |  |  |  | CLP00
602-058-00-8 | α,α-dichlorotoluene; benzylidene chloride; benzal chloride | 202-709-2 | 98-87-3 | Carc. 2 Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 | H351 H331 H302 H335 H315 H318 | GHS06 GHS08 GHS05 Dgr | H351 H331 H302 H335 H315 H318 |  |  |  | CLP00
602-059-00-3 | 1-chlorobutane; butyl chloride | 203-696-6 | 109-69-3 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 |  |  |  | CLP00
602-060-00-9 | bromobenzene | 203-623-8 | 108-86-1 | Flam. Liq. 3 Skin Irrit. 2 Aquatic Chronic 2 | H226 H315 H411 | GHS02 GHS07 GHS09 Wng | H226 H315 H411 |  |  |  | CLP00
602-061-00-4 | hexafluoropropene; hexafluoropropylene | 204-127-4 | 116-15-4 | Press. Gas Acute Tox. 4 * STOT SE 3 | H332 H335 | GHS07 Wng | H332 H335 |  |  | U | CLP00
602-062-00-X | 1,2,3-trichloropropane | 202-486-1 | 96-18-4 | Carc. 1B Repr. 1B Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H350 H360F *** H332 H312 H302 | GHS08 GHS07 Dgr | H350 H360F *** H332 H312 H302 |  |  | D | CLP00
602-063-00-5 | heptachlor epoxide; 2,3-epoxy-1,4,5,6,7,8,8-heptachloro-3a,4,7,7a-tetrahydro-4,7-methanoindane | 213-831-0 | 1024-57-3 | Carc. 2 Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H301 H373 ** H410 |  |  |  | CLP00
602-064-00-0 | 1,3-dichloro-2-propanol | 202-491-9 | 96-23-1 | Carc. 1B Acute Tox. 3 * Acute Tox. 4 * | H350 H301 H312 | GHS06 GHS08 Dgr | H350 H301 H312 |  |  |  | CLP00
602-065-00-6 | hexachlorobenzene | 204-273-9 | 118-74-1 | Carc. 1B STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H372 ** H400 H410 | GHS08 GHS09 Dgr | H350 H372 ** H410 |  |  |  | CLP00
602-066-00-1 | tetrachloro-p-benzoquinone | 204-274-4 | 118-75-2 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H319 H400 H410 | GHS07 GHS09 Wng | H319 H315 H410 |  |  |  | CLP00
602-067-00-7 | 1,3-dichlorbenzene | 208-792-1 | 541-73-1 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
602-068-00-2 | ethylene bis(trichloroacetate) | 219-732-9 | 2514-53-6 | Skin Irrit. 2 | H315 | GHS07 Wng | H315 |  |  |  | CLP00
602-069-00-8 | dichloroacetylene |  | 7572-29-4 | Unst. Expl. Carc. 2 STOT RE 2 * | H200 H351 H373 ** | GHS01 GHS08 Wng | H200 H351 H373 ** |  |  |  | CLP00
602-070-00-3 | 3-chloro-4,5,α, α,α-pentafluorotoluene | 401-930-3 | 77227-99-7 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 | H226 H332 H302 H400 | GHS02 GHS07 GHS09 Wng | H226 H332 H302 H400 |  |  |  | CLP00
602-071-00-9 | bromobenzylbromotoluene, reaction mass of isomers | 402-210-1 | 99688-47-8 | STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H373 ** H317 H410 |  |  |  | CLP00
602-072-00-4 | dichloro [(dichlorophenyl)methyl]methylbenzene, reaction mass of isomers; (dichlorophenyl)(dichlorotolyl)methane, reaction mass of isomers (IUPAC) | 278-404-3 | 76253-60-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
602-073-00-X | 1,4-dichlorobut-2-ene | 212-121-8 | 764-41-0 | Carc. 1B Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H330 H311 H301 H314 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H350 H330 H311 H301 H314 H410 |  | Carc. 1B; H350: C ≥ 0,01 % STOT SE 3; H335: C ≥ 5 % |  | CLP00
602-074-00-5 | pentachlorobenzene | 210-172-0 | 608-93-5 | Flam. Sol. 1 Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H228 H302 H400 H410 | GHS02 GHS07 GHS09 Dgr | H228 H302 H410 |  |  | T | CLP00
602-075-00-0 | 4,4,5,5-tetrachloro-1,3-dioxolan-2-one | 404-060-2 | 22432-68-4 | Acute Tox. 2 * Acute Tox. 4 * Skin Corr. 1B | H330 H302 H314 | GHS06 GHS05 Dgr | H330 H302 H314 |  |  |  | CLP00
602-076-00-6 | 2,3,4-trichlorobut-1-ene | 219-397-9 | 2431-50-7 | Carc. 2 Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H331 H302 H335 H315 H319 H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H331 H302 H319 H335 H315 H410 |  | Carc. 2; H351: C ≥ 0,1 % |  | CLP00/ATP01
602-077-00-1 | dodecachloropentacyclo[5.2.1.02,6.03,9.05,8]decane; mirex | 219-196-6 | 2385-85-5 | Carc. 2 Repr. 2 Lact. Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H361fd H362 H312 H302 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H361fd H362 H312 H302 H410 |  |  |  | CLP00
602-078-00-7 | hexachlorocyclopentadiene | 201-029-3 | 77-47-4 | Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H330 H311 H302 H314 H400 H410 | GHS06 GHS05 GHS09 Dgr | H330 H311 H302 H314 H410 |  |  |  | CLP00
602-079-00-2 | 2,3-dichloropropene; 2,3-dichloropropylene | 201-153-8 | 78-88-6 | Flam. Liq. 2 Muta. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H225 H341 H332 H312 H302 H335 H315 H318 H412 | GHS02 GHS08 GHS05 GHS07 Dgr | H225 H341 H332 H312 H302 H335 H315 H318 H412 |  |  |  | CLP00
602-080-00-8 | alkanes, C10-13, chloro; chlorinated paraffins, C10-13 | 287-476-5 | 85535-84-8 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 | EUH066 |  |  | CLP00/ATP01
602-081-00-3 | 2-chloro-4,5-difluorobenzoic acid | 405-380-5 |  | Acute Tox. 4 * Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 | H312 H302 H318 H317 | GHS05 GHS07 Dgr | H312 H302 H318 H317 |  |  |  | CLP00
602-082-00-9 | 2,2,6,6-tetrakis(bromomethyl)-4-oxaheptane-1,7-diol | 408-020-5 | 109678-33-3 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
602-083-00-4 | diphenyl ether, pentabromo derivative pentabromodiphenyl ether | 251-084-2 | 32534-81-9 | Lact. STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H362 H373 ** H400 H410 | GHS08 GHS09 Wng | H373 ** H362 H410 |  |  |  | CLP00
602-084-00-X | 1,1-dichloro-1-fluoroethane | 404-080-1 | 1717-00-6 | Aquatic Chronic 3 Ozone 1 | H412 H420 | GHS07 Wng | H412 H420 |  |  |  | CLP00/ATP02
602-085-00-5 | 2-bromopropane | 200-855-1 | 75-26-3 | Flam. Liq. 2 Repr. 1A STOT RE 2 * | H225 H360F *** H373 ** | GHS02 GHS08 Dgr | H225 H360F *** H373 ** | EUH066 |  |  | CLP00
602-086-00-0 | trifluoroiodomethane; trifluoromethyl iodide | 219-014-5 | 2314-97-8 | Muta. 2 | H341 | GHS08 Wng | H341 |  |  |  | CLP00
602-087-00-6 | 1,2,4-trichlorobenzene | 204-428-0 | 120-82-1 | Acute Tox. 4 * Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H400 H410 | GHS07 GHS09 Wng | H302 H315 H410 |  |  |  | CLP00
602-088-00-1 | 2,3-dibromopropan-1-ol; 2,3-dibromo-1-propanol | 202-480-9 | 96-13-9 | Carc. 1B Repr. 2 Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 3 | H350 H361f *** H311 H332 H302 H412 | GHS08 GHS06 Dgr | H332 H311 H302 H350 H361f *** H412 |  |  |  | CLP00
602-089-00-7 | 4-bromo-2-chlorofluorobenzene | 405-580-2 | 60811-21-4 | Acute Tox. 4 * Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H400 H410 | GHS07 GHS09 Wng | H302 H315 H410 |  |  |  | CLP00
602-090-00-2 | 1-allyl-3-chloro-4-fluorobenzene | 406-630-6 | 121626-73-1 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
602-091-00-8 | 1,3-dichloro-4-fluorobenzene | 406-160-1 | 1435-48-9 | Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Aquatic Chronic 2 | H302 H373 ** H315 H411 | GHS08 GHS07 Wng | H302 H373 ** H315 H411 |  |  |  | CLP00
602-092-00-3 | 1-bromo-3,4,5-trifluorobenzene | 418-480-9 | 138526-69-9 | Flam. Liq. 3 Carc. 2 Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 2 | H226 H351 H315 H318 H411 | GHS02 GHS08 GHS05 GHS09 Dgr | H226 H351 H315 H318 H411 |  |  |  | CLP00
602-093-00-9 | α, α,α,4-tetrachlorotoluene; p-chlorobenzotrichloride | 226-009-1 | 5216-25-1 | Carc. 1B Repr. 2 Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 STOT RE 1 Skin Irrit. 2 | H350 H361f *** H312 H302 H335 H372 ** H315 | GHS08 GHS07 Dgr | H350 H361f *** H372 ** H312 H302 H335 H315 |  |  |  | CLP00
602-094-00-4 | diphenylether; octabromo derivate | 251-087-9 | 32536-52-0 | Repr. 1B | H360Df | GHS08 Dgr | H360Df |  |  |  | CLP00
602-095-00-X | alkanes, C14-17, chloro; chlorinated paraffins, C14-17 | 287-477-0 | 85535-85-9 | Lact. Aquatic Acute 1 Aquatic Chronic 1 | H362 H400 H410 | GHS09 Wng | H362 H410 | EUH066 |  |  | ATP01
602-096-00-5 | malachite green hydrochloride [1] malachite green oxalate [2] | 209-322-8 [1] 219-441-7 [2] | 569-64-2 [1] 2437-29-8 [2] | Repr. 2 Acute Tox. 4 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H302 H318 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H361d *** H302 H318 H410 |  |  |  | CLP00
602-097-00-0 | 1-bromo-9-(4,4,5,5,5-pentafluoropentylthio)nonane | 422-850-5 | 148757-89-5 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
602-098-00-6 | 2-(3-bromophenoxy)tetrahydro-2H-pyran | 429-030-6 | 57999-49-2 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
602-099-00-1 | 3-(4-fluorophenyl)-2-methylpropionylchloride | 426-370-7 | - | Acute Tox. 4 * Skin Corr. 1A Aquatic Chronic 3 | H302 H314 H412 | GHS05 GHS07 Dgr | H314 H302 H412 | EUH014 EUH029 |  |  | ATP01
602-100-00-5 | reaction mass of: (R,R)-1,1,1,2,2,3,4,5,5,5-decafluoropentane; (S,S)-1,1,1,2,2,3,4,5,5,5-decafluoropentane | 420-640-8 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
602-101-00-0 | 2-chloro-4-fluoro-5-nitrophenyl (isobutyl)carbonate | 427-020-6 | 141772-37-4 | STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H373 ** H317 H410 |  |  |  | ATP01
602-102-00-6 | 1,1,1,3,3-pentafluorobutane | 430-250-1 | 406-58-6 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 |  |  |  | ATP01
602-103-00-1 | 1-(chlorophenylmethyl)-2-methylbenzene | 431-450-1 | 41870-52-4 | Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H400 H410 | GHS07 GHS09 Wng | H315 H410 |  |  |  | ATP01
602-104-00-7 | 1,1,2,2,3,3,4-heptafluorocyclopentane | 430-710-1 | 15290-77-4 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | ATP01
602-105-00-2 | sodium 1,1,2,2,3,3,4,4,4-nonafluoro-1-butanesulfinate | 422-100-7 | 102061-82-5 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
602-106-00-8 | 2-bromo-4,6-difluoroaniline | 429-430-0 | 444-14-4 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | ATP01
602-107-00-3 | 3,3,4,4-tetrafluoro-4-iodo-1-butene | 439-500-2 | 33831-83-3 | Acute Tox. 4 * Skin Irrit. 2 Aquatic Chronic 2 | H302 H315 H411 | GHS07 GHS09 Wng | H302 H315 H411 |  |  |  | ATP01
602-108-00-9 | (2,3,5,6-tetrafluorophenyl)methanol | 443-840-7 | 4084-38-2 | Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 | H302 H319 H317 | GHS07 Wng | H302 H319 H317 |  |  |  | ATP01
602-109-00-4 | Hexabromocyclododecane [1] 1,2,5,6,9,10-hexabromocyclododecane [2] | 247-148-4 [1] 221-695-9 [2] | 25637-99-4 [1] 3194-55-6 [2] | Repr. 2 Lact. | H361 H362 | GHS08 Wng | H361 H362 |  |  |  | ATP03
602-110-00-X | tetrafluoroethylene | 204-126-9 | 116-14-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | ATP17
603-001-00-X | methanol | 200-659-6 | 67-56-1 | Flam. Liq. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT SE 1 | H225 H331 H311 H301 H370 ** | GHS02 GHS06 GHS08 Dgr | H225 H331 H311 H301 H370 ** |  | * STOT SE 1; H370: C ≥ 10 % STOT SE 2; H371: 3 % ≤ C < 10 % |  | CLP00
603-002-00-5 | ethanol; ethyl alcohol | 200-578-6 | 64-17-5 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 |  |  |  | CLP00
603-003-00-0 | propan-1-ol; n-propanol | 200-746-9 | 71-23-8 | Flam. Liq. 2 STOT SE 3 Eye Dam. 1 | H225 H336 H318 | GHS02 GHS05 GHS07 Dgr | H225 H318 H336 |  |  |  | CLP00
603-004-00-6 | butan-1-ol; n-butanol | 200-751-6 | 71-36-3 | Flam. Liq. 3 Acute Tox. 4 * STOT SE 3 STOT SE 3 Skin Irrit. 2 Eye Dam. 1 | H226 H302 H335 H336 H315 H318 | GHS02 GHS05 GHS07 Dgr | H226 H302 H335 H315 H318 H336 |  |  |  | CLP00
603-005-00-1 | 2-methylpropan-2-ol; tert-butyl alcohol | 200-889-7 | 75-65-0 | Flam. Liq. 2 Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 | H225 H332 H335 H319 | GHS02 GHS07 Dgr | H225 H332 H319 H335 |  |  |  | CLP00/ATP01
603-006-00-7 | pentanol isomers, with the exception fo those specified elsewhere in this Annex | 250-378-8 |  | Flam. Liq. 3 Acute Tox. 4 * STOT SE 3 | H226 H332 H335 | GHS02 GHS07 Wng | H226 H332 H335 | EUH066 |  | C | CLP00
603-007-00-2 | 2-methylbutan-2-ol; tert-pentanol | 200-908-9 | 75-85-4 | Flam. Liq. 2 Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 | H225 H332 H335 H315 | GHS02 GHS07 Dgr | H225 H332 H335 H315 |  |  |  | CLP00
603-008-00-8 | 4-methylpentan-2-ol; methyl isobutyl carbinol | 203-551-7 | 108-11-2 | Flam. Liq. 3 STOT SE 3 | H226 H335 | GHS02 GHS07 Wng | H226 H335 |  | STOT SE 3; H335: C ≥ 25 % |  | CLP00
603-009-00-3 | cyclohexanol | 203-630-6 | 108-93-0 | Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 | H332 H302 H335 H315 | GHS07 Wng | H332 H302 H335 H315 |  |  |  | CLP00
603-010-00-9 | 2-methylcyclohexanol, mixed isomers [1] cis-2-methylcyclohexanol [2] trans-2-methylcyclohexanol [3] | 209-512-0 [1] 231-187-9 [2] 231-186-3 [3] | 583-59-5 [1] 7443-70-1 [2] 7443-52-9 [3] | Acute Tox. 4 * | H332 | GHS07 Wng | H332 |  |  | C | CLP00
603-011-00-4 | 2-methoxyethanol; ethylene glycol monomethyl ether | 203-713-7 | 109-86-4 | Flam. Liq. 3 Repr. 1B Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H226 H360FD H332 H312 H302 | GHS02 GHS08 GHS07 Dgr | H226 H360FD H332 H312 H302 |  |  |  | CLP00
603-012-00-X | 2-ethoxyethanol; ethylene glycol monoethyl ether | 203-804-1 | 110-80-5 | Flam. Liq. 3 Repr. 1B Acute Tox. 3 Acute Tox. 4 | H226 H360FD H331 H302 | GHS02 GHS08 GHS06 Dgr | H226 H360FD H331 H302 |  |  |  | CLP00/ATP03
603-013-00-5 | 2-isopropoxyethanol; ethylene glycol monoisopropyl ether | 203-685-6 | 109-59-1 | Acute Tox. 4 * Acute Tox. 4 * Eye Irrit. 2 | H332 H312 H319 | GHS07 Wng | H332 H312 H319 |  |  |  | CLP00
603-014-00-0 | 2-butoxyethanol; ethylene glycol monobutyl ether | 203-905-0 | 111-76-2 | Acute Tox. 3 Acute Tox. 4 Skin Irrit. 2 Eye Irrit. 2 | H331 H302 H315 H319 | GHS06 Dgr | H331 H302 H315 H319 |  | inhalation: ATE = 3 mg/L (Vapours) oral: ATE=1 200 mg/kg bw |  | CLP00/ATP18
603-015-00-6 | allyl alcohol | 203-470-7 | 107-18-6 | Flam. Liq. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 | H225 H331 H311 H301 H335 H315 H319 H400 | GHS02 GHS06 GHS09 Dgr | H225 H331 H311 H301 H319 H335 H315 H400 |  |  |  | CLP00
603-016-00-1 | 4-hydroxy-4-methylpentan-2-one; diacetone alcohol | 204-626-7 | 123-42-2 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  | Eye Irrit. 2; H319: C ≥ 10 % |  | CLP00
603-018-00-2 | furfuryl alcohol | 202-626-1 | 98-00-0 | Carc. 2 Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 STOT RE 2 * Eye Irrit. 2 | H351 H331 H312 H302 H335 H373 ** H319 | GHS06 GHS08 Dgr | H351 H331 H312 H302 H373 ** H319 H335 |  |  |  | CLP00/ATP01
603-019-00-8 | dimethyl ether | 204-065-8 | 115-10-6 | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | U | CLP00
603-020-00-3 | ethyl methyl ether |  | 540-67-0 | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | U | CLP00
603-021-00-9 | methyl vinyl ether | 203-475-4 | 107-25-5 | Flam. Gas 1 Press. Gas | H220 | GHS02 GHS04 Dgr | H220 |  |  | D U | CLP00
603-022-00-4 | diethyl ether; ether | 200-467-2 | 60-29-7 | Flam. Liq. 1 Acute Tox. 4 * STOT SE 3 | H224 H302 H336 | GHS02 GHS07 Dgr | H224 H302 H336 | EUH019 EUH066 |  |  | CLP00
603-023-00-X | ethylene oxide; oxirane | 200-849-9 | 75-21-8 | Flam. Gas 1 Press. Gas Carc. 1B Muta. 1B Repr. 1B Acute Tox. 3 Acute Tox. 3 STOT SE 3 STOT SE 3 STOT RE 1 Skin Corr. 1 Eye Dam. 1 | H220 H350 H340 H360Fd H331 H301 H335 H336 H372 (nervous system) H314 H318 | GHS02 GHS05 GHS06 GHS08 Dgr | H220 H331 H301 H314 H340 H350 H360Fd H335 H336 H372 (nervous system) |  | Inhalation: ATE = 700 ppmV (Gases) Oral: ATE = 100 mg/kg bw | U | CLP00/ATP14
603-024-00-5 | 1,4-dioxane | 204-661-8 | 123-91-1 | Flam. Liq. 2 Carc. 1B STOT SE 3 Eye Irrit. 2 | H225 H350 H335 H319 | GHS02 GHS08 GHS07 Dgr | H225 H319 H350 H335 | EUH019 EUH066 |  | D | CLP00/ATP17
603-025-00-0 | tetrahydrofuran | 203-726-8 | 109-99-9 | Flam. Liq. 2 Carc. 2 STOT SE 3 Eye Irrit. 2 | H225 H351 H335 H319 | GHS02 GHS07 GHS08 Dgr | H225 H351 H319 H335 | EUH019 | STOT SE 3; H335: C ≥ 25 % Eye Irrit.2; H319: C ≥ 25 % |  | CLP00/ATP03
603-026-00-6 | 1-chloro-2,3-epoxypropane; epichlorhydrin | 203-439-8 | 106-89-8 | Flam. Liq. 3 Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Skin Sens. 1 | H226 H350 H331 H311 H301 H314 H317 | GHS02 GHS06 GHS08 GHS05 Dgr | H226 H350 H331 H311 H301 H314 H317 |  | * |  | CLP00
603-027-00-1 | ethanediol; ethylene glycol | 203-473-3 | 107-21-1 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
603-028-00-7 | 2-chloroethanol; ethylene chlorohydrin | 203-459-7 | 107-07-3 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * | H310 H330 H300 | GHS06 Dgr | H330 H310 H300 |  |  |  | CLP00
603-029-00-2 | bis(2-chloroethyl) ether | 203-870-1 | 111-44-4 | Carc. 2 Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * | H351 H310 H330 H300 | GHS06 GHS08 Dgr | H351 H330 H310 H300 |  |  |  | CLP00/ATP01
603-030-00-8 | 2-aminoethanol; ethanolamine | 205-483-3 | 141-43-5 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H332 H312 H302 H314 | GHS05 GHS07 Dgr | H332 H312 H302 H314 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
603-031-00-3 | 1,2-dimethoxyethane; ethylene glycol dimethyl ether; EGDME | 203-794-9 | 110-71-4 | Flam. Liq. 2 Repr. 1B Acute Tox. 4 * | H225 H360FD H332 | GHS02 GHS08 GHS07 Dgr | H225 H360FD H332 | EUH019 |  |  | CLP00
603-032-00-9 | ethylene dinitrate; ethylene glycol dinitrate | 211-063-0 | 628-96-6 | Unst. Expl. Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 | H200 H310 H330 H300 H373 ** | GHS01 GHS06 GHS08 Dgr | H200 H330 H310 H300 H373 ** |  |  |  | CLP00/ATP01
603-033-00-4 | oxydiethylene dinitrate; diethylene glycol dinitrate; digol dinitrate | 211-745-8 | 693-21-0 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Chronic 3 Unst. Expl. | H310 H330 H300 H373 ** H412 H200 | GHS01 GHS06 GHS08 Dgr | H200 H330 H310 H300 H373 ** H412 |  |  |  | CLP00
603-033-01-1 | oxydiethylene dinitrate; diethylene glycol dinitrate; digol dinitrate; [>25 % phlegmatiser] | 211-745-8 | 693-21-0 | Expl. 1.1 Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Chronic 3 | H201 H310 H330 H300 H373 ** H412 | GHS01 GHS06 GHS08 Dgr | H201 H330 H310 H300 H373 ** H412 |  |  |  | CLP00
603-034-00-X | glycerol trinitrate; nitroglycerine | 200-240-8 | 55-63-0 | Unst. Expl. Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Chronic 2 | H200 H310 H330 H300 H373 ** H411 | GHS01 GHS06 GHS08 GHS09 Dgr | H200 H330 H310 H300 H373 ** H411 |  |  |  | CLP00
603-034-01-7 | glycerol trinitrate; nitroglycerine; [>40 % phlegmatiser] | 200-240-8 | 55-63-0 | Expl. 1.1 Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Chronic 2 | H201 H310 H330 H300 H373 ** H411 | GHS01 GHS06 GHS08 GHS09 Dgr | H201 H330 H310 H300 H373 ** H411 |  |  |  | CLP00
603-035-00-5 | pentaerythritol tetranitrate; P.E.T.N. | 201-084-3 | 78-11-5 | Unst. Expl. | H200 | GHS01 Dgr | H200 |  |  |  | CLP00
603-035-01-2 | pentaerythritol tetranitrate; pentaerythrite tetranitrate; P.E.T.N.; [>20 % phlegmatiser] | 201-084-3 | 78-11-5 | Expl. 1.1 | H201 | GHS01 Dgr | H201 |  |  | T | CLP00
603-036-00-0 | mannitol hexanitrate; nitromannite | 239-924-6 | 15825-70-4 | Unst. Expl. | H200 | GHS01 Dgr | H200 |  |  |  | CLP00
603-036-01-8 | mannitol hexanitrate; nitromannite; [>40 % phlegmatiser] | 239-924-6 | 15825-70-4 | Expl. 1.1 | H201 | GHS01 Dgr | H201 |  |  |  | CLP00
603-037-00-6 | cellulose nitrate; nitrocellulose | - | - | Expl. 1.1 | H201 | GHS01 Dgr | H201 |  |  | T | CLP00/ATP01
603-038-00-1 | allyl glycidyl ether; allyl 2,3-epoxypropyl ether; prop-2-en-1-yl 2,3-epoxypropyl ether | 203-442-4 | 106-92-3 | Flam. Liq. 3 Carc. 2 Muta. 2 Repr. 2 Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H226 H351 H341 H361f *** H332 H302 H335 H315 H318 H317 H412 | GHS02 GHS08 GHS05 GHS07 Dgr | H226 H351 H341 H361f *** H332 H302 H335 H315 H318 H317 H412 |  |  |  | CLP00
603-039-00-7 | butyl glycidyl ether; butyl 2,3-epoxypropyl ether | 219-376-4 | 2426-08-6 | Flam. Liq. 3 Carc. 2 Muta. 2 Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Sens. 1 Aquatic Chronic 3 | H226 H351 H341 H332 H302 H335 H317 H412 | GHS02 GHS08 GHS07 Wng | H226 H351 H341 H332 H302 H335 H317 H412 |  |  |  | CLP00
603-040-00-2 | sodium methanolate; sodium methoxide [1] potassium methanolate; potassium methoxide [2] lithium methanolate; lithium methoxide [3] | 204-699-5 [1] 212-736-1 [2] 212-737-7 [3] | 124-41-4 [1] 865-33-8 [2] 865-34-9 [3] | Self-heat. 1 Skin Corr. 1B | H251 H314 | GHS02 GHS05 Dgr | H251 H314 | EUH014 |  | T | CLP00
603-041-00-8 | potassium ethanolate; potassium ethoxide [1] sodium ethanolate; sodium ethoxide [2] | 213-029-0 [1] 205-487-5 [2] | 917-58-8 [1] 141-52-6 [2] | Self-heat. 1 Skin Corr. 1B | H251 H314 | GHS02 GHS05 Dgr | H251 H314 | EUH014 |  | T | CLP00
603-042-00-3 | aluminium-tri-isopropoxide | 209-090-8 | 555-31-7 | Flam. Sol. 1 | H228 | GHS02 Dgr | H228 |  |  | T | CLP00
603-043-00-9 | triarimol (ISO); 2,4-dichloro-α-(pyrimidin-5-yl) benzhydryl alcohol |  | 26766-27-8 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
603-044-00-4 | dicofol (ISO); 2,2,2-trichloro-1,1-bis(4-chlorophenyl)ethanol | 204-082-0 | 115-32-2 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H315 H317 H400 H410 | GHS07 GHS09 Wng | H312 H302 H315 H317 H410 |  |  |  | CLP00
603-045-00-X | diisopropyl ether [1] dipropyl ether [2] | 203-560-6 [1] 203-869-6 [2] | 108-20-3 [1] 111-43-3 [2] | Flam. Liq. 2 STOT SE 3 | H225 H336 | GHS02 GHS07 Dgr | H225 H336 | EUH019 EUH066 |  | C | CLP00
603-046-00-5 | bis(chloromethyl) ether; oxybis(chloromethane) | 208-832-8 | 542-88-1 | Flam. Liq. 2 Carc. 1A Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 4 * | H225 H350 H330 H311 H302 | GHS02 GHS06 GHS08 Dgr | H225 H350 H330 H311 H302 |  | Carc. 1A; H350: C ≥ 0,001 % |  | CLP00/ATP01
603-047-00-0 | 2-dimethylaminoethanol; N,N-dimethylethanolamine | 203-542-8 | 108-01-0 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H226 H332 H312 H302 H314 | GHS02 GHS05 GHS07 Dgr | H226 H332 H312 H302 H314 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
603-048-00-6 | 2-diethylaminoethanol; N,N-diethylethanolamine | 202-845-2 | 100-37-8 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H226 H332 H312 H302 H314 | GHS02 GHS05 GHS07 GHS09 Dgr | H226 H332 H312 H302 H314 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
603-049-00-1 | chlorfenethol (ISO); 1,1-bis (4-chlorophenyl) ethanol | 201-246-3 | 80-06-8 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
603-050-00-7 | 1-(2-Butoxypropoxy)propan-2-ol | 246-011-6 | 24083-03-2 | Acute Tox. 4 * Acute Tox. 4 * | H312 H302 | GHS07 Wng | H312 H302 |  |  |  | CLP00
603-051-00-2 | 2-ethylbutan-1-ol | 202-621-4 | 97-95-0 | Acute Tox. 4 * Acute Tox. 4 * | H312 H302 | GHS07 Wng | H312 H302 |  |  |  | CLP00
603-052-00-8 | 3-butoxypropan-2-ol; propylene glycol monobutyl ether | 225-878-4 | 5131-66-8 | Skin Irrit. 2 Eye Irrit. 2 | H315 H319 | GHS07 Wng | H319 H315 |  |  |  | CLP00
603-053-00-3 | 2-methylpentane-2,4-diol | 203-489-0 | 107-41-5 | Skin Irrit. 2 Eye Irrit. 2 | H315 H319 | GHS07 Wng | H319 H315 |  |  |  | CLP00
603-054-00-9 | di-n-butyl ether; dibutyl ether | 205-575-3 | 142-96-1 | Flam. Liq. 3 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 3 | H226 H335 H315 H319 H412 | GHS02 GHS07 Wng | H226 H319 H335 H315 H412 |  | STOT SE 3; H335: C ≥ 10 % |  | CLP00
603-055-00-4 | propylene oxide; 1,2-epoxypropane; methyloxirane | 200-879-2 | 75-56-9 | Flam. Liq. 1 Carc. 1B Muta. 1B Acute Tox. 3 Acute Tox. 3 Acute Tox. 4 STOT SE 3 Eye Irrit. 2 | H224 H350 H340 H331 H311 H302 H335 H319 | GHS02 GHS08 GHS06 Dgr | H224 H331 H311 H302 H319 H340 H350 H335 |  |  |  | CLP00/ATP09
603-056-00-X | [(p-tolyloxy)methyl]oxirane [1] [(m-tolyloxy)methyl]oxirane [2] 2,3-epoxypropyl o-tolyl ether [3] [(tolyloxy)methyl]oxirane; cresyl glycidyl ether [4] | 218-574-8 [1] 218-575-3 [2] 218-645-3 [3] 247-711-4 [4] | 2186-24-5 [1] 2186-25-6 [2] 2210-79-9 [3] 26447-14-3 [4] | Muta. 2 Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H341 H315 H317 H411 | GHS08 GHS07 GHS09 Wng | H341 H315 H317 H411 |  |  | C | CLP00
603-057-00-5 | benzyl alcohol | 202-859-9 | 100-51-6 | Acute Tox. 4 Eye Irrit. 2 Skin Sens. 1B | H302 H319 H317 | GHS07 Wng | H302 H319 H317 |  | oral: ATE = 1200 mg/kg bw |  | CLP00/ATP21
603-058-00-0 | 1,3-propylene oxide | 207-964-3 | 503-30-0 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H225 H332 H312 H302 | GHS02 GHS07 Dgr | H225 H332 H312 H302 |  |  |  | CLP00
603-059-00-6 | hexan-1-ol | 203-852-3 | 111-27-3 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
603-060-00-1 | 2,2'-bioxirane; 1,2:3,4-diepoxybutane | 215-979-1 | 1464-53-5 | Carc. 1B Muta. 1B Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B | H350 H340 H330 H311 H301 H314 | GHS06 GHS08 GHS05 Dgr | H350 H340 H330 H311 H301 H314 |  |  |  | CLP00
603-061-00-7 | tetrahydro-2-furylmethanol; tetrahydrofurfuryl alcohol | 202-625-6 | 97-99-4 | Repr. 1B Eye Irrit. 2 | H360Df H319 | GHS08 GHS07 Dgr | H319 H360Df |  |  |  | CLP00/ATP06
603-062-00-2 | tetrahydrofuran-2,5-diyldimethanol | 203-239-0 | 104-80-3 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H335 H315 H319 | GHS07 Wng | H319 H335 H315 |  | STOT SE 3; H335: C ≥ 10 % |  | CLP00
603-063-00-8 | 2,3-epoxypropan-1-ol; glycidol; oxiranemethanol | 209-128-3 | 556-52-5 | Carc. 1B Muta. 2 Repr. 1B Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H350 H341 H360F *** H331 H312 H302 H335 H315 H319 | GHS06 GHS08 Dgr | H350 H341 H360F *** H331 H312 H302 H319 H335 H315 |  |  |  | CLP00
603-064-00-3 | 1-methoxy-2-propanol; monopropylene glycol methyl ether | 203-539-1 | 107-98-2 | Flam. Liq. 3 STOT SE 3 | H226 H336 | GHS02 GHS07 Wng | H226 H336 |  |  |  | CLP00/ATP01
603-065-00-9 | m-bis(2,3-epoxypropoxy)benzene; resorcinol diglycidyl ether | 202-987-5 | 101-90-6 | Carc. 1B Muta. 2 Acute Tox. 3 Acute Tox. 4 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H350 H341 H311 H302 H315 H319 H317 H412 | GHS08 GHS06 Dgr | H311 H302 H315 H319 H317 H341 H350 H412 |  | Dermal: ATE = 300 mg/kg bw Oral: ATE = 500 mg/kg bw |  | CLP00/ATP15
603-066-00-4 | 7-oxa-3-oxiranylbicyclo [4.1.0]heptane; 1,2-epoxy- 4-epoxyethylcyclohexane; 4-vinylcyclohexene diepoxide | 203-437-7 | 106-87-6 | Carc. 1B Muta. 2 Repr. 1B Acute Tox. 3 Acute Tox. 4 | H350 H341 H360F H331 H302 | GHS08 GHS06 Dgr | H331 H302 H341 H350 H360F |  | Inhalation: ATE = 0.5 mg/L (dusts/mists) Oral: ATE = 847 mg/kg bw |  | CLP00/ATP17
603-067-00-X | phenyl glycidyl ether; 2,3-epoxypropyl phenyl ether; 1,2-epoxy-3-phenoxypropane | 204-557-2 | 122-60-1 | Carc. 1B Muta. 2 Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H350 H341 H332 H335 H315 H317 H412 | GHS08 GHS07 Dgr | H350 H341 H332 H335 H315 H317 H412 |  |  |  | CLP00
603-068-00-5 | 2,3-epoxypropyl-2-ethylcyclohexyl ether; ethylcyclohexylglycidyl ether |  | 130014-35-6 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H315 H319 H317 | GHS07 Wng | H319 H315 H317 |  |  |  | CLP00
603-069-00-0 | 2,4,6-tris(dimethylaminomethyl)phenol | 202-013-9 | 90-72-2 | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 | H302 H315 H319 | GHS07 Wng | H302 H319 H315 |  |  |  | CLP00
603-070-00-6 | 2-amino-2-methylpropanol | 204-709-8 | 124-68-5 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 3 | H315 H319 H412 | GHS07 Wng | H319 H315 H412 |  |  |  | CLP00
603-071-00-1 | 2,2'-iminodiethanol; diethanolamine | 203-868-0 | 111-42-2 | Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 | H302 H373 ** H315 H318 | GHS08 GHS05 GHS07 Dgr | H302 H373 ** H315 H318 |  |  |  | CLP00
603-072-00-7 | 1,4-bis(2,3 epoxypropoxy)butane; butanedioldiglycidyl ether | 219-371-7 | 2425-79-8 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H332 H312 H315 H319 H317 | GHS07 Wng | H332 H312 H319 H315 H317 |  |  |  | CLP00
603-073-00-2 | bis-[4-(2,3-epoxipropoxi)phenyl]propane | 216-823-5 | 1675-54-3 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H315 H319 H317 | GHS07 Wng | H319 H315 H317 |  | Eye Irrit. 2; H319: C ≥ 5 % Skin Irrit. 2; H315: C ≥ 5 % |  | CLP00
603-074-00-8 | reaction product: bisphenol-A-(epichlorhydrin); epoxy resin (number average molecular weight ≤ 700) | 500-033-5 | 25068-38-6 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H315 H319 H317 H411 | GHS07 GHS09 Wng | H319 H315 H317 H411 |  | Eye Irrit. 2; H319: C ≥ 5 % Skin Irrit. 2; H315: C ≥ 5 % |  | CLP00
603-075-00-3 | chlormethyl methyl ether; chlorodimethyl ether | 203-480-1 | 107-30-2 | Flam. Liq. 2 Carc. 1A Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H225 H350 H332 H312 H302 | GHS02 GHS08 GHS07 Dgr | H225 H350 H332 H312 H302 |  |  |  | CLP00
603-076-00-9 | but-2-yne-1,4-diol; 2-butyne-1,4-diol | 203-788-6 | 110-65-6 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Skin Sens. 1 | H331 H301 H312 H373 ** H314 H317 | GHS06 GHS05 GHS08 Dgr | H314 H331 H301 H312 H373 ** H317 |  | Skin Corr. 1B; H314: C ≥ 50 % Skin Irrit. 2; H315: 25 % ≤ C < 50 % Eye Irrit. 2; H319: 25 % ≤ C < 50 % | D | CLP00
603-077-00-4 | 1-dimethylaminopropan-2-ol; dimepranol (INN) | 203-556-4 | 108-16-7 | Flam. Liq. 3 Acute Tox. 4 * Skin Corr. 1B | H226 H302 H314 | GHS02 GHS05 GHS07 Dgr | H226 H302 H314 |  |  |  | CLP00
603-078-00-X | prop-2-yn-1-ol; propargyl alcohol | 203-471-2 | 107-19-7 | Flam. Liq. 3 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Aquatic Chronic 2 | H226 H331 H311 H301 H314 H411 | GHS02 GHS06 GHS05 GHS09 Dgr | H226 H331 H311 H301 H314 H411 |  |  |  | CLP00
603-079-00-5 | 2,2'-(methylimino)diethanol; N-methyldiethanolamine | 203-312-7 | 105-59-9 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
603-080-00-0 | 2-methylaminoethanol; N-methylethanolamine; N-methyl-2-ethanolamine; N-methyl-2-amino ethanol; 2-(methylamino)ethanol | 203-710-0 | 109-83-1 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H312 H302 H314 | GHS05 GHS07 Dgr | H312 H302 H314 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
603-081-00-6 | 2,2'-thiodiethanol; thiodiglycol | 203-874-3 | 111-48-8 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
603-082-00-1 | 1-aminopropan-2-ol; isopropanolamine | 201-162-7 | 78-96-6 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
603-083-00-7 | 1,1'-iminodipropan-2-ol; di-isopropanolamine | 203-820-9 | 110-97-4 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
603-084-00-2 | styrene oxide; (epoxyethyl)benzene; phenyloxirane | 202-476-7 | 96-09-3 | Carc. 1B Acute Tox. 4 * Eye Irrit. 2 | H350 H312 H319 | GHS08 GHS07 Dgr | H350 H312 H319 |  |  |  | CLP00
603-085-00-8 | bronopol (INN); 2-bromo-2-nitropropane-1,3-diol | 200-143-0 | 52-51-7 | Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 | H312 H302 H335 H315 H318 H400 | GHS05 GHS07 GHS09 Dgr | H312 H302 H335 H315 H318 H400 |  | M=10 |  | CLP00/ATP01
603-086-00-3 | ethirimol (ISO); 5-butyl-2-ethylamino-6-methylpyrimidin-4-ol | 245-949-3 | 23947-60-6 | Acute Tox. 4 * | H312 | GHS07 Wng | H312 |  |  |  | CLP00
603-087-00-9 | 2-ethylhexane-1,3-diol; octylene glycol; ethoexadiol | 202-377-9 | 94-96-2 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
603-088-00-4 | 2-(octylthio)ethanol; 2-hydroxyethyl octyl sulphide | 222-598-4 | 3547-33-9 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
603-089-00-X | 7,7-dimethyl-3-oxa-6-azaoctan-1-ol | 400-390-6 |  | Acute Tox. 4 * Skin Corr. 1A | H302 H314 | GHS05 GHS07 Dgr | H314 H302 |  |  |  | CLP00
603-090-00-5 | 2-(2-bromoethoxy)anisole | 402-010-4 | 4463-59-6 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
603-091-00-0 | exo-1-methyl-4-(1-methylethyl)-7-oxabicyclo[2.2.1]heptan-2-ol | 402-470-6 | 87172-89-2 | Acute Tox. 4 * Eye Dam. 1 | H302 H318 | GHS05 GHS07 Dgr | H302 H318 |  |  |  | CLP00
603-092-00-6 | 2-methyl-4-phenylpentanol | 402-770-7 | 92585-24-5 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
603-093-00-1 | cinmethylin (ISO); exo-(±)-1-methyl-2-(2-methylbenzyloxy)-4-isopropyl-7-oxabicyclo(2.2.1)heptane | 402-410-9 | 87818-31-3 | Acute Tox. 4 * Aquatic Chronic 2 | H332 H411 | GHS07 GHS09 Dgr | H332 H411 |  |  |  | CLP00
603-094-00-7 | 1,3-bis(2,3-epoxypropoxy)-2,2-dimethylpropane | 241-536-7 | 17557-23-2 | Skin Irrit. 2 Skin Sens. 1 | H315 H317 | GHS07 Wng | H315 H317 |  |  |  | CLP00
603-095-00-2 | 2-(propyloxy)ethanol; EGPE | 220-548-6 | 2807-30-9 | Acute Tox. 4 * Eye Irrit. 2 | H312 H319 | GHS07 Wng | H312 H319 |  |  |  | CLP00
603-096-00-8 | 2-(2-butoxyethoxy)ethanol; diethylene glycol monobutyl ether | 203-961-6 | 112-34-5 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
603-097-00-3 | 1,1',1'-nitrilotripropan-2-ol; triisopropanolamine | 204-528-4 | 122-20-3 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00/ATP05
603-098-00-9 | 2-phenoxyethanol | 204-589-7 | 122-99-6 | Acute Tox. 4 STOT SE 3 Eye Dam. 1 | H302 H335 H318 | GHS05 GHS07 Dgr | H302 H318 H335 |  | Oral: ATE = 1394 mg/kg bw |  | CLP00/ATP17
603-099-00-4 | 3-(N-methyl-N-(4-methylamino-3-nitrophenyl)amino)propane-1,2-diol hydrochloride | 403-440-5 | 93633-79-5 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
603-100-00-8 | 1,2-dimethoxypropane | 404-630-0 | 7778-85-0 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 | EUH019 |  |  | CLP00
603-101-00-3 | tetrahydro-2-isobutyl-4-methylpyran-4-ol, mixed isomers (cis and trans) | 405-040-6 |  | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
603-102-00-9 | 1,2-epoxybutane | 203-438-2 | 106-88-7 | Flam. Liq. 2 Carc. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H225 H351 H332 H312 H302 H335 H315 H319 | GHS02 GHS08 GHS07 Dgr | H225 H302 H312 H332 H315 H319 H351 H335 |  |  |  | CLP00/ATP07
603-103-00-4 | oxirane, mono[(C12-14-alkyloxy)methyl] derivs. | 271-846-8 | 68609-97-2 | Skin Irrit. 2 Skin Sens. 1 | H315 H317 | GHS07 Wng | H315 H317 |  |  |  | CLP00
603-104-00-X | fenarimol (ISO); 2,4'-dichloro-α-(pyrimidin-5-yl)benzhydryl alcohol | 262-095-7 | 60168-88-9 | Repr. 2 Lact. Aquatic Chronic 2 | H361fd H362 H411 | GHS08 GHS09 Wng | H361fd H362 H411 |  |  |  | CLP00
603-105-00-5 | furan | 203-727-3 | 110-00-9 | Flam. Liq. 1 Carc. 1B Muta. 2 Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Aquatic Chronic 3 | H224 H350 H341 H332 H302 H373 ** H315 H412 | GHS02 GHS08 GHS07 Dgr | H224 H350 H341 H332 H302 H373 ** H315 H412 | EUH019 |  |  | CLP00
603-106-00-0 | 2-methoxypropanol | 216-455-5 | 1589-47-5 | Flam. Liq. 3 Repr. 1B STOT SE 3 Skin Irrit. 2 Eye Dam. 1 | H226 H360D *** H335 H315 H318 | GHS02 GHS08 GHS05 GHS07 Dgr | H226 H360D *** H335 H315 H318 |  |  |  | CLP00
603-107-00-6 | 2-(2-methoxyethoxy)ethanol; diethylene glycol monomethyl ether | 203-906-6 | 111-77-3 | Repr. 1B | H360D | GHS08 Dgr | H360D |  | Repr. 1B; H360D: C ≥ 3 % |  | CLP00/ATP18
603-108-00-1 | 2-methylpropan-1-ol; iso-butanol | 201-148-0 | 78-83-1 | Flam. Liq. 3 STOT SE 3 STOT SE 3 Skin Irrit. 2 Eye Dam. 1 | H226 H335 H336 H315 H318 | GHS02 GHS05 GHS07 Dgr | H226 H335 H315 H318 H336 |  |  |  | CLP00
603-109-00-7 | reaction mass of: 1-ethoxy-1,1,2,3,3,3-hexafluoro-2-(trifluoromethyl)propane; 1-ethoxy-1,1,2,2,3,3,4,4,4-nonafluorobutane | 425-340-0 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
603-110-00-2 | reaction mass of: cis-2-isobutyl-5-methyl 1,3-dioxane; trans-2-isobutyl-5-methyl 1,3-dioxane | 426-130-1 | 166301-21-9 | Skin Irrit. 2 Aquatic Chronic 3 | H315 H412 | GHS07 Wng | H315 H412 |  |  |  | ATP01
603-111-00-8 | reaction mass of: 1-(1,1-dimethylpropyl)-4-ethoxy-cis-cyclohexane; 1-(1,1-dimethylpropyl)-4-ethoxy-trans-cyclohexane | 426-530-6 | - | Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H400 H410 | GHS07 GHS09 Wng | H315 H410 |  |  |  | ATP01
603-112-00-3 | cyclopentyl 2-phenylethyl ether | 428-340-9 | - | Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H400 H410 | GHS07 GHS09 Wng | H315 H410 |  |  |  | ATP01
603-113-00-9 | 6-glycidyloxynapht-1-yl oxymethyloxirane | 429-960-2 | 27610-48-6 | Muta. 2 Acute Tox. 4 * Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H341 H312 H315 H317 H412 | GHS08 GHS07 Wng | H341 H312 H315 H317 H412 |  |  |  | ATP01
603-114-00-4 | 9-(2-propenyloxy)tricyclo[5.2.1.0(2,6)]dec-3(or-4-)-ene | 430-830-2 | 26912-64-1 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | ATP01
603-115-00-X | reaction mass of: O,O',O''-(methylsilanetriyl)tris(4-methyl-2-pentanone oxime) (3 stereoisomers) | 423-580-0 | - | STOT RE 2 * Aquatic Chronic 4 | H373 ** H413 | GHS08 Wng | H373 ** H413 |  |  |  | ATP01
603-116-00-5 | (Z)-(2,4-difluorophenyl)piperidin-4-ylmethanone oxime monohydrochloride | 424-740-2 | 138271-16-6 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | ATP01
603-117-00-0 | propan-2-ol; isopropyl alcohol; isopropanol | 200-661-7 | 67-63-0 | Flam. Liq. 2 STOT SE 3 Eye Irrit. 2 | H225 H336 H319 | GHS02 GHS07 Dgr | H225 H319 H336 |  |  |  | CLP00
603-118-00-6 | 6-dimethylaminohexan-1-ol | 404-680-3 | 1862-07-3 | Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 3 | H302 H314 H412 | GHS05 GHS07 Dgr | H302 H314 H412 |  |  |  | CLP00
603-119-00-1 | 1,1'-(1,3-phenylenedioxy)bis(3-(2-(prop-2-enyl)phenoxy)propan-2-ol) | 405-840-5 |  | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
603-120-00-7 | 2-methyl-5-phenylpentanol | 405-890-8 | 25634-93-9 | Skin Irrit. 2 Eye Irrit. 2 | H315 H319 | GHS07 Wng | H319 H315 |  |  |  | CLP00
603-121-00-2 | 4-[4-(1,3-dihydroxyprop-2-yl)phenylamino]-1,8-dihydroxy-5-nitroanthraquinone | 406-057-1 | 114565-66-1 | Carc. 2 Skin Sens. 1 Aquatic Chronic 4 | H351 H317 H413 | GHS08 GHS07 Wng | H351 H317 H413 |  |  |  | CLP00
603-122-00-8 | sodium 2-ethylhexanolate | 406-150-7 | 38411-13-1 | Flam. Sol. 1 Skin Corr. 1B Aquatic Chronic 3 | H228 H314 H412 | GHS02 GHS05 Dgr | H228 H314 H412 |  |  | T | CLP00
603-123-00-3 | 4-methyl-8-methylenetricyclo[3.3.1.13,7]decan-2-ol | 406-330-5 | 122760-84-3 | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H315 H317 H411 | GHS07 GHS09 Wng | H315 H317 H411 |  |  |  | CLP00
603-124-00-9 | 1,4-bis[2-(vinyloxy)ethoxy]benzene | 406-900-3 | 84563-49-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
603-125-00-4 | 2-(2,4-dichlorophenyl)-1-(1H-1,2,4-triazol-1-yl)pent-4-en-2-ol | 407-850-5 | 89544-40-1 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | CLP00
603-126-00-X | 2-((4-methyl-2-nitrophenyl)amino)ethanol | 408-090-7 | 100418-33-5 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 3 | H302 H317 H412 | GHS07 Wng | H302 H317 H412 |  |  |  | CLP00
603-127-00-5 | butan-2-ol [1] (S)-butan-2-ol [2] (R)-butan-2-ol [3] (±)-butan-2-ol [4] | 201-158-5 [1] 224-168-1 [2] 238-967-8 [3] 240-029-8 [4] | 78-92-2 [1] 4221-99-2 [2] 14898-79-4 [3] 15892-23-6 [4] | Flam. Liq. 3 STOT SE 3 STOT SE 3 Eye Irrit. 2 | H226 H335 H336 H319 | GHS02 GHS07 Wng | H226 H319 H335 H336 |  |  | C | CLP00/ATP01
603-128-00-0 | 2-(phenylmethoxy)naphthalene | 405-490-3 | 613-62-7 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
603-129-00-6 | 1-tert-butoxypropan-2-ol | 406-180-0 | 57018-52-7 | Flam. Liq. 3 Eye Dam. 1 | H226 H318 | GHS02 GHS05 Dgr | H226 H318 |  |  |  | CLP00
603-130-00-1 | reaction mass of isomers of: α-((dimethyl)biphenyl)-ω-hydroxypoly(oxyethylene) | 406-325-8 |  | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
603-131-00-7 | reaction mass of: 1-deoxy-1-[methyl-(1-oxododecyl)amino]-D-glucitol; 1-deoxy-1-[methyl-(1-oxotetradecyl)amino]-D-glucitol (3:1) | 407-290-1 |  | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
603-132-00-2 | 2-hydroxymethyl-9-methyl-6-(1-methylethyl)-1,4-dioxaspiro[4.5]decane | 408-200-3 | 63187-91-7 | Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H315 H318 H412 | GHS05 Dgr | H315 H318 H412 |  |  |  | CLP00
603-133-00-8 | reaction mass of: 3-[(4-amino-2-chloro-5-nitrophenyl)amino]-propane-1,2-diol; 3,3'-(2-chloro-5-nitro-1,4-phenylenediimino)bis(propan-1,2-diol) | 408-240-1 |  | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
603-134-00-3 | reaction mass of substituted dodecyl and/or tetradecyl, diphenyl ethers. The substance is produced by the Friedel Crafts reaction. The catalyst is removed from the reaction product. Diphenyl ether is substituted by C1-C10 alkyl groups. The alkyl groups are bonded randomly between C1 and C6. Linear C12 and C14, 50/50 used. | 410-450-3 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
603-135-00-9 | bis[[2,2',2''-nitrilotris-[ethanolato]]-1-N,O]-bis[2-(2-methoxyethoxy)ethoxy]-titanium | 410-500-4 |  | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
603-136-00-4 | 3-((4-(bis(2-hydroxyethyl)amino)-2-nitrophenyl)amino)-1-propanol | 410-910-3 | 104226-19-9 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
603-137-00-X | reaction mass of: 1-deoxy-1-[methyl-(1-oxohexadecyl)amino]-D-glucitol; 1-deoxy-1-[methyl-(1-oxooctadecyl)amino]-D-glucitol | 411-130-6 |  | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
603-138-00-5 | 3-(2,2-dimethyl-3-hydroxypropyl)toluene; (alt.): 2,2-dimethyl-3-(3-methylphenyl)propanol | 403-140-4 | 103694-68-4 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
603-139-00-0 | bis(2-methoxyethyl) ether | 203-924-4 | 111-96-6 | Flam. Liq. 3 Repr. 1B | H226 H360FD | GHS02 GHS08 Dgr | H226 H360FD | EUH019 |  |  | CLP00
603-140-00-6 | 2,2' -oxybisethanol; diethylene glycol | 203-872-2 | 111-46-6 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
603-141-00-1 | reaction mass of: dodecyloxy-1-methyl-1-[oxy-poly-(2-hydroxymethylethanoxy)]pentadecane; dodecyloxy-1-methyl-1-[oxy-poly-(2-hydroxymethylethanoxy)]heptadecane | 413-780-6 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
603-142-00-7 | 2-(2-(2-hydroxyethoxy)ethyl)-2-aza-bicyclo[2.2.1]heptane | 407-360-1 | 116230-20-7 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 | H312 H302 H373 ** H315 H318 | GHS06 GHS08 GHS05 Dgr | H312 H302 H373 ** H315 H318 |  |  |  | CLP00
603-143-00-2 | R-2,3-epoxy-1-propanol | 404-660-4 | 57044-25-4 | Self-react. C **** Carc. 1B Muta. 2 Repr. 1B Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H242 H350 H341 H360F *** H331 H312 H302 H314 | GHS02 GHS06 GHS08 GHS05 Dgr | H242 H350 H341 H360F *** H331 H312 H302 H314 |  |  |  | CLP00
603-144-00-8 | reaction mass of: 2,6,9-trimethyl-2,5,9-cyclododecatrien-1-ol; 6,9-dimethyl-2-methylen-5,9-cyclododecadien-1-ol | 413-530-6 | 111850-00-1 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
603-145-00-3 | 2-isopropyl-2-(1-methylbutyl)-1,3-dimethoxypropane | 406-970-5 | 129228-11-1 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
603-146-00-9 | 2-[(2-[2-(dimethylamino)ethoxy]ethyl)methylamino]ethanol | 406-080-7 | 83016-70-0 | Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 3 | H302 H314 H412 | GHS05 GHS07 Dgr | H302 H314 H412 |  |  |  | CLP00
603-147-00-4 | (-)-trans-4-(4'-fluorophenyl)-3-hydroxymethyl-N-methylpiperidine | 406-030-4 | 105812-81-5 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | CLP00
603-148-00-X | 1,4-bis[(vinyloxy)methyl]cyclohexane | 413-370-7 | 17351-75-6 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
603-149-00-5 | reaction mass of diastereoisomers of 1-(1-hydroxyethyl)-4-(1-methylethyl)cyclohexane | 407-640-3 | 63767-86-2 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H315 H319 H411 | GHS07 GHS09 Wng | H319 H315 H411 |  |  |  | CLP00
603-150-00-0 | (±) trans-3,3-dimethyl-5-(2,2,3-trimethyl-cyclopent-3-en-1-yl)-pent-4-en-2-ol | 411-580-3 | 107898-54-4 | Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H400 H410 | GHS07 GHS09 Wng | H315 H410 |  |  |  | CLP00
603-151-00-6 | (±)-2-(2,4-dichlorophenyl)-3-(1H-1,2,4-triazol-1-yl)propan-1-ol | 413-570-4 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
603-152-00-1 | 2-(4-tert-butylphenyl)ethanol | 410-020-5 | 5406-86-0 | Repr. 2 STOT RE 2 * Eye Dam. 1 Aquatic Chronic 2 | H361f *** H373 ** H318 H411 | GHS08 GHS05 GHS09 Dgr | H361f *** H373 ** H318 H411 |  |  |  | CLP00
603-153-00-7 | 3-((2-nitro-4-(trifluoromethyl)phenyl)amino)propane-1,2-diol | 410-010-0 | 104333-00-8 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
603-154-00-2 | 1-[(2-tert-butyl)cyclohexyloxy]-2-butanol | 412-300-2 | 139504-68-0 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
603-156-00-3 | 2-(2,4-dichlorophenyl)-2-(2-propenyl)oxirane | 411-210-0 | 89544-48-9 | Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H317 H400 H410 | GHS07 GHS09 Wng | H315 H317 H410 |  |  |  | CLP00
603-157-00-9 | 6,9-bis(hexadecyloxymethyl)-4,7-dioxanonane-1,2,9-triol | 411-450-6 | 143747-72-2 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
603-158-00-4 | reaction mass of 4 diastereoisomers of 2,7-dimethyl-10-(1-methylethyl)-1-oxaspiro[4.5]deca-3,6-diene | 412-460-3 |  | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
603-159-00-X | 2-cyclododecylpropan-1-ol | 411-410-8 | 118562-73-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
603-160-00-5 | 1,2-diethoxypropane | 412-180-1 | 10221-57-5 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 | EUH019 |  |  | CLP00
603-161-00-0 | 1,3-diethoxypropane | 413-140-6 | 3459-83-4 | Flam. Liq. 3 | H226 | GHS02 Wng | H226 |  |  |  | CLP00
603-162-00-6 | α[2-[[[(2-hydroxyethyl)methylamino]acetyl]amino]propyl]-ω-(nonylphenoxy)poly[oxo(methyl-1,2-ethanediyl)] | 413-420-8 | 144736-29-8 | Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H314 H317 H411 | GHS05 GHS07 GHS09 Dgr | H314 H317 H411 |  |  |  | CLP00
603-163-00-1 | 2-phenyl-1,3-propanediol | 411-810-2 | 1570-95-2 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
603-164-00-7 | 2-butyl-4-chloro-4,5-dihydro-5-hydroxymethyl-1-[2'-(2-triphenylmethyl-1,2,3,4-2H-tetrazol-5-yl)-1,1'-biphenyl-4-methyl]-1H-imidazole | 412-420-5 | 133909-99-6 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
603-165-00-2 | reaction mass of: 4-allyl-2,6-bis(2,3-epoxypropyl)phenol; 4-allyl-6-[3-[6-[3-[6-[3-(4-allyl-2,6-bis(2,3-epoxypropyl)phenoxy)-2-hydroxypropyl]-4-allyl-2-(2,3-epoxypropyl)phenoxy]-2-hydroxypropyl]-4-allyl-2-(2,3-epoxypropyl)phenoxy]-2-hydroxypropyl]-2-(2,3-epoxypropyl)phenol; 4-allyl-6-[3-(4-allyl-2,6-bis(2,3-epoxypropyl)phenoxy)-2-hydroxypropyl]-2-(2,3-epoxypropyl)phenol; 4-allyl-6-[3-[6-[3-(4-allyl-2,6-bis(2,3-epoxypropyl)phenoxy)-2-hydroxypropyl]-4-allyl-2-(2,3-epoxypropyl)phenoxy]-2-hydroxypropyl]-2-(2,3-epoxypropyl)phenol | 417-470-1 |  | Muta. 2 Skin Sens. 1 | H341 H317 | GHS08 GHS07 Wng | H341 H317 |  |  |  | CLP00
603-166-00-8 | R-1-chloro-2,3-epoxypropane | 424-280-2 | 51594-55-9 | Flam. Liq. 3 Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Skin Sens. 1 | H226 H350 H331 H311 H301 H314 H317 | GHS02 GHS06 GHS08 GHS05 Dgr | H226 H350 H331 H311 H301 H314 H317 |  |  |  | CLP00
603-167-00-3 | 3,3',5,5'-tetra-tert-butylbiphenyl-2,2'-diol | 407-920-5 | 6390-69-8 | Aquatic Chronic 4 | H413 | GHS05 Dgr | H413 |  |  |  | CLP00
603-168-00-9 | 3-(2-ethylhexyloxy)propane-1,2-diol | 408-080-2 | 70445-33-9 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
603-169-00-4 | (±)-trans-4-(4-fluorophenyl)-3-hydroxymethyl-N-methylpiperidine | 415-550-0 | 109887-53-8 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | CLP00
603-170-00-X | reaction mass of: 2-methyl-1-(6-methylbicyclo[2.2.1]hept-5-en-2-yl)pent-1-en-3-ol; 2-methyl-1-(1-methylbicyclo[2.2.1]hept-5-en-2-yl)-pent-1-en-3-ol; 2-methyl-1-(5-methylbicyclo[2.2.1]hept-5-en-2-yl)pent-1-en-3-ol | 415-990-3 | 67739-11-1 | Eye Irrit. 2 Aquatic Chronic 2 | H319 H411 | GHS07 GHS09 Wng | H319 H411 |  |  |  | CLP00
603-171-00-5 | 5-thiazolylmethanol | 414-780-9 | 38585-74-9 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
603-172-00-0 | mono-2-[2-(4-dibenzo[b,f][1,4]thiazepin-11-yl)piperazinium-1-yl]ethoxy)ethanol trans-butenedioate | 415-180-1 | 773058-82-5 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | CLP00
603-173-00-6 | 4,4-dimethyl-3,5,8-trioxabicyclo[5.1.0]octane | 421-750-9 | 57280-22-5 | Eye Irrit. 2 Skin Sens. 1 | H319 H317 | GHS07 Wng | H319 H317 |  |  |  | CLP00
603-174-00-1 | 4-cyclohexyl-2-methyl-2-butanol | 420-630-3 | 83926-73-2 | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
603-175-00-7 | 2-(2-hexyloxyethoxy)ethanol; DEGHE; diethylene glycol monohexyl ether; 3,6-dioxa-1-dodecanol; hexyl carbitol; 3,6-dioxadodecan-1-ol | 203-988-3 | 112-59-4 | Acute Tox. 4 * Eye Dam. 1 | H312 H318 | GHS05 GHS07 Dgr | H312 H318 |  |  |  | CLP00
603-176-00-2 | 1,2-bis(2-methoxyethoxy)ethane; TEGDME; triethylene glycol dimethyl ether; triglyme | 203-977-3 | 112-49-2 | Repr. 1B | H360Df | GHS08 Dgr | H360Df | EUH019 |  |  | CLP00
603-177-00-8 | 1-ethoxypropan-2-ol; 2PG1EE; 1-ethoxy-2-propanol; propylene glycol monoethyl ether [1] 2-ethoxy-1-methylethyl acetate; 2PG1EEA [2] | 216-374-5 [1] 259-370-9 [2] | 1569-02-4 [1] 54839-24-6 [2] | Flam. Liq. 3 STOT SE 3 | H226 H336 | GHS02 GHS07 Wng | H226 H336 |  |  |  | CLP00
603-178-00-3 | 2-hexyloxyethanol; ethylene glycol monohexyl ether; n-hexylglycol | 203-951-1 | 112-25-4 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H312 H302 H314 | GHS05 GHS07 Dgr | H312 H302 H314 |  |  |  | CLP00
603-179-00-9 | ergocalciferol (ISO); Vitamin D2 | 200-014-9 | 50-14-6 | Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 | H330 H311 H301 H372 ** | GHS06 GHS08 Dgr | H330 H311 H301 H372 ** |  |  |  | CLP00
603-180-00-4 | colecalciferol; cholecalciferol; Vitamin D3 | 200-673-2 | 67-97-0 | Acute Tox. 2 Acute Tox. 2 Acute Tox. 2 STOT RE 1 | H330 H310 H300 H372 | GHS06 GHS08 Dgr | H330 H310 H300 H372 |  | Inhalation: ATE = 0.05 mg/L (dusts/mists) Dermal: ATE = 50 mg/kg bw Oral: ATE = 35 mg/kg bw STOT RE 1; H372: C ≥ 3 % STOT RE 2; H373: 0,3 % ≤ C < 3 % |  | CLP00/ATP13
603-181-00-X | tert-butyl methyl ether; MTBE; 2-methoxy-2-methylpropane | 216-653-1 | 1634-04-4 | Flam. Liq. 2 Skin Irrit. 2 | H225 H315 | GHS02 GHS07 Dgr | H225 H315 |  |  |  | CLP00
603-182-00-5 | reaction product of: saturated, monounsaturated and multiple unsaturated long-chained partly estrified alcohols of vegetable origin (Brassica napus L., Brassica rapa L., Helianthus annuus L., Glycine hispida, Gossypium hirsutum L., Cocos nucifera L., Elaeis guineensis) with O,O-diisobutyldithiophosphate and 2-ethylhexylamine and hydrogen peroxide | 428-630-5 | - | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
603-183-00-0 | 2-[2-(2-butoxyethoxy)ethoxy]ethanol; TEGBE; triethylene glycol monobutyl ether; butoxytriethylene glycol | 205-592-6 | 143-22-6 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  | Eye Dam. 1; H318: C ≥ 30 % Eye Irrit. 2; H319: 20 % ≤ C < 30 % |  | CLP00
603-184-00-6 | 2-(hydroxymethyl)-2-[[2-hydroxy-3-(isooctadecyloxy)propoxy]methyl]-1,3-propanediol | 416-380-1 | 146925-83-9 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
603-185-00-1 | 2,4-dichloro-3-ethyl-6-nitrophenol | 420-740-1 | 99817-36-4 | Acute Tox. 3 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H318 H317 H400 H410 | GHS06 GHS05 GHS09 Dgr | H301 H318 H317 H410 |  |  |  | CLP00
603-186-00-7 | trans-(5RS,6SR)-6-amino-2,2-dimethyl-1,3-dioxepan-5-ol | 419-050-3 | 79944-37-9 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
603-187-00-2 | 2-((4,6-bis(4-(2-(1-methylpyridinium-4-yl)vinyl)phenylamino)-1,3,5-triazin-2-yl)(2-hydroxyethyl)amino)ethanol dichloride | 419-360-9 | 163661-77-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
603-188-00-8 | reaction mass of: 6,7-epoxy-1,2,3,4,5,6,7,8-octahydro-1,1,2,4,4,7-hexamethylnaphthalene; 7,8-epoxy-1,2,3,4,6,7,8,8a-octahydro-1,1,2,4,4,7-hexamethylnaphthalene | 426-970-9 | - | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
603-189-00-3 | reaction mass of complexes of: titanium, 2,2'-oxydiethanol, ammonium lactate, nitrilotris(2-propanol) and ethylene glycol | 405-250-8 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
603-190-00-9 | 8,8-dimethyl-7-isopropyl-6,10-dioxaspiro[4.5]decane | 424-030-2 | 62406-73-9 | Skin Irrit. 2 Aquatic Chronic 3 | H315 H412 | GHS07 Wng | H315 H412 |  |  |  | ATP01
603-191-00-4 | 2-(4,6-bis(2,4-dimethylphenyl)-1,3,5-triazin-2-yl)-5-(3-((2-ethylhexyl)oxy)-2-hydroxypropoxy)phenol | 419-740-4 | 137658-79-8 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
603-192-00-X | (E,E)-3,7,11-trimethyldodeca-1,4,6,10-tetraen-3-ol | 423-240-1 | 125474-34-2 | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H315 H318 H317 H410 |  |  |  | ATP01
603-193-00-5 | disodium 9,10-anthracenedioxide | 426-030-8 | 46492-07-3 | Skin Corr. 1A | H314 | GHS05 Dgr | H314 |  |  |  | ATP01
603-194-00-0 | 2-(2-aminoethylamino)ethanol; (AEEA) | 203-867-5 | 111-41-1 | Repr. 1B Skin Corr. 1B Skin Sens. 1 | H360Df H314 H317 | GHS05 GHS08 GHS07 Dgr | H314 H317 H360Df |  | STOT SE 3; H335: C ≥ 5 % |  | ATP01/ATP01corr
603-195-00-6 | 2-[4-(4-methoxyphenyl)-6-phenyl-1,3,5-triazin-2-yl]-phenol | 430-810-3 | 154825-62-4 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
603-196-00-1 | 2-(7-ethyl-1H-indol-3-yl)ethanol | 431-020-1 | 41340-36-7 | Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 2 | H302 H373 ** H411 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H411 |  |  |  | CLP00
603-197-00-7 | tebuconazole (ISO); 1-(4-chlorophenyl)-4,4-dimethyl-3-(1,2,4-triazol-1-ylmethyl)pentan-3-ol | 403-640-2 | 107534-96-3 | Repr. 2 Acute Tox. 4 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H302 H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H361d *** H410 |  | M=1 M=10 |  | CLP00/ATP07
603-199-00-8 | etoxazol (ISO); (RS)-5-tert-butyl-2-[2-(2,6-difluorophenyl)-4,5-dihydro-1,3-oxazol-4-yl]phenetole |  | 153233-91-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=100 |  | CLP00
603-200-00-1 | 1-pentanol [1] 3-pentanol [2] | 200-752-1 [1] 209-526-7 [2] | 71-41-0 [1] 584-02-1 [2] | Flam. Liq. 3 Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 | H226 H332 H335 H315 | GHS02 GHS07 Wng | H226 H332 H335 H315 |  |  |  | ATP01
603-201-00-7 | (E)-(7R,11R)-3,7,11,15-tetramethylhexadec-2-ene-1-ol | 416-120-5 | - | Skin Irrit. 2 Aquatic Chronic 4 | H315 H413 | GHS07 Wng | H315 H413 |  |  |  | ATP01
603-202-00-2 | 4,4,5,5,5-pentafluoropentan-1-ol | 421-360-9 | 148043-73-6 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | ATP01
603-203-00-8 | (1R,3S,7R,8R,10R,13R)-5,5,7,9,9,13-hexamethyl-4,6-dioxatetracyclo[6.5.1.01,10.03,7]tetradecane | 427-580-1 | - | Skin Irrit. 2 | H315 | GHS07 Wng | H315 |  |  |  | ATP01
603-204-00-3 | reaction mass of: 2,2'-(heptane-1,7-diyl)bis-1,3-dioxolane; 2,2'-(heptane-1,6-diyl)bis-1,3-dioxolane | 428-110-8 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
603-205-00-9 | (1S-cis)-4-(2-amino-6-chloro-9H-purin-9-yl)-2-cyclopentene-1-methanol hydrochloride | 426-200-1 | 172015-79-1 | Acute Tox. 4 * STOT RE 1 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H302 H372 ** H318 H317 H412 | GHS05 GHS08 GHS07 Dgr | H372 ** H302 H318 H317 H412 |  |  |  | ATP01
603-206-00-4 | 2,2-dichloro-1,3-benzodioxol | 426-850-6 | 2032-75-9 | Flam. Liq. 3 Acute Tox. 4 * Skin Corr. 1A Skin Sens. 1 | H226 H302 H314 H317 | GHS02 GHS05 GHS07 Dgr | H226 H314 H302 H317 | EUH014 |  |  | ATP01
603-207-00-X | 2-isobutyl-2-isopropyl-1,3-dimethoxypropane | 430-800-9 | 129228-21-3 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | ATP01
603-208-00-5 | 1,2-diethoxyethane | 211-076-1 | 629-14-1 | Flam. Liq. 2 Repr. 1A Eye Irrit. 2 | H225 H360Df H319 | GHS02 GHS08 GHS07 Dgr | H225 H360Df H319 | EUH019 |  |  | ATP01
603-209-00-0 | spinosad (ISO) (reaction mass of spinosyn A and spinosyn D in ratios between 95:5 to 50:50); reaction mass of 50-95% of (2R,3aS,5aR,5bS,9S,13S,14R,16aS,16bR)-2-(6-deoxy-2,3,4-tri-O-methyl-α-.sc.l.sc.-mannopyranosyloxy)-13-(4-dimethylamino-2,3,4,6-tetradeoxy-β-.sc.d.sc.-erythropyranosyloxy)-9-ethyl-2,3,3a,5a,5b,6,7,9,10,11,12,13,14,15,16a,16b-hexadecahydro-14-methyl-1H-8-oxacyclododeca[b]as-indacene-7,15-dione and 50-5% (2S,3aR,5aS,5bS,9S,13S,14R,16aS,16bS)-2-(6-deoxy-2,3,4-tri-O-methyl-α-.sc.l.sc.-mannopyranosyloxy)-13-(4-dimethylamino-2,3,4,6-tetradeoxy-β-.sc.d.sc.-erythropyranosyloxy)-9-ethyl-2,3,3a,5a,5b,6,7,9,10,11,12,13,14,15,16a,16b-hexadecahydro-4,14-dimethyl-1H-8-oxacyclododeca[b]as-indacene-7,15-dione [1] spinosyn A [2] spinosyn D [3] |  | 131929-60-7 [2] 131929-63-0 [3] | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=10 |  | ATP01
603-210-00-6 | 2,4-diethyl-1,5-pentanediol | 429-310-8 | 57987-55-0 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
603-211-00-1 | 2,3-epoxypropyltrimethylammonium chloride ...%; glycidyl trimethylammonium chloride ...% | 221-221-0 | 3033-77-0 | Carc. 1B Muta. 2 Repr. 2 Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H350 H341 H361f *** H312 H302 H373 ** H318 H317 H412 | GHS05 GHS08 GHS07 Dgr | H350 H341 H361f *** H312 H302 H373 ** H318 H317 H412 |  |  | B | ATP01
603-212-00-7 | 1,3,4,6,7,8-hexahydro-4,6,6,7,8,8-hexamethylindeno[5,6-c]pyran; galaxolide; (HHCB) | 214-946-9 | 1222-05-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
603-213-00-2 | 2-methoxy-2-methylbutane; tert-amyl methyl ether | 213-611-4 | 994-05-8 | Flam. Liq. 2 Acute Tox. 4 * STOT SE 3 | H225 H302 H336 | GHS02 GHS07 Dgr | H225 H302 H336 |  |  |  | ATP01
603-214-00-8 | 1,1-diisopropoxycyclohexane | 413-740-8 | 1132-95-2 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  |  |  | ATP01
603-215-00-3 | 1-hydroxy-4-fluoro-1,4-diazoniabicyclo[2.2.2]octane bis(tetrafluoroborate) | 418-330-2 | 162241-33-0 | Expl. 1.1 **** Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H201 H302 H373 ** H318 H317 H400 H410 | GHS01 GHS05 GHS08 GHS07 GHS09 Dgr | H201 H302 H373 ** H318 H317 H410 |  |  |  | ATP01
603-216-00-9 | cis-1-amino-2,3-dihydro-1H-inden-2-ol | 422-660-2 | 7480-35-5 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | ATP01
603-217-00-4 | 2,4,6-tri-tert-butylphenyl 2-butyl-2-ethyl-1,3-propanediolphosphite | 423-560-1 | 161717-32-4 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
603-220-00-0 | 1-{benzyl[2-(2-methoxyphenoxy)ethyl]amino}-3-(9H-carbazol-4-yloxy)propan-2-ol | 432-890-5 | 72955-94-3 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
603-221-00-6 | 1-(2-amino-5-chlorophenyl)-2,2,2-trifluoro-1,1-ethanediol, hydrochloride; [containing < 0.1 % 4-chloroaniline (EC No 203-401-0)] | 433-580-2 | 214353-17-0 | Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 2 | H302 H314 H411 | GHS05 GHS07 GHS09 Dgr | H302 H314 H411 |  |  |  | ATP01
603-221-01-3 | 1-(2-amino-5-chlorophenyl)-2,2,2-trifluoro-1,1-ethanediol, hydrochloride; [containing ≥ 0.1 % 4-chloroaniline (EC No 203-401-0)] | 433-580-2 | 214353-17-0 | Carc. 1B Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 2 | H350 H302 H314 H411 | GHS05 GHS08 GHS07 GHS09 Dgr | H350 H302 H314 H411 |  |  |  | ATP01
603-222-00-1 | (2R,3S,4R,5R,7R,9R,10R,11S,12S,13R)-10-[(4-dimethylamino-3-hydroxy-6-methyltetrahydropyran-2-yl)oxy]-2-ethyl-3,4,12-trihydroxy-9-methoxy-3,5,7,9,11,13-hexamethyl-6,14-dioxo-1-oxacyclotetradecane | 433-820-6 | 118058-74-5 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | ATP01
603-223-00-7 | 2-cyclopentylidene cyclopentanol; 1,1'-bi(cyclopentyliden)-2-ol | 434-270-1 | 6261-30-9 | Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H315 H318 H412 | GHS05 Dgr | H315 H318 H412 |  |  |  | ATP01
603-224-00-2 | 3-ethoxy-1,1,1,2,3,4,4,5,5,6,6,6-dodecafluoro-2-(trifluoromethyl)-hexane | 435-790-1 | 297730-93-9 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
603-225-00-8 | erythromycin A9-oxime (E); (3R,4S,5S,6R,7R,9R,11R,12R,13S,14R)-4-((2,6-didesoxy-3-C-methyl-3-O-methyl-α-L-ribo-hexopiranosyl)oxy)-14-ethyl-7,12,13-trihydroxy-3,5,7,9,11,13-hexamethyl-6-((3,4,6-tridesoxy-3-dimethylamino-β-.sc.d.sc.-xylohexapiranosyl)oxy)oxacyclotetradecan-2-ona-10-oxime (E) | 437-070-0 | 13127-18-9 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
603-226-00-3 | 4,4'(4-(4-methoxyphenyl)-1,3,5-triazin-2,4-diyl)bisbenzene-1,3-diol | 444-500-0 | 1440-00-2 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
603-227-00-9 | α-hydro-ω-[[[(1,1-dimethylethyl)dioxy]carbonyl]oxy]-poly[oxy(methyl-1,2-ethanediyl)] ether with 2,2-bis(hydroxymethyl)-1,3-propanediol (4:1); reaction product of: α-hydro-ω-((chlorocarbonyl)oxy)-poly(oxy(methyl-1,2-ethanediyl)) ether with 2,2-bis(hydroxymethyl)-1,3-propanediol with potassium 1,1-dimethylethylperoxalate | 445-060-2 | 203574-04-3 | **** Aquatic Acute 1 Aquatic Chronic 1 | **** H400 H410 | **** GHS09 Wng | **** H410 |  |  |  | ATP01
603-228-00-4 | (+/-)-(R*,R*)-6-fluoro-3,4-dihydro-2-oxiranyl-2H-1-benzopyran; 6-fluoro-2-(2-oxiranyl)chromane | 419-620-1 | - | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
603-229-00-X | sodium (Z)-3-chloro-3-(4-chlorophenyl)-1-hydroxy-2-propene-1-sulfonate | 420-800-7 | - | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H315 H318 H317 H410 |  |  |  | ATP01
603-230-00-5 | 2,6,6,7,8,8-hexamethyldecahydro-2H-indeno[4,5-b]furan | 440-030-5 | - | Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 4 | H315 H318 H413 | GHS05 Dgr | H315 H318 H413 |  |  |  | ATP01
603-231-00-0 | (S)-1,1-diphenyl-1,2-propanediol | 443-220-6 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
603-232-00-6 | 3,3,8,8,10,10-hexamethyl-9-[1-(4-oxiranylmethoxy-phenyl)-ethoxy]-1,5-dioxa-9-aza-spiro[5.5]undecane | 444-420-6 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
603-233-00-1 | reaction mass of: 4-(1,3a,4,6,7,7a-hexahydro-4,7-methanoinden-5-ylidene)-3-methylbutan-2-ol; 4-(3,3a,4,6,7,7a-hexahydro-4,7-methanoinden-5-ylidene)-3-methylbutan-2-ol; 1-(1,3a,4,6,7,7a-hexahydro-4,7-methanoinden-5-ylidene)pentan-3-ol; 1-(3,3a,4,6,7,7a-hexahydro-4,7-methanoinden-5-ylidene)pentan-3-ol; (E)-4-(3a,4,5,6,7,7a-hexahydro-1H-4,7-methanoinden-5-yl)-3-methylbut-3-en-2-ol; (E)-4-(3a,4,5,6,7,7a-hexahydro-3H-4,7-methanoinden-5-yl)-3-methylbut-3-en-2-ol | 444-430-0 | - | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
603-234-00-7 | (1R,4R)-4-methoxy-2,2,7,7-tetramethyltricyclo(6.2.1.0(1,6))undec-5-ene | 444-480-3 | - | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | ATP01
603-235-00-2 | linalool; 3,7-dimethyl-1,6-octadien-3-ol; dl-linalool [1] coriandrol; (S)-3,7-dimethyl-1,6-octadien-3-ol; d-linalool [2] licareol; (R)-3,7-dimethyl-1,6-octadien-3-ol; l-linalool [3] | 201-134-4 [1] 204-810-7 [2] 204-811-2 [3] | 78-70-6 [1] 126-90-9 [2] 126-91-0 [3] | Skin Sens. 1B | H317 | GHS07 Wng | H317 |  |  |  | ATP10
603-236-00-8 | ethanol, 2,2'-iminobis-, N-(C13-15-branched and linear alkyl) derivs. | 308-208-6 | 97925-95-6 | Repr. 1B | H360D | GHS08 Dgr | H360D |  |  |  | ATP14
603-237-00-3 | ipconazole (ISO); (1RS,2SR,5RS;1RS,2S­R,5SR)-2-(4-chloro­benzyl)-5-isopropyl-1- (1H-1,2,4-triazol-1-yl­methyl)cyclopentanol |  | 125225-28-7 | Repr. 1B Acute Tox. 4 STOT RE 2 Aquatic Chronic 1 | H360D H302 H373 (eyes, skin, liver) H410 | GHS08 GHS07 GHS09 Dgr | H302 H360D H373 (eyes, skin, liver) H410 |  | Oral: ATE = 500 mg/kg bw M=100 |  | ATP15
603-238-00-9 | bis(2-(2-methoxyethoxy)ethyl)ether; tetraglyme | 205-594-7 | 143-24-8 | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  |  | ATP15
603-239-00-4 | paclobutrazol (ISO); (2RS,3RS)-1-(4-chloro­phenyl)-4,4-dimethyl- 2-(1H-1,2,4-triazol-1-yl)pentan-3-ol |  | 76738-62-0 | Repr. 2 Acute Tox. 4 Acute Tox. 4 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H361d H332 H302 H319 H400 H410 | GHS08 GHS07 GHS09 Wng | H332 H302 H319 H361d H410 |  | Inhalation: ATE = 3.13 mg/L (dusts/mists) Oral: ATE = 490 mg/kg bw M=10 M=10 |  | ATP15
603-240-00-X | 2,2-bis(bromomethyl) propane-1,3-diol | 221-967-7 | 3296-90-0 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H340 H350 |  |  |  | ATP15
603-241-00-5 | geraniol; (2E)-3,7-dimethylocta-2,6-dien-1-ol | 203-377-1 | 106-24-1 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP15
603-243-00-6 | 2,2-dimethylpropan-1-ol, tribromo derivative; 3-bromo-2,2-bis(bromomethyl)propan-1-ol | 253-057-0 | 36483-57-5; 1522-92-5 | Carc. 1B Muta. 2 | H350 H341 | GHS08 Dgr | H350 H341 |  |  |  | ATP18
603-244-00-1 | reaction mass of 1-(2,3-epoxypropoxy)–2,2-bis ((2,3-epoxypropoxy)methyl) butane and 1-(2,3-epoxypropoxy)-2-((2,3-epoxypropoxy)methyl)-2-hydroxymethyl butane |  |  | Muta. 2 Repr. 1B | H341 H360F | GHS08 Dgr | H341 H360F |  |  |  | ATP21
603-245-00-7 | 2,2'-[[3-methyl-4-[(4-nitrophenyl)azo]phenyl]imino]bisethanol | 221-665-5 | 3179-89-3 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP21
603-246-00-2 | 3,3,4,4,5,5,6,6,7,7,8,8,8-tridecafluorooctan-1-ol | 211-477-1 | 647-42-7 | STOT RE 2 Aquatic Chronic 1 | H373 (teeth, bones) H410 | GHS08 GHS09 Wng | H373 (teeth, bones) H410 |  | M = 1 |  | ATP21
603-247-00-8 | reaction mass of 1,3-dioxan-5-ol and 1,3-dioxolan-4-ylmethanol | - | - | Repr. 1B | H360Df | GHS08 Dgr | H360Df |  |  |  | ATP22
604-001-00-2 | phenol; carbolic acid; monohydroxybenzene; phenylalcohol | 203-632-7 | 108-95-2 | Muta. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Skin Corr. 1B | H341 H331 H311 H301 H373 ** H314 | GHS06 GHS08 GHS05 Dgr | H341 H331 H311 H301 H373 ** H314 |  | * Skin Corr. 1B; H314: C ≥ 3 % Skin Irrit. 2; H315: 1 % ≤ C < 3 % Eye Irrit. 2; H319: 1 % ≤ C < 3 % |  | CLP00
604-002-00-8 | pentachlorophenol | 201-778-6 | 87-86-5 | Carc. 2 Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H330 H311 H301 H335 H315 H319 H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H330 H311 H301 H319 H335 H315 H410 |  |  |  | CLP00
604-003-00-3 | sodium pentachlorophenolate [1] potassium pentachlorophenolate [2] | 205-025-2 [1] 231-911-3 [2] | 131-52-2 [1] 7778-73-6 [2] | Carc. 2 Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H330 H311 H301 H335 H315 H319 H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H330 H311 H301 H319 H335 H315 H410 |  |  |  | CLP00
604-004-00-9 | m-cresol [1] o-cresol [2] p-cresol [3] mix-cresol [4] | 203-577-9 [1] 202-423-8 [2] 203-398-6 [3] 215-293-2 [4] | 108-39-4 [1] 95-48-7 [2] 106-44-5 [3] 1319-77-3 [4] | Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B | H311 H301 H314 | GHS06 GHS05 Dgr | H311 H301 H314 |  | * | C | CLP00
604-005-00-4 | 1,4-dihydroxybenzene; hydroquinone; quinol | 204-617-8 | 123-31-9 | Carc. 2 Muta. 2 Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 | H351 H341 H302 H318 H317 H400 | GHS05 GHS08 GHS07 GHS09 Dgr | H351 H341 H302 H318 H317 H400 |  | M=10 |  | CLP00/ATP01
604-006-00-X | 3,4-xylenol [1] 2,5-xylenol [2] 2,4-xylenol [3] 2,3-xylenol [4] 2,6-xylenol [5] xylenol [6] 2,4(or 2,5)-xylenol [7] | 202-439-5 [1] 202-461-5 [2] 203-321-6 [3] 208-395-3 [4] 209-400-1 [5] 215-089-3 [6] 276-245-4 [7] | 95-65-8 [1] 95-87-4 [2] 105-67-9 [3] 526-75-0 [4] 576-26-1 [5] 1300-71-6 [6] 71975-58-1 [7] | Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Aquatic Chronic 2 | H311 H301 H314 H411 | GHS06 GHS05 GHS09 Dgr | H311 H301 H314 H411 |  |  | C | CLP00
604-007-00-5 | 2-naphthol | 205-182-7 | 135-19-3 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 | H332 H302 H400 | GHS07 GHS09 Wng | H332 H302 H400 |  |  |  | CLP00
604-008-00-0 | 2-chlorophenol [1] 4-chlorophenol [2] 3-chlorophenol [3] chlorophenol [4] | 202-433-2 [1] 203-402-6 [2] 203-582-6 [3] 246-691-4 [4] | 95-57-8 [1] 106-48-9 [2] 108-43-0 [3] 25167-80-0 [4] | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 2 | H332 H312 H302 H411 | GHS07 GHS09 Wng | H332 H312 H302 H411 |  |  | C | CLP00
604-009-00-6 | pyrogallol; 1,2,3-trihydroxybenzene | 201-762-9 | 87-66-1 | Muta. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 3 | H341 H332 H312 H302 H412 | GHS08 GHS07 Wng | H341 H332 H312 H302 H412 |  | * |  | CLP00
604-010-00-1 | resorcinol; 1,3-benzenediol | 203-585-2 | 108-46-3 | Acute Tox. 4 STOT SE 1 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1B Aquatic Acute 1 | H302 H370 (nervous system) H315 H319 H317 H400 | GHS08 GHS07 GHS09 Dgr | H302 H370 (nervous system) H315 H319 H317 H400 |  | oral: ATE = 500 mg/kg bw M = 1 |  | CLP00/ATP21
604-011-00-7 | 2,4-dichlorophenol | 204-429-6 | 120-83-2 | Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 2 | H311 H302 H314 H411 | GHS06 GHS05 GHS09 Dgr | H311 H302 H314 H411 |  |  |  | CLP00
604-012-00-2 | 4-chloro-o-cresol; 4-chloro-2-methyl phenol | 216-381-3 | 1570-64-5 | Acute Tox. 3 * Skin Corr. 1A Aquatic Acute 1 | H331 H314 H400 | GHS06 GHS05 GHS09 Dgr | H331 H314 H400 |  | STOT SE 3; H335: C ≥ 1 % |  | CLP00
604-013-00-8 | 2,3,4,6-tetrachlorophenol | 200-402-8 | 58-90-2 | Acute Tox. 3 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H301 H315 H319 H400 H410 | GHS06 GHS09 Dgr | H301 H319 H315 H410 |  | * Eye Irrit. 2; H319: C ≥ 5 % Skin Irrit. 2; H315: C ≥ 5 % |  | CLP00
604-014-00-3 | chlorocresol; 4-chloro-m-cresol; 4-chloro-3-methylphenol | 200-431-6 | 59-50-7 | Acute Tox. 4 STOT SE 3 Skin Corr. 1C Eye Dam. 1 Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 3 | H302 H335 H314 H318 H317 H400 H412 | GHS07 GHS05 GHS09 Dgr | H302 H314 H317 H335 H410 |  | M=1 |  | CLP00/ATP13
604-015-00-9 | 2,2'-methylenebis-(3,4,6-trichlorophenol); hexachlorophene | 200-733-8 | 70-30-4 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H311 H301 H400 H410 | GHS06 GHS09 Dgr | H311 H301 H410 |  | * |  | CLP00
604-016-00-4 | 1,2-dihydroxybenzene; pyrocatechol | 204-427-5 | 120-80-9 | Carc. 1B Muta. 2 Acute Tox. 3 Acute Tox. 3 Skin Irrit. 2 Eye Irrit. 2 | H350 H341 H311 H301 H315 H319 | GHS08 GHS06 Dgr | H311 H301 H315 H319 H341 H350 |  | Dermal: ATE = 600 mg/kg bw Oral: ATE = 300 mg/kg bw |  | CLP00/ATP13
604-017-00-X | 2,4,5-trichlorophenol | 202-467-8 | 95-95-4 | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H315 H410 |  | * Eye Irrit. 2; H319: C ≥ 5 % Skin Irrit. 2; H315: C ≥ 5 % |  | CLP00
604-018-00-5 | 2,4,6-trichlorophenol | 201-795-9 | 88-06-2 | Carc. 2 Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H315 H319 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H302 H319 H315 H410 |  |  |  | CLP00
604-019-00-0 | dichlorophen (ISO) | 202-567-1 | 97-23-4 | Acute Tox. 4 * Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H410 |  |  |  | CLP00
604-020-00-6 | biphenyl-2-ol; 2-phenylphenol; 2-hydroxybiphenyl | 201-993-5 | 90-43-7 | Carc. 2 Skin Corr. 1 Eye Dam. 1 Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 1 | H351 H314 H318 H317 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H351 H314 H317 H410 |  | M=1 M=1 |  | CLP00/ATP22
604-021-00-1 | sodium 2-biphenylate; 2-phenylphenol, sodium salt | 205-055-6 | 132-27-4 | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 | H302 H335 H315 H318 H400 | GHS05 GHS07 GHS09 Wng | H302 H335 H315 H318 H400 |  |  |  | CLP00
604-022-00-7 | 2,2-dimethyl-1,3-benzodioxol-4-ol | 400-900-7 | 22961-82-6 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
604-023-00-2 | 2,4-dichloro-3-ethylphenol | 401-060-4 |  | Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H314 H400 H410 | GHS05 GHS09 Dgr | H314 H410 |  |  |  | CLP00
604-024-00-8 | 4,4-isobutylethylidenediphenol | 401-720-1 | 6807-17-6 | Repr. 1B Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H360F *** H319 H400 H410 | GHS08 GHS09 Dgr | H360F *** H319 H410 |  |  |  | CLP00
604-025-00-3 | 2,5-bis(1,1-dimethylbutyl)hydroquinone | 400-220-0 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
604-026-00-9 | 2,2-spirobi(6-hydroxy-4,4,7-trimethylchromane) | 400-270-3 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
604-027-00-4 | 2-methyl-5-(1,1,3,3-tetramethylbutyl)hydroquinone | 400-530-6 |  | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H318 H317 H411 |  |  |  | CLP00
604-028-00-X | 4-amino-3-fluorophenol | 402-230-0 | 399-95-1 | Carc. 1B Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H350 H302 H317 H411 | GHS08 GHS07 GHS09 Dgr | H350 H302 H317 H411 |  |  |  | CLP00
604-029-00-5 | 1-naphtol | 201-969-4 | 90-15-3 | Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 | H312 H302 H335 H315 H318 | GHS05 GHS07 Dgr | H312 H302 H335 H315 H318 |  |  |  | CLP00
604-030-00-0 | 4,4'-isopropylidenediphenol; bisphenol A | 201-245-8 | 80-05-7 | Repr. 1B STOT SE 3 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H360F H335 H318 H317 H400 H410 | GHS08 GHS07 GHS05 GHS09 Dgr | H360F H335 H318 H317 H410 |  | M = 1 M = 10 |  | CLP00/ATP18
604-031-00-6 | guaiacol | 201-964-7 | 90-05-1 | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 | H302 H315 H319 | GHS07 Wng | H302 H319 H315 |  |  |  | CLP00
604-032-00-1 | thymol | 201-944-8 | 89-83-8 | Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 2 | H302 H314 H411 | GHS05 GHS07 GHS09 Dgr | H302 H314 H411 |  |  |  | CLP00
604-033-00-7 | isobutyl but-3-enoate | 401-170-2 | 24342-03-8 | Flam. Liq. 3 | H226 | GHS02 Wng | H226 |  |  |  | CLP00
604-034-00-2 | 4,4'-thiodi-o-cresol | 403-330-7 | 24197-34-0 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | CLP00
604-035-00-8 | 4-nonylphenol, reaction products with formaldehyde and dodecane-1-thiol | 404-160-6 |  | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
604-036-00-3 | 4,4'-oxybis(ethylenethio)diphenol | 404-590-4 | 90884-29-0 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
604-037-00-9 | 3,5-xylenol; 3,5-dimethylphenol | 203-606-5 | 108-68-9 | Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B | H311 H301 H314 | GHS06 GHS05 Dgr | H311 H301 H314 |  |  |  | CLP00
604-038-00-4 | 4-chloro-3,5-dimethylphenol [1] chloroxylenol [2] | 201-793-8 [1] 215-316-6 [2] | 88-04-0 [1] 1321-23-9 [2] | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H302 H315 H319 H317 | GHS07 Wng | H302 H319 H315 H317 |  |  |  | CLP00
604-039-00-X | ethyl 2-[4-[(6-chlorobenzoxazol-2-yl)oxy]phenoxy]propionate; fenoxaprop-ethyl | 266-362-9 | 66441-23-4 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
604-040-00-5 | fomesafen (ISO); 5-[2-chloro-4-(trifluoromethyl)phenoxy]-N-(methylsulphonyl)-2-nitrobenzamide | 276-439-9 | 72178-02-0 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
604-041-00-0 | acifluorfen (ISO); 5-[2-chloro-4-(trifluoromethyl)phenoxy]-2-nitrobenzoic acid [1] sodium 5-[2-chloro-4-(trifluoromethyl) phenoxy]-2-nitrobenzoate; acifluorfen-sodium [2] | 256-634-5 [1] 263-560-7 [2] | 50594-66-6 [1] 62476-59-9 [2] | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H315 H318 H410 |  |  |  | CLP00
604-042-00-6 | 4-nitrosophenol | 203-251-6 | 104-91-6 | Muta. 2 Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H341 H302 H318 H411 | GHS08 GHS05 GHS07 GHS09 Dgr | H341 H302 H318 H411 |  |  |  | CLP00
604-043-00-1 | monobenzone; 4-hydroxyphenyl benzyl ether; hydroquinone monobenzyl ether | 203-083-3 | 103-16-2 | Eye Irrit. 2 Skin Sens. 1 | H319 H317 | GHS07 Wng | H319 H317 |  |  |  | CLP00
604-044-00-7 | mequinol; 4-methoxyphenol; hydroquinone monomethyl ether | 205-769-8 | 150-76-5 | Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 | H302 H319 H317 | GHS07 Wng | H302 H319 H317 |  |  |  | CLP00
604-045-00-2 | 2,3,5-trimethylhydroquinone | 211-838-3 | 700-13-0 | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H332 H335 H315 H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H332 H335 H315 H318 H317 H410 |  |  |  | CLP00
604-046-00-8 | 4-(4-isopropoxyphenylsulfonyl)phenol | 405-520-5 | 95235-30-6 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
604-047-00-3 | 4-(4-tolyloxy)biphenyl | 405-730-7 | 51601-57-1 | STOT RE 2 * Aquatic Chronic 4 | H373 ** H413 | GHS08 Wng | H373 ** H413 |  |  |  | CLP00
604-048-00-9 | 4,4',4''-(ethan-1,1,1-triyl)triphenol | 405-800-7 | 27955-94-8 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
604-049-00-4 | 4-4'-methylenebis(oxyethylenethio)diphenol | 407-480-4 | 93589-69-6 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
604-051-00-5 | 3,5-bis((3,5-di-tert-butyl-4-hydroxy)benzyl)-2,4,6-trimethylphenol | 401-110-5 | 87113-78-8 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
604-052-00-0 | 2,2'-methylenebis(6-(2H-benzotriazol-2-yl)-4-(1,1,3,3-tetramethylbutyl)phenol) | 403-800-1 | 103597-45-1 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
604-053-00-6 | 2-methyl-4-(1,1-dimethylethyl)-6-(1-methyl-pentadecyl)-phenol | 410-760-9 | 157661-93-3 | Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H317 H400 H410 | GHS07 GHS09 Wng | H315 H317 H410 |  |  |  | CLP00
604-054-00-1 | reaction mass of: 2-methoxy-4-(tetrahydro-4-methylene-2H-pyran-2-yl)-phenol; 4-(3,6-dihydro-4-methyl-2H-pyran-2-yl)-2-methoxyphenol | 412-020-0 |  | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
604-055-00-7 | 2,2'-((3,3',5,5'-tetramethyl-(1,1'-biphenyl)-4,4'-diyl)-bis(oxymethylene))-bis-oxirane | 413-900-7 | 85954-11-6 | Carc. 2 Skin Sens. 1 | H351 H317 | GHS08 GHS07 Wng | H351 H317 |  |  |  | CLP00/ATP01
604-056-00-2 | 2-(2-hydroxy-3,5-dinitroanilino)ethanol | 412-520-9 | 99610-72-7 | Flam. Sol. 2 Repr. 2 Acute Tox. 4 * | H228 H361f *** H302 | GHS02 GHS07 GHS08 Dgr | H228 H361f *** H302 |  |  |  | CLP00
604-057-00-8 | reaction mass of: isomers of 2-(2H-benzotriazol-2-yl)-4-methyl-(n)-dodecylphenol; isomers of 2-(2H-benzotriazol-2-yl)-4-methyl-(n)-tetracosylphenol; isomers of 2-(2H-benzotriazol-2-yl)-4-methyl-5,6-didodecyl-phenol. n = 5 or 6 | 401-680-5 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00/ATP10
604-058-00-3 | 1,2-bis(3-methylphenoxy)ethane | 402-730-9 | 54914-85-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
604-059-00-9 | 2-n-hexadecylhydroquinone | 406-400-5 |  | STOT RE 2 * Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 4 | H373 ** H315 H317 H413 | GHS08 GHS07 Wng | H373 ** H315 H317 H413 |  |  |  | CLP00
604-060-00-4 | 9,9-bis(4-hydroxyphenyl)fluorene | 406-950-6 | 3236-71-3 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H319 H400 H410 | GHS07 GHS09 Wng | H319 H315 H410 |  |  |  | CLP00
604-061-00-X | reaction mass of: 2-chloro-5-sec-tetradecylhydroquinones where sec-tetradecyl = 1-methyltridecyl; 1-ethyldodecyl; 1-propylundecyl; 1-butyldecyl; 1-pentylnonyl; 1-hexyloctyl | 407-740-7 |  | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H315 H317 H412 | GHS07 Wng | H315 H317 H412 |  |  |  | CLP00
604-062-00-5 | 2,4-dimethyl-6-(1-methyl-pentadecyl)phenol | 411-220-5 |  | Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H317 H400 H410 | GHS07 GHS09 Wng | H315 H317 H410 |  |  |  | CLP00
604-063-00-0 | 5,6-dihydroxyindole | 412-130-9 | 3131-52-0 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | CLP00
604-064-00-6 | 2-(4,6-diphenyl-1,3,5-triazin-2-yl)-5-((hexyl)oxy)-phenol | 411-380-6 | 147315-50-2 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
604-065-00-1 | 4,4',4''-(1-methylpropan-1-yl-3-ylidene)tris(2-cyclohexyl-5-methylphenol) | 407-460-5 | 111850-25-0 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
604-066-00-7 | reaction mass of: phenol, 6-(1,1-dimethylethyl)-4-tetrapropyl-2-[(2-hydroxy-5-tetra-propylphenyl)methyl (C41-compound) and methane, 2,2'-bis[6-(1,1-dimethyl-ethyl)-1-hydroxy-4-tetrapropyl-phenyl)]- (C45-compound); 2,6-bis(1,1-dimethylethyl)-4-tetra-propyl-phenol and 2-(1,1-dimethylethyl)-4-tetrapropyl-phenol; 2,6-bis[(6-(1,1-dimethylethyl)-1-hydroxy-4-tetrapropylphenyl)methyl]-4-(tetrapropyl)phenol and 2-[(6-(1,1-dimethylethyl)-1-hydroxy-4-tetrapropylphenylmethyl]-6-[1-hydroxy-4-tetrapropylphenyl)methyl]-4-(tetrapropyl)phenol | 414-550-8 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
604-067-00-2 | reaction mass of: 2,2'-[[(2-hydroxyethyl)imino]bis(methylene)bis[4-dodecylphenol]; formaldehyde, oligomer with 4-dodecyl phenol and 2-aminoethanol(n = 2); formaldehyde, oligomer with 4-dodecyl phenol and 2-aminoethanol(n = 3, 4 and higher) | 414-520-4 |  | Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H400 H410 | GHS05 GHS09 Dgr | H315 H318 H410 |  |  |  | CLP00
604-068-00-8 | (±)-4-[2-[[3-(4-hydroxyphenyl)-1-methylpropyl]amino]-1-hydroxyethyl]phenol hydrochloride | 415-170-5 | 90274-24-1 | Acute Tox. 4 * Acute Tox. 4 * Skin Sens. 1 | H332 H302 H317 | GHS07 Wng | H332 H302 H317 |  |  |  | CLP00
604-069-00-3 | 2-(1-methylpropyl)-4-tert-butylphenol | 421-740-4 | 51390-14-8 | Skin Corr. 1B Aquatic Chronic 2 | H314 H411 | GHS05 GHS09 Dgr | H314 H411 |  |  |  | CLP00
604-070-00-9 | triclosan; 2,4,4'-trichloro-2'-hydroxy-diphenyl-ether; 5-chloro-2-(2,4-dichlorophenoxy)phenol | 222-182-2 | 3380-34-5 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H319 H400 H410 | GHS07 GHS09 Wng | H319 H315 H410 |  | M=100 |  | CLP00
604-071-00-4 | 4,4'-(1-{4-[1-(4-hydroxyphenyl)-1-methylethyl]phenyl}ethylidene)diphenol | 425-600-3 | 110726-28-8 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
604-072-00-X | 1,2-bis(phenoxymethyl)benzene | 428-620-0 | 10403-74-4 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
604-073-00-5 | (E)-3-[1-[4-[2-(dimethylamino)ethoxy]phenyl]-2-phenylbut-1-enyl]phenol | 428-010-4 | 82413-20-5 | Carc. 2 Repr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H360F *** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H351 H360F *** H317 H410 |  |  |  | ATP01
604-074-00-0 | 2,2',6,6'-tetrabromo–4,4'-isopropylidenediphenol; tetrabromobisphenol-A | 201-236-9 | 79-94-7 | Carc. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H400 H410 | GHS08 GHS09 Dgr | H350 H410 |  |  |  | ATP01/ATP21
604-075-00-6 | 4-(1,1,3,3-tetramethylbutyl)phenol; 4-tert-octylphenol | 205-426-2 | 140-66-9 | Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H400 H410 | GHS05 GHS09 Dgr | H315 H318 H410 |  | M=10 |  | ATP01
604-076-00-1 | phenolphthalein | 201-004-7 | 77-09-8 | Carc. 1B Muta. 2 Repr. 2 | H350 H341 H361f *** | GHS08 Dgr | H341 H350 H361f *** |  | Carc. 1B; H350: C ≥ 1 % |  | ATP01/ATP01corr
604-077-00-7 | 2-benzotriazol-2-yl-4-methyl-6-(2-methylallyl)phenol | 419-750-9 | 98809-58-6 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
604-079-00-8 | 4,4'-(1,3-phenylene-bis(1-methylethylidene))bis-phenol | 428-970-4 | 13595-25-0 | Repr. 2 Skin Sens. 1 Aquatic Chronic 2 | H361f *** H317 H411 | GHS08 GHS07 GHS09 Wng | H317 H361f *** H411 |  |  |  | ATP01/ATP01corr
604-080-00-3 | 4-fluoro-3-trifluoromethylphenol | 432-560-0 | 61721-07-1 | Acute Tox. 4 * Skin Corr. 1A Skin Sens. 1 Aquatic Chronic 2 | H332 H314 H317 H411 | GHS05 GHS07 GHS09 Dgr | H332 H314 H317 H411 |  |  |  | ATP01
604-081-00-9 | 1,1-bis(4-hydroxyphenyl)-1-phenylethane | 433-130-5 | 1571-75-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
604-082-00-4 | 2-chloro-6-fluoro-phenol | 433-890-8 | 2040-90-6 | Muta. 1B Repr. 2 Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H340 H361f *** H302 H314 H317 H411 | GHS05 GHS08 GHS07 GHS09 Dgr | H340 H361f *** H302 H314 H317 H411 |  |  |  | ATP01
604-084-00-5 | 1-ethoxy-2,3-difluorobenzene | 441-000-4 | 121219-07-6 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | ATP01
604-087-00-1 | reaction mass of: 1,2-naphthoquinonediazide-5-sulfonylchloride (or sulfonic acid)monoester with 4,4'-(1-(4-(1-(4-hydroxyphenyl)-1-methylethyl)phenyl)ethylidene)bisphenol; 1,2-naphthoquinonediazide-5-sulfonylchloride (or sulfonic acid)diester with 4,4'-(1-(4-(1-(4-hydroxyphenyl)-1-methylethyl)phenyl)ethylidene)bisphenol; 1,2-naphthoquinonediazide-5-sulfonylchloride (or sulfonic acid)triester with 4,4'-(1-(4-(1-(4-hydroxyphenyl)-1-methylethyl)phenyl)ethylidene)bisphenol | 433-640-8 | - | Pyr. Sol. 1 Aquatic Chronic 4 | H250 H413 | GHS02 Dgr | H250 H413 | EUH044 |  |  | ATP01
604-089-00-2 | 2-methyl-5-tert-butylthiophenol | 444-970-7 | - | Flam. Liq. 3 Repr. 2 Asp. Tox. 1 STOT SE 3 STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H226 H361d *** H304 H336 H373 ** H315 H319 H317 H400 H410 | GHS02 GHS08 GHS07 GHS09 Dgr | H226 H361d *** H373 ** H304 H319 H315 H317 H336 H410 |  |  |  | ATP01
604-090-00-8 | 4-tert-butylphenol | 202-679-0 | 98-54-4 | Repr. 2 Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 1 | H361f H315 H318 H410 | GHS08 GHS05 GHS09 Dgr | H315 H318 H361f H410 |  | M=1 |  | ATP06/ATP13
604-091-00-3 | etofenprox (ISO); 2-(4-ethoxyphenyl)-2-methylpropyl 3-phenoxybenzyl ether | 407-980-2 | 80844-07-1 | Lact. Aquatic Acute 1 Aquatic Chronic 1 | H362 H400 H410 | GHS09 Wng | H362 H410 |  | M=100 M=1000 |  | ATP06
604-092-00-9 | phenol, dodecyl-, branched [1] phenol, 2-dodecyl-, branched [2] phenol, 3-dodecyl-, branched [3] phenol, 4-dodecyl-, branched [4] phenol, (tetrapropenyl) derivatives [5] | 310-154-3 [1] | 121158-58-5 [1] 210555-94-5 [4] 74499-35-7 [5] | Repr. 1B Skin Corr. 1C Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H360F H314 H318 H400 H410 | GHS08 GHS05 GHS09 Dgr | H314 H360F H410 |  | M=10 M=10 |  | ATP07/ATP09
604-093-00-4 | clorofene; chlorophene; 2-benzyl-4-chlorophenol | 204-385-8 | 120-32-1 | Carc. 2 Repr. 2 Acute Tox. 4 STOT RE 2 Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H361f H332 H373 (kidney) H315 H318 H317 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H332 H315 H318 H317 H351 H361f H373 (kidney) H410 |  | M=1 M=100 |  | ATP10
604-094-00-X | isoeugenol [1] (E)-2-methoxy-4-(prop-1-enyl) phenol [2] (Z)-2-methoxy-4-(prop-1-enyl) phenol [3] | 202-590-7 [1] 227-678-2 [2] 227-633-7 [3] | 97-54-1 [1] 5932-68-3 [2] 5912-86-7 [3] | Skin Sens. 1A | H317 | GHS07 Wng | H317 |  | Skin Sens. 1A; H317: C ≥ 0,01 % |  | ATP13
604-095-00-5 | 6,6'-di-tert-butyl-2,2'-methylenedi-p-cresol; [DBMC] | 204-327-1 | 119-47-1 | Repr. 1B | H360F | GHS08 Dgr | H360F |  |  |  | ATP17
604-096-00-0 | piperonyl butoxide (ISO); 2-(2-butoxyethoxy)ethyl 6-propylpiperonyl ether | 200-076-7 | 51-03-6 | STOT SE 3 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H335 H319 H400 H410 | GHS07 GHS09 Wng | H335 H319 H410 | EUH066 | M=1 M=1 |  | ATP18
604-097-00-6 | 2,4,6-tri-tert-butylphenol | 211-989-5 | 732-26-3 | Repr. 1B Acute Tox. 4 STOT RE 2 Skin Sens. 1B | H360D H302 H373 (liver) H317 | GHS08 GHS07 Dgr | H360D H302 H373 (liver) H317 |  | oral: ATE = 500 mg/kg bw |  | ATP18
604-098-00-1 | 4,4’-sulphonyldiphenol; bisphenol S | 201-250-5 | 80-09-1 | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  |  | ATP18
604-099-00-7 | 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]diphenol; bisphenol AF | 216-036-7 | 1478-61-1 | Repr. 1B | H360F | GHS08 Dgr | H360F |  |  |  | ATP21
604-100-00-0 | nonylphenol, branched and linear, ethoxylated (with average molecular weight ≤ 1 540 g/mol) [includes ortho-, meta-, para- isomers or any combination thereof] | 500-315-8 500-024-6 500-045-0 500-209-1 248-762-5 243-816-4 248-291-5 230-770-5 248-743-1 247-555-7 248-293-6 and others | 127087-87-0 9016-45-9 26027-38-3 68412-54-4 27986-36-3 20427-84-3 27176-93-8 1119449-38-5 7311-27-5 27942-27-4 26264-02-8 27177-05-5 14409-72-4 and others | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M = 1 M = 10 |  | ATP21
605-001-00-5 | formaldehyde …% | 200-001-8 | 50-00-0 | Carc. 1B Muta. 2 Acute Tox. 2 Acute Tox. 4 Skin Corr. 1B Skin Sens. 1A | H350 H341 H330 H302 H314 H317 | GHS08 GHS06 GHS05 Dgr | H350 H341 H330 H302 H314 H317 | EUH071 | inhalation: ATE = 100 ppmV (gases) oral: ATE = 500 mg/kg bw STOT SE 3; H335: C ≥ 5 % Skin Corr. 1B; H314: C ≥ 25 % Skin Irrit. 2; H315: 5 % ≤ C < 25 % Eye Irrit. 2; H319: 5 % ≤ C < 25 % | B D F | CLP00/ATP22
605-002-00-0 | 1,3,5-trioxan; trioxymethylene | 203-812-5 | 110-88-3 | Flam. Sol. 1 Repr. 2 STOT SE 3 | H228 H361d *** H335 | GHS02 GHS08 GHS07 Dgr | H228 H361d *** H335 |  |  | T | CLP00
605-003-00-6 | acetaldehyde; ethanal | 200-836-8 | 75-07-0 | Flam. Liq. 1 Carc. 1B Muta. 2 STOT SE 3 Eye Irrit. 2 | H224 H350 H341 H335 H319 | GHS02 GHS08 GHS07 Dgr | H224 H319 H341 H350 H335 |  |  |  | CLP00/ATP13
605-004-00-1 | 2,4,6-trimethyl-1,3,5-trioxane; paraldehyde | 204-639-8 | 123-63-7 | Flam. Liq. 3 | H226 | GHS02 Wng | H226 |  |  |  | CLP00/ATP01
605-005-00-7 | metaldehyde (ISO); 2,4,6,8-tetramethyl- 1,3,5,7-tetraoxacyclooctane | 203-600-2 | 108-62-3 | Flam. Sol. 2 Repr. 2 Acute Tox. 3 Aquatic Chronic 3 | H228 H361f H301 H412 | GHS02 GHS06 GHS08 Dgr | H228 H301 H361f H412 |  | Oral: ATE = 283 mg/kg bw |  | CLP00/ATP14
605-006-00-2 | butyraldehyde | 204-646-6 | 123-72-8 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 |  |  |  | CLP00
605-007-00-8 | 1,1-dimethoxyethane; dimethyl acetal | 208-589-8 | 534-15-6 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 |  |  |  | CLP00
605-008-00-3 | acrolein; prop-2-enal; acrylaldehyde | 203-453-4 | 107-02-8 | Flam. Liq. 2 Acute Tox. 1 Acute Tox. 2 Acute Tox. 3 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H225 H330 H300 H311 H314 H400 H410 | GHS02 GHS06 GHS05 GHS09 Dgr | H225 H300 H311 H330 H314 H410 | EUH071 | Skin Corr. 1B; H314: C ≥ 0,1 % M=100 M=1 | D | CLP00/ATP06
605-009-00-9 | crotonaldehyde; 2-butenal [1] (E)-2-butenal; (E)-crotonaldehyde [2] | 224-030-0 [1] 204-647-1 [2] | 4170-30-3 [1] 123-73-9 [2] | Flam. Liq. 2 Muta. 2 Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 | H225 H341 H330 H311 H301 H335 H373 ** H315 H318 H400 | GHS02 GHS06 GHS08 GHS05 GHS09 Dgr | H225 H341 H330 H311 H301 H373 ** H335 H315 H318 H400 |  |  |  | CLP00
605-010-00-4 | 2-furaldehyde | 202-627-7 | 98-01-1 | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H351 H331 H301 H312 H335 H315 H319 | GHS06 GHS08 Dgr | H351 H331 H301 H312 H319 H335 H315 |  |  |  | CLP00/ATP01
605-011-00-X | 2-chlorobenzaldehyde; o-chlorobenzaldehyde | 201-956-3 | 89-98-5 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
605-012-00-5 | benzaldehyde | 202-860-4 | 100-52-7 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
605-013-00-0 | chloralose (INN); (R)-1,2-O-(2,2,2-trichloroethylidene)-α-D-glucofuranose; glucochloralose; anhydroglucochloral | 240-016-7 | 15879-93-3 | Acute Tox. 3 Acute Tox. 4 * STOT SE 3 Aquatic Acute 1 Aquatic Chronic 1 | H301 H332 H336 H400 H410 | GHS06 GHS09 Dgr | H332 H301 H336 H410 |  | M=10 M=10 | C | CLP00/ATP09
605-014-00-6 | chloral hydrate; 2,2,2-trichloroethane-1,1-diol | 206-117-5 | 302-17-0 | Acute Tox. 3 * Skin Irrit. 2 Eye Irrit. 2 | H301 H315 H319 | GHS06 Dgr | H301 H319 H315 |  |  |  | CLP00
605-015-00-1 | 1,1-diethoxyethane; acetal | 203-310-6 | 105-57-7 | Flam. Liq. 2 Skin Irrit. 2 Eye Irrit. 2 | H225 H315 H319 | GHS02 GHS07 Dgr | H225 H319 H315 |  |  |  | CLP00
605-016-00-7 | glyoxal … %; ethandial … % | 203-474-9 | 107-22-2 | Muta. 2 Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H341 H332 H315 H319 H317 | GHS07 GHS08 Wng | H341 H332 H319 H315 H317 |  | * | B | CLP00
605-017-00-2 | 1,3-dioxolane | 211-463-5 | 646-06-0 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 |  |  |  | CLP00
605-018-00-8 | propanal; propionaldehyde | 204-623-0 | 123-38-6 | Flam. Liq. 2 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H225 H335 H315 H319 | GHS02 GHS07 Dgr | H225 H319 H335 H315 |  |  |  | CLP00
605-019-00-3 | citral | 226-394-6 | 5392-40-5 | Skin Irrit. 2 Skin Sens. 1 | H315 H317 | GHS07 Wng | H315 H317 |  |  |  | CLP00
605-020-00-9 | safrole; 5-allyl-1,3-benzodioxole | 202-345-4 | 94-59-7 | Carc. 1B Muta. 2 Acute Tox. 4 * | H350 H341 H302 | GHS08 GHS07 Dgr | H350 H341 H302 |  |  |  | CLP00
605-021-00-4 | formaldehyde, reaction products with butylphenol | 294-145-9 | 91673-30-2 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
605-022-00-X | glutaral; glutaraldehyde; 1,5-pentanedial | 203-856-5 | 111-30-8 | Acute Tox. 2 Acute Tox. 3 STOT SE 3 Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 2 | H330 H301 H335 H314 H334 H317 H400 H411 | GHS06 GHS05 GHS08 GHS09 Dgr | H330 H301 H314 H334 H317 H335 H410 | EUH071 | STOT SE 3; H335: ,5 % ≤ C < 5 % M=1 |  | CLP00/ATP09
605-023-00-5 | 5-chloro-2-(4-chlorophenoxy)phenol; [DCPP] | 429-290-0 | 3380-30-1 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  | M=10 M=10 |  | ATP01/ATP10
605-024-00-0 | 2-bromo-5-hydroxy-4-methoxybenzaldehyde | 426-540-0 | 2973-59-3 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
605-025-00-6 | chloroacetaldehyde | 203-472-8 | 107-20-0 | Carc. 2 Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Aquatic Acute 1 | H351 H330 H311 H301 H314 H400 | GHS06 GHS08 GHS05 GHS09 Dgr | H351 H330 H311 H301 H314 H400 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
605-026-00-1 | 2,5,7,7-tetramethyloctanal | 405-690-0 | 114119-97-0 | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H315 H317 H411 | GHS07 GHS09 Wng | H315 H317 H411 |  |  |  | CLP00
605-027-00-7 | reaction mass of: 3a,4,5,6,7,7a-hexahydro-4,7-methano-1H-indene-6-carboxaldehyde; 3a,4,5,6,7,7a-hexahydro-4,7-methano-1H-indene-5-carboxaldehyde | 410-480-7 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
605-028-00-2 | β-methyl-3-(1-methylethyl)-benzenepropanal | 412-050-4 | 125109-85-5 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
605-029-00-8 | 2-cyclohexylpropanal | 412-270-0 | 2109-22-0 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
605-030-00-3 | 1-(p-methoxyphenyl)acetaldehyde oxime | 411-510-1 | 3353-51-3 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
605-031-00-9 | reaction mass of: 2,2-dimethoxyethanal [(this component is considered to be anhydrous in terms of identity, structure and composition. However, 2,2-dimethoxyethanal will exist in a hydrated form. 60 % anhydrous is equivalent to 70.4 % hydrate; water(Including free water and water in hydrated 2,2-dimethoxyethanal)] | 421-890-0 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
605-032-00-4 | 3-[3-(4-fluorophenyl)-1-(1-methylethyl)-1H-indol-2-yl]-(E)-2-propenal | 425-370-4 | 93957-50-7 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | ATP01
605-033-00-X | reaction mass of: 3,7,11-trimethyl-cis-6,10-dodecadienal; 3,7,11-trimethyl-trans-6,10-dodecadienal | 425-910-9 | 32480-08-3 | Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H400 H410 | GHS07 GHS09 Wng | H315 H410 |  |  |  | ATP01
605-034-00-5 | reaction mass of: (1RS,2RS,3SR,6RS,9SR)-9-methoxytricyclo[5.2.1.0(2,6)]decane-3-carbaldehyde; (1RS,2RS,3RS,6RS,8SR)-8-methoxytricyclo[5.2.1.0(2,6)]decane-3-carbaldehyde; (1RS,2RS,4SR,6RS,8SR)-8-methoxytricyclo[5.2.1.0(2,6)]decane-4-carbaldehyde | 429-860-9 | - | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
605-035-00-0 | (E)-3-(4-(4-fluorophenyl)-5-methoxymethyl-2,6-bis(1-methoxymethyl)pyridin-3-yl)prop-2-enal | 426-330-9 | 177964-68-0 | Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 4 | H319 H317 H413 | GHS07 Wng | H319 H317 H413 |  |  |  | ATP01
605-036-00-6 | 2-bromomalonaldehyde | 430-470-6 | 2065-75-0 | Acute Tox. 4 * Eye Dam. 1 | H302 H318 | GHS05 GHS07 Dgr | H302 H318 |  |  |  | ATP01
605-037-00-1 | trans-3-[2-(7-chloro-2-quinolinyl)vinyl]benzaldehyde; 3-[(E)-2-(7-chloro-2-quinolinyl)vinyl]benzaldehyde | 421-800-1 | 120578-03-2 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
605-038-00-7 | 3-methyl-5-phenylpentan-1-al | 433-900-0 | 55066-49-4 | Acute Tox. 4 * Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H302 H315 H317 H411 | GHS07 GHS09 Wng | H302 H315 H317 H411 |  |  |  | ATP01
605-039-00-2 | 3,4-dihydroxy-5-nitrobenzaldehyde | 441-810-8 | 116313-85-0 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 | H302 H318 H317 | GHS05 GHS07 Dgr | H302 H318 H317 |  |  |  | ATP01
605-040-00-8 | hydroxyisohexyl 3-cyclohexene carboxaldehyde (INCI); reaction mass of 4-(4-hydroxy-4-methylpentyl)cyclohex-3-ene-1-carbaldehyde and 3-(4-hydroxy-4-methylpentyl)cyclohex-3-ene-1-carbaldehyde [1] 4-(4-hydroxy-4-methylpentyl)cyclohex-3-ene-1-carbaldehyde [2] 3-(4-hydroxy-4-methylpentyl)cyclohex-3-ene-1-carbaldehyde [3] | 250-863-4 [2] 257-187-9 [3] | 130066-44-3 [1] 31906-04-4 [2] 51414-25-6 [3] | Skin Sens. 1A | H317 | GHS07 Wng | H317 |  |  |  | ATP09
605-041-00-3 | 2-(4-tert-butylbenzyl) propionaldehyde | 201-289-8 | 80-54-6 | Repr. 1B | H360Fd | GHS08 Dgr | H360Fd |  |  |  | ATP15
605-042-00-9 | α-methyl-1,3-benzodioxole-5-propionaldehyde [1] (S)-α-methyl-1,3-benzodioxole-5-propionaldehyde; (2S)-3-(1,3-benzodioxol-5-yl)-2-methylpropanal [2] (R)-α-methyl-1,3-benzodioxole-5-propionaldehyde; (2R)-3-(1,3-benzodioxol-5-yl)-2-methylpropanal [3] | 214-881-6 [1] - [2] - [3] | 1205-17-0 [1] 737776-68-0 [2] 737776-59-9 [3] | Skin Sens. 1B | H317 | GHS07 Wng | H317 |  |  |  | ATP22
605-043-00-4 | 2,4-dimethylcyclohex-3-ene-1-carbaldehyde [1] (1α,2α,5α)-2,5-dimethylcyclohex-3-ene-1-carbaldehyde [2] 2,6-dimethylcyclohex-3-ene-1-carbaldehyde [3] 3,5-dimethylcyclohex-3-ene-1-carbaldehyde [4] 3,6-dimethylcyclohex-3-ene-1-carbaldehyde [5] 4,6-dimethylcyclohex-3-ene-1-carbaldehyde [6] reaction mass of 3,5-dimethylcyclohex-3-ene-1-carbaldehyde and 2,4-dimethylcyclohex-3-ene-1-carbaldehyde [7] dimethylcyclohex-3-ene-1-carbaldehyde [8] Dimethylcyclohex-3-ene-1-carbaldehyde [9] 1,2,4(or 1,3,5)-trimethylcyclohex-3-ene-1-carbaldehyde [10] 1,3,4-trimethylcyclohex-3-ene-1-carbaldehyde [11] 2,2,4-trimethylcyclohex-3-ene-1-carbaldehyde [12] 2,4,6-trimethylcyclohex-3-enecarbaldehyde [13] isocyclocitral [14] 3,5,6-trimethylcyclohex-3-ene-1-carbaldehyde [15] 4,6,6-trimethylcyclohex-3-ene-1-carbaldehyde [16] | 268-264-1 [1] 252-395-6 [2] - [3] 268-263-6 [4] 267-186-5 [5] 253-139-6 [6] - [7] 248-742-6 [8] 272-113-5 [9] 276-055-1 [10] - [11] - [12] 215-833-7 [13] 215-638-7 [14] 266-810-3 [15] - [16] | 68039-49-6 [1] 35145-02-9 [2] 6975-94-6 [3] 68039-48-5 [4] 67801-65-4 [5] 36635-35-5 [6] - [7] 27939-60-2 [8] 68737-61-1 [9] 71832-78-5 [10] 40702-26-9 [11] 1726-47-2 [12] 1423-46-7 [13] 1335-66-6 [14] 67634-07-5 [15] 6754-27-4 [16] | Skin Sens. 1B | H317 | GHS07 Wng | H317 |  |  |  | ATP22
606-001-00-8 | acetone; propan-2-one; propanone | 200-662-2 | 67-64-1 | Flam. Liq. 2 STOT SE 3 Eye Irrit. 2 | H225 H336 H319 | GHS02 GHS07 Dgr | H225 H319 H336 | EUH066 |  |  | CLP00
606-002-00-3 | butanone; ethyl methyl ketone | 201-159-0 | 78-93-3 | Flam. Liq. 2 STOT SE 3 Eye Irrit. 2 | H225 H336 H319 | GHS02 GHS07 Dgr | H225 H319 H336 | EUH066 |  |  | CLP00
606-003-00-9 | heptan-3-one; butyl ethyl ketone | 203-388-1 | 106-35-4 | Flam. Liq. 3 Acute Tox. 4 * Eye Irrit. 2 | H226 H332 H319 | GHS02 GHS07 Wng | H226 H332 H319 |  |  |  | CLP00
606-004-00-4 | 4-methylpentan-2-one; isobutyl methyl ketone | 203-550-1 | 108-10-1 | Flam. Liq. 2 Carc. 2 Acute Tox. 4 STOT SE 3 Eye Irrit. 2 | H225 H351 H332 H336 H319 | GHS02 GHS07 GHS08 Dgr | H225 H332 H319 H351 H336 | EUH066 | Inhalation: ATE = 11 mg/L (Vapours) |  | CLP00/ATP17
606-005-00-X | 2,6-dimethylheptan-4-one; di-isobutyl ketone | 203-620-1 | 108-83-8 | Flam. Liq. 3 STOT SE 3 | H226 H335 | GHS02 GHS07 Wng | H226 H335 |  | STOT SE 3; H335: C ≥ 10 % |  | CLP00
606-006-00-5 | pentan-3-one; diethyl ketone | 202-490-3 | 96-22-0 | Flam. Liq. 2 STOT SE 3 STOT SE 3 | H225 H335 H336 | GHS02 GHS07 Dgr | H225 H335 H336 | EUH066 |  |  | CLP00
606-007-00-0 | 3-methylbutan-2-one; methyl isopropyl ketone | 209-264-3 | 563-80-4 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 |  |  |  | CLP00
606-009-00-1 | 4-methylpent-3-en-2-one; mesityl oxide | 205-502-5 | 141-79-7 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H226 H332 H312 H302 | GHS02 GHS07 Wng | H226 H332 H312 H302 |  | * |  | CLP00
606-010-00-7 | cyclohexanone | 203-631-1 | 108-94-1 | Flam. Liq. 3 Acute Tox. 4 * | H226 H332 | GHS02 GHS07 Wng | H226 H332 |  |  |  | CLP00
606-011-00-2 | 2-methylcyclohexanone | 209-513-6 | 583-60-8 | Flam. Liq. 3 Acute Tox. 4 * | H226 H332 | GHS02 GHS07 Wng | H226 H332 |  |  |  | CLP00
606-012-00-8 | 3,5,5-trimethylcyclohex-2-enone; isophorone | 201-126-0 | 78-59-1 | Carc. 2 Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 | H351 H312 H302 H335 H319 | GHS08 GHS07 Wng | H351 H312 H302 H319 H335 |  | STOT SE 3; H335: C ≥ 10 % |  | CLP00
606-013-00-3 | p-benzoquinone; quinone | 203-405-2 | 106-51-4 | Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 | H331 H301 H335 H315 H319 H400 | GHS06 GHS09 Dgr | H331 H301 H319 H335 H315 H400 |  | M=10 |  | CLP00/ATP01
606-014-00-9 | chlorophacinone (ISO); 2-[(4-chlorophenyl)(phenyl)acetyl]-1H-indene-1,3(2H)-dione | 223-003-0 | 3691-35-8 | Repr. 1B Acute Tox. 1 Acute Tox. 1 Acute Tox. 1 STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D H330 H310 H300 H372 (blood) H400 H410 | GHS08 GHS06 GHS09 Dgr | H330 H310 H300 H360D H372 (blood) H410 |  | Repr. 1B; H360D: C ≥ 0,003 % STOT RE 1; H372: C ≥ 0,1 % STOT RE 2; H373: 0,01 % ≤ C < 0,1 % M=1 M=1 |  | CLP00/ATP09
606-016-00-X | pindone (ISO); 2-pivaloylindan-1,3-dione | 201-462-8 | 83-26-1 | Acute Tox. 3 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H372 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H301 H372 ** H410 |  |  |  | CLP00
606-017-00-5 | diketene; diketen | 211-617-1 | 674-82-8 | Flam. Liq. 3 Acute Tox. 4 * | H226 H332 | GHS02 GHS07 Wng | H226 H332 |  |  | D | CLP00
606-018-00-0 | dichlone (ISO); 2,3-dichloro-1,4-naphthoquinone | 204-210-5 | 117-80-6 | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H315 H410 |  |  |  | CLP00
606-019-00-6 | chlordecone (ISO); perchloropentacyclo[5,3,0,02,6,03,9,04,8]decan-5-one; decachloropentacyclo[5,2,1,02,6,03,9,05,8]decan-4-one | 205-601-3 | 143-50-0 | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H311 H301 H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H311 H301 H410 |  |  |  | CLP00
606-020-00-1 | 5-methylheptan-3-one | 208-793-7 | 541-85-5 | Flam. Liq. 3 STOT SE 3 Eye Irrit. 2 | H226 H335 H319 | GHS02 GHS07 Wng | H226 H319 H335 |  | STOT SE 3; H335: C ≥ 10 % |  | CLP00
606-021-00-7 | N-methyl-2-pyrrolidone; 1-methyl-2-pyrrolidone | 212-828-1 | 872-50-4 | Repr. 1B STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H360D *** H335 H315 H319 | GHS08 GHS07 Dgr | H360D *** H319 H335 H315 |  | STOT SE 3; H335: C ≥ 10 % |  | CLP00/ATP09
606-022-00-2 | 1-phenyl-3-pyrazolidone | 202-155-1 | 92-43-3 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
606-023-00-8 | 4-methoxy-4-methylpentan-2-one | 203-512-4 | 107-70-0 | Flam. Liq. 3 Acute Tox. 4 * | H226 H332 | GHS02 GHS07 Wng | H226 H332 |  |  |  | CLP00
606-024-00-3 | heptan-2-one; methyl amyl ketone | 203-767-1 | 110-43-0 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * | H226 H332 H302 | GHS02 GHS07 Wng | H226 H332 H302 |  |  |  | CLP00
606-025-00-9 | cyclopentanone | 204-435-9 | 120-92-3 | Flam. Liq. 3 Skin Irrit. 2 Eye Irrit. 2 | H226 H315 H319 | GHS02 GHS07 Wng | H226 H319 H315 |  |  |  | CLP00
606-026-00-4 | 5-methylhexan-2-one; isoamyl methyl ketone | 203-737-8 | 110-12-3 | Flam. Liq. 3 Acute Tox. 4 * | H226 H332 | GHS02 GHS07 Wng | H226 H332 |  |  |  | CLP00
606-027-00-X | heptan-4-one; di-n-propyl ketone | 204-608-9 | 123-19-3 | Flam. Liq. 3 Acute Tox. 4 * | H226 H332 | GHS02 GHS07 Wng | H226 H332 |  |  |  | CLP00
606-028-00-5 | 2,4-dimethylpentan-3-one; di-isopropyl ketone | 209-294-7 | 565-80-0 | Flam. Liq. 2 Acute Tox. 4 * | H225 H332 | GHS02 GHS07 Dgr | H225 H332 |  |  |  | CLP00
606-029-00-0 | pentane-2,4-dione; acetylacetone | 204-634-0 | 123-54-6 | Flam. Liq. 3 Acute Tox. 4 * | H226 H302 | GHS02 GHS07 Wng | H226 H302 |  |  |  | CLP00
606-030-00-6 | hexan-2-one; methyl butyl ketone; butyl methyl ketone; methyl-n-butyl ketone | 209-731-1 | 591-78-6 | Flam. Liq. 3 Repr. 2 STOT SE 3 STOT RE 1 | H226 H361f *** H336 H372 ** | GHS02 GHS08 GHS07 Dgr | H226 H361f *** H372 ** H336 |  |  |  | CLP00
606-031-00-1 | 3-propanolide; 1,3-propiolactone | 200-340-1 | 57-57-8 | Carc. 1B Acute Tox. 2 * Skin Irrit. 2 Eye Irrit. 2 | H350 H330 H315 H319 | GHS06 GHS08 Dgr | H350 H330 H319 H315 |  |  |  | CLP00
606-032-00-7 | hexachloroacetone | 204-129-5 | 116-16-5 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
606-033-00-2 | 2-(3,4-dichlorophenyl)-4-methyl-1,2,4-oxadiazolidinedione; methazole | 243-761-6 | 20354-26-1 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H312 H302 H315 H319 H411 | GHS07 GHS09 Wng | H312 H302 H319 H315 H411 |  |  |  | CLP00
606-034-00-8 | metribuzin (ISO); 4-amino-6-tert-butyl-3-methylthio-1,2,4-triazin-5(4H)-one; 4-amino-4,5-dihydro-6-(1,1-dimethylethyl)-3-methylthio-1,2,4-triazin-5-one | 244-209-7 | 21087-64-9 | Acute Tox. 4 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 (blood system) H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 (blood system) H410 |  | oral: ATE = 320 mg/kg bw M = 10 M = 10 |  | CLP00/ATP21
606-035-00-3 | chloridazon (ISO); 5-amino-4-chloro-2-phenylpyridazine-3-(2H)-one; pyrazon | 216-920-2 | 1698-60-8 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
606-036-00-9 | quinomethionate; chinomethionat (ISO); 6-methyl-1,3-dithiolo(4,5-b)quinoxalin-2-one | 219-455-3 | 2439-01-2 | Repr. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361f *** H332 H312 H302 H373 ** H319 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H361f *** H332 H312 H302 H373 ** H319 H317 H410 |  |  |  | CLP00
606-037-00-4 | triadimefon (ISO); 1-(4-chlorophenoxy)-3,3-dimethyl-1-(1,2,4-triazol-1-yl)butanone | 256-103-8 | 43121-43-3 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | CLP00
606-038-00-X | diphacinone (ISO); 2-diphenylacetylindan-1,3-dione | 201-434-5 | 82-66-6 | Acute Tox. 2 * STOT RE 1 | H300 H372 ** | GHS06 GHS08 Dgr | H300 H372 ** |  |  |  | CLP00
606-039-00-5 | 5(or 6)-tert-butyl-2'-chloro-6'-ethylamino-3',7'-dimethylspiro(isobenzofuran-1(1H),9'-xanthene)-3-one | 400-680-2 |  | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H400 H410 | GHS07 GHS09 Wng | H332 H410 |  |  |  | CLP00
606-040-00-0 | (N-benzyl-N-ethyl)amino-3-hydroxyacetophenone hydrochloride | 401-840-4 | 55845-90-4 | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
606-041-00-6 | 2-methyl-1-(4-methylthiophenyl)-2-morpholinopropan-1-one | 400-600-6 | 71868-10-5 | Repr. 1B Acute Tox. 4 * Aquatic Chronic 2 | H360FD H302 H411 | GHS08 GHS07 GHS09 Dgr | H302 H360FD H411 |  |  |  | CLP00/ATP10
606-042-00-1 | acetophenone | 202-708-7 | 98-86-2 | Acute Tox. 4 * Eye Irrit. 2 | H302 H319 | GHS07 Wng | H302 H319 |  |  |  | CLP00
606-043-00-7 | 2,4-di-tert-butylcyclohexanone | 405-340-7 | 13019-04-0 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
606-044-00-2 | 2,4,6-trimethylbenzophenone | 403-150-9 | 954-16-5 | Acute Tox. 4 * Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H410 |  |  |  | CLP00
606-045-00-8 | oxadiazon (ISO); 3-[2,4-dichloro-5-(1-methylethoxy)phenyl]-5-(1,1-dimethylethyl)-1,3,4-oxadiazol-2(3H)-one | 243-215-7 | 19666-30-9 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
606-046-00-3 | reaction mass of cis- and trans-cyclohexadec-8-en-1-one | 401-700-2 | 3100-36-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
606-047-00-9 | 2-benzyl-2-dimethylamino-4'-morpholinobutyrophenone | 404-360-3 | 119313-12-1 | Repr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H360D H400 H410 | GHS08 GHS09 Dgr | H360D H410 |  |  |  | CLP00/ATP14
606-048-00-4 | 2'-anilino-3'-methyl-6'-dipentylaminospiro(isobenzofuran-1(1H),9'-xanthen)-3-one | 406-480-1 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
606-049-00-X | 4-(trans-4-propylcyclohexyl)acetophenone | 406-700-6 | 78531-61-0 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
606-050-00-5 | 6-anilino-1-benzoyl-4-(4-tert-pentylphenoxy)naphto[1,2,3-de]quinoline-2,7-(3H)-dione | 412-480-2 | 72453-58-8 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
606-051-00-0 | 4-pentylcyclohexanone | 406-670-4 | 61203-83-6 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
606-052-00-6 | 4-(N,N-dibutylamino)-2-hydroxy-2'-carboxybenzophenone | 410-410-5 | 54574-82-2 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
606-053-00-1 | flurtamone (ISO); (RS)-5-methylamino-2-phenyl-4-(α, α,α-trifluoro-m-tolyl)furan-3(2H)-one |  | 96525-23-4 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
606-054-00-7 | isoxaflutole (ISO); 5-cyclopropyl-1,2-oxazol-4-yl α,α,α-trifluoro-2-mesyl-p-tolyl ketone |  | 141112-29-0 | Repr. 2 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H400 H410 | GHS08 GHS09 Wng | H361d *** H410 |  | M=10 M=100 |  | CLP00/ATP07
606-055-00-2 | 1-(2,3-dihydro-1,3,3,6-tetramethyl-1-(1-methylethyl)-1H-inden-5-yl)ethanone | 411-180-9 | 92836-10-7 | Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 2 | H302 H373 ** H411 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H411 |  |  |  | CLP00
606-056-00-8 | 4-chloro-3',4'-dimethoxybenzophenone | 404-610-1 | 116412-83-0 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
606-057-00-3 | 4-propylcyclohexanone | 406-810-4 | 40649-36-3 | Skin Irrit. 2 Aquatic Chronic 3 | H315 H412 | GHS07 Wng | H315 H412 |  |  |  | CLP00
606-058-00-9 | 4'-fluoro-2,2-dimethoxyacetophenone | 407-500-1 | 21983-80-2 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
606-059-00-4 | 2,4-difluoro-α-(1H-1,2,4-triazol-1-yl)acetophenone hydrochloride | 412-390-3 | 86386-75-6 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 | H302 H318 H317 | GHS05 GHS07 Dgr | H302 H318 H317 |  |  |  | CLP00
606-060-00-X | reaction mass of: trans-2,4-dimethyl-2-(5,6,7,8-tetrahydro-5,5,8,8-tetramethyl-naphthalene-2-yl)-1,3-dioxolane; cis-2,4-dimethyl-2-(5,6,7,8-tetrahydro-5,5,8,8-tetramethyl-naphthalene-2-yl)-1,3-dioxolane | 412-950-7 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
606-061-00-5 | (3-chlorophenyl)-(4-methoxy-3-nitrophenyl)methanone | 423-290-4 | 66938-41-8 | Muta. 2 Aquatic Acute 1 Aquatic Chronic 1 | H341 H400 H410 | GHS08 GHS09 Wng | H341 H410 |  |  |  | CLP00
606-062-00-0 | tetrahydrothiopyran-3-carboxaldehyde | 407-330-8 | 61571-06-0 | Repr. 1B Eye Dam. 1 Aquatic Chronic 3 | H360D *** H318 H412 | GHS08 GHS05 Dgr | H360D *** H318 H412 |  |  |  | CLP00
606-063-00-6 | (E)-3-(2-chlorophenyl)-2-(4-fluorophenyl)propenal | 410-980-5 | 112704-51-5 | Eye Irrit. 2 Skin Sens. 1 | H319 H317 | GHS07 Wng | H319 H317 |  |  |  | CLP00
606-064-00-1 | pregn-5-ene-3,20-dione bis(ethylene ketal) | 407-450-0 | 7093-55-2 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
606-065-00-7 | 1-(4-morpholinophenyl)butan-1-one | 413-790-0 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
606-066-00-2 | (E)-5[(4-chlorophenyl)methylene]-2,2-dimethylcyclopentanone | 410-440-9 | 164058-20-2 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
606-067-00-8 | reaction mass of: 1-(2,3,6,7,8,9-hexahydro-1,1-dimethyl-1H-benz(g)inden-4-yl)ethanone; 1-(2,3,5,6,7,8-hexahydro-1,1-dimethyl-1H-benz(f)inden-4-yl)ethanone; 1-(2,3,6,7,8,9-hexahydro-1,1-dimethyl-1H-benz(g)inden-5-yl)ethanone; 1-(2,3,6,7,8,9-hexahydro-3,3-dimethyl-1H-benz(g)inden-5-yl)ethanone | 414-870-8 | 96792-67-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
606-068-00-3 | 2,7,11-trimethyl-13-(2,6,6-trimethylcyclohex-1-en-1-yl)tridecahexaen-2,4,6,8,10,12-al | 415-770-7 | 1638-05-7 | STOT RE 2 * Skin Sens. 1 Aquatic Chronic 3 | H373 ** H317 H412 | GHS08 GHS07 Wng | H373 ** H317 H412 |  |  |  | CLP00
606-069-00-9 | spiro[1,3-dioxolane-2,5'-(4',4',8',8'-tetramethyl-hexahydro-3',9'-methanonaphthalene)] | 415-460-1 | 154171-76-3 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
606-070-00-4 | butroxydim (ISO); 5-(3-butyryl-2,4,6-trimethylphenyl)-2-[1-(ethoxyimino)propyl]-3-hydroxycyclohex-2-en-1-one | 414-790-3 | 138164-12-2 | Repr. 2 Acute Tox. 4 * Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H361fd H302 H315 H400 H410 | GHS08 GHS07 GHS09 Wng | H361fd H302 H315 H410 |  |  |  | CLP00
606-071-00-X | 17-spiro(5,5-dimethyl-1,3-dioxan-2-yl)androsta-1,4-diene-3-one | 421-050-3 | 13258-43-0 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
606-072-00-5 | 3-acetyl-1-phenyl-pyrrolidine-2,4-dione | 421-600-2 | 719-86-8 | STOT RE 2 * Aquatic Chronic 2 | H373 ** H411 | GHS08 GHS09 Wng | H373 ** H411 |  |  |  | CLP00
606-073-00-0 | 4,4'-bis(dimethylamino)benzophenone; Michler's ketone | 202-027-5 | 90-94-8 | Carc. 1B Muta. 2 Eye Dam. 1 | H350 H341 H318 | GHS08 GHS05 Dgr | H350 H341 H318 |  |  |  | CLP00
606-074-00-6 | reaction mass of: (1R*,2S*)-2-acetyl-1,2,3,4,5,6,7,8-octahydro-1,2,8,8-tetramethylnaphthalene; (2R*,3S*)-2-acetyl-1,2,3,4,5,6,7,8-octahydro-2,3,8,8-tetramethylnaphthalene | 425-570-1 | - | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
606-075-00-1 | 1-benzyl-5-ethoxyimidazolidine-2,4-dione | 417-340-4 | 65855-02-9 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
606-076-00-7 | 1-((2-quinolinyl-carbonyl)oxy)-2,5-pyrrolidinedione | 418-630-3 | 136465-99-1 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
606-077-00-2 | (3S,4S)-3-hexyl-4-[(R)-2-hydroxytridecyl]-2-oxetanone | 418-650-2 | 104872-06-2 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
606-078-00-8 | 1-octylazepin-2-one | 420-040-6 | 59227-88-2 | Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H314 H317 H411 | GHS05 GHS07 GHS09 Dgr | H314 H317 H411 |  |  |  | CLP00
606-079-00-3 | 2-n-butyl-benzo[d]isothiazol-3-one | 420-590-7 | 4299-07-4 | Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H314 H317 H410 |  |  |  | CLP00
606-081-00-4 | (3β, 5α, 6β)-3-(acetyloxy)-5-bromo-6-hydroxy-androstan-17-one | 419-790-7 | 4229-69-0 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
606-082-00-X | reaction mass of: butan-2-one oxime; syn-O,O'-di(butan-2-one oxime)diethoxysilane | 406-930-7 |  | STOT RE 1 Skin Sens. 1 Aquatic Chronic 3 | H372 ** H317 H412 | GHS08 GHS07 Dgr | H372 ** H317 H412 |  |  |  | CLP00
606-083-00-5 | 2-chloro-5-sec-hexadecylhydroquinone | 407-750-1 | 137193-60-3 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H315 H319 H317 H412 | GHS07 Wng | H319 H315 H317 H412 |  |  |  | CLP00
606-084-00-0 | 1-(4-methoxy-5-benzofuranyl)-3-phenyl-1,3-propanedione | 414-540-3 | 484-33-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
606-085-00-6 | (1R,4S)-2-azabicyclo[2.2.1]hept-5-en-3-one | 418-530-1 | 79200-56-9 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 | H302 H318 H317 | GHS05 GHS07 Dgr | H302 H318 H317 |  |  |  | CLP00
606-086-00-1 | 1-(3,3-dimethylcyclohexyl)pent-4-en-1-one | 422-330-8 | 56973-87-6 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
606-087-00-7 | 6-ethyl-5-fluoro-4(3H)-pyrimidone | 422-460-5 | 137234-87-8 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
606-088-00-2 | 2,4,4,7-tetramethyl-6-octen-3-one | 422-520-0 | 74338-72-0 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
606-089-00-8 | reaction mass of: 1,4-diamino-2-chloro-3-phenoxyanthraquinone; 1,4-diamino-2,3-bis-phenoxyanthraquinone | 423-220-2 | 12223-77-7 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
606-090-00-3 | 1-[3-[(dimethylamino)methyl]-4-hydroxyphenyl]ethanone | 430-920-1 | 73096-98-7 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | ATP01
606-091-00-9 | 6-chloro-5-(2-chloroethyl)-1,3-dihydroindol-2-one | 421-320-0 | 118289-55-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
606-092-00-4 | reaction mass of: (E)-oxacyclohexadec-12-en-2-one; (E)-oxacyclohexadec-13-en-2-one; a) (Z)-oxacyclohexadec-(12)-en-2-one and b) (Z)-oxacyclohexadec-(13)-en-2-one | 422-320-3 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
606-093-00-X | 5-ethyl-2,4-dihydro-4-(2-phenoxyethyl)-3H-1,2,4-triazol-3-one | 414-470-3 | 95885-13-5 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | ATP01
606-094-00-5 | N-[ethyl(3-methylbutyl)amino]-3-methyl-1-phenyl-spiro[[1]benzo-pyrano[2,3-c]pyrazole-4(1H),1'(3'H)-isobenzofuran]-3'-one | 417-460-7 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
606-095-00-0 | (R,S)-2-azabicyclo[2.2.1]hept-5-en-3-one | 421-830-3 | 49805-30-3 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  |  |  | ATP01
606-096-00-6 | 3-(6-O-(6-desoxy-α-.sc.l.sc.-mannopyranosyl-O-(α-.sc.d.sc.-glucopyranosyl)-(β-.sc.d.sc.-glucopyranosyl)oxy)-2-(3,4-dihydroxyphenyl)-5,7-dihydroxy-4H-1-benzopyran-4-one | 424-170-4 | 130603-71-3 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
606-097-00-1 | 2,2''-dihydroxy-4,4''-(2-hydroxy-propane-1,3-diyldioxy)dibenzophenone | 424-210-0 | 23911-85-5 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
606-098-00-7 | 1-benzyl-5-(hexadecyloxy)-2,4-imidazolidinedione | 431-220-9 | 158574-65-3 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
606-099-00-2 | 5-methoxy-4'-(trifluoromethyl)valerophenone | 425-000-1 | 61718-80-7 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
606-100-00-6 | 2-butyryl-3-hydroxy-5-thiocyclohexan-3-yl-cyclohex-2-en-1-one | 425-150-8 | 94723-86-1 | Repr. 1B Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 3 | H360F *** H302 H317 H412 | GHS08 GHS07 Dgr | H360F *** H302 H317 H412 |  |  |  | ATP01
606-101-00-1 | reaction mass of: 1,5-bis[(2-ethylhexyl)amino]-9,10-anthracenedione; 1-[(2-ethylhexyl)amino]-5-[3-[(2-ethylhexyl)oxy]propyl]amino-9,10-anthracenedione; 1,5-bis[3-[(2-ethylhexyl)oxy]propyl]amino-9,10-anthracenedione; 1-[(2-ethylhexyl)amino]-5-[(3-methoxypropyl)amino]-9,10-anthracene dione; 1-[3-[(2-ethylhexyl)oxy]propyl]amino-5-[(3-methoxypropyl)amino]-9,10-anthracenedione; 1,5-bis[(3-methyloxypropyl)amino]-9,10-anthracenedione | 426-050-7 | 165038-51-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
606-102-00-7 | 4-(3-triethoxysilylpropoxy)-2-hydroxybenzophenone | 431-490-8 | 79876-59-8 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
606-103-00-2 | 1-(4-(trans-4-ethylcyclohexyl)phenyl)ethanone | 426-460-6 | - | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
606-104-00-8 | 1-(4-(trans-4-pentylcyclohexyl)phenyl)ethanone | 426-830-7 | 78531-59-6 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
606-105-00-3 | 3,4,3',4'-tetraphenyl-1,1'-ethandiylbispyrol-2,5-dione | 431-500-0 | 226065-73-2 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
606-106-00-9 | 1-(4-(trans-4-butylcyclohexyl)phenyl)ethanone | 427-320-7 | 83626-30-6 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
606-107-00-4 | 8-azaspiro[4.5]decane-7,9-dione | 427-770-4 | 1075-89-4 | Acute Tox. 3 * Aquatic Chronic 2 | H301 H411 | GHS06 GHS09 Dgr | H301 H411 |  |  |  | ATP01
606-108-00-X | 1,1,1,2,2,4,5,5,5-nonafluoro-4-(trifluoromethyl )-3-pentanone | 436-710-6 | 756-13-8 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
606-109-00-5 | 2-(4-methyl-3-pentenyl)anthraquinone | 428-320-1 | 71308-16-2 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 4 | H302 H317 H413 | GHS07 Wng | H302 H317 H413 |  |  |  | ATP01/ATP01corr
606-110-00-0 | 5-ethoxy-5H-furan-2-one | 428-330-4 | 2833-30-9 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Skin Sens. 1 | H312 H302 H373 ** H314 H317 | GHS05 GHS08 GHS07 Dgr | H314 H312 H302 H373 ** H317 |  |  |  | ATP01
606-111-00-6 | 5-amino-6-methyl-1,3-dihydrobenzoimidazol-2-one | 428-410-9 | 67014-36-2 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | ATP01
606-112-00-1 | (4aR*,8aR*)-4a,5,9,10,11,12-hexahydro-3-methoxy-11-methyl-6H-benzofuro[3a,3,2-ef][2]benzazepin-6-one | 428-690-2 | 1668-86-6 | Acute Tox. 4 * Eye Irrit. 2 Aquatic Chronic 3 | H302 H319 H412 | GHS07 Wng | H302 H319 H412 |  |  |  | ATP01
606-113-00-7 | 1-[4-(4-benzoylphenylsulfanyl)phenyl]-2-methyl-2-(4-methylphenylsulfonyl)propan-1-one | 429-040-0 | 272460-97-6 | Eye Dam. 1 Aquatic Chronic 4 | H318 H413 | GHS05 Dgr | H318 H413 |  |  |  | ATP01/ATP01corr
606-114-00-2 | 4,4',5,5',6,6',7,7'-octachloro-(2,2')biisoindolyl-1,1',3,3'-tetraone | 429-150-9 | 67887-47-2 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
606-115-00-8 | profoxydim (ISO); 2-{(EZ)-1-[(2RS)-2-(4-chlorophenoxy)propoxyimino]butyl}-3-hydroxy-5-(thian-3-yl)cyclohex-2-en-1-one | - | 139001-49-3 | Carc. 2 Repr. 2 Skin Sens. 1 | H351 H361d H317 | GHS08 GHS07 Wng | H351 H361d H317 |  |  |  | ATP01
606-116-00-3 | tepraloxydim (ISO); (RS)-(EZ)-2-{1-[(2E)-3-chloroallyloxyimino]propyl}-3-hydroxy-5-perhydropyran-4-ylcyclohex-2-en-1-one | - | 149979-41-9 | Carc. 2 Repr. 2 | H351 H361fd | GHS08 Wng | H351 H361fd |  |  |  | ATP01
606-117-00-9 | 2,6-bis(1,1-dimethylethyl)-4-(phenylenemethylene)cyclohexa-2,5-dien-1-one | 429-460-4 | 7078-98-0 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
606-118-00-4 | N-(1,3-dimethylbutyl)-N'-(phenyl)-1,4-benzoquinonediimine | 429-640-2 | 52870-46-9 | Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H319 H400 H410 | GHS07 GHS09 Wng | H319 H410 |  |  |  | ATP01
606-119-00-X | (E)-3-methyl-5-cyclopentadecen-1-one | 429-900-5 | - | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | ATP01
606-120-00-5 | 2,5-dihydroxy-5-methyl-3-(morpholin-4-yl)-2-cyclopenten-1-one | 430-170-5 | 114625-74-0 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | ATP01
606-121-00-0 | (+)-(1S,2S,3S,5R)-2,6,6-trimethylbicyclo[3.1.1]heptane-3-spiro-1'-(cyclohex-2'-en-4'-one) | 430-460-1 | 133636-82-5 | Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H314 H317 H410 |  |  |  | ATP01
606-122-00-6 | 3-(2-bromopropionoyl)-4,4-dimethyl-1,3-oxazolan-2-one | 430-820-8 | 114341-88-7 | Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H315 H318 H317 H400 H410 | GHS05 GHS08 GHS07 GHS09 Dgr | H302 H373 ** H315 H318 H317 H410 |  |  |  | ATP01
606-123-00-1 | 4-hexadecyl-1-phenylpyrazolidin-3-one | 430-840-7 | - | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
606-124-00-7 | 1-cyclopropyl-3-(2-methylthio-4-trifluoromethylphenyl)-1,3-propanedione | 421-080-7 | 161462-35-7 | STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H400 H410 | GHS08 GHS09 Wng | H373 ** H410 |  |  |  | ATP01
606-125-00-2 | 1-benzylimidazolidine-2,4-dione | 421-340-1 | 6777-05-5 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | ATP01
606-126-00-8 | 1,4-bis(2,3-dihydroxypropylamino)anthraquinone | 421-470-7 | 99788-75-7 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
606-128-00-9 | 2,2'-(1,3-phenylene)bis[5-chloro-1H-isoindole]-1,3(2H)-dione | 422-650-8 | 148935-94-8 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
606-129-00-4 | 5-amino-[2S-di(methylphenyl)amino]-1,6-diphenyl-4Z-hexen-3-one; (2S,4Z)-5-amino-2-(dibenzylamino)-1,6-diphenylhex-4-en-3-one | 423-090-7 | 156732-13-7 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
606-130-00-X | 4-(1,4-dioxa-spiro[4.5]dec-8-yl)-cyclohexanone | 423-860-2 | 56309-94-5 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
606-131-00-5 | cyclic 3-(1,2-ethanediylacetale)-estra-5(10),9(11)-diene-3,17-dione | 427-230-8 | 5571-36-8 | Repr. 1B STOT RE 2 * Aquatic Chronic 2 | H360F *** H373 ** H411 | GHS08 GHS09 Dgr | H360F *** H373 ** H411 |  |  |  | ATP01
606-132-00-0 | (6β)-6,19-epoxyandrost-4-ene-3,17-dione | 433-490-3 | 6563-83-3 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
606-134-00-1 | androsta-1,4,9(11)-triene-3,17-dione | 433-560-3 | 15375-21-0 | Repr. 2 | H361f *** | GHS08 Wng | H361f *** |  |  |  | ATP01
606-135-00-7 | cyclohexadecanone | 438-930-8 | 2550-52-9 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
606-136-00-2 | (3S,6R,9S,12R,15S,18R,21S,24R)-6,18-dibenzyl-3,9,15,21-tetraisobutyl-4,10,12,16,22,24-hexamethyl-1,7,13,19-tetraoxa-4,10,16,22-tetraazacyclo-tetracosane-2,5,8,11,14,17,20,23-octaone | 444-350-6 | 133413-70-4 | Eye Irrit. 2 Aquatic Chronic 4 | H319 H413 | GHS07 Wng | H319 H413 |  |  |  | ATP01
606-137-00-8 | trans-7,7'-dimethyl-(4H,4H')-(2,2')bi[benzo[1,4]thiazinylidene]-3,3'-dione | 444-750-0 | 211387-26-7 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
606-138-00-3 | (2-butyl-5-nitrobenzofuran-3-yl)[4-(3-dibutylaminopropoxy)phenyl]methanone | 444-800-1 | 141645-23-0 | Flam. Liq. 3 Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H226 H302 H373 ** H315 H318 H317 H400 H410 | GHS02 GHS05 GHS08 GHS07 GHS09 Dgr | H226 H302 H373 ** H315 H318 H317 H410 |  | M=10 |  | ATP01/ATP01corr
606-139-00-9 | (S)-4-(3,4-dichlorophenyl)-3,4-dihydro-2H-naphthalen-1-one | 444-830-5 | 124379-29-9 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
606-140-00-4 | 2-hydroxy-1-(4-(4-(2-hydroxy-2-methylpropionyl)benzyl)phenyl)-2-methylpropan-1-one | 444-860-9 | 474510-57-1 | STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H400 H410 | GHS08 GHS09 Wng | H373 ** H410 |  |  |  | ATP01
606-141-00-X | sodium 3-(methoxycarbonyl)-4-oxo-3,4,5,6-tetrahydro-2-pyridinolate | 418-410-7 | - | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | ATP01
606-142-00-5 | reaction mass of: (1RS,2SR,7SR,8SR,E) 9 and 10-ethylidene-3-oxatricyclo[6.2.1.0(2,7)]undecan-4-one; (1RS,2SR,7SR,8SR,Z)-10-ethylidene-3-oxatricyclo[6.2.1.0(2,7)]undecan-4-one; (1RS,2SR,7SR,8SR,Z)-9-ethylidene-3-oxatricyclo[6.2.1.0(2,7)]undecan-4-one | 434-290-9 | - | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | ATP01
606-143-00-0 | abamectin (combination of avermectin B1a and avermectin B1b) (ISO) [1] avermectin B1a (purity ≥80 %) [2] | _ [1] 265-610-3 [2] | 71751-41-2 [1] 65195-55-3 [2] | Repr. 2 Acute Tox. 1 Acute Tox. 2 STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d H330 H300 H372 (nervous system) H400 H410 | GHS06 GHS08 GHS09 Dgr | H361d H300 H330 H372 (nervous system) H410 |  | STOT RE 1; H372: C ≥ 5 % STOT RE 2; H373: 0,5 % ≤C<5 % M=10000 |  | ATP03
606-144-00-6 | acequinocyl (ISO); 3-dodecyl-1,4-dioxo-1,4-dihydronaphthalen-2-yl acetate | — | 57960-19-7 | STOT SE 1 STOT RE 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H370 (lung) (Inhalation) H373 (blood system) H317 H400 H410 | GHS07 GHS08 GHS09 Dgr | H317 H370 (lung) (Inhalation) H373 (blood system) H410 |  | M=1000 |  | ATP03
606-145-00-1 | sulcotrione (ISO); 2-[2-chloro-4-(methylsulfonyl)benzoyl]cyclohexane-1,3-dione |  | 99105-77-8 | Repr. 2 STOT RE 2 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 1 | H361d H373 (kidneys) H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H317 H361d H373 (kidneys) H410 |  | M=1 M=10 |  | ATP05
606-146-00-7 | tralkoxydim (ISO); 2-(N-ethoxypropanimidoyl)-3-hydroxy-5-mesitylcyclohex-2-en-1-one |  | 87820-88-0 | Carc. 2 Acute Tox. 4 Aquatic Chronic 2 | H351 H302 H411 | GHS08 GHS07 GHS09 Wng | H302 H351 H411 |  |  |  | ATP06
606-147-00-2 | cycloxydim (ISO); 2-(N-ethoxybutanimidoyl)-3-hydroxy-5-(tetrahydro-2H-thiopyran-3-yl)cyclohex-2-en-1-one | 405-230-9 | 101205-02-1 | Repr. 2 | H361d | GHS08 Wng | H361d |  |  |  | ATP06
606-148-00-8 | carvone (ISO); 2-methyl-5-(prop-1-en-2-yl)cyclohex-2-en-1-one [1] d-carvone; (5S)-2-methyl-5-(prop-1-en-2-yl)cyclohex-2-en-1-one [2] l-carvone; (5R)-2-methyl-5-(prop-1-en-2-yl)cyclohex-2-en-1-one [3] | 202-759-5 [1] 218-827-2 [2] 229-352-5 [3] | 99-49-0 [1] 2244-16-8 [2] 6485-40-1 [3] | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP07
606-149-00-3 | tembotrione (ISO); 2-{2-chloro-4-(methylsulfonyl)-3-[(2,2,2-trifluoroethoxy)methyl]benzoyl}cyclohexane-1,3-dione |  | 335104-84-2 | Repr. 2 STOT RE 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d H373 (eyes, kidneys, liver) H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H317 H361d H373 (eyes, kidneys, liver) H410 |  | M=100 M=10 |  | ATP07
606-150-00-9 | clethodim (ISO); (5RS)-2-{(1EZ)-1-[(2E)-3-chloroallyloxyimino]propyl}-5-[(2RS)-2-(ethylthio)propyl]-3-hydroxycyclohex-2-en-1-one |  | 99129-21-2 | Acute Tox. 4 Skin Sens. 1 Aquatic Chronic 3 | H302 H317 H412 | GHS07 Wng | H302 H317 H412 | EUH066 |  |  | ATP10
606-151-00-4 | anthraquinone | 201-549-0 | 84-65-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | ATP10
606-152-00-X | (5-chloro-2-methoxy-4-methyl-3-pyridyl)(4,5,6-trimethoxy-o-tolyl)methanone; pyriofenone |  | 688046-61-9 | Carc. 2 Aquatic Chronic 1 | H351 H410 | GHS08 GHS09 Wng | H351 H410 |  | M=1 |  | ATP17
606-153-00-5 | benzophenone | 204-337-6 | 119-61-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | ATP18
606-154-00-0 | quinoclamine (ISO); 2-amino-3-chloro-1,4-naphthoquinone | 220-529-2 | 2797-51-5 | Carc. 2 Repr. 2 Acute Tox. 4 STOT RE 2 Eye Irrit. 2 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 1 | H351 H361d H302 H373 (blood system, kidney) H319 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H361d H302 H373 (blood system, kidney) H319 H317 H410 |  | oral: ATE = 500 mg/kg bw M=10 M=10 |  | ATP18
606-155-00-6 | cinnamaldehyde; 3-phenylprop-2-enal; cinnamic aldehyde; cinnamal;[1] (2E)-3-phenylprop-2-enal[2] | 203-213-9[1] -[2] | 104-55-2[1] 14371-10-9[2] | Skin Sens. 1A | H317 | GHS07 Wng | H317 |  | Skin Sens. 1; H317:C >= 0.01 % |  | ATP21
606-156-00-1 | acetone oxime | 204-820-1 | 127-06-0 | Carc. 1B Acute Tox. 4 STOT SE 3 STOT RE 2 Eye Dam. 1 Skin Sens. 1 | H350 H312 H336 H373 (blood system) H318 H317 | GHS08 GHS07 GHS05 Dgr | H350 H312 H336 H373 (blood system) H318 H317 |  | dermal: ATE = 1 100 mg/kg bw |  | ATP22
606-157-00-7 | (3E)-dec-3-en-2-one | - | 18402-84-1 | Acute Tox. 4 Asp. Tox. 1 Skin Irrit. 2 Aquatic Chronic 2 | H332 H304 H315 H411 | GHS07 GHS08 GHS09 Dgr | H332 H304 H315 H411 | EUH071 | inhalation: ATE = 1,5 mg/L (dusts or mists) |  | ATP22
606-158-00-2 | 2-(dimethylamino)- 2-[(4-methylphenyl) methyl]-1-[4-(morpholin- 4-yl)phenyl]butan-1-one | 438-340-0 | 119344-86-4 | Repr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H360Df H400 H410 | GHS08 GHS09 Dgr | H360Df H410 |  | M=1 M=1 |  | ATP22
607-001-00-0 | formic acid … % | 200-579-1 | 64-18-6 | Flam. Liq. 3 Met. Corr. 1 Acute Tox. 3 Acute Tox. 4 Skin Corr. 1A Eye Dam. 1 | H226 H290 H331 H302 H314 H318 | GHS02 GHS05 GHS06 Dgr | H226 H290 H331 H302 H314 | EUH071 | inhalation: ATE = 7,4 mg/L (vapours) oral: ATE = 500 mg/kg bw Flam. Liq. 3; H226: C > 85 % Skin Corr. 1A; H314: C ≥ 90 % Skin Corr. 1B; 314: 10 % ≤ C < 90 % Skin Irrit. 2; H315: 2 % ≤ C < 10 % Eye Dam. 1; H318: C ≥ 10 % Eye Irrit. 2; H319: 2 % ≤ C < 10 % | B | CLP00/ATP22
607-002-00-6 | acetic acid … % | 200-580-7 | 64-19-7 | Flam. Liq. 3 Skin Corr. 1A | H226 H314 | GHS02 GHS05 Dgr | H226 H314 |  | Skin Corr. 1A; H314: C ≥ 90 % Skin Corr. 1B; H314: 25 % ≤ C < 90 % Skin Irrit. 2; H315: 10 % ≤ C < 25 % Eye Irrit. 2; H319: 10 % ≤ C < 25 % | B | CLP00
607-003-00-1 | chloroacetic acid | 201-178-4 | 79-11-8 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Aquatic Acute 1 | H331 H311 H301 H314 H400 | GHS06 GHS05 GHS09 Dgr | H331 H311 H301 H314 H400 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00/ATP01
607-004-00-7 | TCA (ISO); trichloroacetic acid | 200-927-2 | 76-03-9 | Skin Corr. 1A Aquatic Acute 1 Aquatic Chronic 1 | H314 H400 H410 | GHS05 GHS09 Dgr | H314 H410 |  | STOT SE 3; H335: C ≥ 1 % |  | CLP00
607-005-00-2 | TCA-sodium (ISO); sodium trichloroacetate | 211-479-2 | 650-51-1 | STOT SE 3 Aquatic Acute 1 Aquatic Chronic 1 | H335 H400 H410 | GHS07 GHS09 Wng | H335 H410 |  |  |  | CLP00
607-006-00-8 | oxalic acid | 205-634-3 | 144-62-7 | Acute Tox. 4 * Acute Tox. 4 * | H312 H302 | GHS07 Wng | H312 H302 |  | * |  | CLP00
607-007-00-3 | salts of oxalic acid with the exception of those specified elsewhere in this Annex | - | - | Acute Tox. 4 * Acute Tox. 4 * | H312 H302 | GHS07 Wng | H312 H302 |  | * | A | CLP00/ATP01
607-008-00-9 | acetic anhydride | 203-564-8 | 108-24-7 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H226 H332 H302 H314 | GHS02 GHS05 GHS07 Dgr | H226 H332 H302 H314 |  | Skin Corr. 1B; H314: C ≥ 25 % Skin Irrit. 2; H315: 5 % ≤ C < 25 % Eye Dam. 1; H318: 5 % ≤ C < 25 % Eye Irrit. 2; H319: 1 % ≤ C < 5 % STOT SE 3; H335: C ≥ 5 % |  | CLP00
607-009-00-4 | phthalic anhydride | 201-607-5 | 85-44-9 | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H302 H335 H315 H318 H334 H317 | GHS08 GHS05 GHS07 Dgr | H302 H335 H315 H318 H334 H317 |  |  |  | CLP00
607-010-00-X | propionic anhydride | 204-638-2 | 123-62-6 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  | Skin Corr. 1B; H314: C ≥ 25 % Skin Irrit. 2; H315: 10 % ≤ C < 25 % Eye Irrit. 2; H319: 10 % ≤ C < 25 % |  | CLP00
607-011-00-5 | acetyl chloride | 200-865-6 | 75-36-5 | Flam. Liq. 2 Skin Corr. 1B | H225 H314 | GHS02 GHS05 Dgr | H225 H314 | EUH014 |  |  | CLP00
607-012-00-0 | benzoyl chloride | 202-710-8 | 98-88-4 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 | H332 H312 H302 H314 H317 | GHS05 GHS07 Dgr | H332 H312 H302 H314 H317 |  |  |  | CLP00/ATP01
607-013-00-6 | dimethyl carbonate | 210-478-4 | 616-38-6 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 |  |  |  | CLP00
607-014-00-1 | methyl formate | 203-481-7 | 107-31-3 | Flam. Liq. 1 Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 | H224 H332 H302 H335 H319 | GHS02 GHS07 Dgr | H224 H332 H302 H319 H335 |  |  |  | CLP00
607-015-00-7 | ethyl formate | 203-721-0 | 109-94-4 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 | H225 H332 H302 H335 H319 | GHS02 GHS07 Dgr | H225 H332 H302 H319 H335 |  |  |  | CLP00
607-016-00-2 | propyl formate [1] isopropyl formate [2] | 203-798-0 [1] 210-901-2 [2] | 110-74-7 [1] 625-55-8 [2] | Flam. Liq. 2 STOT SE 3 STOT SE 3 Eye Irrit. 2 | H225 H335 H336 H319 | GHS02 GHS07 Dgr | H225 H319 H335 H336 |  |  | C | CLP00
607-017-00-8 | butyl formate [1] tert-butyl formate [2] isobutyl formate [3] | 209-772-5 [1] 212-105-0 [2] 208-818-1 [3] | 592-84-7 [1] 762-75-4 [2] 542-55-2 [3] | Flam. Liq. 2 STOT SE 3 Eye Irrit. 2 | H225 H335 H319 | GHS02 GHS07 Dgr | H225 H319 H335 |  |  | C | CLP00
607-018-00-3 | isopentyl formate [1] pentyl formate [2] | 203-769-2 [1] 252-343-2 [2] | 110-45-2 [1] 35073-27-9 [2] | Flam. Liq. 2 STOT SE 3 Eye Irrit. 2 | H225 H335 H319 | GHS02 GHS07 Dgr | H225 H319 H335 |  |  | C | CLP00
607-019-00-9 | methyl chloroformate | 201-187-3 | 79-22-1 | Flam. Liq. 2 Acute Tox. 2 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H225 H330 H312 H302 H314 | GHS02 GHS06 GHS05 Dgr | H225 H330 H312 H302 H314 |  |  |  | CLP00
607-020-00-4 | ethyl chloroformate | 208-778-5 | 541-41-3 | Flam. Liq. 2 Acute Tox. 2 * Acute Tox. 4 * Skin Corr. 1B | H225 H330 H302 H314 | GHS02 GHS06 GHS05 Dgr | H225 H330 H302 H314 |  |  |  | CLP00
607-021-00-X | methyl acetate | 201-185-2 | 79-20-9 | Flam. Liq. 2 STOT SE 3 Eye Irrit. 2 | H225 H336 H319 | GHS02 GHS07 Dgr | H225 H319 H336 | EUH066 |  |  | CLP00
607-022-00-5 | ethyl acetate | 205-500-4 | 141-78-6 | Flam. Liq. 2 STOT SE 3 Eye Irrit. 2 | H225 H336 H319 | GHS02 GHS07 Dgr | H225 H319 H336 | EUH066 |  |  | CLP00
607-023-00-0 | vinyl acetate | 203-545-4 | 108-05-4 | Flam. Liq. 2 Carc. 2 Acute Tox. 4 STOT SE 3 | H225 H351 H332 H335 | GHS02 GHS08 GHS07 Dgr | H225 H332 H351 H335 |  |  | D | CLP00/ATP05
607-024-00-6 | propyl acetate [1] isopropyl acetate [2] | 203-686-1 [1] 203-561-1 [2] | 109-60-4 [1] 108-21-4 [2] | Flam. Liq. 2 STOT SE 3 Eye Irrit. 2 | H225 H336 H319 | GHS02 GHS07 Dgr | H225 H319 H336 | EUH066 |  | C | CLP00
607-025-00-1 | n-butyl acetate | 204-658-1 | 123-86-4 | Flam. Liq. 3 STOT SE 3 | H226 H336 | GHS02 GHS07 Wng | H226 H336 | EUH066 |  |  | CLP00
607-026-00-7 | sec-butyl acetate [1] isobutyl acetate [2] tert-butyl acetate [3] | 203-300-1 [1] 203-745-1 [2] 208-760-7 [3] | 105-46-4 [1] 110-19-0 [2] 540-88-5 [3] | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 | EUH066 |  | C | CLP00
607-027-00-2 | methyl propionate | 209-060-4 | 554-12-1 | Flam. Liq. 2 Acute Tox. 4 * | H225 H332 | GHS02 GHS07 Dgr | H225 H332 |  |  |  | CLP00
607-028-00-8 | ethyl propionate | 203-291-4 | 105-37-3 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 |  |  |  | CLP00
607-029-00-3 | n-butyl propionate [1] sec-butyl propionate [2] tert-butyl propionate [3] | 209-669-5 [1] 208-746-0 [3] | 590-01-2 [1] 591-34-4 [2] 540-42-1 [3] | Flam. Liq. 3 | H226 | GHS02 Wng | H226 |  |  | C | CLP00
607-030-00-9 | propyl propionate | 203-389-7 | 106-36-5 | Flam. Liq. 3 Acute Tox. 4 * | H226 H332 | GHS02 GHS07 Wng | H226 H332 |  |  |  | CLP00
607-031-00-4 | butyl butyrate | 203-656-8 | 109-21-7 | Flam. Liq. 3 | H226 | GHS02 Wng | H226 |  |  | C | CLP00
607-032-00-X | ethyl acrylate | 205-438-8 | 140-88-5 | Flam. Liq. 2 Acute Tox. 3 Acute Tox. 4 Acute Tox. 4 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H225 H331 H312 H302 H335 H315 H319 H317 | GHS02 GHS06 Dgr | H225 H331 H312 H302 H335 H315 H319 H317 |  | inhalation: ATE = 9 mg/L(vapours) oral: ATE = 1120 mg/kg bw dermal: ATE = 1800 mg/kg bw STOT SE 3; H335: C >= 5 % Skin Irrit. 2; H315: C >= 5 % Eye Irrit. 2; H319: C >= 5 % | D | CLP00/ATP21
607-033-00-5 | n-butyl methacrylate | 202-615-1 | 97-88-1 | Flam. Liq. 3 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H226 H335 H315 H319 H317 | GHS02 GHS07 Wng | H226 H319 H335 H315 H317 |  |  | D | CLP00
607-034-00-0 | methyl acrylate; methyl propenoate | 202-500-6 | 96-33-3 | Flam. Liq. 2 Acute Tox. 3 Acute Tox. 4 Acute Tox. 4 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H225 H331 H312 H302 H335 H315 H319 H317 | GHS02 GHS06 Dgr | H225 H331 H312 H302 H335 H315 H319 H317 |  | inhalation: ATE = 3 mg/L(vapours) oral: ATE = 500 mg/kg bw dermal: ATE = 1100 mg/kg bw | D | CLP00/ATP21
607-035-00-6 | methyl methacrylate; methyl 2-methylprop-2-enoate; methyl 2-methylpropenoate | 201-297-1 | 80-62-6 | Flam. Liq. 2 STOT SE 3 Skin Irrit. 2 Skin Sens. 1 | H225 H335 H315 H317 | GHS02 GHS07 Dgr | H225 H335 H315 H317 |  |  | D | CLP00
607-036-00-1 | 2-methoxyethyl acetate; methylglycol acetate | 203-772-9 | 110-49-6 | Repr. 1B Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H360FD H332 H312 H302 | GHS08 GHS07 Dgr | H360FD H332 H312 H302 |  |  |  | CLP00
607-037-00-7 | 2-ethoxyethyl acetate; ethylglycol acetate | 203-839-2 | 111-15-9 | Flam. Liq. 3 Repr. 1B Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H226 H360FD H332 H312 H302 | GHS02 GHS08 GHS07 Dgr | H226 H360FD H332 H312 H302 |  |  |  | CLP00/ATP01
607-038-00-2 | 2-butoxyethyl acetate; butylglycol acetate | 203-933-3 | 112-07-2 | Acute Tox. 4 * Acute Tox. 4 * | H332 H312 | GHS07 Wng | H332 H312 |  |  |  | CLP00
607-039-00-8 | 2,4-D (ISO); 2,4-dichlorophenoxyacetic acid | 202-361-1 | 94-75-7 | Acute Tox. 4 * STOT SE 3 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H302 H335 H318 H317 H412 | GHS05 GHS07 Dgr | H302 H335 H318 H317 H412 |  |  |  | CLP00
607-040-00-3 | salts of 2,4-D |  |  | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H317 H411 |  |  | A | CLP00
607-041-00-9 | 2,4,5-T (ISO); 2,4,5-trichlorophenoxy acetic acid | 202-273-3 | 93-76-5 | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H335 H315 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H335 H315 H410 |  |  |  | CLP00
607-042-00-4 | salts and esters of 2,4,5-T; salts and esters of 2,4,5-trichlorophenoxy acetic acid |  |  | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H335 H315 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H335 H315 H410 |  |  | A | CLP00
607-043-00-X | dicamba (ISO); 2,5-dichloro-6-methoxybenzoic acid; 3,6-dichloro-2-methoxybenzoic acid | 217-635-6 | 1918-00-9 | Acute Tox. 4 Acute Tox. 4 STOT SE 3 STOT SE 3 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 2 | H332 H302 H335 H336 H318 H400 H411 | GHS07 GHS05 GHS09 Dgr | H332 H302 H335 H336 H318 H410 |  | inhalation: ATE = 4,0 mg/L (dusts or mists) oral: ATE = 1 500 mg/kg bw M = 1 |  | CLP00/ATP22
607-044-00-5 | 3,6-dichloro-o-anisic acid, compound with dimethylamine (1:1) [1] potassium 3,6-dichloro-o-anisate [2] | 218-951-7 [1] 233-002-7 [2] | 2300-66-5 [1] 10007-85-9 [2] | Eye Irrit. 2 Aquatic Chronic 3 | H319 H412 | GHS07 Wng | H319 H412 |  |  |  | CLP00
607-045-00-0 | dichlorprop (ISO); 2-(2,4-dichlorophenoxy) propionic acid | 204-390-5 | 120-36-5 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 | H312 H302 H315 H318 | GHS05 GHS07 Dgr | H312 H302 H315 H318 |  |  |  | CLP00
607-046-00-6 | salts of dichlorprop |  |  | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H332 H312 H302 | GHS07 Wng | H332 H312 H302 |  |  | A | CLP00
607-047-00-1 | fenoprop (ISO); 2-(2,4,5-trichlorophenoxy)propionic acid | 202-271-2 | 93-72-1 | Acute Tox. 4 * Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H400 H410 | GHS07 GHS09 Wng | H302 H315 H410 |  |  |  | CLP00
607-048-00-7 | salts of fenoprop; salts of 2-(2,4,5-trichlorophenoxy)propionic acid |  |  | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  |  | A | CLP00
607-049-00-2 | mecoprop (ISO); 2-(4-chloro-o-tolyloxy) propionic acid; (RS)-2-(4-chloro-o-tolyloxy)propionic acid [1] 2-(4-chloro-2-methylphenoxy)propionic acid [2] | 230-386-8 [1] 202-264-4 [2] | 7085-19-0 [1] | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H315 H318 H410 |  | M=100 |  | CLP00
607-050-00-8 | salts of mecoprop |  |  | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H315 H318 H410 |  |  | A | CLP00
607-051-00-3 | MCPA (ISO); 4-chloro-o-tolyloxyacetic acid | 202-360-6 | 94-74-6 | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H315 H318 H410 |  |  |  | CLP00/ATP01
607-052-00-9 | salts and esters of MCPA | - | - | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  |  | A | CLP00/ATP01
607-053-00-4 | MCPB (ISO); 4-(4-chloro-o-tolyloxy) butyric acid | 202-365-3 | 94-81-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-054-00-X | salts and esters of MCPB |  |  | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  | A | CLP00
607-055-00-5 | endothal-sodium (ISO); disodium 7-oxabicyclo(2,2,1)heptane-2,3-dicarboxylate | 204-959-8 | 129-67-9 | Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H301 H312 H335 H315 H319 | GHS06 Dgr | H301 H312 H319 H335 H315 |  |  |  | CLP00
607-056-00-0 | warfarin (ISO); 4-hydroxy-3-(3-oxo-1-phenylbutyl)-2H-chromen-2-one [1] (S)-4-hydroxy-3-(3-oxo-1-phenylbutyl)-2-benzopyrone [2] (R)-4-hydroxy-3-(3-oxo-1-phenylbutyl)-2-benzopyrone [3] | 201-377-6 [1] 226-907-3 [2] 226-908-9 [3] | 81-81-2 [1] 5543-57-7 [2] 5543-58-8 [3] | Repr. 1A Acute Tox. 1 Acute Tox. 1 Acute Tox. 2 STOT RE 1 Aquatic Chronic 2 | H360D H330 H310 H300 H372 (blood) H411 | GHS08 GHS06 GHS09 Dgr | H330 H310 H300 H360D H372 (blood) H411 |  | Repr. 1A; H360D: C ≥ 0,003 % STOT RE 1; H372: C ≥ 0,5 % STOT RE 2; H373: ,05 % ≤ C < 0,5 % |  | CLP00/ATP09
607-057-00-6 | coumachlor (ISO); 3-[1-(4-chlorophenyl)-3-oxobutyl]-4-hydroxycoumarin | 201-378-1 | 81-82-3 | STOT RE 2 * Aquatic Chronic 3 | H373 ** H412 | GHS08 Wng | H373 ** H412 |  |  |  | CLP00
607-058-00-1 | coumafuryl (ISO); fumarin; (RS)-3-(1-(2-furyl)-3-oxobutyl)4-hydroxycoumarin; 4-hydroxy-3-[3-oxo-1-(2-furyl) butyl]coumarin | 204-195-5 | 117-52-2 | Acute Tox. 3 * STOT RE 1 Aquatic Chronic 3 | H301 H372 ** H412 | GHS06 GHS08 Dgr | H301 H372 ** H412 |  |  |  | CLP00
607-059-00-7 | coumatetralyl (ISO); 4-hydroxy-3-(1,2,3,4-tetrahydro-1-naphthyl)coumarin | 227-424-0 | 5836-29-3 | Repr. 1B Acute Tox. 2 Acute Tox. 2 Acute Tox. 3 STOT RE 1 Aquatic Chronic 1 | H360D H330 H300 H311 H372 (blood) H410 | GHS08 GHS06 GHS09 Dgr | H330 H311 H300 H360D H372 (blood) H410 |  | Repr. 1B; H360D: C ≥ 0,003 % STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,1 % ≤ C < 1 % M=10 |  | CLP00/ATP09
607-060-00-2 | dicoumarol; 4,4'-dihydroxy-3,3'-methylenebis(2H-chromen-2-one) | 200-632-9 | 66-76-2 | Acute Tox. 4 * STOT RE 1 Aquatic Chronic 2 | H302 H372 ** H411 | GHS08 GHS07 GHS09 Dgr | H372 ** H302 H411 |  |  |  | CLP00
607-061-00-8 | acrylic acid; prop-2-enoic acid | 201-177-9 | 79-10-7 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A Aquatic Acute 1 | H226 H332 H312 H302 H314 H400 | GHS02 GHS05 GHS07 GHS09 Dgr | H226 H332 H312 H302 H314 H400 |  | STOT SE 3; H335: C ≥ 1 % | D | CLP00
607-062-00-3 | n-butyl acrylate | 205-480-7 | 141-32-2 | Flam. Liq. 3 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H226 H335 H315 H319 H317 | GHS02 GHS07 Wng | H226 H319 H335 H315 H317 |  |  | D | CLP00
607-063-00-9 | isobutyric acid | 201-195-7 | 79-31-2 | Acute Tox. 4 * Acute Tox. 4 * | H312 H302 | GHS07 Wng | H312 H302 |  |  |  | CLP00
607-064-00-4 | benzyl chloroformate | 207-925-0 | 501-53-1 | Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H314 H400 H410 | GHS05 GHS09 Dgr | H314 H410 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
607-065-00-X | bromoacetic acid | 201-175-8 | 79-08-3 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1A Skin Sens. 1 Aquatic Acute 1 | H331 H311 H301 H314 H317 H400 | GHS06 GHS05 GHS09 Dgr | H331 H311 H301 H314 H317 H400 |  |  |  | CLP00/ATP01
607-066-00-5 | dichloroacetic acid | 201-207-0 | 79-43-6 | Skin Corr. 1A Aquatic Acute 1 | H314 H400 | GHS05 GHS09 Dgr | H314 H400 |  |  |  | CLP00
607-067-00-0 | dichloroacetyl chloride | 201-199-9 | 79-36-7 | Skin Corr. 1A Aquatic Acute 1 | H314 H400 | GHS05 GHS09 Dgr | H314 H400 |  |  |  | CLP00
607-068-00-6 | iodoacetic acid | 200-590-1 | 64-69-7 | Acute Tox. 3 * Skin Corr. 1A | H301 H314 | GHS06 GHS05 Dgr | H301 H314 |  |  |  | CLP00
607-069-00-1 | ethyl bromoacetate | 203-290-9 | 105-36-2 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * | H310 H330 H300 | GHS06 Dgr | H330 H310 H300 |  |  |  | CLP00
607-070-00-7 | ethyl chloroacetate | 203-294-0 | 105-39-5 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 | H331 H311 H301 H400 | GHS06 GHS09 Dgr | H331 H311 H301 H400 |  |  |  | CLP00
607-071-00-2 | ethyl methacrylate | 202-597-5 | 97-63-2 | Flam. Liq. 2 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H225 H335 H315 H319 H317 | GHS02 GHS07 Dgr | H225 H319 H335 H315 H317 |  |  | D | CLP00
607-072-00-8 | 2-hydroxyethyl acrylate | 212-454-9 | 818-61-1 | Acute Tox. 3 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 | H311 H314 H317 H400 | GHS06 GHS05 GHS09 Dgr | H311 H314 H317 H400 |  | * Skin Sens. 1; H317: C ≥ 0,2 % | D | CLP00
607-073-00-3 | 4-CPA (ISO); 4-chlorophenoxyacetic acid | 204-581-3 | 122-88-3 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
607-074-00-9 | chlorfenac (ISO); 2,3,6-trichlorophenylacetic acid | 201-599-3 | 85-34-7 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
607-075-00-4 | chlorfenprop-methyl; methyl 2-chloro-3-(4-chlorophenyl)propionate | 238-413-5 | 14437-17-3 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H400 H410 | GHS07 GHS09 Wng | H312 H302 H410 |  |  |  | CLP00
607-076-00-X | dodine (ISO); dodecylguanidinium acetate | 219-459-5 | 2439-10-3 | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H315 H410 |  |  |  | CLP00
607-077-00-5 | erbon (ISO); 2-(2,4,5-trichlorophenoxy)ethyl 2,2-dichloropropionate |  | 136-25-4 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
607-078-00-0 | fluenetil (ISO); 2-fluoroethyl biphenyl-4-ylacetate |  | 4301-50-2 | Acute Tox. 1 Acute Tox. 2 * | H310 H300 | GHS06 Dgr | H310 H300 |  |  |  | CLP00
607-079-00-6 | kelevan (ISO); ethyl 5-(perchloro-5-hydroxypentacyclo[5,3,0,02,6,03,9,04,8]decan-5-yl)-4-oxopentanoate; ethyl 5-(1,2,3,5,6,7,8,9,10,10-decachloro-4-hydroxypentacyclo(5,2,1,02,6,03,9,05,8)dec-4-yl)-4-oxovalerate |  | 4234-79-1 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Chronic 2 | H311 H302 H411 | GHS06 GHS09 Dgr | H311 H302 H411 |  |  |  | CLP00
607-080-00-1 | chloroacetyl chloride | 201-171-6 | 79-04-9 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 Skin Corr. 1A Aquatic Acute 1 | H331 H311 H301 H372 ** H314 H400 | GHS06 GHS08 GHS05 GHS09 Dgr | H331 H311 H301 H372 ** H314 H400 | EUH014 EUH029 |  |  | CLP00
607-081-00-7 | fluoroacetic acid | 205-631-7 | 144-49-0 | Acute Tox. 2 * Aquatic Acute 1 | H300 H400 | GHS06 GHS09 Dgr | H300 H400 |  |  |  | CLP00
607-082-00-2 | fluoroacetates, soluble |  |  | Acute Tox. 2 * Aquatic Acute 1 | H300 H400 | GHS06 GHS09 Dgr | H300 H400 |  |  | A | CLP00
607-083-00-8 | 2,4-DB (ISO); 4-(2,4-dichlorophenoxy)butyric acid | 202-366-9 | 94-82-6 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
607-084-00-3 | salts of 2,4-DB |  |  | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  | A | CLP00
607-085-00-9 | benzyl benzoate | 204-402-9 | 120-51-4 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00/ATP01
607-086-00-4 | diallyl phthalate | 205-016-3 | 131-17-9 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
607-088-00-5 | methacrylic acid; 2-methylpropenoic acid | 201-204-4 | 79-41-4 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A | H312 H302 H314 | GHS05 GHS07 Dgr | H312 H302 H314 |  | STOT SE 3; H335: C ≥ 1 % | D | CLP00
607-089-00-0 | propionic acid … % | 201-176-3 | 79-09-4 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  | STOT SE 3; H335: C ≥ 10 % Skin Corr. 1B; H314: C ≥ 25 % Skin Irrit. 2; H315: 10 % ≤ C < 25 % Eye Irrit. 2; H319: 10 % ≤ C < 25 % | B | CLP00
607-090-00-6 | thioglycolic acid | 200-677-4 | 68-11-1 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B | H331 H311 H301 H314 | GHS06 GHS05 Dgr | H331 H311 H301 H314 |  | * |  | CLP00
607-091-00-1 | trifluoroacetic acid . . . % | 200-929-3 | 76-05-1 | Acute Tox. 4 * Skin Corr. 1A Aquatic Chronic 3 | H332 H314 H412 | GHS05 GHS07 Dgr | H332 H314 H412 |  | * | B | CLP00
607-092-00-7 | methyl lactate [1] methyl (±)-lactate [2] methyl (R)-lactate [3] methyl (S)-(-)-lactate [4] | 208-930-0 [1] 218-449-8 [2] 241-420-6 [3] 248-704-9 [4] | 547-64-8 [1] 2155-30-8 [2] 17392-83-5 [3] 27871-49-4 [4] | Flam. Liq. 3 STOT SE 3 Eye Irrit. 2 | H226 H335 H319 | GHS02 GHS07 Wng | H226 H319 H335 |  |  | C | CLP00
607-093-00-2 | propionyl chloride | 201-170-0 | 79-03-8 | Flam. Liq. 2 Skin Corr. 1B | H225 H314 | GHS02 GHS05 Dgr | H225 H314 | EUH014 |  | B D | CLP00
607-094-00-8 | peracetic acid . . . % | 201-186-8 | 79-21-0 | Org. Perox. D Acute Tox. 2 Acute Tox. 2 Acute Tox. 3 Skin Corr. 1A Aquatic Acute 1 Aquatic Chronic 1 | H242 H330 H310 H301 H314 H400 H410 | GHS02 GHS06 GHS05 GHS09 Dgr | H242 H330 H310 H301 H314 H410 | EUH071 | inhalation: ATE = 0,2 mg/L (dusts or mists) dermal: ATE = 60 mg/kg bw oral: ATE = 80 mg/kg bw STOT SE 3; H335: C ≥ 1 % M = 10 M = 100 | B D T | CLP00/ATP22
607-095-00-3 | maleic acid | 203-742-5 | 110-16-7 | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H302 H335 H315 H319 H317 | GHS07 Wng | H302 H319 H335 H315 H317 |  | Skin Sens. 1; H317: C ≥ 0,1 % |  | CLP00/ATP01
607-096-00-9 | maleic anhydride | 203-571-6 | 108-31-6 | Acute Tox. 4 STOT RE 1 Skin Corr. 1B Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1A | H302 H372 (respiratory system, inhalation) H314 H318 H334 H317 | GHS07 GHS08 GHS05 Dgr | H302 H314 H334 H317 H372 (respiratory system) (inhalation) | EUH071 | Skin Sens. 1A; H317: C ≥ 0,001 % |  | CLP00/ATP13
607-097-00-4 | benzene-1,2,4-tricarboxylic acid 1,2-anhydride; trimellitic anhydride | 209-008-0 | 552-30-7 | STOT SE 3 Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H335 H318 H334 H317 | GHS08 GHS05 GHS07 Dgr | H335 H318 H334 H317 |  |  |  | CLP00
607-098-00-X | benzene-1,2:4,5-tetracarboxylic dianhydride; benzene-1,2:4,5-tetracarboxylic dianhydride; pyromellitic dianhydride | 201-898-9 | 89-32-7 | Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H318 H334 H317 | GHS08 GHS05 Dgr | H318 H334 H317 |  |  |  | CLP00
607-099-00-5 | 1,2,3,6-tetrahydrophthalic anhydride [1] cis-1,2,3,6-tetrahydrophthalic anhydride [2] 3,4,5,6-tetrahydrophthalic anhydride [3] tetrahydrophthalic anhydride [4] | 201-605-4 [1] 213-308-7 [2] 219-374-3 [3] 247-570-9 [4] | 85-43-8 [1] 935-79-5 [2] 2426-02-0 [3] 26266-63-7 [4] | Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H334 H317 H412 | GHS08 GHS05 Dgr | H318 H334 H317 H412 |  |  | C | CLP00
607-100-00-9 | benzophenone-3,3',4,4'-tetracarboxylic dianhydride; 4,4'-carbonyldi(phthalic anhydride) | 219-348-1 | 2421-28-5 | STOT SE 3 Eye Irrit. 2 | H335 H319 | GHS07 Wng | H319 H335 |  | Eye Irrit. 2; H319: C ≥ 1 % STOT SE 3; H335: C ≥ 1 % |  | CLP00
607-101-00-4 | 1,4,5,6,7,7-hexachlorobicyclo [2,2,1]hept-5-ene-2,3-dicarboxylic anhydride chlorendic anhydride | 204-077-3 | 115-27-5 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H335 H315 H319 | GHS07 Wng | H319 H335 H315 |  | Skin Irrit. 2; H315: C ≥ 1 % Eye Irrit. 2; H319: C ≥ 1 % STOT SE 3; H335: C ≥ 1 % |  | CLP00
607-102-00-X | cyclohexane-1,2-dicarboxylic anhydride [1] cis-cyclohexane-1,2-dicarboxylic anhydride [2] trans-cyclohexane-1,2-dicarboxylic anhydride [3] | 201-604-9 [1] 236-086-3 [2] 238-009-9 [3] | 85-42-7 [1] 13149-00-3 [2] 14166-21-3 [3] | Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H318 H334 H317 | GHS08 GHS05 Dgr | H318 H334 H317 |  |  | C | CLP00
607-103-00-5 | succinic anhydride | 203-570-0 | 108-30-5 | Acute Tox. 4 Skin Corr. 1 Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H302 H314 H318 H334 H317 | GHS07 GHS05 GHS08 Dgr | H302 H314 H334 H317 | EUH071 |  |  | CLP00/ATP13
607-104-00-0 | cyclopentane-1,2,3,4-tetracarboxylic dianhydride | 227-964-7 | 6053-68-5 | STOT SE 3 Eye Irrit. 2 | H335 H319 | GHS07 Wng | H319 H335 |  | Eye Irrit. 2; H319: C ≥ 1 % STOT SE 3; H335: C ≥ 1 % |  | CLP00
607-105-00-6 | 8,9,10-trinorborn-5-ene-2,3-dicarboxylic anhydride [1] 1,2,3,6-tetrahydro-3,6-methanophthalic anhydride [2] (1α,2α,3β,6β)-1,2,3,6-tetrahydro-3,6-methanophthalic anhydride [3] | 204-957-7 [1] 212-557-9 [2] 220-384-5 [3] | 129-64-6 [1] 826-62-0 [2] 2746-19-2 [3] | Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H318 H334 H317 | GHS08 GHS05 Dgr | H318 H334 H317 |  |  | C | CLP00
607-106-00-1 | 8,9-dinorborn-5-ene-2,3-dicarboxylic anhydride |  | 123748-85-6 | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H302 H335 H315 H319 H334 | GHS08 GHS07 Dgr | H302 H319 H335 H315 H334 |  | STOT SE 3; H335: C ≥ 10 % | C | CLP00
607-107-00-7 | 2-ethylhexyl acrylate | 203-080-7 | 103-11-7 | STOT SE 3 Skin Irrit. 2 Skin Sens. 1 | H335 H315 H317 | GHS07 Wng | H335 H315 H317 |  |  | D | CLP00
607-108-00-2 | 2-hydroxy-1-methylethylacrylate [1] 2-hydroxypropylacrylate [2] acrylic acid, monoester with propane-1,2-diol [3] | 220-852-9 [1] 213-663-8 [2] 247-118-0 [3] | 2918-23-2 [1] 999-61-1 [2] 25584-83-2 [3] | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Skin Sens. 1 | H331 H311 H301 H314 H317 | GHS06 GHS05 Dgr | H331 H311 H301 H314 H317 |  | * Skin Sens. 1; H317: C ≥ 0,2 % | C D | CLP00
607-109-00-8 | hexamethylene diacrylate; hexane-1,6-diol diacrylate | 235-921-9 | 13048-33-4 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H315 H319 H317 | GHS07 Wng | H319 H315 H317 |  |  | D | CLP00
607-110-00-3 | pentaerythritol triacrylate | 222-540-8 | 3524-68-3 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H315 H319 H317 | GHS07 Wng | H319 H315 H317 |  |  | D | CLP00
607-111-00-9 | 2-ethyl-2-[[(1-oxoallyl)oxy] methyl]-1,3-propanediyl diacrylate; 2,2-bis(acryloyloxymethyl)butyl acrylate; trimethylolpropane triacrylate | 239-701-3 | 15625-89-5 | Carc. 2 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H315 H319 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H315 H319 H317 H410 |  | M = 1 M = 1 | D | CLP00/ATP18
607-112-00-4 | 2,2-dimethyltrimethylene diacrylate; neopentyl glycol diacrylate | 218-741-5 | 2223-82-7 | Acute Tox. 3 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H311 H315 H319 H317 | GHS06 Dgr | H311 H319 H315 H317 |  | * | D | CLP00
607-113-00-X | isobutyl methacrylate | 202-613-0 | 97-86-9 | Flam. Liq. 3 STOT SE 3 Skin Irrit. 2 Skin Sens. 1B | H226 H335 H315 H317 | GHS02 GHS07 Wng | H226 H315 H317 H335 |  |  | D | CLP00/ATP13
607-114-00-5 | ethylene dimethacrylate | 202-617-2 | 97-90-5 | STOT SE 3 Skin Sens. 1 | H335 H317 | GHS07 Wng | H335 H317 |  | STOT SE 3; H335: C ≥ 10 % | D | CLP00
607-115-00-0 | isobutyl acrylate | 203-417-8 | 106-63-8 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Skin Sens. 1 | H226 H332 H312 H315 H317 | GHS02 GHS07 Wng | H226 H332 H312 H315 H317 |  |  | D | CLP00
607-116-00-6 | cyclohexyl acrylate | 221-319-3 | 3066-71-5 | STOT SE 3 Skin Irrit. 2 Aquatic Chronic 2 | H335 H315 H411 | GHS07 GHS09 Wng | H335 H315 H411 |  | STOT SE 3; H335: C ≥ 10 % | D | CLP00
607-117-00-1 | 2,3-epoxypropyl acrylate; glycidyl acrylate | 203-440-3 | 106-90-1 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Skin Sens. 1 | H331 H311 H301 H314 H317 | GHS06 GHS05 Dgr | H331 H311 H301 H314 H317 |  | * Skin Sens. 1; H317: C ≥ 0,2 % | D | CLP00
607-118-00-7 | 1-methyltrimethylene diacrylate; 1,3-butylene glycol diacrylate | 243-105-9 | 19485-03-1 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 | H312 H314 H317 | GHS05 GHS07 Dgr | H312 H314 H317 |  |  | D | CLP00
607-119-00-2 | tetramethylene diacrylate; 1,4-butyleneglycol diacrylate | 213-979-6 | 1070-70-8 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 | H312 H314 H317 | GHS05 GHS07 Dgr | H312 H314 H317 |  |  | D | CLP00
607-120-00-8 | 2,2'-oxydiethyl diacrylate; diethylene glycol diacrylate | 223-791-6 | 4074-88-8 | Acute Tox. 3 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H311 H315 H319 H317 | GHS06 Dgr | H311 H319 H315 H317 |  | * Skin Sens. 1; H317: C ≥ 0,2 % | D | CLP00
607-121-00-3 | 8,9,10-trinorborn-2-yl acrylate |  | 10027-06-2 | Acute Tox. 4 * Skin Irrit. 2 Skin Sens. 1 | H312 H315 H317 | GHS07 Wng | H312 H315 H317 |  |  | D | CLP00
607-122-00-9 | pentaerythritol tetraacrylate | 225-644-1 | 4986-89-4 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H315 H319 H317 | GHS07 Wng | H319 H315 H317 |  |  | D | CLP00
607-123-00-4 | 2,3-epoxypropyl methacrylate; glycidyl methacrylate | 203-441-9 | 106-91-2 | Carc. 1B Muta. 2 Repr. 1B Acute Tox. 3 Acute Tox. 4 STOT SE 3 STOT RE 1 Skin Corr. 1C Eye Dam. 1 Skin Sens. 1 | H350 H341 H360F H311 H302 H335 H372 (respiratory tract) (Inhalation) H314 H318 H317 | GHS08 GHS06 GHS05 Dgr | H311 H302 H314 H317 H341 H350 H360F H335 H372 (respiratory tract) (Inhalation) |  |  | D | CLP00/ATP10
607-124-00-X | 2-hydroxyethyl methacrylate | 212-782-2 | 868-77-9 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H315 H319 H317 | GHS07 Wng | H319 H315 H317 |  |  | D | CLP00
607-125-00-5 | 2-hydroxypropyl methacrylate [1] 3-hydroxypropyl methacrylate [2] | 213-090-3 [1] 220-426-2 [2] | 923-26-2 [1] 2761-09-3 [2] | Eye Irrit. 2 Skin Sens. 1 | H319 H317 | GHS07 Wng | H319 H317 |  |  | C D | CLP00
607-126-00-0 | 2,2'-(ethylenedioxy)diethyl diacrylate; triethylene glycol diacrylate | 216-853-9 | 1680-21-3 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H315 H319 H317 | GHS07 Wng | H319 H315 H317 |  |  | D | CLP00
607-127-00-6 | 2-diethylaminoethyl methacrylate | 203-275-7 | 105-16-8 | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H332 H315 H319 H317 | GHS07 Wng | H332 H319 H315 H317 |  |  | D | CLP00
607-128-00-1 | 2-tert-butylaminoethyl methacrylate | 223-228-4 | 3775-90-4 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H315 H319 H317 | GHS07 Wng | H319 H315 H317 |  |  | D | CLP00
607-129-00-7 | ethyl lactate; ethyl DL-lactate [1] ethyl (S)-2-hydroxypropionate; ethyl L-lactate; ethyl-(S)-lactate [2] | 202-598-0 [1] 211-694-1 [2] | 97-64-3 [1] 687-47-8 [2] | Flam. Liq. 3 STOT SE 3 Eye Dam. 1 | H226 H335 H318 | GHS02 GHS05 GHS07 Dgr | H226 H335 H318 |  |  | C | CLP00
607-130-00-2 | pentyl acetate [1] isopentyl acetate [2] 1-methylbutyl acetate [3] 2-methylbutyl acetat [4] 2(or 3)-methylbutyl acetate [5] | 211-047-3 [1] 204-662-3 [2] 210-946-8 [3] 210-843-8 [4] 282-263-3 [5] | 628-63-7 [1] 123-92-2 [2] 626-38-0 [3] 624-41-9 [4] 84145-37-9 [5] | Flam. Liq. 3 | H226 | GHS02 Wng | H226 | EUH066 |  | C | CLP00
607-131-00-8 | isopentyl propionate [1] pentyl propionate [2] 2-methylbutyl propionate [3] | 203-322-1 [1] 210-852-7 [2] 219-449-0 [3] | 105-68-0 [1] 624-54-4 [2] 2438-20-2 [3] | Flam. Liq. 3 | H226 | GHS02 Wng | H226 |  |  | C | CLP00
607-132-00-3 | 2-dimethylaminoethyl methacrylate | 220-688-8 | 2867-47-2 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H312 H302 H315 H319 H317 | GHS07 Wng | H312 H302 H319 H315 H317 |  |  | D | CLP00
607-133-00-9 | monoalkyl or monoaryl or monoalkylaryl esters of acrylic acid with the exception of those specified elsewhere in this Annex |  |  | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H335 H315 H319 H411 | GHS07 GHS09 Wng | H319 H335 H315 H411 |  | STOT SE 3; H335: C ≥ 10 % | A | CLP00
607-134-00-4 | monoalkyl or monoaryl or monoalkyaryl esters of methacrylic acid with the exception of those specified elsewhere in this Annex |  |  | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H335 H315 H319 | GHS07 Wng | H319 H335 H315 |  | STOT SE 3; H335: C ≥ 10 % | A | CLP00
607-135-00-X | butyric acid | 203-532-3 | 107-92-6 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
607-136-00-5 | butyryl chloride | 205-498-5 | 141-75-3 | Flam. Liq. 2 Skin Corr. 1B | H225 H314 | GHS02 GHS05 Dgr | H225 H314 |  |  |  | CLP00
607-137-00-0 | methyl acetoacetate | 203-299-8 | 105-45-3 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
607-138-00-6 | butyl chloroformate; chloroformic acid butyl ester | 209-750-5 | 592-34-7 | Flam. Liq. 3 Acute Tox. 3 * Skin Corr. 1B | H226 H331 H314 | GHS02 GHS06 GHS05 Dgr | H226 H331 H314 |  |  |  | CLP00
607-139-00-1 | 2-chloropropionic acid | 209-952-3 | 598-78-7 | Acute Tox. 4 * Skin Corr. 1A | H302 H314 | GHS05 GHS07 Dgr | H302 H314 |  |  |  | CLP00
607-140-00-7 | isobutyryl chloride | 201-194-1 | 79-30-1 | Flam. Liq. 2 Skin Corr. 1A | H225 H314 | GHS02 GHS05 Dgr | H225 H314 |  |  |  | CLP00
607-141-00-2 | oxydiethylene bis(chloroformate) | 203-430-9 | 106-75-2 | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 2 | H302 H315 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H315 H318 H411 |  |  |  | CLP00
607-142-00-8 | propyl chloroformate; chloroformic acid propylester; n-propyl chloroformate | 203-687-7 | 109-61-5 | Flam. Liq. 2 Acute Tox. 3 * Skin Corr. 1B | H225 H331 H314 | GHS02 GHS06 GHS05 Dgr | H225 H331 H314 |  |  |  | CLP00/ATP01
607-143-00-3 | valeric acid | 203-677-2 | 109-52-4 | Skin Corr. 1B Aquatic Chronic 3 | H314 H412 | GHS05 Dgr | H314 H412 |  |  |  | CLP00
607-144-00-9 | adipic acid | 204-673-3 | 124-04-9 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
607-145-00-4 | methanesulphonic acid | 200-898-6 | 75-75-2 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
607-146-00-X | fumaric acid | 203-743-0 | 110-17-8 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
607-147-00-5 | oxalic acid diethylester; diethyl oxalate | 202-464-1 | 95-92-1 | Acute Tox. 4 * Eye Irrit. 2 | H302 H319 | GHS07 Wng | H302 H319 |  |  |  | CLP00
607-148-00-0 | guanidinium chloride; guanadine hydrochloride | 200-002-3 | 50-01-1 | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 | H302 H315 H319 | GHS07 Wng | H302 H319 H315 |  |  |  | CLP00
607-149-00-6 | urethane (INN); ethyl carbamate | 200-123-1 | 51-79-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
607-150-00-1 | endothal (ISO); 7-oxabicyclo(2,2,1)heptane-2,3-dicarboxylic acid | 205-660-5 | 145-73-3 | Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H301 H312 H335 H315 H319 | GHS06 Dgr | H301 H312 H319 H335 H315 |  |  |  | CLP00
607-151-00-7 | propargite (ISO); 2-(4-tert-butylphenoxy) cyclohexyl prop-2-ynyl sulphite | 219-006-1 | 2312-35-8 | Carc. 2 Acute Tox. 3 * Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H331 H315 H318 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H351 H331 H315 H318 H410 |  | M=10 |  | CLP00
607-152-00-2 | 2,3,6-TBA (ISO); 2,3,6-trichlorobenzoic acid | 200-026-4 | 50-31-7 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
607-153-00-8 | benazolin (ISO); 4-chloro-2,3-dihydro-2-oxo-1,3-benzothiazol-3-ylacetic acid | 223-297-0 | 3813-05-6 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 3 | H315 H319 H412 | GHS07 Wng | H319 H315 H412 |  |  |  | CLP00
607-154-00-3 | ethyl N-benzoyl-N-(3,4-dichlorophenyl)-DL-alaninate; benzoylprop-ethyl (ISO) | 244-845-5 | 22212-55-1 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
607-155-00-9 | 3-(3-amino-5-(1-methylguanidino)-1-oxopentylamino-6-(4-amino-2-oxo-2,3-dihydro-pyrimidin-1-yl)-2,3-dihydro-(6H)-pyran-2-carboxylic acid; blasticidin-s |  | 2079-00-7 | Acute Tox. 2 * | H300 | GHS06 Dgr | H300 |  |  |  | CLP00
607-156-00-4 | chlorfenson (ISO); 4-chlorophenyl 4-chlorobenzenesulfonate | 201-270-4 | 80-33-1 | Acute Tox. 4 * Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H400 H410 | GHS07 GHS09 Wng | H302 H315 H410 |  |  |  | CLP00
607-157-00-X | difenacoum (ISO); 3-(3-biphenyl-4-yl-1,2,3,4-tetrahydro-1-naphthyl)-4-hydroxycoumarin | 259-978-4 | 56073-07-5 | Repr. 1B Acute Tox. 1 Acute Tox. 1 Acute Tox. 1 STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D H330 H310 H300 H372 (blood) H400 H410 | GHS08 GHS06 GHS09 Dgr | H330 H310 H300 H360D H372 (blood) H410 |  | Repr. 1B; H360D: C ≥ 0,003 % STOT RE 1; H372: C ≥ 0,02 % STOT RE 2; H373: 0,002 % ≤ C < 0,02 % M=10 M=10 |  | CLP00/ATP09
607-158-00-5 | sodium salt of chloroacetic acid; sodium chloroacetate | 223-498-3 | 3926-62-3 | Acute Tox. 3 * Skin Irrit. 2 Aquatic Acute 1 | H301 H315 H400 | GHS06 GHS09 Dgr | H301 H315 H400 |  |  |  | CLP00
607-159-00-0 | chlorobenzilate (ISO); ethyl 2,2-di(4-chlorophenyl)-2-hydroxyacetate; ethyl 4,4'-dichlorobenzilate | 208-110-2 | 510-15-6 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
607-160-00-6 | isobutyl 2-(4-(4-chlorophenoxy)phenoxy)propionate; clofop-isobutyl (ISO) |  | 51337-71-4 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
607-161-00-1 | diethanolamine salt of 4-CPA |  |  | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
607-162-00-7 | dalapon; 2,2-dichloropropionic acid [1] dalapon-sodium; sodium 2,2-dichloropropionate [2] | 200-923-0 [1] 204-828-5 [2] | 75-99-0 [1] 127-20-8 [2] | Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H315 H318 H412 | GHS05 Dgr | H315 H318 H412 |  |  |  | CLP00/ATP01
607-163-00-2 | 3-acetyl-6-methyl-2H-pyran-2,4(3H)-dione; dehydracetic acid | 208-293-9 | 520-45-6 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
607-164-00-8 | sodium 1-(3,4-dihydro-6-methyl-2,4-dioxo-2H-pyran-3-ylidene)ethonolate; sodium dehydracetate | 224-580-1 | 4418-26-2 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
607-165-00-3 | diclofop-methyl (ISO); methyl 2-(4-(2,4-dichlorophenoxy)phenoxy)propionate; methyl (RS)-2-[4-(2,4-dichlorophenoxy)phenoxy]propionate | 257-141-8 | 51338-27-3 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
607-166-00-9 | medinoterb acetate (ISO); 6-tert-butyl-3-methyl-2,4-dinitrophenyl acetate | 219-634-6 | 2487-01-6 | Acute Tox. 3 * Acute Tox. 4 * | H301 H312 | GHS06 Dgr | H301 H312 |  |  |  | CLP00
607-167-00-4 | sodium 3-chloroacrylate |  | 4312-97-4 | Acute Tox. 4 * Acute Tox. 4 * | H312 H302 | GHS07 Wng | H312 H302 |  |  |  | CLP00
607-168-00-X | dipropyl 6,7-methylenedioxy-1,2,3,4-tetrahydro-3-methylnaphthalene-1,2-dicarboxylate; propylisome |  | 83-59-0 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H311 H302 H400 H410 | GHS06 GHS09 Dgr | H311 H302 H410 |  |  |  | CLP00
607-169-00-5 | sodium fluoroacetate | 200-548-2 | 62-74-8 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Aquatic Acute 1 | H310 H330 H300 H400 | GHS06 GHS09 Dgr | H330 H310 H300 H400 |  |  |  | CLP00
607-170-00-0 | bis(1,2,3-trithiacyclohexyldimethylammonium) oxalate; thiocyclam-oxalate | 250-859-2 | 31895-22-4 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H400 H410 | GHS07 GHS09 Wng | H312 H302 H410 |  |  |  | CLP00
607-172-00-1 | brodifacoum (ISO); 4-hydroxy-3-(3-(4'-bromo-4-biphenylyl)-1,2,3,4-tetrahydro-1-naphthyl)coumarin | 259-980-5 | 56073-10-0 | Repr. 1A Acute Tox. 1 Acute Tox. 1 Acute Tox. 1 STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D H330 H310 H300 H372 (blood) H400 H410 | GHS08 GHS06 GHS09 Dgr | H330 H310 H300 H360D H372 (blood) H410 |  | Repr. 1A; H360D: C ≥ 0,003 % STOT RE 1; H372: C ≥ 0,02 % STOT RE 2; H373: 0,002 % ≤ C < 0,02 % M=10 M=10 |  | CLP00/ATP09
607-173-00-7 | dimethyl (3-methyl-4-(5-nitro-3-ethoxycarbonyl-2-thienyl)azo)phenylnitrilodipropionate | 400-460-6 |  | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
607-174-00-2 | reaction mass of dodecyl 3-(2,2,4,4-tetramethyl-21-oxo-7-oxa-3,20-diazadispiro(5,1,11,2)henicosan-20-yl)propionate and tetradecyl 3-(2,2,4,4-tetramethyl-21-oxo-7-oxa-3,20-diazadispiro(5,1,11,2)henicosan-20-yl)propionate | 400-580-9 |  | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
607-175-00-8 | methyl 2-(2-nitrobenzylidene)acetoacetate | 400-650-9 | 39562-27-1 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-176-00-3 | reaction mass of α-3-(3-(2H-benzotriazol-2-yl)-5-tert-butyl-4-hydroxyphenyl)propionyl-ω-hydroxypoly(oxyethylene) and α-3-(3-(2H-benzotriazol-2-yl)-5-tert-butyl-4-hydroxyphenyl)propionyl-ω-3-(3-(2H-benzotriazol-2-yl)-5-tert-butyl-4-hydroxyphenyl)propionyloxypoly(oxyethylene) | 400-830-7 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-177-00-9 | tribenuron-methyl (ISO); methyl 2-[N-(4-methoxy-6-methyl-1,3,5-triazin-2-yl)-N-methylcarbamoylsulfamoyl]benzoate | 401-190-1 | 101200-48-0 | STOT RE 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H317 H373 H410 |  | M=100 M=100 |  | CLP00/ATP15
607-178-00-4 | methyl α-((4,6-dimethoxypyrimidin-2-yl)ureidosulphonyl)-o-toluate | 401-340-6 | 83055-99-6 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-179-00-X | (benzothiazol-2-ylthio)succinic acid | 401-450-4 | 95154-01-1 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-180-00-5 | potassium 2-hydroxycarbazole-1-carboxylate | 401-630-2 | 96566-70-0 | Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 Aquatic Chronic 3 | H302 H335 H319 H412 | GHS07 Wng | H302 H319 H335 H412 |  |  |  | CLP00
607-181-00-0 | 3,5-dichloro-2,4-difluorobenzoyl fluoride | 401-800-6 | 101513-70-6 | Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 3 | H331 H302 H314 H317 H412 | GHS06 GHS05 Dgr | H331 H314 H302 H317 H412 | EUH029 |  |  | CLP00
607-182-00-6 | methyl 3-sulphamoyl-2-thenoate | 402-050-2 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-183-00-1 | zinc 2-hydroxy-5-C13-18alkylbenzoate | 402-280-3 |  | Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H315 H319 H411 | GHS07 GHS09 Wng | H319 H315 H411 |  |  |  | CLP00
607-184-00-7 | S-(3-trimethoxysilyl)propyl 19-isocyanato-11-(6-isocyanatohexyl)-10,12-dioxo-2,9,11,13-tetraazanonadecanethioate | 402-290-8 | 85702-90-5 | Flam. Liq. 3 Resp. Sens. 1 Skin Sens. 1 | H226 H334 H317 | GHS02 GHS08 Dgr | H226 H334 H317 |  |  |  | CLP00
607-185-00-2 | ethyl trans-3-dimethylaminoacrylate | 402-650-4 | 1117-37-9 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-186-00-8 | quinclorac (ISO); 3,7-dichloroquinoline-8-carboxylic acid | 402-780-1 | 84087-01-4 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-187-00-3 | bis(2,2,6,6-tetramethyl-4-piperidyl) succinate | 402-940-0 | 62782-03-0 | Eye Irrit. 2 Aquatic Chronic 3 | H319 H412 | GHS07 Wng | H319 H412 |  |  |  | CLP00
607-188-00-9 | hydrogen sodium N-carboxylatoethyl-N-octadec-9-enylmaleamate | 402-970-4 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-189-00-4 | trimethylenediaminetetraacetic acid | 400-400-9 | 1939-36-2 | Acute Tox. 4 * Eye Dam. 1 | H302 H318 | GHS05 GHS07 Dgr | H302 H318 |  |  |  | CLP00/ATP01
607-190-00-X | methyl acrylamidomethoxyacetate (containing ≥ 0,1 % acrylamid) | 401-890-7 | 77402-03-0 | Carc. 1B Muta. 1B Acute Tox. 4 * Eye Irrit. 2 | H350 H340 H302 H319 | GHS08 GHS07 Dgr | H350 H340 H302 H319 |  |  |  | CLP00
607-191-00-5 | isobutyl 3,4-epoxybutyrate | 401-920-9 | 100181-71-3 | Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H317 H400 H410 | GHS07 GHS09 Wng | H315 H317 H410 |  |  |  | CLP00
607-192-00-0 | disodium N-carboxymethyl-N-(2-(2-hydroxyethoxy)ethyl)glycinate | 402-360-8 | 92511-22-3 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
607-194-00-1 | propylene carbonate | 203-572-1 | 108-32-7 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
607-195-00-7 | 2-methoxy-1-methylethyl acetate | 203-603-9 | 108-65-6 | Flam. Liq. 3 | H226 | GHS02 Wng | H226 |  |  |  | CLP00/ATP01
607-196-00-2 | heptanoic acid | 203-838-7 | 111-14-8 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
607-197-00-8 | nonanoic acid | 203-931-2 | 112-05-0 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 3 | H315 H319 H412 | GHS07 Wng | H315 H319 H412 |  |  |  | CLP00/ATP07
607-198-00-3 | propyl 3,4,5-trihydroxybenzoate | 204-498-2 | 121-79-9 | Acute Tox. 4 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  | oral: ATE = 1 700 mg/kg bw M = 1 M = 1 |  | CLP00/ATP22
607-199-00-9 | octyl 3,4,5-trihydroxybenzoate | 213-853-0 | 1034-01-1 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  |  |  | CLP00
607-200-00-2 | dodecyl 3,4,5-trihydroxybenzoate | 214-620-6 | 1166-52-5 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-201-00-8 | thiocarbonyl chloride | 207-341-6 | 463-71-8 | Acute Tox. 3 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H331 H302 H335 H315 H319 | GHS06 Dgr | H331 H302 H319 H335 H315 |  |  |  | CLP00
607-203-00-9 | 2-ethylhexyl[[[3,5-bis(1,1-dimethylethyl)-4-hydroxyphenyl]methyl]thio]acetate | 279-452-8 | 80387-97-9 | Repr. 1B Skin Sens. 1 Aquatic Chronic 3 | H360D *** H317 H412 | GHS08 GHS07 Dgr | H360D *** H317 H412 |  |  |  | CLP00
607-204-00-4 | (chlorophenyl)(chlorotolyl)methane, mixed isomers | 400-140-6 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-205-00-X | methyl chloroacetate | 202-501-1 | 96-34-4 | Flam. Liq. 3 Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 | H226 H331 H301 H335 H315 H318 | GHS02 GHS06 GHS05 Dgr | H226 H331 H301 H335 H315 H318 |  |  |  | CLP00
607-206-00-5 | isopropyl chloroacetate | 203-301-7 | 105-48-6 | Flam. Liq. 3 Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H226 H301 H335 H315 H319 | GHS02 GHS06 Dgr | H226 H301 H319 H335 H315 |  |  |  | CLP00
607-207-00-0 | haloxyfop-etotyl (ISO); 2-ethoxyethyl 2-(4-(3-chloro-5-trifluoromethyl-2-pyridyloxy)phenoxy)propionate; haloxyfop-(2-ethoxyethyl) | 402-560-5 | 87237-48-7 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
607-208-00-6 | 4,8,12-trimethyltrideca-3,7,11-trienoic acid, mixed isomers | 403-000-2 | 91853-67-7 | Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H400 H410 | GHS07 GHS09 Wng | H315 H410 |  |  |  | CLP00
607-209-00-1 | reaction mass of O,O'-diisopropyl (pentathio)dithioformate and O,O'-diisopropyl (trithio)dithioformate and O,O'-diisopropyl (tetrathio)dithioformate | 403-030-6 |  | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
607-210-00-7 | methyl acrylamidoglycolate (containing ≥ 0,1 % acrylamide) | 403-230-3 | 77402-05-2 | Carc. 1B Muta. 1B Skin Corr. 1B Skin Sens. 1 | H350 H340 H314 H317 | GHS08 GHS05 GHS07 Dgr | H350 H340 H314 H317 |  |  |  | CLP00
607-211-00-2 | methyl 3-(3-tert-butyl-4-hydroxy-5-methylphenyl)propionate | 403-270-1 | 6386-39-6 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
607-212-00-8 | poly(oxypropylenecarbonyl-co-oxy(ethylethylene)carbonyl), containing 27 % hydroxyvalerate | 403-300-3 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-213-00-3 | ethyl 3,3-bis(tert-pentylperoxy)butyrate | 403-320-2 | 67567-23-1 | Flam. Liq. 3 Org. Perox. D **** Aquatic Chronic 2 | H226 H242 H411 | GHS02 GHS09 Dgr | H242 H226 H411 |  |  |  | CLP00/ATP01
607-214-00-9 | N,N-hydrazinodiacetic acid | 403-510-5 | 19247-05-3 | Acute Tox. 3 * STOT RE 2 * Skin Sens. 1 Aquatic Chronic 3 | H301 H373 ** H317 H412 | GHS06 GHS08 Dgr | H301 H373 ** H317 H412 |  |  |  | CLP00
607-215-00-4 | 3-(3-tert-butyl-4-hydroxyphenyl)propionic acid | 403-920-4 | 107551-67-7 | Acute Tox. 4 * Eye Irrit. 2 | H302 H319 | GHS07 Wng | H302 H319 |  |  |  | CLP00
607-216-00-X | glutamic acid, reaction products with N-(C12-14-alkyl)propylenediamine | 403-950-8 | - | Acute Tox. 2 * Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 | H330 H302 H314 H400 | GHS06 GHS05 GHS09 Dgr | H330 H302 H314 H400 |  |  |  | CLP00/ATP01
607-217-00-5 | 2-ethoxyethyl 2-(4-(2,6-dihydro-2,6-dioxo-7-phenyl-1,5-dioxaindacen-3-yl)phenoxy)acetate | 403-960-2 |  | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
607-218-00-0 | dichlorprop-P (ISO); (+)-R-2-(2,4-dichlorophenoxy)propionic acid | 403-980-1 | 15165-67-0 | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H302 H315 H318 H317 | GHS05 GHS07 Dgr | H302 H315 H318 H317 |  |  |  | CLP00
607-219-00-6 | bis(2-ethylhexyl) dithiodiacetate | 404-510-8 | 62268-47-7 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | CLP00
607-221-00-7 | 6-docosyloxy-1-hydroxy-4-(1-(4-hydroxy-3-methylphenanthren-1-yl)-3-oxo-2-oxaphenalen-1-yl)naphthalene-2-carboxylic acid | 404-550-6 |  | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
607-222-00-2 | 6-(2,3-dimethylmaleimido)hexyl methacrylate | 404-870-6 | 63740-41-0 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-223-00-8 | transfluthrin (ISO); 2,3,5,6-tetrafluorobenzyl (1R,3S)-3-(2,2-dichlorovinyl)-2,2-dimethylcyclopropanecarboxylate | 405-060-5 | 118712-89-3 | Carc. 2 Acute Tox. 4 STOT SE 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H370 (nervous system) H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H302 H370 (nervous system) H410 | EUH066 | oral: ATE = 580 mg/kg bw M = 1000 M = 1000 |  | CLP00/ATP21
607-224-00-3 | methyl 2-(3-nitrobenzylidene)acetoacetate | 405-270-7 | 39562-17-9 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
607-225-00-9 | 3-azidosulfonylbenzoic acid | 405-310-3 | 15980-11-7 | STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Self-React. C **** | H373 ** H318 H317 H241 | GHS02 GHS08 GHS05 GHS07 Dgr | H241 H373 ** H318 H317 |  |  |  | CLP00
607-226-00-4 | reaction mass of 2-acryloyloxyethyl hydrogen cyclohexane-1,2-dicarboxylate and 2-methacryloyloxyethyl hydrogen cyclohexane-1,2-dicarboxylate | 405-360-6 |  | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H315 H318 H317 H412 | GHS05 GHS07 Dgr | H315 H318 H317 H412 |  |  |  | CLP00
607-227-00-X | potassium 2-amino-2-methylpropionate octahydrate | 405-560-3 | 120447-91-8 | Acute Tox. 4 * Skin Corr. 1A | H302 H314 | GHS05 GHS07 Dgr | H302 H314 |  |  |  | CLP00
607-228-00-5 | bis(2-methoxyethyl) phthalate | 204-212-6 | 117-82-8 | Repr. 1B | H360Df | GHS08 Dgr | H360Df |  |  |  | CLP00
607-229-00-0 | diethylcarbamoyl chloride | 201-798-5 | 88-10-8 | Carc. 2 Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H351 H332 H302 H335 H315 H319 | GHS08 GHS07 Wng | H351 H332 H302 H319 H335 H315 |  |  |  | CLP00
607-230-00-6 | 2-ethylhexanoic acid and its salts, with the exception of those specified elsewhere in this Annex |  |  | Repr. 1B | H360D | GHS08 Dgr | H360D |  |  | A X 12 | CLP00/ATP20
607-231-00-1 | clopyralid (ISO); 3,6-dichloropyridine-2-carboxylic acid | 216-935-4 | 1702-17-6 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00/ATP01
607-232-00-7 | pyridate (ISO); O-(6-chloro-3-phenylpyridazin-4-yl) S-octyl thiocarbonate | 259-686-7 | 55512-33-9 | Acute Tox. 4 Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H317 H400 H410 | GHS07 GHS09 Wng | H302 H315 H317 H410 |  | Oral: ATE = 500 mg/kg bw M=1 M=10 |  | CLP00/ATP14
607-233-00-2 | hexyl acrylate | 219-698-5 | 2499-95-8 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H335 H315 H319 H317 H411 | GHS07 GHS09 Wng | H319 H335 H315 H317 H411 |  |  |  | CLP00
607-234-00-8 | flurenol (ISO); 9-hydroxy-9H-fluorene-9-carboxylic acid | 207-397-1 | 467-69-6 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-235-00-3 | mecrilate; methyl 2-cyanoacrylate | 205-275-2 | 137-05-3 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H335 H315 H319 | GHS07 Wng | H319 H335 H315 |  | STOT SE 3; H335: C ≥ 10 % |  | CLP00
607-236-00-9 | ethyl 2-cyanoacrylate | 230-391-5 | 7085-85-0 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H335 H315 H319 | GHS07 Wng | H319 H335 H315 |  | STOT SE 3; H335: C ≥ 10 % |  | CLP00
607-237-00-4 | benzyl 2-chloro-4-(trifluoromethyl)thiazole-5-carboxylate; flurazole | 276-942-3 | 72850-64-7 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-238-00-X | tau-fluvalinate (ISO); cyano-(3-phenoxyphenyl)methyl N-[2-chloro-4-(trifluoromethyl)phenyl]-D-valinate |  | 102851-06-9 | Acute Tox. 4 * Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H315 H400 H410 | GHS07 GHS09 Wng | H302 H315 H410 |  |  |  | CLP00
607-239-00-5 | fenpropathrin (ISO); α-cyano-3-phenoxybenzyl 2,2,3,3-tetramethylcyclopropanecarboxylate | 254-485-0 | 39515-41-8 | Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H330 H301 H312 H400 H410 | GHS06 GHS09 Dgr | H330 H301 H312 H410 |  |  |  | CLP00
607-240-00-0 | cis-1,2,3,6-tetrahydro-4-methylphthalic anhydride [1] 1,2,3,6-tetrahydro-4-methylphthalic anhydride [2] 1,2,3,6-tetrahydro-3-methylphthalic anhydride [3] tetrahydromethylphthalic anhydride [4] 1,2,3,6-tetrahydromethylphthalic anhydride [5] tetrahydro-4-methylphthalic anhydride [6] 2,3,5,6-tetrahydro-2-methylphthalic anhydride [7] | 216-906-6 [1] 222-323-8 [2] 226-247-6 [3] 234-290-7 [4] 247-830-1 [5] 251-823-9 [6] 255-853-3 [7] | 1694-82-2 [1] 3425-89-6 [2] 5333-84-6 [3] 11070-44-3 [4] 26590-20-5 [5] 34090-76-1 [6] 42498-58-8 [7] | Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H318 H334 H317 | GHS08 GHS05 Dgr | H318 H334 H317 |  |  | C | CLP00
607-241-00-6 | hexahydro-4-methylphthalic anhydride [1] hexahydromethylphthalic anhydride [2] hexahydro-1-methylphthalic anhydride [3] hexahydro-3-methylphthalic anhydride [4] | 243-072-0 [1] 247-094-1 [2] 256-356-4 [3] 260-566-1 [4] | 19438-60-9 [1] 25550-51-0 [2] 48122-14-1 [3] 57110-29-9 [4] | Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H318 H334 H317 | GHS08 GHS05 Dgr | H318 H334 H317 |  |  | C | CLP00
607-242-00-1 | tetrachlorophthalic anhydride | 204-171-4 | 117-08-8 | Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H334 H317 H400 H410 | GHS08 GHS05 GHS09 Dgr | H318 H334 H317 H410 |  |  |  | CLP00
607-243-00-7 | sodium 3,6-dichloro-o-anisate [1] 3,6-dichloro-o-anisic acid, compound with 2,2'-iminodiethanol (1:1) [2] 3,6-dichloro-o-anisic acid, compound with 2-aminoethanol (1:1) [3] | 217-846-3 [1] 246-590-5 [2] 258-527-9 [3] | 1982-69-0 [1] 25059-78-3 [2] 53404-28-7 [3] | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-244-00-2 | isooctyl acrylate | 249-707-8 | 29590-42-9 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H335 H315 H319 H400 H410 | GHS07 GHS09 Wng | H319 H335 H315 H410 |  | STOT SE 3; H335: C ≥ 10 % |  | CLP00
607-245-00-8 | tert-butyl acrylate | 216-768-7 | 1663-39-4 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H225 H332 H312 H302 H335 H315 H317 H411 | GHS02 GHS07 GHS09 Dgr | H225 H302 H312 H332 H315 H317 H335 H411 |  |  | D | CLP00/ATP01corr
607-246-00-3 | allyl methacrylate; 2-methyl-2-propenoic acid 2-propenyl ester | 202-473-0 | 96-05-9 | Flam. Liq. 3 Acute Tox. 2 Acute Tox. 3 Acute Tox. 4 Aquatic Acute 1 | H226 H330 H311 H302 H400 | GHS02 GHS06 GHS09 Dgr | H226 H330 H311 H302 H400 |  | inhalation: ATE = 1.5 mg/L(vapours) oral: ATE = 400 mg/kg bw dermal: ATE = 300 mg/kg bw |  | CLP00/ATP21
607-247-00-9 | dodecyl methacrylate | 205-570-6 | 142-90-5 | STOT SE 3 | H335 | GHS07 Wng | H335 |  | STOT SE 3; H335: C ≥ 10 % |  | CLP00/ATP14
607-248-00-4 | naptalam-sodium (ISO); sodium N-naphth-1-ylphthalamate | 205-073-4 | 132-67-2 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
607-249-00-X | (1-methyl-1,2-ethanediyl)bis[oxy(methyl-2,1-ethanediyl)] diacrylate | 256-032-2 | 42978-66-5 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H335 H315 H319 H317 H411 | GHS07 GHS09 Wng | H319 H335 H315 H317 H411 |  | STOT SE 3; H335: C ≥ 10 % |  | CLP00
607-250-00-5 | 4H-3,1-benzoxazine-2,4(1H)-dione | 204-255-0 | 118-48-9 | Eye Irrit. 2 Skin Sens. 1 | H319 H317 | GHS07 Wng | H319 H317 |  |  |  | CLP00
607-251-00-0 | 2-methoxypropyl acetate | 274-724-2 | 70657-70-4 | Flam. Liq. 3 Repr. 1B STOT SE 3 | H226 H360D *** H335 | GHS02 GHS08 GHS07 Dgr | H226 H360D *** H335 |  |  |  | CLP00
607-252-00-6 | lambda-cyhalothrin (ISO); reaction mass of (S)-α-cyano-3-phenoxybenzyl(Z)-(1R)-cis-3-(2-chloro-3,3,3-trifluoropropenyl)-2,2-dimethylcyclopropanecarboxylate and (R)-α-cyano-3-phenoxybenzyl (Z)-(1S)-cis-3-(2-chloro-3,3,3-trifluoropropenyl)-2,2-dimethylcyclopropanecarboxylate (1:1) | 415-130-7 | 91465-08-6 | Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H330 H301 H312 H400 H410 | GHS06 GHS09 Dgr | H330 H301 H312 H410 |  | M=10000 |  | CLP00/ATP01
607-253-00-1 | cyfluthrin (ISO); α-cyano-4-fluoro-3-phenoxybenzyl-3-(2,2-dichlorovinyl)-2,2-dimethylcyclopropanecarboxylate | 269-855-7 | 68359-37-5 | Lact. Acute Tox. 2 Acute Tox. 2 STOT SE 1 Aquatic Acute 1 Aquatic Chronic 1 | H362 H330 H300 H370 (nervous system) H400 H410 | GHS06 GHS08 GHS09 Dgr | H362 H330 H300 H370 (nervous system) H410 |  | inhalation: ATE = 0,14 mg/L (dusts or mists) oral: ATE = 14 mg/kg bw M = 1 000 000 M = 1 000 000 |  | CLP00/ATP18
607-254-00-7 | beta-cyfluthrin (ISO); reaction mass of rel-(R)-cyano(4-fluoro-3-phenoxyphenyl)methyl(1S,3S)-3-(2,2-dichloroethenyl)-2,2-dimethylcyclopropane-1-carboxylate and rel-(R)-cyano(4-fluoro-3-phenoxyphenyl)methyl(1S,3R)-3-(2,2-dichloroethenyl)-2,2-dimethylcyclopropane-1-carboxylate |  | 1820573-27-0 | Lact. Acute Tox. 2 Acute Tox. 2 STOT SE 1 Aquatic Acute 1 Aquatic Chronic 1 | H362 H330 H300 H370 (nervous system) H400 H410 | GHS06 GHS08 GHS09 Dgr | H362 H330 H300 H370 (nervous system) H410 |  | inhalation: ATE = 0,081 mg/L (dusts or mists) oral: ATE = 11 mg/kg bw M = 1 000 000 M = 1 000 000 |  | CLP00/ATP18
607-255-00-2 | fluroxypyr (ISO); 4-amino-3,5-dichloro-6-fluoro-2-pyridyloxyacetic acid |  | 69377-81-7 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-256-00-8 | azoxystrobin (ISO); methyl (E)-2-{2-[6-(2-cyanophenoxy)pyrimi­din-4-yloxy]phenyl}-3-methoxyacrylate |  | 131860-33-8 | Acute Tox. 3 Aquatic Acute 1 Aquatic Chronic 1 | H331 H400 H410 | GHS06 GHS09 Dgr | H331 H410 |  | Inhalation: ATE = 0.7 mg/L (dusts/mists) M=10 M=10 |  | CLP00/ATP15
607-257-00-3 | isopropyl propionate | 211-300-8 | 637-78-5 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 |  |  |  | CLP00
607-258-00-9 | dodecyl 3-(2-(3-benzyl-4-ethoxy-2,5-dioxoimidazolidin-1-yl)-3-(4-methoxybenzoyl)acetamido)-4-chlorobenzoate | 403-990-6 | 70950-45-7 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-259-00-4 | methyl 2R,3S-(-)-3-(4-methoxyphenyl)oxiranecarboxylate | 404-130-2 | 105560-93-8 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | CLP00
607-260-00-X | ethyl 2-(3-nitrobenzylidene)acetoacetate | 404-490-0 | 39562-16-8 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | CLP00
607-261-00-5 | iso(C10-C14)alkyl (3,5-di-tert-butyl-4-hydroxyphenyl)methylthioacetate | 404-800-4 | 118832-72-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-262-00-0 | 7-chloro-1-cyclopropyl-6-fluoro-1,4-dihydro-4-oxoquinoline-3-carboxylic acid | 405-050-0 | 86393-33-1 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
607-263-00-6 | potassium iron(III) 1,3-propanediamine-N,N,N',N'-tetraacetate hemihydrate | 405-680-6 |  | Self-heat. 2 **** Aquatic Chronic 2 | H252 H411 | GHS02 GHS09 Wng | H252 H411 |  |  |  | CLP00
607-264-00-1 | 2-chloro-4-(methylsulfonyl)benzoic acid | 406-520-8 | 53250-83-2 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
607-265-00-7 | ethyl-2-chloro-2,2-diphenylacetate | 406-580-5 | 52460-86-3 | Skin Irrit. 2 Aquatic Chronic 3 | H315 H412 | GHS07 Wng | H315 H412 |  |  |  | CLP00
607-266-00-2 | reaction mass of: hydroxyaluminium bis[2-hydroxy-3,5-di-tert-butylbenzoate]; 3,5-di-tert-butyl-salicylic acid | 406-890-0 | 130296-87-6 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
607-267-00-8 | tert-butyl (5S,6R,7R)-3-bromomethyl-5,8-dioxo-7-(2-(2-phenylacetamido)-5-thia-1-azabicyclo[4.2.0] oct-2-ene-2-carboxylate | 407-620-4 | 33610-13-8 | Resp. Sens. 1 Skin Sens. 1 Aquatic Chronic 3 | H334 H317 H412 | GHS08 Dgr | H334 H317 H412 |  |  |  | CLP00
607-268-00-3 | 2-methylpropyl (R)-2-hydroxypropanoate | 407-770-0 | 61597-96-4 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
607-269-00-9 | (R)-2-(4-hydroxyphenoxy)propanoic acid | 407-960-3 | 94050-90-5 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
607-270-00-4 | 3,9-bis(2-(3-(3-tert-butyl-4-hydroxy-5-methylphenyl)propionyloxy-1,1-dimethylethyl)-2,4,8,10- tetraoxaspiro[5.5]undecane | 410-730-5 | 90498-90-1 | Acute Tox. 4 * | H312 | GHS07 Wng | H312 |  |  |  | CLP00
607-271-00-X | 2-isopropyl-5-methylcyclohexyloxycarbonyloxy-2-hydroxypropane | 417-420-9 | 156324-82-2 | Eye Irrit. 2 Aquatic Chronic 2 | H319 H411 | GHS07 GHS09 Wng | H319 H411 |  |  |  | CLP00
607-272-00-5 | fluroxypyr-meptyl (ISO); methylheptyl, O-(4-amino-3,5-dichloro-6-fluoro-2-pyridyloxy) acetate [1] fluroxypyr-butometyl (ISO); 2-butoxy-1-methylethyl, O-(4-amino-3,5-dichloro-6-fluoro-2-pyridyloxy) acetate [2] | 279-752-9 [1] | 81406-37-3 [1] 154486-27-8 [2] | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-273-00-0 | ammonium 7-(2,6-dimethyl-8-(2,2-dimethylbutyryloxy)-1,2,6,7,8,8a-hexahydro-1-naphthyl)-3,5-dihydroxyheptanoate | 404-520-2 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-274-00-6 | 2-(N-benzyl-N-methylamino)ethyl 3-amino-2-butenoate | 405-350-1 | 54527-73-0 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-275-00-1 | sodium benzoyloxybenzene-4-sulfonate | 405-450-5 | 66531-87-1 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-276-00-7 | bis[(1-methylimidazol)-(2-ethyl-hexanoate)], zinc complex | 405-635-0 |  | Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H400 H410 | GHS05 GHS09 Dgr | H315 H318 H410 |  |  |  | CLP00
607-277-00-2 | reaction mass of: 2-(hexylthio)ethylamine hydrochloride; sodium propionate | 405-720-2 |  | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H317 H411 |  |  |  | CLP00
607-278-00-8 | reaction mass of isomers of: sodium phenethylnaphthalenesulfonate; sodium naphthylethylbenzenesulfonate | 405-760-0 |  | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | CLP00
607-279-00-3 | reaction mass of n-octadecylaminodiethyl bis(hydrogen maleate); n-octadecylaminodiethyl hydrogen maleate hydrogenphthalate | 405-960-8 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-280-00-9 | sodium 4-chloro-1-hydroxybutane-1-sulfonate | 406-190-5 | 54322-20-2 | Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 | H302 H319 H317 | GHS07 Wng | H302 H319 H317 |  |  |  | CLP00
607-281-00-4 | reaction mass of branched and linear C7-C9 alkyl 3-[3-(2H-benzotriazol-2-yl)-5-(1,1-dimethylethyl)-4-hydroxyphenyl]propionates | 407-000-3 | 127519-17-9 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-282-00-X | 2-acetoxymethyl-4-benzyloxybut-1-yl acetate | 407-140-5 | 131266-10-9 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-283-00-5 | E-ethyl-4-oxo-4-phenylcrotonate | 408-040-4 | 15121-89-8 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H315 H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H312 H302 H315 H318 H317 H410 |  |  |  | CLP00
607-284-00-0 | reaction mass of: sodium 3,3'-(1,4-phenylenebis(carbonylimino-3,1-propanediylimino))bis(10-amino-6,13-dichloro-4,11-triphenodioxazinedisulfonate); lithium 3,3'-(1,4-phenylenebis-(carbonylimino-3,1-propanediyl-imino))bis(10-amino-6,13-dichloro)-4,11-triphenodioxazinedisulfonate (9:1) | 410-040-4 | 136213-76-8 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-285-00-6 | reaction mass of: 7-(((3-aminophenyl)sulfonyl)amino)-naphthalene-1,3-disulfonic acid; sodium 7-(((3-aminophenyl)sulfonyl)amino)-naphthalene-1,3-disulfonate; potassium 7-(((3-aminophenyl)sulfonyl)amino)-naphthalene-1,3-disulfonate | 410-065-0 |  | Skin Sens. 1 | H317 | GHS07 Wng |  |  |  |  | CLP00
607-286-00-1 | reaction mass of: sodium/potassium 7-[[[3-[[4-((2-hydroxy-naphthyl)azo)phenyl]azo]phenyl]sulfonyl]amino]-naphthalene-1,3-disulfonate | 410-070-8 | 141880-36-6 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
607-287-00-7 | O'-methyl O-(1-methyl-2-methacryloyloxy-ethyl)-1,2,3,6-tetrahydrophthalate | 410-140-8 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-288-00-2 | Tetrasodium (c-(3-(1-(3-(e-6-dichloro-5-cyanopyrimidin-f-yl(methyl)amino)propyl)-1,6-dihydro-2-hydroxy-4-methyl-6-oxo-3-pyridylazo)-4-sulfonatophenylsulfamoyl)phthalocyanine-a,b,d-trisulfonato(6-))nickelato II, where a is 1 or 2 or 3 or 4,b is 8 or 9 or 10 or 11,c is 15 or 16 or 17 or 18, d is 22 or 23 or 24 or 25 and where e and f together are 2 and 4 or 4 and 2 respectively | 410-160-7 | 148732-74-5 | Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H319 H317 H412 | GHS07 Wng | H319 H317 H412 |  |  |  | CLP00
607-289-00-8 | 3-(3-(4-(2,4-bis(1,1-dimethylpropyl)phenoxy)butylaminocarbonyl-4-hydroxy-1-naphthalenyl)thio)propanoic acid | 410-370-9 | 105488-33-3 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-290-00-3 | reaction mass (ratio not known) of: ammonium 1-C14-C18-alkyloxycarbonyl-2-(3-allyloxy-2-hydroxypropoxycarbonyl)ethane-1-sulfonate; ammonium 2-C14-C18-alkyloxycarbonyl-1-(3-allyloxy-2-hydroxypropoxycarbonyl)ethane-1-sulfonate | 410-540-2 |  | Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H317 H400 H410 | GHS07 GHS09 Wng | H315 H317 H410 |  |  |  | CLP00
607-291-00-9 | dodecyl-ω-(C5/C6-cycloalkyl)alkyl carboxylate | 410-630-1 | 104051-92-5 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-292-00-4 | reaction mass of: [1-(methoxymethyl)-2-(C12-alkoxy)-ethoxy]acetic acid; [1-(methoxymethyl)-2-(C14-alkoxy)-ethoxy]acetic acid | 410-640-6 |  | Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H400 H410 | GHS05 GHS09 Dgr | H315 H318 H410 |  |  |  | CLP00
607-293-00-X | reaction mass of: N-aminoethylpiperazonium mono-2,4,6-trimethylnonyldiphenyl ether di-sulfonate; N-aminoethylpiperazonium di-2,4,6-trimethylnonyldiphenyl ether di-sulfonate | 410-650-0 |  | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H318 H317 H411 |  |  |  | CLP00
607-294-00-5 | sodium 2-benzoyloxy-1-hydroxyethane-sulfonate | 410-680-4 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-295-00-0 | reaction mass of: tetrasodium phosphonoethane-1,2-dicarboxylate; hexasodium phosphonobutane-1,2,3,4-tetracarboxylate | 410-800-5 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-296-00-6 | reaction mass of: pentaerythriol tetraesters with heptanoic acid and 2-ethylhexanoic acid | 410-830-9 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-297-00-1 | (E-E)-3,3'-(1,4-phenylenedimethylidene)bis(2-oxobornane-10-sulfonic acid) | 410-960-6 | 92761-26-7 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
607-298-00-7 | 2-(trimethylammonium)ethoxycarboxybenzene-4-sulfonate | 411-010-3 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-299-00-2 | methyl 3-(acetylthio)-2-methyl-propanoate | 411-040-7 | 97101-46-7 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
607-300-00-6 | trisodium [2-(5-chloro-2,6-difluoropyrimidin-4-ylamino)-5-(b-sulfamoyl-c, d-sulfonatophthalocyanin-a-yl-K4,N29,N30,N31,N32-sulfonylamino)benzoato(5-)]cuprate(II) where a = 1,2,3,4 b = 8,9,10,11 c = 15,16,17,18 d = 22,23,24,25 | 411-430-7 |  | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
607-301-00-1 | reaction mass of: dodecanoic acid; poly(1-7)lactate esters of dodecanoic acid | 411-860-5 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-302-00-7 | reaction mass of: tetradecanoic acid; poly(1-7)lactate esters of tetradecanoic acid | 411-910-6 |  | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H315 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H315 H318 H317 H411 |  |  |  | CLP00
607-303-00-2 | 1-cyclopropyl-6,7-difluoro-1,4-dihydro-4-oxoquinoline-3-carboxylic acid | 413-760-7 | 93107-30-3 | Repr. 2 Aquatic Chronic 3 | H361f *** H412 | GHS08 Wng | H361f *** H412 |  |  |  | CLP00
607-304-00-8 | fluazifop-butyl (ISO); butyl (RS)-2-[4-(5-trifluoromethyl-2-pyridyloxy)phenoxy]propionate | 274-125-6 | 69806-50-4 | Repr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H360D *** H400 H410 | GHS08 GHS09 Dgr | H360D *** H410 |  |  |  | CLP00
607-305-00-3 | fluazifop-P-butyl (ISO); butyl (R)-2-[4-(5-trifluoromethyl-2-pyridyloxy)phenoxy]propionate |  | 79241-46-6 | Repr. 2 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H400 H410 | GHS08 GHS09 Wng | H361d *** H410 |  |  |  | CLP00
607-306-00-9 | chlozolinate (ISO); ethyl (RS)-3-(3,5-dichlorophenyl)-5-methyl-2,4-dioxo-oxazolidine-5-carboxylate | 282-714-4 | 84332-86-5 | Carc. 2 Aquatic Chronic 2 | H351 H411 | GHS08 GHS09 Wng | H351 H411 |  |  |  | CLP00
607-307-00-4 | vinclozolin (ISO); N-3,5-dichlorophenyl-5-methyl-5-vinyl-1,3-oxazolidine-2,4-dione | 256-599-6 | 50471-44-8 | Carc. 2 Repr. 1B Skin Sens. 1 Aquatic Chronic 2 | H351 H360FD H317 H411 | GHS08 GHS07 GHS09 Dgr | H351 H360FD H317 H411 |  |  |  | CLP00
607-308-00-X | esters of 2,4-D |  |  | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  | A | CLP00
607-309-00-5 | carfentrazone-ethyl (ISO); ethyl (RS)-2-chloro-3-[2-chloro-4-fluoro-5-[4-difluoromethyl-4,5-dihydro-3-methyl-5-oxo-1H-1,2,4-triazol-1-yl]phenyl]propionate |  | 128639-02-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-310-00-0 | kresoxim-methyl (ISO); methyl (E)-2-methoxyimino-[2-(o-tolyloxymethyl)phenyl]acetate |  | 143390-89-0 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  |  |  | CLP00
607-311-00-6 | benazolin-ethyl; ethyl 4-chloro-2-oxo-2H-benzothiazole-3-acetate | 246-591-0 | 25059-80-7 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-312-00-1 | methoxyacetic acid | 210-894-6 | 625-45-6 | Repr. 1B Acute Tox. 4 * Skin Corr. 1B | H360FD H302 H314 | GHS08 GHS05 GHS07 Dgr | H360FD H302 H314 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
607-313-00-7 | neodecanoyl chloride | 254-875-0 | 40292-82-8 | Acute Tox. 2 * Acute Tox. 4 * Skin Corr. 1B | H330 H302 H314 | GHS06 GHS06 Dgr | H330 H302 H314 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
607-314-00-2 | ethofumesate (ISO); (RS)-2-ethoxy-2,3-di­hydro-3,3-dimethyl­benzofuran-5-yl methanesulfonate | 247-525-3 | 26225-79-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1 M=1 |  | CLP00/ATP15
607-315-00-8 | glyphosate (ISO); N-(phosphonomethyl)glycine | 213-997-4 | 1071-83-6 | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00/ATP22
607-316-00-3 | glyphosate-trimesium; glyphosate-trimethylsulfonium |  | 81591-81-3 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
607-317-00-9 | bis(2-ethylhexyl) phthalate; di-(2-ethylhexyl) phthalate; DEHP | 204-211-0 | 117-81-7 | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  |  | CLP00
607-318-00-4 | dibutyl phthalate; DBP | 201-557-4 | 84-74-2 | Repr. 1B Aquatic Acute 1 | H360Df H400 | GHS08 GHS09 Dgr | H360Df H400 |  |  |  | CLP00
607-319-00-X | deltamethrin (ISO); (S)-α-cyano-3-phenoxybenzyl (1R, 3R)-3-(2,2-dibromovinyl)-2,2-dimethylcyclopropanecarboxylate | 258-256-6 | 52918-63-5 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H301 H410 |  | M=1000000 |  | CLP00/ATP01
607-320-00-5 | bis[4-(ethenyloxy)butyl] 1,3-benzenedicarboxylate | 413-930-0 | 130066-57-8 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
607-321-00-0 | (S)-methyl-2-chloropropionate | 412-470-8 | 73246-45-4 | Flam. Liq. 3 STOT RE 2 * Eye Irrit. 2 | H226 H373 ** H319 | GHS02 GHS08 Wng | H226 H373 ** H319 |  |  |  | CLP00
607-322-00-6 | 4-(4,4-dimethyl-3-oxo-pyrazolidin-1-yl)-benzoic acid | 413-120-7 | 107144-30-9 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
607-323-00-1 | 2-(1-(2-hydroxy-3,5-di-tert-pentyl-phenyl)ethyl)-4,6-di-tert-pentylphenyl acrylate | 413-850-6 | 123968-25-2 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-324-00-7 | reaction mass of: N,N-di(hydrogenated alkyl C14-C18)phtalamic acid; dihydrogenated alkyl (C14-C18)amine | 413-800-3 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-325-00-2 | (S)-2-chloropropionic acid | 411-150-5 | 29617-66-1 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A | H312 H302 H314 | GHS05 GHS07 Dgr | H312 H302 H314 |  |  |  | CLP00
607-326-00-8 | reaction mass of: isobutyl hydrogen 2-(α-2,4,6-trimethylnon-2-enyl)succinate; isobutyl hydrogen 2-(ß-2,4,6-trimetyhylnon-2-enyl)succinate | 410-720-0 | 141847-13-4 | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
607-327-00-3 | 2-(2-iodoethyl)-1,3-propanediol diacetate | 411-780-0 | 127047-77-2 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
607-328-00-9 | methyl 4-bromomethyl-3-methoxybenzoate | 410-310-1 | 70264-94-7 | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H315 H318 H317 H410 |  |  |  | CLP00
607-329-00-4 | reaction mass of: sodium 2-(C12-18-n-alkyl)amino-1,4-butandioate; sodium 2-octadecenyl-amino-1,4-butandioate | 411-250-9 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-330-00-X | (S)-2,3-dihydro-1H-indole-2-carboxylic acid | 410-860-2 | 79815-20-6 | Repr. 2 STOT RE 2 * Skin Sens. 1 | H361f *** H373 ** H317 | GHS08 GHS07 Wng | H361f *** H373 ** H317 |  |  |  | CLP00
607-332-00-0 | cyclopentyl chloroformate | 411-460-0 | 50715-28-1 | Flam. Liq. 3 Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 | H226 H331 H302 H373 ** H318 H317 | GHS02 GHS06 GHS08 GHS05 Dgr | H226 H331 H302 H373 ** H318 H317 |  |  |  | CLP00
607-333-00-6 | reaction mass of: dodecyl N-(2,2,6,6-tetramethylpiperidin-4-yl)-β-alaninate; tetradecyl N-(2,2,6,6-tetramethylpiperidin-4-yl)-β-alaninate | 405-670-1 |  | Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H314 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H302 H373 ** H314 H410 |  |  |  | CLP00
607-334-00-1 | ethyl 1-ethyl-6,7,8-trifluoro-1,4-dihydro-4-oxoquinoline-3-carboxylate | 405-880-3 | 100501-62-0 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
607-335-00-7 | methyl (R)-2-(4-(3-chloro-5-trifluoromethyl-2-pyridyloxy)phenoxy)propionate | 406-250-0 | 72619-32-0 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
607-336-00-2 | 4-methyl-8-methylenetricyclo[3.3.1.13,7]dec-2-yl acetate | 406-560-6 | 122760-85-4 | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H315 H317 H411 | GHS07 GHS09 Wng | H315 H317 H411 |  |  |  | CLP00
607-337-00-8 | di-tert-(C12-14)-alkylammonium 2-benzothiazolylthiosuccinate | 406-052-4 | 125078-60-6 | Flam. Liq. 3 Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 2 | H226 H302 H315 H318 H411 | GHS02 GHS05 GHS07 GHS09 Dgr | H226 H302 H315 H318 H411 |  |  |  | CLP00
607-338-00-3 | 2-methylpropyl 2-hydroxy-2-methylbut-3-enoate | 406-235-9 | 72531-53-4 | Skin Irrit. 2 Eye Irrit. 2 | H315 H319 | GHS07 Wng | H319 H315 |  |  |  | CLP00
607-339-00-9 | 2,3,4,5-tetrachlorobenzoylchloride | 406-760-3 | 42221-52-3 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 | H302 H314 H317 | GHS05 GHS07 Dgr | H302 H314 H317 |  |  |  | CLP00
607-340-00-4 | 1,3-bis(4-benzoyl-3-hydroxyphenoxy)prop-2-yl acetate | 406-990-4 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-341-00-X | (9S)-9-amino-9-deoxyerythromycin | 406-790-7 | 26116-56-3 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | CLP00
607-342-00-5 | 4-chlorobutyl veratrate | 410-950-1 | 69788-75-6 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-343-00-0 | 4,7-methanooctahydro-1H-indene-diyldimethyl bis(2-carboxybenzoate) | 407-410-2 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-344-00-6 | reaction mass of: 3-(N-(3-dimethylaminopropyl)-(C4-8)perfluoroalkylsulfonamido)propionic acid; N-[dimethyl-3-(C4-8-perfluoroalkylsulfonamido)propylammonium propionate; 3-(N-(3-dimethyl-propylammonium)-(C4-8)perfluoroalkylsulfonamido)propionic acid propionate | 407-810-7 |  | STOT RE 2 * | H373 ** | GHS08 Wng | H373 ** |  |  |  | CLP00
607-345-00-1 | potassium 2-(2,4-dichlorophenoxy)-(R)-propionate | 413-580-9 | 113963-87-4 | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H302 H315 H318 H317 | GHS05 GHS07 Dgr | H302 H315 H318 H317 |  |  |  | CLP00
607-346-00-7 | 3-icosyl-4-henicosylidene-2-oxetanone | 401-210-9 | 83708-14-9 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-347-00-2 | sodium (R)-2-(2,4-dichlorophenoxy)propionate | 413-340-3 | 119299-10-4 | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H302 H315 H318 H317 | GHS05 GHS07 Dgr | H302 H315 H318 H317 |  |  |  | CLP00
607-348-00-8 | magnesium bis((R)-2-(2,4-dichlorophenoxy)propionate) | 413-360-2 |  | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H302 H315 H318 H317 | GHS05 GHS07 Dgr | H302 H315 H318 H317 |  |  |  | CLP00
607-349-00-3 | mono-(tetrapropylammonium) hydrogen 2,2'-dithiobisbenzoate | 411-270-8 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-350-00-9 | bis(4-(1,2-bis(ethoxycarbonyl)ethylamino)-3-methylcyclohexyl)methane | 412-060-9 | 136210-32-7 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
607-351-00-4 | methyl O-(4-amino-3,5-dichloro-6-fluoropyridin-2-yloxy)acetate | 407-550-4 | 69184-17-4 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-352-00-X | 4,4'-oxydiphthalic anhydride | 412-830-4 | 1823-59-2 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-353-00-5 | reaction mass of: ethyl exo-tricyclo[5.2.1.02,6]decane-endo-2-carboxylate; ethyl endo-tricyclo[5.2.1.02,6]decane-exo-2-carboxylate | 407-520-0 | 80657-64-3 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
607-354-00-0 | ethyl 2-cyclohexylpropionate | 412-280-5 | 2511-00-4 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-355-00-6 | p-tolyl 4-chlorobenzoate | 411-530-0 | 15024-10-9 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
607-356-00-1 | ethyl trans-2,2,6-trimethylcyclohexanecarboxylate | 412-540-8 |  | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
607-357-00-7 | reaction mass of: trans-4-acetoxy-4-methyl-2-propyl-tetrahydro-2H-pyran; cis-4-acetoxy-4-methyl-2-propyl-tetrahydro-2H-pyran | 412-450-9 | 131766-73-9 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-358-00-2 | (1S,3S,5R,6R)-(4-nitrophenylmethyl)-1-dioxo-6-phenylacetamido-penam-3-carboxylate | 412-670-5 | 54275-93-3 | Resp. Sens. 1 | H334 | GHS08 Dgr | H334 |  |  |  | CLP00
607-359-00-8 | (1S,4R,6R,7R)-(4-nitrophenylmethyl)3-methylene-1-oxo-7-phenylacetamido-cepham-4-carboxylateido-penam-3-carboxylate | 412-800-0 | 76109-32-5 | Resp. Sens. 1 | H334 | GHS08 Dgr | H334 |  |  |  | CLP00
607-360-00-3 | sodium 3-acetoacetylamino-4-methoxytolyl-6-sulfonate | 411-680-7 | 133167-77-8 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-361-00-9 | methyl (R)-2-(4-hydroxyphenoxy)propionate | 411-950-4 | 96562-58-2 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
607-362-00-4 | reaction mass of: (3-methoxy)propylammonium/[tris-(2-hydroxyethyl)]ammonium 2-(2-(bis(2-hydroxyethyl)amino)ethoxycarbonylmethyl)hexadec-4-enoate; (3-methoxy)propylammonium/[tris-(2-hydroxyethyl)]ammonium 2-(2-(bis(2-hydroxyethyl)amino)ethoxycarbonylmethyl)tetradec-4-enoate; (3-methoxy)propylammonium/[tris-(2-hydroxyethyl)]ammonium 2-(3-methoxypropylcarbamoylmethyl)hexadec-4-enoate; (3-methoxy)propylammonium/[tris-(2-hydroxyethyl)]ammonium 2-(3-methoxypropylcarbamoylmethyl)tetradec-4-enoate | 413-500-2 |  | Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 2 | H315 H318 H411 | GHS05 GHS09 Dgr | H315 H318 H411 |  |  |  | CLP00
607-363-00-X | methyl-3-methoxyacrylate | 412-900-4 | 5788-17-0 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-364-00-5 | 3-phenyl-7-[4-(tetrahydrofurfuryloxy)phenyl]-1,5-dioxa-s-indacen-2,6-dione | 413-330-9 | 134724-55-3 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-365-00-0 | 2-(2-amino-1,3-thiazol-4-yl)-(Z)-2-methoxyiminoacetyl chloride hydrochloride | 410-620-7 | 119154-86-8 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 | H302 H314 H317 | GHS05 GHS07 Dgr | H302 H314 H317 |  |  |  | CLP00
607-366-00-6 | 3,5-dimethylbenzoyl chloride | 413-010-9 | 6613-44-1 | Skin Corr. 1B Skin Sens. 1 | H314 H317 | GHS05 GHS07 Dgr | H314 H317 |  |  |  | CLP00
607-367-00-1 | potassium bis(N-carboxymethyl)-N-methyl-glycinato-(2-)N,O,O,N)-ferrate-(1-) monohydrate | 411-640-9 | 153352-59-1 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
607-368-00-7 | 1-(N,N-dimethylcarbamoyl)-3-tert-butyl-5-carbethoxymethylthio-1H-1,2,4-triazole | 411-650-3 | 110895-43-7 | Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H301 H410 |  |  |  | CLP00
607-369-00-2 | reaction mass of: trans-(2R)-5-acetoxy-1,3-oxathiolane-2-carboxylic acid; cis-(2R)-5-acetoxy-1,3-oxathiolane-2-carboxylic acid | 411-660-8 | 147027-04-1 | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H302 H315 H318 H317 | GHS05 GHS07 Dgr | H302 H315 H318 H317 |  |  |  | CLP00
607-370-00-8 | 2-[[2-(acetyloxy)-3-(1,1-dimethyl-ethyl)-5-methylphenyl]methyl]-6-(1,1-dimethylethyl)-4-methylphenol | 412-210-3 | 41620-33-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-371-00-3 | 3-ethyl 5-methyl 4-(2-chlorophenyl)-1,4-dihydro-2-[2-(1,3-dihydro-1,3-dioxo-(2H)isoindol-2-yl)-ethoxymethyl]-6-methyl-3,5-pyridinedicarboxylate | 413-410-3 | 88150-62-3 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-372-00-9 | ethoxylated bis phenol A di-(norbornene carboxylate) | 412-410-0 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-373-00-4 | quizalofop-P-tefuryl (ISO); (+/-) tetrahydrofurfuryl (R)-2-[4-(6-chloroquinoxalin-2-yloxy)phenyloxy]propionate | 414-200-4 | 200509-41-7 | Carc. 2 Repr. 2 Acute Tox. 4 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H361fd H302 H373 H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H351 H361fd H373 H410 |  | M=1 M=1 |  | CLP00/ATP13
607-374-00-X | 5-amino-2,4,6-triiodo-1,3-benzenedicarbonyldichloride | 417-220-1 | 37441-29-5 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-375-00-5 | flocoumafen (ISO); reaction mass of: cis-4-hydroxy-3-(1,2,3,4-tetrahydro-3-(4-(4-trifluoromethylbenzyloxy)phenyl)-1-naphthyl)coumarin and trans-4-hydroxy-3-(1,2,3,4-tetrahydro-3-(4-(4-trifluoromethylbenzyloxy)phenyl)-1-naphthyl)coumarin | 421-960-0 | 90035-08-8 | Repr. 1B Acute Tox. 1 Acute Tox. 1 Acute Tox. 1 STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D H330 H310 H300 H372 (blood) H400 H410 | GHS08 GHS06 GHS09 Dgr | H330 H310 H300 H360D H372 (blood) H410 |  | Repr. 1B; H360D: C ≥ 0,003 % STOT RE 1; H372: C ≥ 0,05 % STOT RE 2; H373: 0,005 % ≤ C < 0,05 % M=10 M=10 |  | CLP00/ATP09
607-376-00-0 | benzyl 2,4-dibromobutanoate | 420-710-8 | 23085-60-1 | Repr. 2 Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361f *** H315 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H361f *** H315 H317 H410 |  |  |  | CLP00
607-377-00-6 | trans-4-cyclohexyl-L-proline monohydrochloride | 419-160-1 | 90657-55-9 | Repr. 2 Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H361f *** H302 H315 H318 H317 | GHS08 GHS05 GHS07 Dgr | H361f *** H302 H315 H318 H317 |  |  |  | CLP00
607-378-00-1 | ammonium (Z)-α-methoxyimino-2-furylacetate | 405-990-1 | 97148-39-5 | Flam. Sol. 2 | H228 | GHS02 Dgr | H228 |  |  | T | CLP00
607-379-00-7 | reaction mass of: 2-[N-(2-hydroxyethyl)stearamido]ethyl stearate; sodium [bis[2-(stearoyloxy)ethyl]amino]methylsulfonate; sodium [bis(2-hydroxyethyl)amino]methylsulfonate; N,N-bis(2-hydroxyethyl)stearamide | 401-230-8 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-380-00-2 | reaction mass of: ammonium-1,2-bis(hexyloxycarbonyl)ethanesulfonate; ammonium-1-hexyloxycarbonyl-2-octyloxycarbonylethanesulfonate; ammonium-2-hexyloxycarbonyl-1-octyloxycarbonylethanesulfonate | 407-320-3 |  | Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H315 H318 H412 | GHS05 Dgr | H315 H318 H412 |  |  |  | CLP00
607-381-00-8 | reaction mass of triesters of 2,2-bis(hydroxymethyl)butanol with C7-alkanoic acids and 2-ethylhexanoic acid | 413-710-4 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-382-00-3 | 2-((4-amino-2-nitrophenyl)amino)benzoic acid | 411-260-3 | 117907-43-4 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | CLP00
607-383-00-9 | reaction mass of: 2,2,6,6-tetramethylpiperidin-4-yl-hexadecanoate; 2,2,6,6-tetramethylpiperidin-4-yl-octadecanoate | 415-430-8 | 86403-32-9 | Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H318 H317 H410 |  |  |  | CLP00
607-384-00-4 | reaction mass of: esters of C14-C15 branched alcohols with 3,5-di-t-butyl-4-hydroxyphenyl propionic acid; C15 branched and linear alkyl 3,5-bis(1,1-dimethylethyl)-4-hydroxybenzenepropanoate; C13 branched and linear alkyl 3,5-bis(1,1-dimethylethyl)-4-hydroxybenzenepropanoate | 413-750-2 | 171090-93-0 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-385-00-X | Copolymer of vinyl-alcohol and vinyl acetate partially acetilized with 4-(2-(4-formylphenyl)ethenyl)-1-methylpyridinium methylsulfate | 414-590-6 | 125229-74-5 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-386-00-5 | reaction mass of: tetradecanoic acid (42.5-47.5 %); poly(1-7)lactate esters of tetradecanoic acid (52.5-57.5 %) | 412-580-6 | 174591-51-6 | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H315 H318 H317 H410 |  |  |  | CLP00
607-387-00-0 | reaction mass of: dodecanoic acid (35-40 %); poly(1-7)lactate esters of dodecanoic acid (60-65 %) | 412-590-0 | 58856-63-6 | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H315 H318 H317 H410 |  |  |  | CLP00
607-388-00-6 | 4-ethylamino-3-nitrobenzoic acid | 412-090-2 | 2788-74-1 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 3 | H302 H317 H412 | GHS07 Wng | H302 H317 H412 |  |  |  | CLP00
607-389-00-1 | trisodium N,N-bis(carboxymethyl)-3-amino-2-hydroxypropionate | 414-130-4 | 119710-96-2 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
607-390-00-7 | 1,2,3,4-tetrahydro-6-nitro-quinoxaline | 414-270-6 | 41959-35-7 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
607-391-00-2 | dimethylcyclopropane-1,1-dicarboxylate | 414-240-2 | 6914-71-2 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-392-00-8 | 2-phenoxyethyl 4-((5-cyano-1,6-dihydro-2-hydroxy-1,4-dimethyl-6-oxo-3-pyridinyl)azo)benzoate | 414-260-1 | 88938-37-8 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-393-00-3 | 3-(cis-1-propenyl)-7-amino-8-oxo-5-thia-1-azabicyclo[4.2.0]oct-2-ene-2-carboxylic acid | 415-750-8 | 106447-44-3 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-394-00-9 | 5-methylpyrazine-2-carboxylic acid | 413-260-9 | 5521-55-1 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
607-395-00-4 | reaction mass of: sodium 1-tridecyl-4-allyl-(2 or 3)-sulfobutanedioate; sodium 1-dodecyl-4-allyl-(2 or 3)-sulfobutanedioate | 410-230-7 |  | Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H314 H317 H411 | GHS05 GHS07 GHS09 Dgr | H314 H317 H411 |  |  |  | CLP00
607-396-00-X | bis(1,2,2,6,6-pentamethyl-4-piperidinyl) 2-(4-methoxybenzylidene)malonate | 414-840-4 | 147783-69-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-397-00-5 | reaction mass of: Ca salicylates (branched C10-14 and C18-30 alkylated); Ca phenates (branched C10-14 and C18-30 alkylated); Ca sulfurised phenates (branched C10-14 and C18-30 alkylated) | 415-930-6 | - | Repr. 2 Skin Sens. 1 | H361f *** H317 | GHS08 GHS07 Wng | H361f *** H317 |  |  |  | CLP00/ATP01
607-398-00-0 | ethyl N-(5-chloro-3-(4-(diethylamino)-2-methylphenylimino)-4-methyl-6-oxo-1,4-cyclohexadienyl)carbamate | 414-820-5 | 125630-94-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-399-00-6 | 2,2-dimethyl 3-methyl-3-butenyl propanoate | 415-610-6 | 104468-21-5 | Skin Irrit. 2 Aquatic Chronic 3 | H315 H412 | GHS07 Wng | H315 H412 |  |  |  | CLP00
607-400-00-X | methyl 3-[[(dibutylamino)thioxomethyl]thio]propanoate | 414-400-1 | 32750-89-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-401-00-5 | ethyl 3-hydroxy-5-oxo-3-cyclohexene-1-carboxylate | 414-450-4 | 88805-65-6 | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H315 H318 H317 | GHS05 GHS07 Dgr | H315 H318 H317 |  |  |  | CLP00
607-402-00-0 | methyl N-(phenoxycarbonyl)-L-valinate | 414-500-5 | 153441-77-1 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-403-00-6 | reaction mass of: bis(1S,2S,4S)-(1-benzyl-4-tert-butoxycarboxamido-2-hydroxy-5-phenyl)pentylammonium succinate; isopropyl alcohol | 414-810-0 |  | STOT RE 2 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H318 H400 H410 | GHS08 GHS05 GHS09 Dgr | H373 ** H318 H410 |  |  |  | CLP00
607-404-00-1 | reaction mass of: ((Z)-3,7-dimethyl-2,6-octadienyl)oxycarbonylpropanoic acid; di-((E)-3,7-dimethyl-2,6-octadienyl) butandioate; di-((Z)-3,7-dimethyl-2,6-octadienyl) butandioate; (Z)-3,7-dimethyl-2,6-octadienyl butandioate; ((E)-3,7-dimethyl-2,6-octadienyl)oxycarbonylpropanoic acid | 415-190-4 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-405-00-7 | 2-hexyldecyl-p-hydroxybenzoate | 415-380-7 | 148348-12-3 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-406-00-2 | potassium 2,5-dichlorobenzoate | 415-700-5 | 184637-62-5 | Acute Tox. 4 * Eye Dam. 1 | H302 H318 | GHS05 GHS07 Dgr | H302 H318 |  |  |  | CLP00
607-407-00-8 | ethyl 2-carboxy-3-(2-thienyl)propionate | 415-680-8 | 143468-96-6 | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H315 H318 H317 | GHS05 GHS07 Dgr | H315 H318 H317 |  |  |  | CLP00
607-408-00-3 | potassium N-(4-fluorophenyl)glycinate | 415-710-1 | 184637-63-6 | STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H373 ** H318 H317 H412 | GHS08 GHS05 GHS07 Dgr | H373 ** H318 H317 H412 |  |  |  | CLP00
607-409-00-9 | reaction mass of: (3R)-[1S-(1α, 2α, 6β-((2S)-2-methyl-1-oxo-butoxy)-8aγ)hexahydro-2,6-dimethyl-1-naphthalene]-3,5-dihydroxyheptanoic acid; inert biomass from Aspergillus terreus | 415-840-7 |  | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
607-410-00-4 | mono[2-(dimethylamino)ethyl]monohydrogen-2-(hexadec-2-enyl)butanedioate and/or mono[2-(dimethylamino)ethyl]monohydrogen-3-(hexadec-2-enyl)butanedioate | 415-880-5 | 779343-34-9 | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H315 H318 H317 H410 |  |  |  | CLP00
607-411-00-X | oxiranemethanol, 4-methylbenzene-sulfonate, (S)- | 417-210-7 | 70987-78-9 | Carc. 1B Muta. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H350 H341 H318 H317 H411 | GHS08 GHS05 GHS07 GHS09 Dgr | H350 H341 H318 H317 H411 |  |  |  | CLP00
607-412-00-5 | ethyl 2-(1-cyanocyclohexyl)acetate | 415-970-4 | 133481-10-4 | Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 3 | H302 H373 ** H412 | GHS08 GHS07 Wng | H302 H373 ** H412 |  |  |  | CLP00
607-413-00-0 | trans-4-phenyl-L-proline | 416-020-1 | 96314-26-0 | Repr. 2 Skin Sens. 1 | H361f *** H317 | GHS08 GHS07 Wng | H361f *** H317 |  |  |  | CLP00
607-415-00-1 | poly-(methyl methacrylate)-co-(butylmethacrylate)-co-(4-acryloxybutyl-isopropenyl-α, α-dimethylbenzyl carbamate)-co-(maleicanhydride) | 419-590-1 |  | Flam. Sol. 1 Skin Sens. 1 | H228 H317 | GHS02 GHS07 Dgr | H228 H317 |  |  | T | CLP00
607-416-00-7 | 4-(2-carboxymethylthio)ethoxy-1-hydroxy-5-isobutyloxycarbonylamino-N-(3-dodecyloxypropyl)-2-naphthamide | 420-730-7 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-417-00-2 | 3-chloropropyl chloroformiate | 425-770-9 | 628-11-5 | Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H331 H302 H373 ** H315 H318 H317 | GHS06 GHS05 GHS08 Dgr | H331 H302 H373 ** H315 H318 H317 |  |  |  | ATP01
607-418-00-8 | 2-ethylhexyl 4-aminobenzoate | 420-170-3 | 26218-04-2 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-419-00-3 | (3'-carboxymethyl-5-(2-(3-ethyl-3H-benzothiazol-2-ylidene)-1-methyl-ethylidene)-4,4'-dioxo-2'-thioxo-(2,5')bithiazolidinyliden-3-yl)-acetic acid | 422-240-9 | 166596-68-5 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
607-420-00-9 | 2,2-bis(hydroxymethyl)butanoic acid | 424-090-1 | 10097-02-6 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
607-421-00-4 | cypermethrin (ISO); α-cyano-3-phenoxybenzyl 3-(2,2-dichlorovi-nyl)-2,2-dimethylcyclopro-panecarboxylate; cypermethrin cis/trans +/- 40/60 | 257-842-9 | 52315-07-8 | Acute Tox. 4 Acute Tox. 4 STOT SE 3 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H335 H373 (nervous system) H400 H410 | GHS07 GHS08 GHS09 Wng | H332 H302 H335 H373 (nervous system) H410 |  | Inhalation: ATE = 3.3 mg/L (dusts/mists) Oral: ATE = 500 mg/kg bw M=100000 M=100000 |  | CLP00/ATP17
607-422-00-X | α-cypermethrin (ISO); racemate comprising (R)-α-cyano-3-phenoxybenzyl (1S,3S)-3-(2,2-dichlorovinyl)-2,2-dimethylcyclopropanecarboxylate; (S)-α-cyano-3-phenoxybenzyl (1R,3R)-3-(2,2-dichlorovinyl)-2,2-dimethylcyclopropanecarboxylate | 257-842-9 | 67375-30-8 | Acute Tox. 3 * STOT SE 3 STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H335 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H301 H373 ** H335 H410 |  | M=1000 |  | CLP00/ATP01
607-423-00-5 | esters of mecoprop and of mecoprop-P |  |  | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  | A | CLP00
607-424-00-0 | trifloxystrobin (ISO); methyl (E)-methoxyimino- {(E)-α-[1-(α,α,α-trifluoro-m- tolyl) ethylideneaminooxy]-o- tolyl}acetate |  | 141517-21-7 | Lact. Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H362 H317 H400 H410 | GHS07 GHS09 Wng | H317 H362 H410 |  | M=100 M=10 |  | CLP00/ATP17
607-425-00-6 | metalaxyl (ISO); methyl-N-(2,6-dimethylphenyl)-N-(methoxyacetyl)-DL-alaninate | 260-979-7 | 57837-19-1 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 3 | H302 H317 H412 | GHS07 Wng | H302 H317 H412 |  |  |  | CLP00
607-426-00-1 | 1,2-benzenedicarboxylic acid, dipentylester, branched and linear [1] n-pentyl-isopentylphthalate [2] di-n-pentyl phthalate [3] diisopentylphthalate [4] | 284-032-2 [1] 205-017-9 [3] 210-088-4 [4] | 84777-06-0 [1] 131-18-0 [3] 605-50-5 [4] | Repr. 1B Aquatic Acute 1 | H360FD H400 | GHS08 GHS09 Dgr | H360FD H400 |  |  |  | CLP00
607-427-00-7 | bromoxynil heptanoate (ISO); 2,6-dibromo-4-cyanophenyl heptanoate | 260-300-4 | 56634-95-8 | Repr. 2 Acute Tox. 4 * Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H332 H302 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H361d *** H332 H302 H317 H410 |  |  |  | CLP00
607-428-00-2 | tetrasodium ethylene diamine tetraacetate | 200-573-9 | 64-02-8 | Acute Tox. 4 * Eye Dam. 1 | H302 H318 | GHS05 GHS07 Dgr | H302 H318 |  |  |  | ATP01
607-429-00-8 | edetic acid; (EDTA) | 200-449-4 | 60-00-4 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | ATP01
607-430-00-3 | BBP; benzyl butyl phthalate | 201-622-7 | 85-68-7 | Repr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H360Df H400 H410 | GHS08 GHS09 Dgr | H360Df H410 |  |  |  | CLP00
607-431-00-9 | prallethrin (ISO); ETOC; 2-methyl-4-oxo-3-(prop-2-ynyl)cyclopent-2-en-1-yl 2,2-dimethyl-3-(2-methylprop-1-enyl)cyclopropanecarboxylate | 245-387-9 | 23031-36-9 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H302 H400 H410 | GHS06 GHS09 Dgr | H331 H302 H410 |  |  |  | CLP00
607-432-00-4 | S-metolachlor (ISO); 2-chloro-N-(2-ethyl-6-methylphenyl)-N-[(2S)-1-methoxypropan-2-yl]acetamide; (R a S a)-2-chloro-N-(6-ethyl-o-tolyl)-N-[(1S)-2-methoxy-1-methylethyl]acetamide [contains 80-100 % 2-chloro-N-(2-ethyl-6-methylphenyl)-N-[(2S)-1-methoxypropan-2-yl]acetamide and 0-20 % 2-chloro-N-(2-ethyl-6-methylphenyl)-N-[(2R)-1-methoxypropan-2-yl]acetamide] | - | 87392-12-9 | Carc. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H317 H410 | EUH066 | M = 10 M = 10 |  | CLP00/ATP22
607-433-00-X | cypermethrin cis/trans +/- 80/20; (RS)-α-cyano-3-phenoxybenzyl (1RS; 3RS; 1RS, 3SR)-3-(2,2-dichlorovinyl)-2,2-dimethylcyclopropanecarboxylate | 257-842-9 | 52315-07-8 | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H335 H315 H317 H400 H410 | GHS07 GHS09 Wng | H302 H335 H315 H317 H410 |  |  |  | CLP00
607-434-00-5 | mecoprop-P [1] and its salts;; (R)-2-(4-chloro-2-methylphenoxy)propionic acid and its salts | 240-539-0 | 16484-77-8 | Acute Tox. 4 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H318 H400 H410 | GHS07 GHS05 GHS09 Dgr | H302 H318 H410 |  | Oral: ATE = 431 mg/kg bw M=10 M=10 |  | CLP00/ATP17
607-435-00-0 | 2S-isopropyl-5R-methyl-1R-cyclohexyl 2,2-dihydroxyacetate | 416-810-6 | 111969-64-3 | STOT RE 2 * Eye Dam. 1 Aquatic Chronic 2 | H373 ** H318 H411 | GHS08 GHS05 GHS09 Dgr | H373 ** H318 H411 |  |  |  | CLP00
607-436-00-6 | 2-hydroxy-3-(2-ethyl-4-methylimidazoyl)propyl neodecanoate | 417-350-9 |  | Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H400 H410 | GHS05 GHS09 Dgr | H315 H318 H410 |  |  |  | CLP00
607-437-00-1 | 3-(4-aminophenyl)-2-cyano-2-propenoic acid | 417-480-6 | 252977-62-1 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-438-00-7 | methyl-2-[(aminosulfonyl)methyl]benzoate | 419-010-5 | 112941-26-1 | Acute Tox. 4 * Eye Irrit. 2 | H302 H319 | GHS07 Wng | H302 H319 |  |  |  | CLP00
607-439-00-2 | methyl tetrahydro-2-furancarboxylate | 420-670-1 | 37443-42-8 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
607-440-00-8 | methyl 2-aminosulfonyl-6-(trifluoromethyl)pyridine-3-c arboxylate | 421-220-7 | 144740-59-0 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
607-441-00-3 | 3-[3-(2-dodecyloxy-5-methylphenylcarbamoyl)-4-hydroxy-1-naphthylthio]propionic acid | 421-490-6 | 167684-63-1 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-442-00-9 | benzyl [hydroxy-(4-phenylbutyl)phosphinyl] acetate | 416-050-5 | 87460-09-1 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
607-444-00-X | reaction mass of: cis-1,4-dimethylcyclohexyl dibenzoate; trans-1,4-dimethylcyclohexyl dibenzoate | 416-230-3 | 35541-81-2 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-445-00-5 | Iron (III) tris(4-methylbenzenesulfonate) | 420-960-8 | 77214-82-5 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
607-446-00-0 | methyl 2-[4-(2-chloro-4-nitrophenylazo)-3-(1-oxopropyl)amino]phenylaminopropionate | 416-240-8 | 155522-12-6 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
607-447-00-6 | sodium 4-[4-(4-hydroxyphenylazo)phenylamino]-3-nitrobenzenesulfonate | 416-370-5 | 156738-27-1 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
607-448-00-1 | 2,3,5,6-tetrafluorobenzoic acid | 416-800-1 | 652-18-6 | Skin Irrit. 2 Eye Dam. 1 | H315 H318 | GHS05 Dgr | H315 H318 |  |  |  | CLP00
607-449-00-7 | reaction mass of: 4,4',4''-[(2,4,6-trioxo-1,3,5(2H,4H,6H)-triazine-1,3,5-triyl)tris[methylene(3,5,5-trimethyl-3,1-cyclohexanediyl)iminocarbonyloxy-2,1-ethanediyl(ethyl)amino]]trisbenzenediazoniumtri[bis(2-methylpropyl)naphthalenesulfonate]; 4,4',4'',4'''-[[5,5'-[carbonylbis[imino(1,5,5-trimethyl-3,1-cyclohexanediyl)methylene]]-2,4,6-trioxo-1,3,5(2H,4H,6H)-triazine-1,1',3,3'-tetrayl]tetrakis[methylene(3,5,5-trimethyl-3,1-cyclohexanediyl)iminocarbonyloxy-2,1-ethanediyl(ethyl)amino]]tetrakisbenzenediazoniumtetra[bis(2-methylpropyl)naphthalenesulfonate] | 417-080-1 |  | Self-react. D **** Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H242 H317 H400 H410 | GHS02 GHS07 GHS09 Dgr | H242 H317 H410 |  |  |  | CLP00
607-450-00-2 | 2-mercaptobenzothiazolyl-(Z)-(2-aminothiazol-4-yl)-2-(tert-butoxycarbonyl) isopropoxyiminoacetate | 419-040-9 | 89604-92-2 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-451-00-8 | 4-[4-amino-5-hydroxy-3-(4-(2-sulfoxyethylsulfonyl)phenylazo)-2,7-disulfonapht-6-ylazo]-6-[3-(4-amino-5-hydroxy-3-(4-(2-sulfoxyethylsulfonyl)phenylazo)-2,7-disulfonapht-6-ylazo]phenylcarbonylamino]benzenesulfonic acid, sodium salt | 417-640-5 | 161935-19-9 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
607-453-00-9 | 4-benzyl-2,6-dihydroxy-4-aza-heptylene bis(2,2-dimethyloctanoate) | 418-100-1 | 172964-15-7 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
607-454-00-4 | reaction mass of: trans-2-(1-methylethyl)-1,3-dioxane-5-carboxylic acid; cis-2-(1-methylethyl)-1,3-dioxane-5-carboxylic acid | 418-170-3 | 116193-72-7 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
607-455-00-X | 1-amino-4-(3-[4-chloro-6-(2,5-di-sulfophenylamino)-1,3,5-triazin-2-ylamino]-2,2-dimethyl-propylamino)-anthraquinone-2-sulfonic acid, sodium/lithiumsalt | 419-520-8 | 172890-93-6 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-456-00-5 | 3-amino-4-chlorobenzoic acid, hexadecyl ester | 419-700-6 | 143269-74-3 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-457-00-0 | tetrasodium dihydrogen 1,1''-dihydroxy-8,8''-[p-phenylbis(imino-{}{6-[4-(2-aminoethyl)piperazin-1-yl]}}-1,3,5-triazine-4,2-diyl-imino)]bis(2,2'-azonaphthalene-1',3,6-trisulfonate) | 420-350-1 | 172277-97-3 | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
607-458-00-6 | reaction mass of: 2-ethyl-[2,6-dibromo-4-[1-[3,5-dibromo-4-(2-hydroxyethoxy)phenyl]-1-methylethyl]phenoxy]propenoate; 2,2'-diethyl-[4,4'-bis(2,6-dibromophenoxy)-1-methylethylidene] dipropenoate; 2,2'-[(1-methylethylidene)bis[[2,6-dibromo-4,1-phenylene)oxy]ethanol]] | 420-850-1 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-459-00-1 | isopentyl 4-{}{2-[5-cyano-1,2,3,6-tetrahydro-1-(2-isopropoxyethoxy-carbonylmethyl)-4-methyl-2,6-dioxo-3-pyridylidene]hydrazino}}benzoate | 418-930-4 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-460-00-7 | 3-tridecyloxy-propyl-ammonium 9-octadecenoate | 418-990-1 | 778577-53-0 | STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H315 H319 H400 H410 | GHS08 GHS07 GHS09 Wng | H373 ** H319 H315 H410 |  |  |  | CLP00
607-461-00-2 | reaction mass of: pentasodium 2-{}{4-{}{3-methyl-4-[6-sulfonato-4-(2-sulfonato-phenylazo)-naphthalen-1-ylazo]-phenylamino}}-6-[3-(2-sulfato-ethanesulfonyl)-phenylamino]-1,3,5-triazin-2-ylamino}}-benzene-1,4-disulfonate; pentasodium 2-{}{4-{}{3-methyl-4-[7-sulfonato-4-(2-sulfonato-phenylazo)-naphthalen-1-ylazo]-phenylamino}}-6-[3-(2-sulfato-ethanesulfonyl)-phenylamino]-1,3,5-triazin-2-ylamino}}-benzene-1,4-disulfonate | 421-160-1 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-462-00-8 | reaction mass of: 1-hexyl acetate; 2-methyl-1-pentyl acetate; 3-methyl-1-pentyl acetate; 4-methyl-1-pentyl acetate; other mixed linear and branched C6-alkyl acetates | 421-230-1 | 88230-35-7 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-463-00-3 | 3-(phenothiazin-10-yl)propionic acid | 421-260-5 | 362-03-8 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-464-00-9 | reaction mass of: 7-chloro-1-ethyl-6-fluoro-1,4-dihydro-4-oxo-quinoline-3-carboxylic acid; 5-chloro-1-ethyl-6-fluoro-1,4-dihydro-4-oxo-quinoline-3-carboxylic acid | 421-280-4 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-465-00-4 | tris(2-hydroxyethyl)ammonium 7-{}{4-[4-(2-cyanoamino-4-hydroxy-6-oxidopyrimidin-5-ylazo)benzamido]-2-ethoxy-phenylazo}}naphthalene-1,3-disulfonate | 421-440-3 | 778583-04-3 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-466-00-X | reaction mass of: phenyl 1-(1-[2-chloro-5-(hexadecyloxycarbonyl)phenylcarbamoyl]-3,3-dimethyl-2-oxobutyl)-1H-2,3,3a,7a-tetrahydrobenzotriazole-5-carboxylate; phenyl 2-(1-(2-chloro-5-(hexadecyloxycarbonyl)phenylcarbamoyl)-3,3-dimethyl-2-oxobutyl)-1H-2,3,3a,7a-tetrahydrobenzotriazole-5-carboxylate; phenyl 3-(1-(2-chloro-5-(hexadecyloxycarbonyl)phenylcarbamoyl)-3,3-dimethyl-2-oxobutyl)-1H-2,3,3a,7a-tetrahydrobenzotriazole-5-carboxylate | 421-480-1 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-467-00-5 | 1,1,3,3-tetrabutyl-1,3-ditinoxydicaprylate | 419-430-9 | 56533-00-7 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H373 ** H314 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H312 H302 H373 ** H314 H410 |  |  |  | CLP00
607-468-00-0 | reaction mass of: monosodium 4-((4-(5-sulfonato-2-methoxyphenylamino)-6-chloro-1,3,5-triazine-2-yl)amino)-2-((1,4-dimethyl-6-oxido-2-oxo-5-sulfonatomethyl-1,2-dihydropyridine-3-yl)azo)benzenesulfonate; disodium 4-((4-(5-sulfonato-2-methoxyphenylamino)-6-chloro-1,3,5-triazine-2-yl)amino)-2-((1,4-dimethyl-6-oxido-2-oxo-5-sulfonatomethyl-1,2-dihydropyridine-3-yl)azo)benzenesulfonate; trisodium 4-((4-(5-sulfonato-2-methoxyphenylamino)-6-chloro-1,3,5-triazine-2-yl)amino)-2-((1,4-dimethyl-6-oxido-2-oxo-5-sulfonatomethyl-1,2-dihydropyridine-3-yl)azo)benzenesulfonate; tetrasodium 4-((4-(5-sulfonato-2-methoxyphenylamino)-6-chloro-1,3,5-triazine-2-yl)amino)-2-((1,4-dimethyl-6-oxido-2-oxo-5-sulfonatomethyl-1,2-dihydropyridine-3-yl)azo)benzenesulfonate | 419-450-8 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-469-00-6 | disodium 7-((4,6-bis(3-diethylaminopropylamino)-1,3,5-triazine-2-yl)amino)-4-hydroxy-3-(4-(4-sulfonatophenylazo)phenylazo)-2-naphthalene sulfonate | 419-460-2 | 120029-06-3 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-470-00-1 | potassium sodium 6,13-dichloro-3,10-bis{}{2-[4-[3-(2-hydroxysulphonyloxyethanesulfonyl)phenylamino]-6-(2,5-disulfonatophenylamino)-1,3,5-triazin-2-ylamino]ethylamino}}benzo[5,6][1,4]oxazino[2,3-b]phenoxazine-4,11-disulfonate | 414-100-0 | 154336-20-6 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
607-471-00-7 | 1,6-bis((dibenzylthiocarbamoyl)disulfanyl)hexane | 429-280-6 | 151900-44-6 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-473-00-8 | pentaerythritol, dipentaerythritol, fatty acids, C6-10, mixed esters with adipic acid, heptanoic acid and isostearic acid | 426-590-3 | 187412-41-5 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-474-00-3 | (4-(4-(4-dimethylaminobenzyliden-1-yl)-3-methyl-5-oxo-2-pyrazolin-1-yl)benzoic acid | 410-430-4 | 117573-89-4 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-475-00-9 | reaction mass of: tetrasodium 7-(4-[4-chloro-6-[methyl-(3-sulfonatophenyl)amino]-1,3,5-triazin-2-ylamino]-2-ureidophenylazo)naphthalene-1,3,6-trisulfonate; tetrasodium 7-(4-[4-chloro-6-[methyl-(4-sulfonatophenyl)amino]-1,3,5-triazin-2-ylamino]-2-ureidophenylazo)naphthalene-1,3,6-trisulfonate (1:1) | 412-940-2 | 148878-18-6 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-476-00-4 | trisodium N,N-bis(carboxymethyl)-β-alanine | 414-070-9 | 129050-62-0 | Skin Corr. 1B Aquatic Chronic 3 | H314 H412 | GHS05 Dgr | H314 H412 |  |  |  | CLP00
607-477-00-X | (1α5α6α)-6-nitro-3-benzyl-3-azabicyclo[3.1.0]hexane methanesulfonate salt | 426-740-8 | - | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | ATP01
607-478-00-5 | tetramethylammonium hydrogen phthalate | 416-900-5 | 79723-02-7 | Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 | H301 H373 ** H400 | GHS06 GHS08 GHS09 Dgr | H301 H373 ** H400 |  |  |  | CLP00
607-479-00-0 | hexadecyl 4-chloro-3-[2-(5,5-dimethyl-2,4-dioxo-1,3-oxazolidin-3-yl)-4,4-dimethyl-3-oxopentamido]benzoate | 418-550-9 | 168689-49-4 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-480-00-6 | 1,2-benzenedicarboxylic acid; di-C7-11-branched and linear alkylesters | 271-084-6 | 68515-42-4 | Repr. 1B | H360Df | GHS08 Dgr | H360Df |  |  |  | CLP00
607-481-00-1 | reaction mass of: trihexyl citrate; dihexyloctyl citrate; dioctylhexyl citrate; dihexyldecyl citrate | 430-290-8 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-482-00-7 | N-[1-(S)-ethoxycarbonyl-3-phenylpropyl]-.sc.l.sc.-alanyl-N-carboxyanhydride | 430-360-8 | 84793-24-8 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
607-483-00-2 | 1,2-benzenedicarboxylic acid; di-C6-8-branched alkylesters, C7-rich | 276-158-1 | 71888-89-6 | Repr. 1B | H360D *** | GHS08 Dgr | H360D *** |  |  |  | ATP01
607-484-00-8 | ethyl 2-{[3-acetylamino-4-(6-bromo-2-methyl-1,3-dioxo-2,3-dihydro-1H-isoindol-5-ylazo)phenyl]ethylamino}propionate | 430-480-0 | 221452-67-1 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-485-00-3 | (3S-trans)-phenyl-3-[(1,3-benzodioxol-5-yloxy)methyl]-4-(4-fluorophenyl)-1-piperidinecarboxylate | 430-510-2 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-486-00-9 | potassium sodium 5'-(6-chloro-4-(2-(2-vinylsulfonylethoxy)ethylamino)-1,3,5-triazin-2-ylamino)-4'-hydroxy-2,3'-azodinaphthalene-1,2',5,7'-disulfonate | 402-110-8 | 110081-40-8 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
607-487-00-4 | reaction mass of: disodium 4-(3-ethoxycarbonyl-4-(5-(3-ethoxycarbonyl-5-hydroxy-1-(4-sulfonatophenyl)pyrazol-4-yl)penta-2,4-dienylidene)-4,5-dihydro-5-oxopyrazol-1-yl)benzenesulfonate; trisodium 4-(3-ethoxycarbonyl-4-(5-(3-ethoxycarbonyl-5-oxido-1-(4-sulfonatophenyl)pyrazol-4-yl)penta-2,4-dienylidene)-4,5-dihydro-5-oxopyrazol-1-yl)benzenesulfonate | 402-660-9 |  | Repr. 1B Aquatic Chronic 3 | H360D *** H412 | GHS08 Dgr | H360D *** H412 |  |  |  | CLP00
607-488-00-X | ethyl (2-acetylamino-5-fluoro-4-isothiocyanatophenoxy)acetate | 414-210-9 | 147379-38-2 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-489-00-5 | reaction mass of: 2-ethylhexyl linolenate, linoleate and oleate; 2-ethylhexyl epoxyoleate; 2-ethylhexyl diepoxylinoleate; 2-ethylhexyl triepoxylinolenate | 414-890-7 | 71302-79-9 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-490-00-0 | N-[2-hydroxy-3-(C12-16-alkyloxy)propyl]-N-methyl glycinate | 415-060-7 |  | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
607-491-00-6 | reaction mass of: diester of 4,4'-methylenebis[2-(2-hydroxy-5-methylbenzyl)-3,6-dimethylphenol] and 6-diazo-5,6-dihydro-5-oxonaphthalene-1-sulfonic acid (1:2); triester of 4,4'-methylenebis[2-(2-hydroxy-5-methylbenzyl)-3,6-dimethylphenol] and 6-diazo-5,6-dihydro-5-oxonaphthalene-1-sulfonic acid (1:3) | 427-140-9 | - | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | ATP01
607-492-00-1 | 2-(1-(3',3'-dimethyl-1'-cyclohexyl)ethoxy)-2-methyl propyl propanoate | 415-490-5 | 141773-73-1 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-493-00-7 | methyl (3aR,4R,7aR)-2-methyl-4-(1S,2R,3-triacetoxypropyl)-3a,7a-dihydro-4H-pyrano[3,4-d]oxazole-6-carboxylate | 415-670-3 | 78850-37-0 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
607-494-00-2 | bis(2-ethylhexyl)octylphosphonate | 417-170-0 | 52894-02-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-495-00-8 | sodium 4-sulfophenyl-6-((1-oxononyl)amino)hexanoate | 417-550-6 | 168151-92-6 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
607-496-00-3 | 2,2'-methylenebis(4,6-di-tert-butyl-phenyl)-2-ethylhexyl phosphite | 418-310-3 | 126050-54-2 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-497-00-9 | cerium oxide isostearate | 419-760-3 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-498-00-4 | (E)-3,7-dimethyl-2,6-octadienylhexadecanoate | 421-370-3 | 3681-73-0 | Skin Irrit. 2 Aquatic Chronic 4 | H315 H413 | GHS07 Wng | H315 H413 |  |  |  | CLP00
607-499-00-X | bis(dimethyl-(2-hydroxyethyl)ammonium) 1,2-ethanediyl-bis(2-hexadecenylsuccinate) | 421-660-1 |  | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H318 H317 H411 |  |  |  | CLP00
607-500-00-3 | calcium 2,2,bis[(5-tetrapropylene-2-hydroxy)phenyl]ethanoate | 421-670-4 |  | Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H400 H410 | GHS07 GHS09 Wng | H315 H410 |  |  |  | CLP00
607-501-00-9 | reaction mass of: triphenylthiophosphate and tertiary butylated phenyl derivatives | 421-820-9 | 192268-65-8 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
607-502-00-4 | (N-benzyl-N,N,N-tributyl)ammonium 4-dodecylbenzenesulfonate | 422-200-0 | 178277-55-9 | Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 2 | H302 H314 H411 | GHS05 GHS07 GHS09 Dgr | H314 H302 H411 |  |  |  | CLP00
607-503-00-X | 2,4,6-tri-n-propyl-2,4,6-trioxo-1,3,5,2,4,6-trioxatriphosphorinane | 422-210-5 | 68957-94-8 | Skin Corr. 1B | H314 | GHS05 Dgr | H314 |  |  |  | CLP00
607-504-00-5 | diammonium 1-hydroxy-2-(4-(4-carboxyphenylazo)-2,5-dimethoxyphenylazo)-7-amino-3-naphthalenesulfonate | 422-670-7 |  | Repr. 2 Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H361f H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H301 H361f H373 ** H410 |  |  |  | ATP01/ATP01corr
607-505-00-0 | pentasodium 7-(4-(4-(5-amino-4-sulfonato-2-(4-((2-(sulfonato-ethoxy)sulfonyl)phenylazo)phenylamino)-6-chloro-1,3,5-triazin-2-yl)amino-2-ureidophenylazo)naphtalene-1,3,6-trisulfonate | 422-930-1 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-506-00-6 | reaction mass of: strontium (4-chloro-2-((4,5-dihydro-3-methyl-5-oxo-1-(3-sulfonatophenyl)-1H-pyrazol-4-yl)azo)-5-methyl)benzenesulfonate; disodium (4-chloro-2-((4,5-dihydro-3-methyl-5-oxo-1-(3-sulfonatophenyl)-1H-pyrazol-4-yl)azo)-5-methyl)benzenesulfonate | 422-970-8 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
607-507-00-1 | potassium,sodium 2,4-diamino-3-[4-(2-sulfonatoethoxysulfonyl)phenylazo]-5-[4-(2-sulfonatoethoxysulfonyl)-2-sulfonatophenylazo]-benzenesulfonate | 422-980-2 | 187026-95-5 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
607-508-00-7 | disodium 3,3'-[iminobis[sulfonyl-4,1-phenylene-(5-hydroxy-3-methylpyrazole-1,4-diyl)azo-4,1-phenylenesulfonylimino-(4-amino-6-hydroxypyrimidine-2,5-diyl)azo-4,1-phenylenesulfonylimino(4-amino-6-hydroxypyrimidine-2,5-diyl)azo]bis(benzenesulfonate)] | 423-110-4 |  | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
607-509-00-2 | 2-phenoxyethyl 4-aminobenzoate | 430-880-5 | 88938-23-2 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
607-510-00-8 | (2S,5R)-6,6-dibromo-3,3-dimethyl-7-oxo-4-thia-1-azabicyclo[3.2.0]heptane-2-carboxylic acid 4,4-dioxide | 427-200-4 | 76646-91-8 | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H302 H315 H318 H317 | GHS05 GHS07 Dgr | H302 H315 H318 H317 |  |  |  | ATP01
607-511-00-3 | reaction mass of: 4-[(3-decyloxypropyl)(3-isobutoxy-1-isobutoxycarbonyl-3-oxopropyl)amino]-4-oxobutyric acid; 4-[(3-isobutoxy-1-isobutoxycarbonyl-3-oxopropyl)(3-octyloxypropyl)amino]-4-oxobutyric acid | 423-750-4 | - | Eye Irrit. 2 Aquatic Chronic 2 | H319 H411 | GHS07 GHS09 Wng | H319 H411 |  |  |  | ATP01
607-512-00-9 | trisodium 2,4-diamino-3,5-bis-[4-(2-sulfonatoethoxy)sulfonyl)phenylazo]benzenesulfonate | 423-970-0 | 182926-43-8 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
607-513-00-4 | reaction mass of: Trisodium 4-benzoylamino-6-(6-ethenesulfonyl-1-sulfato-naphthalen-2-ylazo)-5-hydroxynaphthalene-2,7-disulfonate; 5-(benzoylamino)-4-hydroxy-3-((1-sulfo-6-((2-(sulfooxy)ethyl)sulfonyl)-2-naphthyl)azo)naphthalene-2,7-disulfonic acid sodium salt; 5-(benzoylamino)-4-hydroxy-3-((1-sulfo-6-((2-(sulfooxy)ethyl)sulfonyl)-2-naphthyl)azo)naphthalene-2,7-disulfonic acid | 423-200-3 |  | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | CLP00
607-514-00-X | potassium N-(1-methoxy-1-oxobut-2-en-3-yl)valinate | 427-240-2 | 134841-35-3 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-515-00-5 | reaction mass of: disodium hexyldiphenyl ether disulphonate; disodium dihexyldiphenyl ether disulphonate | 429-650-7 | 147732-60-3 | Eye Irrit. 2 Aquatic Chronic 2 | H319 H411 | GHS07 GHS09 Wng | H319 H411 |  |  |  | CLP00
607-516-00-0 | N,N'-bis(trifluoroacetyl)-S,S'-bis L-homocysteine | 429-670-6 | 105996-54-1 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
607-517-00-6 | (S)-α-(acetylthio)benzenepropanoic acid | 430-300-0 | 76932-17-7 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 | H302 H318 H317 | GHS05 GHS07 Dgr | H302 H318 H317 |  |  |  | CLP00
607-518-00-1 | 3-oxoandrost-4-ene-17-ß-carboxylic acid | 414-990-0 | 302-97-6 | Repr. 2 Aquatic Chronic 4 | H361f H413 | GHS08 Wng | H361f H413 |  |  |  | ATP01/ATP01corr
607-519-00-7 | poly-[((4-((4-ethyl-ethylene)amino)phenyl)-((4-(ethyl-(2-oxyethylene)amino)phenyl)methinyl)cyclohexa-2,5-dienylidene)-N-ethyl-N-(2-hydroxyethyl)ammonium acetate] | 427-280-0 | 176429-27-9 | STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H335 H315 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H335 H315 H318 H410 |  |  |  | ATP01
607-520-00-2 | reaction mass of: sodium 4,5-dihydro-2-[(propionato)(C6-18)alkyl]-3H-imidazolium-N-ethylphosphate; disodium 4,5-dihydro-2-[(dipropionato)(C6-18)alkyl]-3H-imidazolium-N-ethylphosphate | 427-740-0 | - | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
607-521-00-8 | tetraethyl N,N'-(methylenedicyclohexane-4,1-diyl)bis-.sc.dl.sc.-aspartate | 429-270-1 | 136210-30-5 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
607-522-00-3 | sodium salt of the polymer of: sodium 2-methyl-buta-1,3-diene-1-sulfonate with acrylic acid and 2-hydroxyethyl-2-methylacrylate | 429-720-7 | 184246-86-4 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
607-523-00-9 | reaction mass of mono to tetra(lithium and/or sodium)3-amino-10-[4-(4-amino-3-sulfonatoanilino)-6-[methyl-(2-sulfonatoethyl)amino]-1,3,5-triazin-2-ylamino]-6-13-dichlorobenzo[1,2-B:4,5-B']di[1,4]benzoxazine-4,11-disulfonate; mono to tetra(lithium and/or sodium)3-amino-10-[4,6-bis(4-amino-3-sulfonatoanilino)-1,3,5-triazin-2-ylamino]-6-13-dichlorobenzo[1,2-B:4,5-B']di[1,4]benzoxazine-4,11-disulfonate; mono to penta(lithium and/or sodium)10,10´-diamino-6,6',13,13´-tetrachloro-3,3'-[6-[methyl-(2-sulfonatoethyl)amino]-1,3,5-triazin-2,4-diyldiimino]bis[benzo[1,2-B:4,5-B']di[1,4]benzoxazine-4,11-disulfonate; mono to hepta(lithium and/or sodium)10-amino-6,6',13,13'-tetrachloro-10´[4-(4-amino-3-sulfonatoanilino)-[6-methyl-(2-sulfonatoethyl)amino]-1,3,5-triazin-2,4-diimino]bis[benzo[1,2-B:4,5-B']di[1,4]benzoxazine-4,11-disulfonate; mono to hepta(lithium and/or sodium)10,10'-diamino-6,6',3,3'[(2-sulfonato)-1,4-phenylenediiminobis[6-methyl-(2-sulfonatoethyl)amino]-1,3,5-triazin-2,4-diyldiimino]bis[benzo[1,2-B:4,5-B']di[1,4]benzoxazine-4,11-disulfonate | 430-200-7 | - | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
607-524-00-4 | tall oil 2-[(tetrahydro-2H-pyran-2-yl) thio]ethyl esters | 430-310-5 | - | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | ATP01
607-525-00-X | (Z)-2-methoxymino-2-[2-(tritylamino)thiazol-4-yl]acetic acid | 431-520-1 | 64485-90-1 | Flam. Sol. 1 **** Carc. 2 Aquatic Chronic 3 | H228 H351 H412 | GHS02 GHS08 Dgr | H228 H351 H412 |  |  |  | ATP01
607-526-00-5 | cartap (ISO); 1,3-bis(carbamoylthio)-2-(dimethylamino)propane |  | 15263-53-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
607-527-00-0 | reaction mass of: 1-(1'H,1'H,2'H,2'H-tridecafluorooctyl)-12-(1''H,1''H,2''H,2''H-tridecafluorooctyl)dodecanedioate; 1-(1'H,1'H,2'H,2'H-tridecafluorooctyl)-12-(1''H,1''H,2''H,2''H-heptdecafluorodecyl)dodecanedioate; 1-(1'H,1'H,2'H,2'H-tridecafluorooctyl)-12-(1''H,1''H,2''H,2''H-heneicosafluorododecyl)dodecanedioate; 1-(1'H,1'H,2'H,2'H-tridecafluorooctyl)-12-(1''H,1''H,2''H,2''H-pentacosafluorotetradecyl)dodecanedioate; 1-(1'H,1'H,2'H,2'H-heptadecafluorodecyl)-12-(1''H,1''H,2''H,2''H-heptadecafluorodecyl)dodecanedioate; 1-(1'H,1'H,2'H,2'H-heptadecafluorodecyl)-12-(1''H,1''H,2''H,2''H-heneicosafluorododecyl)dodecanedioate | 423-180-6 |  | STOT RE 2 * | H373 ** | GHS08 Wng | H373 ** |  |  |  | CLP00
607-528-00-6 | (S)-3-methyl-2-(2-oxotetrahydropyrimidine-1-yl)butyric acid | 430-900-2 | 192725-50-1 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-529-00-1 | benzyl cis-4-ammonium-4'-toluenesulfonato-1-cyclohexanecarboxylate | 426-070-6 | 67299-45-0 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
607-530-00-7 | reaction mass of isomers of: C7-9-alkyl 3-(3,5-di-tert-butyl-4-hydroxyphenyl)propionate | 406-040-9 | 125643-61-0 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-531-00-2 | methyl 3-amino-4,6-dibromo-2-methyl-benzoate | 425-190-6 | 119916-05-1 | STOT RE 2 * Aquatic Chronic 2 | H373 ** H411 | GHS08 GHS09 Wng | H373 ** H411 |  |  |  | ATP01
607-532-00-8 | (S)-1-[2-tert-butoxycarbonyl-3-(2-methoxyethoxy)propyl]-1-cyclopentanecarboxylic acid, cyclohexylamine salt | 425-510-4 | 167944-94-7 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
607-533-00-3 | pentasodium monohydrogen 6-chloro-3,10-bis[2-[4-chloro-6-(2,4-disulfophenylamino)-1,3,5-triazin-2-yl-amino]ethylamino]-13-ethylbenzo[5.6][1.4]oxazino[2,3-b]phenoxazine-4,11-disulfonate | 414-910-4 | - | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
607-534-00-9 | ethyl 2-(3-benzoylphenyl)propanoate | 414-920-9 | 60658-04-0 | Acute Tox. 3 * STOT RE 1 Skin Sens. 1 Aquatic Chronic 2 | H301 H372 ** H317 H411 | GHS06 GHS08 GHS09 Dgr | H301 H372 ** H317 H411 |  |  |  | ATP01
607-535-00-4 | potassium 4-iodo-2-sulfonato-benzoic acid | 426-620-5 | - | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
607-536-00-X | (2,6-xylyloxy) acetic acid | 430-910-7 | 13335-71-2 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | ATP01
607-537-00-5 | isopropylammonium 2-(3-benzoylphenyl)propionate | 417-970-1 | - | Acute Tox. 3 * Acute Tox. 4 * STOT RE 1 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H312 H372 ** H318 H400 H410 | GHS06 GHS05 GHS08 GHS09 Dgr | H301 H312 H372 ** H318 H410 |  |  |  | ATP01
607-539-00-6 | propyl((4-(5-oxo-3-propylisoxazolidin-4-ylidenmethin)phenyl)propoxycarbonylmethyleneamino)acetate | 431-000-2 | 198705-81-6 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-540-00-1 | 1-(mercaptomethyl)cyclopropylacetic acid | 420-240-3 | 162515-68-6 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H312 H302 H314 H317 H411 | GHS05 GHS07 GHS09 Dgr | H314 H312 H302 H317 H411 |  |  |  | ATP01
607-541-00-7 | [(1-methyl-1,2-ethanediyl)bis[nitrilobis(methylene)]]tetrakis(phosphonic acid) | 421-940-1 | 28698-31-9 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | ATP01
607-542-00-2 | methyl 2-(4-butanesulfonamidophenoxy)tetradecanoate | 422-110-1 | - | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
607-543-00-8 | poly-[((4-((4-(ethyl-ethylene)amino)phenyl)-(4-(ethyl-(2-oxyethylene)amino)phenyl)methinyl)-3-methylcyclohexa-2,5-dienylidene)-N-ethyl-N-(2-hydroxyethyl)ammonium acetate] | 427-480-8 | 176429-22-4 | STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H335 H315 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H335 H315 H318 H410 |  |  |  | ATP01
607-544-00-3 | ethyl 6,8-difluoro-1-(formylmethylamino)-1,4-dihydro-7-(4-methyl)piperazin-1-yl)-4-oxo-quinoline-3-carboxylate | 427-490-2 | 158585-86-5 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
607-545-00-9 | 1,2-dimethyl-3-(1-methylethenyl)cyclopentyl acetate | 424-070-0 | 94346-09-5 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | ATP01
607-546-00-4 | reaction mass of: methyl {[5-acetylamino-4-(2-chloro-4-nitrophenylazo)phenyl]methoxycarbonylmethylamino}acetate; methyl {[5-acetylamino-4-(2-chloro-4-nitrophenylazo)phenyl]ethoxycarbonylmethylamino}acetate | 424-290-7 | 188070-47-5 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-547-00-X | 18-methylnonadecyl 2,2-dimethylpropanoate | 424-370-1 | 125496-22-2 | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 4 | H315 H317 H413 | GHS07 Wng | H315 H317 H413 |  |  |  | ATP01/ATP01corr
607-548-00-5 | 1-(2,4-dichlorophenyl)-2-(1H-imidazol-1-yl)ethanone methanesulfonate | 431-010-7 | 154486-26-7 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | ATP01
607-549-00-0 | methyl (E)-2((3-(1,3-benzodioxol-5-yl)-2-methyl-1-propenyl)amino)benzoate | 424-430-7 | 125778-19-0 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
607-550-00-6 | 2-amino-4-bromo-5-chlorobenzoic acid | 424-700-4 |  | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01/ATP01corr
607-551-00-1 | tetrabutylammonium 2-amino-6-iodopurinate | 424-710-9 | 156126-48-6 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H312 H302 H373 ** H315 H318 H317 H411 | GHS05 GHS08 GHS07 GHS09 Dgr | H312 H302 H373 ** H315 H318 H317 H411 |  |  |  | ATP01
607-552-00-7 | hexadecyl 3-amino-4-isopropoxybenzoate | 424-830-1 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-553-00-2 | 7-amino-4-hydroxy-2-naphthalenesulfonic acid, coupled with 5 (or 8) -amino-8 (or 5)-[[4-[[4-[[4-amino-6(or 7)-sulfo-1-naphthyl]azo]phenyl]amino]-3-sulfophenyl]azo]-2-naphthalenesulfonic acid and 4-hydroxy-7-(phenylamino)-2-naphthalenesulfonic acid, sodium salt | 424-850-0 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-554-00-8 | 2,4-diamino-5-[4-[(2-sulfoxyl ethyl)sulfonyl]phenylazo]benzenesulfonic acid | 424-870-1 | 27624-67-5 | Expl. 1.1 Eye Dam. 1 Aquatic Chronic 3 | H201 H318 H412 | GHS01 GHS05 Dgr | H201 H318 H412 |  |  |  | ATP01
607-555-00-3 | 1,1,3,3-tetramethylbutylperoxypivalate | 424-980-8 | 22288-41-1 | Flam. Liq. 2 Org. Perox. D Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H225 H242 H315 H317 H411 | GHS02 GHS07 GHS09 Dgr | H225 H242 H315 H317 H411 |  |  |  | ATP01
607-556-00-9 | 2-acetoxymethylene-4-acetylphenylacetate | 425-160-2 | 24085-06-1 | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H318 H317 H400 H410 | GHS05 GHS08 GHS07 GHS09 Dgr | H302 H373 ** H318 H317 H410 |  |  |  | ATP01
607-557-00-4 | salt of: (1S-cis)-1-amino-2,3-dihydro-1H-inden-2-ol and [R-[R*R*]]-2,3-dihydroxybutanedioic acid | 425-210-3 | 169939-84-8 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-558-00-X | 2S-isopropyl-5R-methyl-1R-cyclohexyl (2R,5S)-5-(4-amino-2-oxo-2H-pyrimidin-1-yl)-[1.3]-oxathiolane-2-carboxylate | 425-250-1 | 147027-10-9 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
607-559-00-5 | coconut oil, reaction products with glycerol esters of 3,5-bis(1,1-dimethylethyl)-4-hydroxybenzenepropanoic acid | 425-400-6 | 179986-09-5 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-560-00-0 | (R,S)-2-butyloctanedioic acid | 431-210-4 | 50905-10-7 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-561-00-6 | sodium 4-hydroxy-3-(N'-(2-(2-hydroxyethylenesulfonyl)ethylene)ureido)-5-nitrobenzenesulfonate | 425-460-3 | - | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
607-562-00-1 | reaction mass of: (2R,3R)-3-(2-ethoxyphenoxy)-2-hydroxy-3-phenylpropylammonium methanesulfonate; (2S,3S)-3-(2-ethoxyphenoxy)-2-hydroxy-3-phenylpropylammonium methanesulfonate | 425-530-3 | 98769-75-6 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | ATP01
607-563-00-7 | 5,7-dichloro-4-hydroxyquinoline-3-carboxylic acid | 431-250-2 | 171850-30-9 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
607-564-00-2 | 1,6-hexanediammonium, sodium 5-sulfato-1,3-benzenedicarboxylate | 425-730-0 | 51178-75-7 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-565-00-8 | 3-ethyl 5-methyl 2-(2-aminoethoxymethyl)-4-(2-chlorophenyl)-1,4-dihydro-6-methyl-3,5-pyridinedicarboxylate | 425-820-1 | 88150-42-9 | Acute Tox. 3 * STOT RE 2 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H373 ** H318 H400 H410 | GHS06 GHS05 GHS08 GHS09 Dgr | H301 H373 ** H318 H410 |  |  |  | ATP01
607-566-00-3 | reaction mass of: dodecylphenyl dodecylhydroxybenzenecarboxylate; bis(dodecylphenyl)dodecyl hydroxybenzenedicarboxylate | 426-140-6 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-567-00-9 | potassium 3-iodo-6-methylbenzenesulfonate | 426-300-5 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-568-00-4 | potassium 2-chloro-3-(benzyloxy)propionate | 426-350-8 | 138666-92-9 | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 | H302 H373 ** H318 H317 | GHS05 GHS08 GHS07 Dgr | H302 H373 ** H318 H317 |  |  |  | ATP01
607-569-00-X | reaction mass of: sodium 2-amino-4-(2,6-difluoropyrimidin-4-ylamino)benzenesulfonate; sodium 2-amino-4-(4,6-difluoropyrimidin-4-ylamino)benzenesulfonate | 426-470-0 | - | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-570-00-5 | sodium (6R-trans)-7-amino-8-oxo-3-[[[1-(sulfomethyl)-1H-tetrazol-5-yl]thio]methyl]-5-thia-1-azabicyclo[4.2.0]oct-2-ene-2-carboxylate monohydrate | 426-520-1 | 71420-85-4 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-571-00-0 | 2-cyclopentene-1-acetic acid, 3-hydroxy-2-pentyl-, methyl ester acetate | 431-400-7 | 57374-49-9 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
607-572-00-6 | diethyl thiophosphoryl (Z)-(2-aminothiazol-4-yl)methoxyimino acetate | 426-790-0 | 162208-27-7 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H312 H302 H373 ** H317 H410 |  |  |  | ATP01
607-573-00-1 | reaction mass of: disodium 7-(2,4-difluoropyrimidin-6-ylamino)-4-hydroxy-3-(4-methoxy-2-sulfonatophenylazo)naphthalene-2-sulfonate; disodium 7-(4,6-difluoropyrimidin-2-ylamino)-4-hydroxy-3-(4-methoxy-2-sulfonatophenylazo)naphthalene-2-sulfonate | 426-840-1 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-574-00-7 | [1R-(1-α,2β,5α)]-mono[5-methyl-2-(1-methylethyl)cyclohexyl]butanedioate | 426-890-4 | 77341-67-4 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-575-00-2 | 4-(5-(5-[1-(4-carboxyphenyl)hexahydro-2,4,6-trioxopyrimidin-5-ylidene]penta-1,3-dienyl)-1,2,3,4-tetrahydro-6-hydroxy-2,4-dioxopyrimidin-1-yl)benzoic acid-triethylamine salt | 426-900-7 | - | STOT SE 3 Aquatic Chronic 3 | H335 H412 | GHS07 Wng | H335 H412 |  |  |  | ATP01
607-576-00-8 | branched, octyl 3-[3,5-di(tert-butyl)-4-hydroxyphenyl]propanoate | 427-030-0 | - | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
607-577-00-3 | (2R*,3S*)-2-(2,4-difluorophenyl)-3-(5-fluoro-4-pyrimidinyl)-1-(1H-1,2,4-triazol-1-yl)butan-2-ol (1R)-10-camphorsulfonate | 427-100-0 | - | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H302 H318 H317 H412 | GHS05 GHS07 Dgr | H302 H318 H317 H412 |  |  |  | ATP01
607-578-00-9 | ethyl 4-((4-diethylamino-2-methylphenyl)imino)-4,5-dihydro-1-isopropyl-5-oxo-1H-pyrazole-3-carboxylate | 427-110-5 | - | Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 4 | H302 H373 ** H413 | GHS08 GHS07 Wng | H302 H373 ** H413 |  |  |  | ATP01
607-579-00-4 | diethyl[(p-ethoxyanilino)methylene]malonate | 431-430-0 | 103976-28-9 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | ATP01
607-580-00-X | ethyl 7-chloro-1-(2,4-difluorophenyl)-6-fluoro-1,4-dihydro-4-oxo-1,8-naphthyridine-3-carboxylate | 422-360-1 | 100491-29-0 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
607-581-00-5 | ethyl 2-ethoxy-4-carboxymethylbenzoate | 427-630-2 | 99469-99-5 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-582-00-0 | reaction mass of: tetrasodium 7-(4-(4-fluoro-6-(4-(2-sulfonatoethylsulfonyl)phenylamino)-1,3,5-triazin-2-ylamino)-2-ureidophenylazo)naphthalene-1,3,6-trisulfonate; tetrasodium 7-(4-(4-hydroxy-6-(4-(2-sulfonatoethylsulfonyl)phenylamino)-1,3,5-triazin-2-ylamino)-2-ureidophenylazo)naphthalene-1,3,6-trisulfonate | 427-650-1 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
607-583-00-6 | 4-amino-3-[[4-[[2-(sulfooxy)ethyl]sulfonyl]phenyl]azo]-1-naphthalene sulfonic acid | 427-680-5 | 188907-52-0 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | ATP01
607-584-00-1 | trisodium 3-[2-acetylamino-4-[4-chloro-6-[4-(2-sulfonatoxyethylsulfonyl)phenylamino]-1,3,5-triazine-2-ylamino]phenylazo]naphthalene-1,5-disulfonate | 427-710-7 | 215612-56-9 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | ATP01
607-585-00-7 | strontium 2-[(2-hydroxy-6-sulfonato-1-naphthyl)azo]naphthalene-1-sulfonate | 427-930-3 | - | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-586-00-2 | dodecyl 3-amino-4-chlorobenzoate | 428-020-9 | 6195-20-6 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
607-587-00-8 | ethyl cis-4-[4-[[2-(2,4-dichlorophenyl)-2-(1H-imidazol-1-ylmethyl)-1,3-dioxolan-4-yl]methoxy]phenyl]piperazine-1-carboxylate | 428-030-3 | 67914-69-6 | Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H410 |  |  |  | ATP01
607-588-00-3 | reaction mass of: 2-ethylhexyl 2,3,4,5-tetrabromobenzoate; bis(2-ethylhexyl) 3,4,5,6-tetrabromophthalate | 428-050-2 | - | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | ATP01
607-589-00-9 | tetrakis(1,2,2,6,6-pentamethyl-4-piperidyl)-1,2,3,4-butanetetracarboxylate | 428-070-1 | 91788-83-9 | Acute Tox. 4 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H372 ** H400 H410 | GHS08 GHS07 GHS09 Dgr | H372 ** H302 H410 |  |  |  | ATP01
607-590-00-4 | hexadecyl 3-[2-(5,5-dimethyl-2,4-dioxo-1,3-oxazolidin-3-yl)-4,4-dimethyl-3-oxovaleramido]-4-isopropoxybenzoate | 428-140-1 | 210706-50-6 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-591-00-X | reaction mass of: trisodium 5-(4-fluoro-6-morpholin-4-yl-1,3,5-triazin-2-ylamino)-4-hydroxy-3-(4-(2-sulfooxyethanesulfonyl)phenylazo)naphthalene-2,7-disulfonate; disodium 3-(4-ethenesulfonylphenylazo)-5-(4-fluoro-6-morpholin-4-yl-1,3,5-triazin-2-ylamino)-4-hydroxynaphthalene-2,7-disulfonate | 428-400-4 |  | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01/ATP01corr
607-592-00-5 | di(C9-11-alkyl) cyclohexane-1,4-dicarboxylate | 428-870-0 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-593-00-0 | 4-(2-methylacryloyloxy)phenyl 4-allyloxybenzoate | 429-000-2 | 159235-16-2 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
607-594-00-6 | ethyl (1S,5R,6S)-5-(1-ethylpropoxy)-7-oxabicyclo[4.1.0]hept-3-ene-3-carboxylate | 429-020-1 | 204254-96-6 | STOT RE 2 * Skin Sens. 1 | H373 ** H317 | GHS08 GHS07 Wng | H373 ** H317 |  |  |  | ATP01
607-595-00-1 | N-amidino-N-methylglycine-2-oxopropionate | 429-120-5 | 208535-04-0 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-596-00-7 | ethyl 2-(4-phenoxyphenyl)lactate | 429-220-9 | 132584-17-9 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | ATP01
607-597-00-2 | tetrasodium 4,4'-bis{4-[4-(2-hydroxyethylamino)-6-(4-sulfonatoanilino)-1,3,5-triazin-2-ylamino]phenylazo}stilbene-2,2'-disulfonate | 429-230-3 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-598-00-8 | trisodium 3-amino-4-[4-[4-(2-(2-ethenylsulfonylethoxy)ethylamino)-6-fluoro-1,3,5-triazine-2-ylamino]-2-sulfophenylazo]-5-hydroxynaphthalene-2,7-disulfonate | 429-240-8 | 212652-59-0 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-599-00-3 | 1,1-dimethylpropyl 3,5,5-trimethylperoxyhexanoate | 431-610-9 | 68860-54-8 | Org. Perox. D Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H242 H317 H400 H410 | GHS02 GHS07 GHS09 Dgr | H242 H317 H410 |  |  |  | ATP01
607-600-00-7 | (1S,1'R)-[1-(3',3'-dimethyl-1'-cyclohexyl)ethoxycarbonyl]methyl propanoate | 431-700-8 | - | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
607-601-00-2 | 1,4-dihydroxy-2,2,6,6-tetramethyl piperidinium-2-hydroxy-1,2,3-propanetricarboxylate | 429-370-5 | 220410-74-2 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | ATP01
607-602-00-8 | ethyl (3-cyanomethyl-3,4-dihydro-4-oxophthalazin-1-yl)acetate | 429-680-0 | 122665-86-5 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
607-603-00-3 | lithium sodium 4,4',4''-(nitrilotris(ethane-2,1-diylimino(6-chloro-1,3,5-triazine-4,2-diyl)imino))tris(5-hydroxy-6-(1-sulfonaphthalene-2-ylazo)-2,7-naphthalene)disulfonate | 429-730-1 | 193562-37-7 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
607-604-00-9 | guanidinium benzoate | 429-820-0 | 26739-54-8 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | ATP01
607-605-00-4 | methyl 4-iodo-2-(3-(4-methoxy-6-methyl-1,3,5-triazine-2-yl)ureidosulfonyl)benzoate | 429-890-2 | 144550-06-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
607-606-00-X | (Z)-2-(2-t-butoxycarbonylamino-4-thiazolyl)pent-2-enoic acid | 430-100-3 | 86978-24-7 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | ATP01
607-607-00-5 | reaction mass of: calcium bis(C10-14 branched alkyl salicylate); calcium bis(C18-30-alkyl salicylate); calcium C10-14 branched alkylsalicylato-C18-30-alkyl salicylate; calcium bis (C10-14 branched alkyl phenolate); calcium bis (C18-30-alkyl phenolate); calcium C10-14 branched alkylphenolato-C18-30-alkyl phenolate; C10-14 branched alkyl phenol; C18-30-alkyl phenol | 430-180-1 | - | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | ATP01
607-608-00-0 | pentapotassium 2-(4-{5-[1-(2,5-disulfophenyl)-4,5-dihydro-3-methylcarbamoyl-5-oxopyrazol-4-ylidene]-3-(2-pyrrolidinone-1-yl)-1,3-pentadienyl}-3-methylcarbamoyl-5-oxopyrazol-1-yl)benzene-1,4-disulfonate | 430-210-1 | - | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
607-609-00-6 | ethyl (3R)-4-cyano-3-hydroxybutanoate | 430-220-6 | 141942-85-0 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | ATP01
607-610-00-1 | trisodium 4-hydroxy-6-(sulfonatomethylamino)-5-(2-(2-sulfatoethylsulfonyl)phenylazo)naphthalene-2-sulfonate | 430-280-3 | - | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-611-00-7 | methyl 3-amino-2,2,3-trimethylbutyrate | 431-720-7 | 90886-53-6 | Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 3 | H302 H314 H412 | GHS05 GHS07 Dgr | H314 H302 H412 |  |  |  | ATP01
607-612-00-2 | reaction mass of: 3,3,4,4,5,5,6,6,7,7,8,8,8-tridecafluoro-1-octanesulfonic acid; ammonium 3,3,4,4,5,5,6,6,7,7,8,8,8-tridecafluoro-1-octanesulfonate | 432-190-1 | 182176-52-9 | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 | H302 H373 ** H318 | GHS05 GHS08 GHS07 Dgr | H302 H373 ** H318 |  |  |  | ATP01
607-613-00-8 | reaction mass of: succinic acidmonopersuccinic aciddipersuccinic acidmonomethyl ester of succinic acidmonomethyl ester of persuccinic aciddimethyl succinate glutaric acidmonoperglutaric aciddiperglutaric acidmonomethyl ester of glutaric acidmonomethyl ester of perglutaric aciddimethyl glutarate adipic acidmonoperadipic aciddiperadipic acidmonomethyl ester of adipic acidmonomethyl ester of peradipic aciddimethyl adipatehydrogen peroxidemethanolwater | 432-790-1 |  | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 2 Skin Corr. 1B | H332 H312 H302 H371 H314 | GHS07 GHS05 GHS08 Dgr | H302 H312 H332 H314 H371 |  |  |  | ATP01/ATP05
607-614-00-3 | 2-(10-oxo-10H-9-oxa-10-phosphaphenanthren-10-ylmethyl)succinic acid | 426-480-5 | 63562-33-4 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
607-615-00-9 | reaction product of thioglycerol and mercaptoacetic acid consisting mainly of 3-mercapto-1,2-bismercaptoacetoxypropane and oligomers of this substance | 431-120-5 | - | Acute Tox. 3 * Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 | H331 H302 H319 H317 | GHS06 Dgr | H331 H302 H319 H317 |  |  |  | ATP01
607-616-00-4 | 2,4-dichloro-5-fluorobenzoylchloride | 428-390-1 | 86393-34-2 | STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H335 H315 H318 H317 H412 | GHS05 GHS07 Dgr | H335 H315 H318 H317 H412 |  |  |  | ATP01
607-617-00-X | bis(2-ethylhexyl)-4,5-epoxycyclohexane-1,2-dicarboxylate | 430-700-5 | 10138-36-0 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-618-00-5 | menadione sodium bisulfite; 2-naphthalenesulfonic acid,1,2,3,4-tetrahydro-2-methyl-1,4-dioxo-, sodium salt | 204-987-0 | 130-37-0 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H319 H400 H410 | GHS07 GHS09 Wng | H319 H315 H410 |  |  |  | ATP01
607-619-00-0 | menadione nicotinamide bisulfite; 1,2,3,4-tetrahydro-2-methyl-1,4-dioxonaphthalene-2-sulfonic acid, compound with nicotin-3-amide (1:1) | 277-543-7 | 73581-79-0 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H319 H400 H410 | GHS07 GHS09 Wng | H319 H315 H410 |  |  |  | ATP01
607-620-00-6 | trisodium nitrilotriacetate | 225-768-6 | 5064-31-3 | Carc. 2 Acute Tox. 4 * Eye Irrit. 2 | H351 H302 H319 | GHS08 GHS07 Wng | H351 H302 H319 |  | Carc. 2; H351: C ≥ 5 % |  | ATP01
607-621-00-1 | milbemectin (ISO); [reaction mass of milbemycin A3 (CAS No 51596-10-2) and milbemycin A4 (CAS No 51596-11-3) (30:70)] | - | - | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H400 H410 | GHS07 GHS09 Wng | H332 H302 H410 |  | M=100 |  | ATP01
607-622-00-7 | 2-ethylhexyl-2-ethylhexanoate | 231-057-1 | 7425-14-1 | Repr. 2 | H361d *** | GHS08 Wng | H361d *** |  |  |  | ATP01
607-623-00-2 | diisobutyl phthalate | 201-553-2 | 84-69-5 | Repr. 1B | H360Df | GHS08 Dgr | H360Df |  |  |  | ATP01/ATP09
607-624-00-8 | perfluorooctane sulfonic acid; heptadecafluorooctane-1-sulfonic acid [1] potassium perfluorooctanesulfonate; potassium heptadecafluorooctane-1-sulfonate [2] diethanolamine perfluorooctane sulfonate [3] ammonium perfluorooctane sulfonate; ammonium heptadecafluorooctanesulfonate [4] lithium perfluorooctane sulfonate; lithium heptadecafluorooctanesulfonate [5] | 217-179-8 [1] 220-527-1 [2] 274-460-8 [3] 249-415-0 [4] 249-644-6 [5] | 1763-23-1 [1] 2795-39-3 [2] 70225-14-8 [3] 29081-56-9 [4] 29457-72-5 [5] | Carc. 2 Repr. 1B Lact. Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Aquatic Chronic 2 | H351 H360D *** H362 H332 H302 H372 ** H411 | GHS08 GHS07 GHS09 Dgr | H351 H360D *** H372 ** H332 H302 H362 H411 |  |  |  | ATP01
607-625-00-3 | clodinafop-propargyl (ISO) | - | 105512-06-9 | Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H317 H410 |  | Skin Sens. 1; H317: C ≥ 0,001 % M=1 |  | ATP01
607-626-00-9 | ethyl 1-(2,4-dichlorophenyl)-5-(trichloromethyl)-1H-1,2,4-triazole-3-carboxylate | 401-290-5 | 103112-35-2 | Carc. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H400 H410 | GHS08 GHS09 Dgr | H350 H410 |  |  |  | ATP01
607-627-00-4 | [(4S,5S)-4-benzyl-2-oxo-5-oxazolidinyl]methyl 4-nitrobenzenesulfonate | 416-360-0 | 162221-28-5 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-628-00-X | 4-oxo-4-(p-tolyl)butyric acid adduct with 4-ethylmorpholine | 419-240-6 | 171054-89-0 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-629-00-5 | [[2-methyl-1-(1-oxopropoxy)propoxy](4-phenylbutyl)phosphinyl] acetic acid | 419-270-1 | 123599-82-6 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | ATP01
607-630-00-0 | acrylic acid, 3-(trimethoxysilyl)propyl ester | 419-560-6 | 4369-14-6 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 3 | H332 H314 H317 H412 | GHS05 GHS07 Dgr | H332 H314 H317 H412 |  |  |  | ATP01
607-631-00-6 | reaction mass of: 2-(2-((oxo(phenyl)acetyl)oxy)ethoxy)ethyl oxo(phenyl)acetate; (2-(2-hydroxyethoxy)ethyl) oxo(phenyl)acetate | 442-300-8 | - | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-632-00-1 | N-[3-(2,4-di-(1,1-dimethyl-propyl)phenoxy)-propyl]-1-hydroxy-5-(2-methylpropyl-oxycarbonylamino)-naphthamide | 420-210-1 | 111244-14-5 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-633-00-7 | trisodium 5-{[4-chloro-6-(1-naphthylamino)-1,3,5-triazin-2-yl]amino}-4-hydroxy-3-[(E)-(4-methoxy-2-sulfonatophenyl)diazenyl]-2,7-naphthalenedisulfonate | 440-480-2 | 341026-59-3 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
607-634-00-2 | (S)-(-)-2-acetoxypropionylchloride; (1S)-2-chloro-1-methyl-2-oxoethyl acetate | 420-610-4 | 36394-75-9 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 | H302 H314 H317 | GHS05 GHS07 Dgr | H302 H314 H317 |  |  |  | ATP01
607-635-00-8 | trisodium N-(3-propionato)-.sc.l.sc.-aspartate | 422-090-4 | 172737-80-3 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-636-00-3 | 1-bromo-2-methylpropyl propionate | 422-900-6 | 158894-67-8 | Flam. Liq. 3 Carc. 2 Skin Corr. 1B Skin Sens. 1 | H226 H351 H314 H317 | GHS02 GHS05 GHS08 GHS07 Dgr | H226 H351 H314 H317 |  |  |  | ATP01
607-637-00-9 | disodium 8-amino-5-{4-[2-(sulfonatoethoxy)sulfonyl]phenylazo}naphthalene-2-sulfonate | 423-730-5 | 250688-43-8 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-638-00-4 | 2-hydroxybenzoic acid 2-butyloctyl ester | 431-090-3 | 190085-41-7 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-639-00-X | 2-(2-oxo-5-(1,1,3,3-tetramethylbutyl)-2,3-dihydro-1-benzofuran-3-yl)-4-(1,1,3,3-tetramethylbutyl)phenyl acetate | 431-770-1 | 216698-07-6 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-641-00-0 | 2-(formylamino)-3-thiophenecarboxylic acid; 2-formamido-3-thiophenecarboxylic acid | 431-930-9 | 43028-69-9 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  |  |  | ATP01
607-642-00-6 | 3,6,9-trithiaundecamethylene-1,11-dimethacrylate | 432-210-7 | 141631-22-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
607-643-00-1 | dimethyl (2S)-2-hydroxysuccinate | 432-310-0 | 617-55-0 | Flam. Liq. 3 Eye Dam. 1 Skin Sens. 1 | H226 H318 H317 | GHS02 GHS05 GHS07 Dgr | H226 H318 H317 |  |  |  | ATP01
607-644-00-7 | methyl 2,2-dimethyl-6-methylenecyclohexanecarboxylate | 432-350-9 | 81752-87-6 | Skin Irrit. 2 | H315 | GHS07 Wng | H315 |  |  |  | ATP01
607-645-00-2 | tetrasodium 2-(4-fluoro-6-(methyl-(2-(sulfatoethylsulfonyl)ethyl)amino)-1,3,5-triazin-2-ylamino)-5-hydroxy-6-(4-methyl-2-sulfonatophenylazo)naphthalene-1,7-disulfonate | 432-550-6 | 243858-01-7 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-646-00-8 | .sc.d.sc.-erythro-hexanoic acid 2,4-dideoxy-3,5-O-(1-methylethylidene)-1,1-dimethylethylester; tert-butyl 2-[(4R,6S)-6-(hydroxymethyl)-2,2-dimethyl-1,3-dioxan-4-yl]acetate | 432-960-5 | 124655-09-0 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | ATP01
607-647-00-3 | 5-acetoxy-2-(R,S)butyryloxymethyl-1,3-oxathiolane | 433-530-1 | 143446-73-5 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 | H302 H317 H400 | GHS07 GHS09 Wng | H302 H317 H400 |  |  |  | ATP01
607-649-00-4 | [3-(chlorocarbonyl)-2-methylphenyl]acetate | 433-690-0 | 167678-46-8 | Skin Corr. 1A Skin Sens. 1 | H314 H317 | GHS05 GHS07 Dgr | H314 H317 |  |  |  | ATP01
607-650-00-X | 2-methyl-1,5-pentanediamine-1,3-benzenedicarboxylate | 433-910-5 | 145153-52-2 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-651-00-5 | sodium 2-(nonanoyloxy)benzenesulfonate | 434-360-9 | 91125-43-8 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
607-652-00-0 | ethyl N2-dodecanoyl-.sc.l.sc.-argininate hydrochloride | 434-630-6 | 60372-77-2 | Eye Dam. 1 Aquatic Acute 1 | H318 H400 | GHS05 GHS09 Dgr | H318 H400 |  |  |  | ATP01
607-653-00-6 | tetrakis(bis(2-hydroxyethyl)methylammonium) 3-(4-(7-acetylamino-1-hydroxy-3-sulfonatonaphthalen-2-ylazo)-5-methoxy-2-sulfonatophenylazo)-7-(4-amino-3-sulfonatophenylamino)-4-hydroxynaphthalene-2-sulfonate | 434-840-8 | 225786-91-4 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
607-654-00-1 | (S)-3-hydroxy-γ-butyrolactone | 434-990-4 | 7331-52-4 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-655-00-7 | ethyl 6,8-dichlorooctanoate | 435-080-1 | 1070-64-0 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
607-656-00-2 | sodium salt of 4-amino-3,6-bis[[5-[[4-chloro-6-[(2-methyl-4-sulfophenyl)amino]-1,3,5-triazin-2-yl]amino]-2-sulfophenyl]azo]-5-hydroxy-2,7-naphthalenedisulfonic acid | 435-350-7 | 141250-43-3 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
607-657-00-8 | pentasodium 7-(4-(4-(3-(2-sulfatoethanesulfonyl)phenylamino)-6-(4-(2-sulfatoethanesulfonyl)phenylamino)-1,3,5-triazin-2-ylamino)-2-ureidophenylazo)naphthalene-1,3,6-trisulfonate | 436-920-8 | 172399-10-9 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-658-00-3 | 3,10-diamino-6,13-dichloro-2-((6-(((4-(1,1-dimethylethyl)phenyl)sulfonyl)amino)-2-naphthalenyl)sulfonyl)-4,11-triphenodioxazinedisulfonic acid, lithium potassium sodium salt | 440-770-9 | 371921-63-0 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
607-659-00-9 | pentasodium N-[5-[[4-[[3-[(aminocarbonyl)amino]-4-[(3,6,8-trisulfonatonaphthalen-2-yl)azo]phenyl]amino]-6-chloro-1,3,5-triazin-2-yl]amino]-2-sulfonato-4-[[4-[[-2-(oxysulfonato)ethyl] sulfonyl]phenyl]azo]phenyl]-3-aminopropanoic acid | 442-030-0 | 321912-47-4 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-660-00-4 | 2-{4-[4-[4-fluoro-6-(2-(2-vinylsulfonylethoxy)ethylamino)-1,3,5-triazin-2-ylamino]phenylazo]phenylazo}naphthalene-4,6,8-trisulfonate, trisodium salt | 442-230-8 | 321679-52-1 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
607-661-00-X | 1,1-dimethylethyl 4'-(bromomethyl)biphenyl-2-carboxylate | 442-850-9 | 114772-40-6 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
607-662-00-5 | methyl 2-(acetylamino)-3-chloropropionate | 442-860-3 | 87333-22-0 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | ATP01
607-663-00-0 | bis(2-ethylhexyl) naphthalene-2,6-dicarboxylate | 442-980-6 | 127474-91-3 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-664-00-6 | methyl 2-chlorosulfonyl-4-(methanesulfonylaminomethyl) benzoate | 443-120-2 | 393509-79-0 | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | ATP01
607-665-00-1 | trans-methyl-2-ethyl-but-2-enoate | 443-150-6 | 101226-85-1 | Flam. Liq. 3 | H226 | GHS02 Wng | H226 |  |  |  | ATP01
607-666-00-7 | (2S)-5-(benzyloxy)-2-(1,3-dioxo-1,3-dihydro-2H-isoindol-2-yl)-5-oxopentanoic acid | 443-560-5 | 88784-33-2 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | ATP01
607-667-00-2 | chloro-1-ethylcyclohexyl carbonate | 444-950-8 | 99464-83-2 | Muta. 2 Skin Sens. 1 | H341 H317 | GHS08 GHS07 Wng | H341 H317 |  |  |  | ATP01
607-668-00-8 | trans-2-isopropyl-5-carboxy-1,3-dioxane | 445-770-2 | 42031-28-7 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
607-669-00-3 | methyl (9-acetoxy-3,8,10-triethyl-7,8,10-trimethyl-1,5-dioxa-9-aza-spiro[5.5]undec-3-yl)octadecanoate | 445-990-9 | 376588-17-9 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
607-670-00-9 | dibutyl-3-(4-(5-ammonio-2-butyl)benzofuran-3-yl)carbonyl)phenoxy)propyl ammonium oxalate; (5-amino-2-butylbenzofuran-3-yl) [4-(3-dibutylaminopropoxy)phenyl]methanone, dioxalate | 448-700-9 | 500791-70-8 | STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H318 H317 H400 H410 | GHS05 GHS08 GHS07 GHS09 Dgr | H373 ** H318 H317 H410 |  | M=10 |  | ATP01
607-671-00-4 | diethyl 1,4-cyclohexanedicarboxylate | 417-310-0 | 72903-27-6 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
607-672-00-X | reaction mass of: 2-hydroxy-3-(methacryloyloxy)propyl (2-benzoyl)benzoate; 1-hydroxymethyl-2-(methacryloyloxy)ethyl (2-benzoyl)benzoate; x-hydroxy-y-(methacryloyloxy)propyl(or -ethyl) (2-benzoyl)benzoate | 419-000-0 | - | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
607-673-00-5 | 1-ethyl-5,6,7,8-tetrahydroquinolinium tosylate | 419-570-0 | - | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | ATP01
607-675-00-6 | reaction mass of: cis-9-octadecenedioic acid; cis-9-cis-12-octadecadienedioic acid; hexadecanedioic acid; octadecanedioic acid | 422-260-8 | - | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | ATP01
607-676-00-1 | reaction mass of: 2-methylnonanedioic acid; 2,4-dimethyl-4-methoxycarbonylundecanedioic acid; 2,4,6-trimethyl-4,6-dimethoxycarbonyltridecanedioic acid; 8,9-dimethyl-8,9-dimethoxycarbonylhexadecanedioic acid | 423-670-1 | - | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
607-677-00-7 | 2,5-dioxopyrrolidin-1-yl N-{[methyl[[2-(1-methylethyl)-4-thiazolyl]methyl]amino]carbonyl}-.sc.l.sc.-valinate | 424-660-8 | - | STOT RE 2 * Eye Dam. 1 Skin Sens. 1 | H373 ** H318 H317 | GHS05 GHS08 GHS07 Dgr | H373 ** H318 H317 |  |  |  | ATP01
607-678-00-2 | reaction mass of: ethyl (2R,3R)-3-isopropylbicyclo[2.2.1]hept-5-ene-2-carboxylate; ethyl (2S,3S)-3-isopropylbicyclo[2.2.1]hept-5-ene-2-carboxylate | 427-090-8 | - | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
607-679-00-8 | reaction mass of: 3-{5-[3-(4-{1,6-dihydro-2-hydroxy-4-methyl-1-[3-(methylammonio)propyl]-6-oxo-3-pyridylazo}benzamido)phenylazo]-1,2-dihydro-6-hydroxy-4-methyl-2-oxo-1-pyridyl}propyl(methyl)ammonium di(acetate); 3-{5-[4-(3-{1,6-dihydro-2-hydroxy-4-methyl-1-[3-(methylammonio)propyl]-6-oxo-3-pyridylazo}benzamido]phenylazo]-1,2-dihydro-6-hydroxy-4-methyl-2-oxo-1-pyridyl}propyl(dimethyl)ammonium di(acetate); 3-{5-[3-(4-{1-[3-(dimethylammonio)propyl]-1,6-dihydro-2-hydroxy-4-methyl-6-oxo-3-pyridylazo}benzamido)phenylazo]-1,2-dihydro-6-hydroxy-4-methyl-2-oxo-1-pyridyl}propyl(dimethyl)ammonium di(acetate) | 431-440-5 | - | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | ATP01
607-680-00-3 | tert-butyl(6-{2-[4-(4-fluorophenyl)-6-isopropyl-2-[methyl(methylsulfonyl)amino]pyrimidin-5-yl]vinyl}(4S,6S)-2,2-dimethyl[1,3]dioxan-4-yl)acetate | 432-810-9 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-681-00-9 | reaction mass of: 9-nonyl-10-octyl-19-carbonyloxyhexadecylnonadecanoic acid; 9-nonyl-10-octyl-19-carbonyloxyoctadecylnonadecanoic acid; dihexadecyl 9-nonyl-10-octylnonadecandioate; 1-octadecyl,19-hexadecyl 9-nonyl-10-octylnonadecandioate; dioctadecyl 9-nonyl-10-octylnonadecandioate | 432-910-2 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-682-00-4 | complex reaction mass of Chinese gum rosin post reacted with acrylic acid | 434-230-1 | 144413-22-9 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-683-00-X | reaction mass of: methyl 3-((1E)-2-methylprop-1-enyl)-2,2-dimethylcyclopropanecarboxylate; methyl 3-((1Z)-2-methylprop-1-enyl)-2,2-dimethylcyclopropanecarboxylate (20:80) | 435-450-0 | - | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
607-684-00-5 | alkenes, C12-14, hydroformylation products, distn. residues, C-(hydrogen sulfobutanedioates), disodium salts | 435-660-2 | 243662-67-1 | Skin Irrit. 2 Skin Sens. 1 | H315 H317 | GHS07 Wng | H315 H317 |  |  |  | ATP01
607-685-00-0 | ammonium 2-cocoyloxyethanesulfonate | 441-050-7 | - | Skin Irrit. 2 Eye Dam. 1 | H315 H318 | GHS05 Dgr | H315 H318 |  |  |  | ATP01
607-686-00-6 | 6,6'-bis(diazo-5,5',6,6'-tetrahydro-5,5'-dioxo)[methylene-bis(5-(6-diazo-5,6-dihydro-5-oxo-1-naphthylsulphonyloxy)-6-methyl-2-phenylene]di(naphthalene-1-sulfonate) | 441-550-5 | - | Self-react. C **** Carc. 2 | H242 H351 | GHS02 GHS08 Dgr | H242 H351 |  |  |  | ATP01
607-687-00-1 | reaction mass of: 2-{3,6-bis-[(2-ethylphenyl)-methylamino]-xanthylium-9-yl}-benzenesulfonate (2-10%); 2-{3,6-bis-[(2,3-dimethylphenyl)-methylamino]-xanthylium-9-yl}-benzenesulfonate (2-10%); 2-{3,6-bis-[(2,4-dimethylphenyl)-methylamino]-xanthylium-9-yl}-benzenesulfonate (2-10%); 2-{3,6-bis-[(2,5-dimethylphenyl)-methylamino]-xanthylium-9-yl}-benzenesulfonate (2-10%); 2-{3-[(2,3-dimethylphenyl)-methylamino]-6-[(2-ethylphenyl)-methylamino]-xanthylium-9-yl}-benzenesulfonate (7-20%); 2-{3-[(2,4-dimethylphenyl)-methylamino]-6-[(2-ethylphenyl)-methylamino]-xanthylium-9-yl}-benzenesulfonate (7-20%); 2-{3-[(2,5-dimethylphenyl)-methylamino]-6-[(2-ethylphenyl)-methylamino]-xanthylium-9-yl}-benzenesulfonate (7-20%); 2-{3-[(2,3-dimethylphenyl)-methylamino]-6-[(2,4-dimethylphenyl)-methylamino]-xanthylium-9-yl}-benzenesulfonate (7-20%); 2-{3-[(2,3-dimethylphenyl)-methylamino]-6-[(2,5-dimethylphenyl)-methylamino]-xanthylium-9-yl}-benzenesulfonate (7-20%); 2-{3-[(2,4-dimethylphenyl)-methylamino]-6-[(2,5-dimethylphenyl)-methylamino]-xanthylium-9-yl}-benzenesulfonate (7-20%) | 442-800-6 | - | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | ATP01
607-688-00-7 | (R)-1-cyclohexa-1,4-dienyl-1-methoxycarbonyl-methylammoniumchloride | 444-320-2 | - | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | ATP01
607-689-00-2 | reaction mass of: methyl 1,4-dimethylcyclohexanecarboxylate ("para-isomer" including cis- and trans- isomers); methyl 1,3-dimethylcyclohexanecarboxylate ("meta-isomer" including cis- and trans-isomers) | 444-920-4 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
607-690-00-8 | dimethyl[2S,2S']-6,6,6'6'-tetramethoxy-2,2'-[N,N'-bis(trifluoracetyl)-S,S'-bi(L-homocysteinyl) diimino]dihexanoate | 432-860-1 | 255387-46-3 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
607-691-00-3 | magnesium salts, fatty acids, C16-18 and C18 unsaturated, branched and linear | 448-690-6 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-692-00-9 | zinc salts, fatty acids, C16-18 and C18 unsaturated, branched and linear | 446-470-4 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
607-694-00-X | ethyl 5,5-diphenyl-2-isoxazoline-3-carboxylate | 443-870-0 | 163520-33-0 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | ATP01
607-696-00-0 | pentyl formate | 211-340-6 | 638-49-3 | Flam. Liq. 3 STOT SE 3 Eye Irrit. 2 | H226 H335 H319 | GHS02 GHS07 Dgr | H226 H319 H335 |  |  | C | CLP00
607-697-00-6 | tert-butyl propionate |  | 20487-40-5 | Flam. Liq. 2 | H225 | GHS02 Dgr | H225 |  |  | C | CLP00
607-698-00-1 | 4-tert-butylbenzoic acid | 202-696-3 | 98-73-7 | Repr. 1B Acute Tox. 4 STOT RE 1 | H360F H302 H372 | GHS07 GHS08 Dgr | H360F H372 H302 |  |  |  | ATP03
607-699-00-7 | bifenthrin (ISO); (2-methylbiphenyl-3-yl)methyl rel-(1R,3R)-3-[(1Z)-2-chloro-3,3,3-trifluoroprop-1-en-1-yl]-2,2-dimethylcyclopropanecarboxylate |  | 82657-04-3 | Carc. 2 Acute Tox. 2 Acute Tox. 3 STOT RE 1 Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 1 | H351 H300 H331 H372 (nervous system) H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H300 H331 H317 H351 H372 (nervous system) H410 |  | M=10000 M=100000 |  | ATP05
607-700-00-0 | indoxacarb (ISO); methyl (4aS)-7-chloro-2-{(methoxycarbonyl)[4-(trifluoromethoxy)phenyl]carbamoyl}-2,5-dihydroindeno[1,2-e][1,3,4]oxadiazine-4a(3H)-carboxylate [1] reaction mass of (S)- Indoxacarb and (R)- Indoxacarb 75:25; methyl 7-chloro-2-{(methoxycarbonyl)[4-(trifluoromethoxy)phenyl]carbamoyl}-2,5-dihydroindeno[1,2-e][1,3,4]oxadiazine-4a(3H)-carboxylate [2] |  | 173584-44-6 [1] 144171-61-9 [2] | Acute Tox. 3 Acute Tox. 4 STOT RE 1 Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 1 | H301 H332 H372 (blood, nervous system, heart) H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H301 H332 H317 H372 (blood, nervous system, heart) H410 |  | M=1 M=1 |  | ATP05
607-702-00-1 | dihexyl phthalate | 201-559-5 | 84-75-3 | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  |  | ATP05
607-703-00-7 | ammoniumpentadeca- fluorooctanoate | 223-320-4 | 3825-26-1 | Carc. 2 Repr. 1B Lact. Acute Tox. 4 Acute Tox. 4 STOT RE 1 Eye Dam. 1 | H351 H360D H362 H332 H302 H372 (liver) H318 | GHS08 GHS07 GHS05 Dgr | H302 H332 H318 H351 H360D H362 H372 (liver) |  |  |  | ATP05
607-704-00-2 | perfluorooctanoic acid | 206-397-9 | 335-67-1 | Carc. 2 Repr. 1B Lact. Acute Tox. 4 Acute Tox. 4 STOT RE 1 Eye Dam. 1 | H351 H360D H362 H332 H302 H372 (liver) H318 | GHS08 GHS07 GHS05 Dgr | H302 H332 H318 H351 H360D H362 H372 (liver) |  |  |  | ATP05
607-705-00-8 | benzoic acid | 200-618-2 | 65-85-0 | STOT RE 1 Skin Irrit. 2 Eye Dam. 1 | H372 (lungs) (Inhalation) H315 H318 | GHS08 GHS05 Dgr | H315 H318 H372 (lungs) (Inhalation) |  |  |  | ATP06
607-706-00-3 | methyl 2,5-dichlorobenzoate | 220-815-7 | 2905-69-3 | Acute Tox. 4 STOT SE 3 Aquatic Chronic 2 | H302 H336 H411 | GHS07 GHS09 Wng | H302 H336 H411 |  |  |  | ATP06
607-707-00-9 | fenoxaprop-P-ethyl (ISO); ethyl (2R)-2-{4-[(6-chloro-1,3-benzoxazol-2-yl)oxy]phenoxy}propanoate |  | 71283-80-2 | STOT RE 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 (kidneys) H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H317 H373 (kidneys) H410 |  | M=1 M=1 |  | ATP07
607-708-00-4 | octanoic acid | 204-677-5 | 124-07-2 | Skin Corr. 1C Aquatic Chronic 3 | H314 H412 | GHS05 Dgr | H314 H412 |  |  |  | ATP07
607-709-00-X | decanoic acid | 206-376-4 | 334-48-5 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 3 | H315 H319 H412 | GHS07 Wng | H315 H319 H412 |  |  |  | ATP07
607-710-00-5 | 1,2-benzenedicarboxylic acid, dihexyl ester, branched and linear | 271-093-5 | 68515-50-4 | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  |  | ATP07
607-711-00-0 | spirotetramat (ISO); (5s,8s)-3-(2,5-dimethylphenyl)-8-methoxy-2-oxo-1-azaspiro[4,5]dec-3-en-4-yl ethyl carbonate |  | 203313-25-1 | Repr. 2 STOT SE 3 Eye Irrit. 2 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 1 | H361fd H335 H319 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H319 H317 H361fd H335 H410 |  | M=1 M=1 |  | ATP07
607-712-00-6 | dodemorph acetate; 4-cyclododecyl-2,6-dimethylmorpholin-4-ium acetate | 250-778-2 | 31717-87-0 | Repr. 2 STOT RE 2 Skin Corr. 1C Skin Sens. 1A Aquatic Chronic 1 | H361d H373 (liver) H314 H317 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H314 H317 H361d H373 (liver) H410 | EUH071 | M=1 |  | ATP07
607-713-00-1 | fenpyroximate (ISO); tert-butyl 4-[({(E)-[(1,3-dimethyl-5-phenoxy-1H-pyrazol-4-yl)methylene]amino}oxy)methyl]benzoate |  | 134098-61-6 | Acute Tox. 2 Acute Tox. 3 Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 1 | H330 H301 H317 H400 H410 | GHS06 GHS09 Dgr | H301 H330 H317 H410 |  | M=100 M=1000 |  | ATP07
607-714-00-7 | triflusulfuron-methyl; methyl 2-({[4-(dimethylamino)-6-(2,2,2- trifluoroethoxy)-1,3,5-triazin-2- yl]carbamoyl}sulfamoyl)-3- methylbenzoate |  | 126535-15-7 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  | M=100 M=10 |  | CLP00/ATP07
607-715-00-2 | bifenazate (ISO); isopropyl 2-(4-methoxybiphenyl-3-yl)hydrazinecarboxylate | 442-820-5 | 149877-41-8 | STOT RE 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H317 H373 H410 |  | M=1 M=1 |  | ATP07
607-716-00-8 | bromadiolone (ISO); 3-[3-(4′-bromobiphenyl-4-yl)-3-hydroxy-1-phenylpropyl]-4-hydroxy-2H-chromen-2-one | 249-205-9 | 28772-56-7 | Repr. 1B Acute Tox. 1 Acute Tox. 1 Acute Tox. 1 STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D H330 H310 H300 H372 (blood) H400 H410 | GHS08 GHS06 GHS09 Dgr | H330 H310 H300 H360D H372 (blood) H410 |  | Repr. 1B; H360D: C ≥ 0,003 % STOT RE 1; H372: C ≥ 0,005 % STOT RE 2; H373: 0,0005 % ≤ C < 0,005 % M=1 M=1 |  | ATP09
607-717-00-3 | difethialone (ISO); 3-[3-(4′-bromobiphenyl-4-yl)-1,2,3,4-tetrahydronaphthalen-1-yl]-4-hydroxy-2H-1-benzothiopyran-2-one |  | 104653-34-1 | Repr. 1B Acute Tox. 1 Acute Tox. 1 Acute Tox. 1 STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D H330 H310 H300 H372 (blood) H400 H410 | GHS08 GHS06 GHS09 Dgr | H330 H310 H300 H360D H372 (blood) H410 | EUH070 | Repr. 1B; H360D: C ≥ 0,003 % STOT RE 1; H372: C ≥ 0,02 % STOT RE 2; H373: 0,002 % ≤ C < 0,02 % M=100 M=100 |  | ATP09
607-718-00-9 | perfluorononan-1-oic acid [1] perfluorononan-1-oic acid sodium salt [2] perfluorononan-1-oic acid ammonium salt [3] | 206-801-3 [1] | 375-95-1 [1] 21049-39-8 [2] 4149-60-4 [3] | Carc. 2 Repr. 1B Lact. Acute Tox. 4 Acute Tox. 4 STOT RE 1 Eye Dam. 1 | H351 H360Df H362 H332 H302 H372 (liver, thymus, spleen) H318 | GHS08 GHS07 GHS05 Dgr | H332 H302 H318 H351 H360Df H362 H372 (liver, thymus, spleen) |  |  |  | ATP09
607-719-00-4 | dicyclohexyl phthalate | 201-545-9 | 84-61-7 | Repr. 1B Skin Sens. 1 | H360D H317 | GHS08 GHS07 Dgr | H317 H360D |  |  |  | ATP09
607-720-00-X | nonadecafluorodecanoic acid [1] ammonium nonadecafluorodecanoate [2] sodium nonadecafluorodecanoate [3] | 206-400-3 [1] 221-470-5 [2] | 335-76-2 [1] 3108-42-7 [2] 3830-45-3 [3] | Carc. 2 Repr. 1B Lact. | H351 H360Df H362 | GHS08 Dgr | H351 H360Df H362 |  |  |  | ATP10
607-721-00-5 | N,N′-methylenedimorpholine; N,N′-methylenebismorpholine; [formaldehyde released from N,N′-methylenebismorpholine]; [MBM] | 227-062-3 | 5625-90-1 | Carc. 1B Muta. 2 Acute Tox. 4 Acute Tox. 4 Acute Tox. 4 STOT RE 2 Skin Corr. 1B Eye Dam. 1 Skin Sens. 1 | H350 H341 H332 H312 H302 H373 (gastrointestinal tract, respiratory tract) H314 H318 H317 | GHS08 GHS07 GHS05 Dgr | H332 H312 H302 H314 H317 H341 H350 H373 (gastrointestinal tract, respiratory tract) | EUH071 |  | 8 9 | ATP10
607-722-00-0 | 2,3,5,6-tetrafluoro-4-(methoxymethyl)benzyl (Z)-(1R,3R)-3-(2-cyanoprop-1-enyl)-2,2-dimethylcyclopropanecarboxylate; epsilon-momfluorothrin |  | 1065124-65-3 | Acute Tox. 4 STOT SE 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H371 (nervous system) H400 H410 | GHS07 GHS08 GHS09 Wng | H302 H371 (nervous system) H410 |  | M=100 M=100 |  | ATP10
607-723-00-6 | tefluthrin (ISO); 2,3,5,6-tetrafluoro-4-methylbenzyl (1RS,3RS)-3-[(Z)-2-chloro-3,3,3-trifluoroprop-1-enyl]-2,2-dimethylcyclopropanecarboxylate |  | 79538-32-2 | Acute Tox. 1 Acute Tox. 2 Acute Tox. 2 Aquatic Acute 1 Aquatic Chronic 1 | H330 H310 H300 H400 H410 | GHS06 GHS09 Dgr | H330 H310 H300 H410 |  | M=10000 M=10000 |  | ATP10
607-724-00-1 | 2,3,5,6-tetrafluoro-4-(methoxymethyl)benzyl (1R,3R)-2,2-dimethyl-3-[(1Z)-prop-1-en-1-yl] cyclopropanecarboxylate; epsilon-metofluthrin |  | 240494-71-7 | Acute Tox. 3 Acute Tox. 4 STOT SE 1 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H301 H332 H370 (nervous system) H373 H400 H410 | GHS06 GHS08 GHS09 Dgr | H332 H301 H370 (nervous system) H373 H410 |  | M=100 M=100 |  | ATP13
607-725-00-7 | isopropyl (2E,4E,7S)-11-methoxy-3,7,11-trimethyldodeca-2,4-dienoate; S-methoprene |  | 65733-16-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1 M=1 |  | ATP13
607-726-00-2 | pinoxaden (ISO); 8-(2,6-diethyl-4-methylphenyl)-7-oxo-1,2,4,5-tetrahydro-7H-pyrazolo[1,2-d][1,4,5]oxadia­zepin-9-yl 2,2-dimethylpropanoate |  | 243973-20-8 | Repr. 2 Acute Tox. 4 Acute Tox. 4 STOT SE 3 Eye Irrit. 2 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 3 | H361d H332 H302 H335 H319 H317 H400 H412 | GHS08 GHS07 GHS09 Wng | H332 H302 H319 H317 H361d H335 H410 |  | Inhalation: ATE = 4.63 mg/L (dusts/mists) Oral: ATE = 500 mg/kg bw M=1 |  | ATP13
607-727-00-8 | tetramethrin (ISO); (1,3-dioxo-1,3,4,5,6,7-hexahydro-2H-isoindol-2-yl)methyl 2,2-dimethyl-3-(2-methylprop-1-en-1-yl)cyclopropanecarboxylate | 231-711-6 | 7696-12-0 | Carc. 2 Acute Tox. 4 STOT SE 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H371 (nervous system) (Inhalation) H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H351 H371 (nervous system) (Inhalation) H410 |  | M=100 M=100 |  | ATP13
607-728-00-3 | (1,3,4,5,6,7-hexahydro-1,3-di­oxo-2H-isoindol-2-yl)methyl (1R-trans)-2,2-dimethyl-3-(2-methylprop-1-enyl)cyclopropanecarboxylate | 214-619-0 | 1166-46-7 | Carc. 2 Acute Tox. 4 STOT SE 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H371 (nervous system) (Inhalation) H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H351 H371 (nervous system) (Inhalation) H410 |  | M=100 M=100 |  | ATP13
607-729-00-9 | mesosulfuron-methyl (ISO); methyl 2-[(4,6-dimethoxypyrimidin-2-ylcarbamoyl)sulfamoyl]-α-(methanesulfonamido)-p-toluate; |  | 208465-21-8 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=100 M=100 |  | ATP13
607-730-00-4 | spirodiclofen (ISO); 3-(2,4-dichlorophenyl)-2-oxo-1-oxaspiro[4.5]dec-3-en-4-yl 2,2-dimethylbutyrate |  | 148477-71-8 | Carc. 1B Repr. 2 STOT RE 2 Skin Sens. 1B Aquatic Chronic 1 | H350 H361f H373 H317 H410 | GHS08 GHS07 GHS09 Dgr | H317 H350 H361f H373 H410 |  | M=10 |  | ATP13
607-731-00-X | sodium methyl [(4-aminophenyl)sulphonyl]carbamate; sodium methyl (EZ)-sulfanilylcarbonimidate; asulam-sodium | 218-953-8 | 2302-17-2 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  | M=1 M=1 |  | ATP13
607-732-00-5 | salicylic acid | 200-712-3 | 69-72-7 | Repr. 2 Acute Tox. 4 Eye Dam. 1 | H361d H302 H318 | GHS08 GHS07 GHS05 Dgr | H302 H318 H361d |  |  |  | ATP13
607-733-00-0 | cyflumetofen (ISO); 2-methoxyethyl (RS)-2-(4- tert-butylphenyl)-2-cyano-3-oxo-3-(α,α,α-trifluoro-o-tolyl)propionate |  | 400882-07-7 | Carc. 2 Skin Sens. 1A | H351 H317 | GHS07 GHS08 Wng | H317 H351 |  |  |  | ATP14
607-734-00-6 | pentapotassium 2,2’,2’’,2’’’,2’’’’-(ethane-1,2-diylnitrilo)pentaacetate | 404-290-3 | 7216-95-7 | Repr. 1B Acute Tox. 4 STOT RE 2 Eye Irrit. 2 | H360D H332 H373 (Inhalation) H319 | GHS07 GHS08 Dgr | H360D H332 H319 H373 (Inhalation) |  | Repr. 1B; H360D: C ≥ 3 % Inhalation: ATE = 1.5 mg/L (dusts/mists) |  | ATP14/ATP18
607-735-00-1 | N-carboxymethyliminobis(ethylenenitrilo)tetra(acetic acid) | 200-652-8 | 67-43-6 | Repr. 1B Acute Tox. 4 STOT RE 2 Eye Irrit. 2 | H360D H332 H373 (Inhalation) H319 | GHS07 GHS08 Dgr | H360D H332 H319 H373 (Inhalation) |  | Repr. 1B; H360D: C ≥ 3 % Inhalation: ATE = 1.5 mg/L (dusts/mists) |  | ATP14/ATP18
607-736-00-7 | pentasodium (carboxylatomethyl)iminobis(ethylenenitrilo)tetraacetate | 205-391-3 | 140-01-2 | Repr. 1B Acute Tox. 4 STOT RE 2 | H360D H332 H373 (Inhalation) | GHS07 GHS08 Dgr | H360D H332 H373 (Inhalation) |  | Repr. 1B; H360D: C ≥ 3 % Inhalation: ATE = 1.5 mg/L (dusts/mists) |  | ATP14/ATP18
607-737-00-2 | diisohexyl phthalate | 276-090-2 | 71850-09-4 | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  |  | ATP14
607-738-00-8 | MCPA-thioethyl (ISO); S-ethyl (4-chloro-2- methylphenoxy)etha­nethioate; S-ethyl 4- chloro-o-tolylox­ythioacetate | 246-831-4 | 25319-90-8 | Acute Tox. 4 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 (liver) H400 H410 | GHS07 GHS08 GHS09 Wng | H302 H373 (liver) H410 |  | Oral: ATE = 450 mg/kg bw M=10 M=10 |  | ATP15
607-740-00-9 | diisooctyl phthalate | 248-523-5 | 27554-26-3 | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  |  | ATP15
607-741-00-4 | 4-{[(6-chloropyridin-3-yl)methyl](2,2-di­fluoroethyl)amino}furan-2(5H)-one; flupyra­difurone |  | 951659-40-8 | Acute Tox. 4 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 (muscle) H400 H410 | GHS07 GHS08 GHS09 Wng | H302 H373 (muscle) H410 |  | Oral: ATE = 500 mg/kg bw M=10 M=10 |  | ATP15
607-742-00-X | thiencarbazone-methyl (ISO); methyl 4-[(4,5-dihydro-3-methoxy-4- methyl-5-oxo-1H-1,2,4-triazol-1-yl)carbonylsulfamoyl]-5- methylthiophene-3- carboxylate |  | 317815-83-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1000 M=1000 |  | ATP15
607-743-00-5 | L-(+)-lactic acid; (2S)-2-hydroxypropanoic acid | 201-196-2 | 79-33-4 | Skin Corr. 1C Eye Dam. 1 | H314 H318 | GHS05 Dgr | H314 | EUH071 |  |  | ATP15
607-744-00-0 | 2-methoxyethyl acrylate | 221-499-3 | 3121-61-7 | Flam. Liq. 3 Muta. 2 Repr. 1B Acute Tox. 3 Acute Tox. 4 Skin Corr. 1C Eye Dam. 1 Skin Sens. 1 | H226 H341 H360FD H331 H302 H314 H318 H317 | GHS02 GHS05 GHS06 GHS08 Dgr | H226 H331 H302 H314 H317 H341 H360FD | EUH071 | Inhalation: ATE = 2.7 mg/L (Vapours) Oral: ATE = 404 mg/kg bw |  | ATP15
607-745-00-6 | glyoxylic acid ...% | 206-058-5 | 298-12-4 | Eye Dam. 1 Skin Sens. 1B | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  | B | ATP15
607-746-00-1 | sodium N-(hydroxymethyl)glycinate; [formaldehyde released from sodium N-(hydroxymethyl)glycinate] | 274-357-8 | 70161-44-3 | Carc. 1B Muta. 2 Acute Tox. 4 Acute Tox. 4 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H350 H341 H332 H302 H335 H315 H319 H317 | GHS08 GHS07 Dgr | H332 H302 H315 H319 H317 H341 H350 H335 |  | Inhalation: ATE = 3 mg/L (dusts/mists) Oral: ATE = 1100 mg/kg bw | 8 9 | ATP15
607-747-00-7 | 2,2-dibromo-2-cyanoacetamide; [DBNPA] | 233-539-7 | 10222-01-2 | Acute Tox. 2 Acute Tox. 3 STOT RE 1 Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H301 H372 (respiratory tract, inhalation) H315 H318 H317 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H330 H301 H315 H318 H317 H372 (respiratory tract) (inhalation) H410 |  | Inhalation: ATE = 0.24 mg/L (dusts/mists) Oral: ATE = 118 mg/kg bw M=1 M=1 |  | ATP17
607-748-00-2 | [S-(Z,E)]-5-(1-hydroxy-2,6,6-trimethyl-4-oxocyclohex-2-en-1-yl)-3-methylpenta-2,4-dienoic acid; S-abscisic acid | 244-319-5 | 21293-29-8 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1 M=1 |  | ATP17
607-749-00-8 | methyl salicylate | 204-317-7 | 119-36-8 | Repr. 2 Acute Tox. 4 Skin Sens. 1B Aquatic Chronic 3 | H361d H302 H317 H412 | GHS07 GHS08 Wng | H302 H317 H361d H412 |  | Oral: ATE = 890 mg/kg bw |  | ATP17
607-750-00-3 | citric acid | 201-069-1 | 77-92-9 | STOT SE 3 Eye Irrit. 2 | H335 H319 | GHS07 Wng | H319 H335 |  |  |  | ATP17
607-751-00-9 | ethametsulfuron-methyl (ISO); methyl 2-({[4-ethoxy-6-(methylamino)-1,3,5-triazin-2-yl]carbamoyl}sulfamoyl)benzoate |  | 97780-06-8 | Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H319 H400 H410 | GHS07 GHS09 Wng | H319 H410 |  | M=1000 M=100 |  | ATP17
607-752-00-4 | trinexapac-ethyl (ISO); ethyl 4-[cyclopropyl(hydroxy)methylene]-3,5-dioxocyclohexanecarboxylate |  | 95266-40-3 | STOT RE 2 Skin Sens. 1B Aquatic Chronic 1 | H373 (gastrointestinal tract) H317 H410 | GHS08 GHS07 GHS09 Wng | H317 H373 (gastrointestinal tract) H410 |  | M=1 |  | ATP17
607-753-00-X | (3aS,5S,6R,7aR,7bS,9aS,10R,12aS,12bS)-10-[(2S,3R,4R,5R)-3,4-dihydroxy-5,6-dimethylheptan-2-yl]-5,6-dihydroxy-7a,9a-dimethylhexadecahydro-3H-benzo[c]indeno[5,4-e]oxepin-3-one; 24-epibrassinolide |  | 78821-43-9 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | ATP17
607-754-00-5 | benzyl salicylate | 204-262-9 | 118-58-1 | Skin Sens. 1B | H317 | GHS07 Wng | H317 |  |  |  | ATP17
607-755-00-0 | (RS)-1-{1-ethyl-4-[4-mesyl-3-(2-methoxyethoxy)-o-toluoyl]pyrazol-5-yloxy}ethyl methyl carbonate; tolpyralate |  | 1101132-67-5 | Carc. 2 Repr. 2 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H361fd H373 (eye) H400 H410 | GHS08 GHS09 Wng | H351 H361fd H373 (eye) H410 |  | M=10 M=100 |  | ATP17
607-756-00-6 | exo-1,7,7-trimethylbicyclo[2.2.1]hept-2-yl acrylate; isobornyl acrylate | 227-561-6 | 5888-33-5 | Skin Sens. 1A | H317 | GHS07 Wng | H317 |  |  |  | ATP18
607-757-00-1 | daminozide (ISO); 4-(2,2-dimethylhydrazino)-4-oxobutanoic acid; N-dimethylaminosuccinamic acid | 216-485-9 | 1596-84-5 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | ATP18
607-758-00-7 | 4,4'-oxydi(benzenesulphonohydrazide) | 201-286-1 | 80-51-3 | Self-react. D Aquatic Acute 1 Aquatic Chronic 1 | H242 H400 H410 | GHS02 GHS09 Dgr | H242 H410 |  | M=1 M=1 |  | ATP18
607-759-00-2 | toluene-4-sulphonohydrazide | 216-407-3 | 1576-35-8 | Self-react. D | H242 | GHS02 Dgr | H242 |  |  |  | ATP18
607-760-00-8 | 2-[N-ethyl-4-[(5-nitrothiazol-2-yl)azo]-m-toluidino]ethyl acetate; C.I. Disperse Blue 124 | 239-203-6 | 15141-18-1 | Skin Sens. 1A | H317 | GHS07 Wng | H317 |  | Skin Sens. 1A; H317: C ≥ 0,001% |  | ATP18
607-761-00-3 | Perfluoroheptanoic acid; tridecafluoroheptanoic acid | 206-798-9 | 375-85-9 | Repr. 1B STOT RE 1 | H360D H372 (liver) | GHS08 Dgr | H360D H372 (liver) |  |  |  | ATP18
607-762-00-9 | methyl N-(isopropoxycarbonyl)-L-valyl-(3RS)-3-(4-chlorophenyl)-β-alaninate; valifenalate | - | 283159-90-0 | Carc. 2 Aquatic Chronic 2 | H351 H411 | GHS08 GHS09 Wng | H351 H411 |  |  |  | ATP18
607-763-00-4 | 6-[C12-18-alkyl-(branched, unsaturated)-2,5-dioxopyrrolidin-1-yl]hexanoic acid, sodium and tris(2-hydroxyethyl)ammonium salts | - | - | Repr. 1B Eye Irrit. 2 | H360FD H319 | GHS08 GHS07 Dgr | H360FD H319 |  |  |  | ATP18
607-764-00-X | 6-[(C10-C13)-alkyl-(branched, unsaturated)-2,5-dioxopyrrolidin-1-yl]hexanoic acid | - | 2156592-54-8 | Repr. 1B Eye Irrit. 2 | H360FD H319 | GHS08 GHS07 Dgr | H360FD H319 |  |  |  | ATP18
607-765-00-5 | 6-[C12-18-alkyl-(branched, unsaturated)-2,5-dioxopyrrolidin-1-yl]hexanoic acid | - | - | Repr. 1B | H360FD | GHS08 Dgr | H360FD |  |  |  | ATP18
607-766-00-0 | tetramethylene dimethacrylate | 218-218-1 | 2082-81-7 | Skin Sens. 1B | H317 | GHS07 Wng | H317 |  |  |  | ATP21
607-767-00-6 | 7,7,9(or 7,9,9)-trimethyl–4,13-dioxo–3,14-dioxa–5,12-diazahexadecane–1,16-diyl bismethacrylate | 276-957-5 | 72869-86-4 | Skin Sens. 1B | H317 | GHS07 Wng | H317 |  |  |  | ATP21
607-768-00-1 | 2,2'-ethylenedioxydiethyl dimethacrylate | 203-652-6 | 109-16-0 | Skin Sens. 1B | H317 | GHS07 Wng | H317 |  |  |  | ATP21
607-769-00-7 | bifenox (ISO); methyl 5-(2,4-dichlorophenoxy)-2-nitrobenzoate | 255-894-7 | 42576-02-3 | Acute Tox. 4 Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  | oral: ATE = 1500 mg/kg bw M = 1000 M = 1000 |  | ATP21
607-770-00-2 | 2,3-epoxypropyl neodecanoate | 247-979-2 | 26761-45-5 | Muta. 2 Skin Sens. 1A | H341 H317 | GHS08 GHS07 Wng | H341 H317 |  | Skin Sens. 1A; H317: C ≥ 0,001 % |  | ATP22
607-771-00-8 | benthiavalicarb-isopropyl (ISO); isopropyl [(S)-1-{[(R)-1-(6-fluoro-1,3-benzothiazol-2-yl)ethyl]carbamoyl}-2-methylpropyl]carbamate | - | 177406-68-7 | Carc. 1B Repr. 2 Skin Sens. 1 Aquatic Chronic 2 | H350 H361fd H317 H411 | GHS08 GHS07 GHS09 Dgr | H350 H361fd H317 H411 |  |  |  | ATP22
607-772-00-3 | hexyl salicylate | 228-408-6 | 6259-76-3 | Repr. 2 Skin Sens. 1 | H361d H317 | GHS08 GHS07 Wng | H361d H317 |  |  |  | ATP22
607-773-00-9 | 7-oxabicyclo[4.1.0]hept-3-ylmethyl 7-oxabicyclo[4.1.0]heptane-3-carboxylate | 219-207-4 | 2386-87-0 | Muta. 2 STOT RE 2 Skin Sens. 1 | H341 H373 (nasal cavity) H317 | GHS08 GHS07 Wng | H341 H373 (nasal cavity) H317 |  |  |  | ATP22
607-774-00-4 | tetrasodium 4-amino-5-hydroxy-3,6-bis[[4-[[2-(sulphonatooxy)ethyl]sulphonyl]phenyl]azo]naphthalene-2,7-disulphonate [1] Reaction products of 4-amino-5-hydroxynaphthalene-2,7-disulfonic acid, coupled twice with diazotized 2-[(4-aminophenyl)sulfonyl]ethyl hydrogen sulfate, sodium salts [2] disodium 4-amino-5-hydroxy-3,6-bis{[4-(vinylsulfonyl) phenyl]diazenyl}naphthalene-2,7-disulfonate [3] | 241-164-5 [1] - [2] - [3] | 17095-24-8 [1] - [2] 100556-82-9 [3] | Resp. Sens. 1A Skin Sens. 1 | H334 H317 | GHS08 Dgr | H334 H317 |  |  |  | ATP22
607-775-00-X | Sodium 3-(allyloxy)-2-hydroxypropanesulphonate | 258-004-5 | 52556-42-0 | Repr. 1B Eye Dam. 1 | H360F H318 | GHS08 GHS05 Dgr | H360F H318 |  |  |  | ATP22
608-001-00-3 | acetonitrile; cyanomethane | 200-835-2 | 75-05-8 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Eye Irrit. 2 | H225 H332 H312 H302 H319 | GHS02 GHS07 Dgr | H225 H332 H312 H302 H319 |  |  |  | CLP00
608-002-00-9 | trichloroacetonitrile | 208-885-7 | 545-06-2 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Chronic 2 | H331 H311 H301 H411 | GHS06 GHS09 Dgr | H331 H311 H301 H411 |  |  |  | CLP00
608-003-00-4 | acrylonitrile | 203-466-5 | 107-13-1 | Flam. Liq. 2 Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H225 H350 H331 H311 H301 H335 H315 H318 H317 H411 | GHS02 GHS06 GHS08 GHS05 GHS09 Dgr | H225 H350 H331 H311 H301 H335 H315 H318 H317 H411 |  | * | D | CLP00
608-004-00-X | 2-hydroxy-2-methylpropionitrile; 2-cyanopropan-2-ol; acetone cyanohydrin | 200-909-4 | 75-86-5 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H400 H410 | GHS06 GHS09 Dgr | H330 H310 H300 H410 |  |  |  | CLP00
608-005-00-5 | n-butyronitrile | 203-700-6 | 109-74-0 | Flam. Liq. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H225 H331 H311 H301 | GHS02 GHS06 Dgr | H225 H331 H311 H301 |  |  |  | CLP00/ATP01
608-006-00-0 | bromoxynil (ISO); 3,5-dibromo-4-hydroxybenzonitrile; bromoxynil phenol | 216-882-7 | 1689-84-5 | Repr. 2 Acute Tox. 2 * Acute Tox. 3 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H330 H301 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H361d *** H330 H301 H317 H410 |  | M=10 |  | CLP00
608-007-00-6 | ioxynil (ISO); 4-hydroxy-3,5-diiodobenzonitrile | 216-881-1 | 1689-83-4 | Repr. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H331 H301 H312 H373 ** H319 H400 H410 | GHS06 GHS08 GHS09 Dgr | H361d *** H331 H301 H312 H373 ** H319 H410 |  | M=10 |  | CLP00
608-008-00-1 | chloroacetonitrile | 203-467-0 | 107-14-2 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Chronic 2 | H331 H311 H301 H411 | GHS06 GHS09 Dgr | H331 H311 H301 H411 |  |  |  | CLP00
608-009-00-7 | malononitrile | 203-703-2 | 109-77-3 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H301 H410 |  |  |  | CLP00
608-010-00-2 | methacrylonitrile; 2-methyl-2-propene nitrile | 204-817-5 | 126-98-7 | Flam. Liq. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Sens. 1 | H225 H331 H311 H301 H317 | GHS02 GHS06 Dgr | H225 H331 H311 H301 H317 |  | * Skin Sens. 1; H317: C ≥ 0,2 % | D | CLP00
608-011-00-8 | oxalonitrile; cyanogen | 207-306-5 | 460-19-5 | Flam. Gas 1 Press. Gas Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H220 H331 H400 H410 | GHS02 GHS04 GHS06 GHS09 Dgr | H220 H331 H410 |  |  | U | CLP00/ATP01corr
608-012-00-3 | benzonitrile | 202-855-7 | 100-47-0 | Acute Tox. 4 * Acute Tox. 4 * | H312 H302 | GHS07 Wng | H312 H302 |  |  |  | CLP00
608-013-00-9 | 2-chlorobenzonitrile | 212-836-5 | 873-32-5 | Acute Tox. 4 * Acute Tox. 4 * Eye Irrit. 2 | H312 H302 H319 | GHS07 Wng | H312 H302 H319 |  |  |  | CLP00
608-014-00-4 | chlorothalonil (ISO); tetrachloroisophthalonitrile | 217-588-1 | 1897-45-6 | Carc. 2 Acute Tox. 2 * STOT SE 3 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H330 H335 H318 H317 H400 H410 | GHS06 GHS05 GHS08 GHS09 Dgr | H351 H330 H335 H318 H317 H410 |  | M=10 |  | CLP00/ATP01
608-015-00-X | dichlobenil (ISO); 2,6-dichlorobenzonitrile | 214-787-5 | 1194-65-6 | Acute Tox. 4 * Aquatic Chronic 2 | H312 H411 | GHS07 GHS09 Wng | H312 H411 |  |  |  | CLP00
608-016-00-5 | 1,4-Dicyano-2,3,5,6-tetra-chloro-benzene | 401-550-8 | 1897-41-2 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
608-017-00-0 | bromoxynil octanoate (ISO); 2,6-dibromo-4-cyanophenyl octanoate | 216-885-3 | 1689-99-2 | Repr. 2 Acute Tox. 3 * Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H331 H302 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H361d *** H331 H302 H317 H410 |  | M=10 |  | CLP00
608-018-00-6 | ioxynil octanoate (ISO); 4-cyano-2,6-diiodophenyl octanoate | 223-375-4 | 3861-47-0 | Repr. 2 Acute Tox. 3 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H301 H319 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H361d *** H301 H319 H317 H410 |  | M=10 |  | CLP00
608-019-00-1 | 2,2'-dimethyl-2,2'-azodipropiononitrile; ADZN | 201-132-3 | 78-67-1 | Self-react. C Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 3 | H242 H332 H302 H412 | GHS02 GHS07 Dgr | H242 H332 H302 H412 |  |  | T | CLP00
608-020-00-7 | diphenoxymethylenecyanamide | 427-300-8 | 79463-77-7 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
608-021-00-2 | 3-(2-(diaminomethyleneamino)thiazol-4-ylmethylthio)propionitrile | 403-710-2 | 76823-93-3 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  |  |  | CLP00
608-022-00-8 | 3,7-dimethyloctanenitrile | 403-620-3 | 40188-41-8 | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H315 H317 H411 | GHS07 GHS09 Wng | H315 H317 H411 |  |  |  | CLP00
608-023-00-3 | fenbuconazole (ISO); 4-(4-chlorophenyl)-2-phenyl-2-[(1H-1,2,4-triazol-1-yl)methyl]butanenitrile | 406-140-2 | 114369-43-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
608-024-00-9 | 2-(4-(N-butyl-N-phenethylamino)phenyl)ethylene-1,1,2-tricarbonitrile | 407-650-8 | 97460-76-9 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
608-025-00-4 | 2-nitro-4,5-bis(benzyloxy)phenylacetonitrile | 410-970-0 | 117568-27-1 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
608-026-00-X | 3-cyano-3,5,5-trimethylcyclohexanone | 411-490-4 | 7027-11-4 | Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Chronic 3 | H302 H373 ** H317 H412 | GHS08 GHS07 Wng | H302 H373 ** H317 H412 |  |  |  | CLP00
608-027-00-5 | reaction mass of: 3-(4-ethylphenyl)-2,2-dimethylpropanenitrile; 3-(2-ethylphenyl)-2,2-dimethylpropanenitrile; 3-(3-ethylphenyl)-2,2-dimethylpropanenitrile | 412-660-0 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
608-028-00-0 | 4-(2-cyano-3-phenylamino-acryloyloxymethyl)-cyclohexyl-methyl 2-cyano-3-phenylamino)-acrylate | 413-510-7 | 147374-67-2 | STOT RE 2 * Skin Sens. 1 Aquatic Chronic 2 | H373 ** H317 H411 | GHS08 GHS09 Wng | H373 ** H317 H411 |  |  |  | CLP00
608-029-00-6 | 1,2-dihydro-6-hydroxy-4-methyl-1-[3-(1-methylethoxy)propyl]-2-oxo-3-pyridinecarbonitrile | 411-990-2 | 68612-94-2 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
608-030-00-1 | N-acetyl-N-[5-cyano-3-(2-dibutylamino-4-phenylthyazol-5-yl-methylene)-4-methyl-2,6-dioxo-1,2,3,6-tetrahydropyridin-1-yl]benzamide | 412-340-0 | 147741-93-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
608-031-00-7 | 2-benzyl-2-methyl-3-butenitrile | 407-870-4 | 97384-48-0 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
608-032-00-2 | acetamiprid (ISO); (1E)-N-[(6-chloropyridin-3-yl) methyl]-N’-cyano-N-methylethanimidamide; (E)-N1-[(6-chloro-3-pyridyl)methyl]-N2-cyano-N1-methylacetamidine | - | 135410-20-7; 160430-64-8 | Repr. 2 Acute Tox. 3 Aquatic Acute 1 Aquatic Chronic 1 | H361d H301 H400 H410 | GHS08 GHS06 GHS09 Dgr | H361d H301 H410 |  | oral: ATE = 140 mg/kg bw M = 10 M = 10 |  | ATP01/ATP18
608-033-00-8 | N-butyl-3-(2-chloro-4-nitrophenylhydrazono)-1-cyano-2-methylprop-1-ene-1,3-dicarboximide | 407-970-8 | 75511-91-0 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
608-034-00-3 | chlorfenapyr (ISO); 4-bromo-2-(4-chlorophenyl)-1-ethoxymethyl-5-trifluoromethylpyrrole-3-carbonitrile | - | 122453-73-0 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H302 H400 H410 | GHS06 GHS09 Dgr | H331 H302 H410 |  | M=100 |  | CLP00/ATP01
608-035-00-9 | (±)-α-[(2-acetyl-5-methylphenyl)-amino]-2,6-dichlorobenzene-aceto-nitrile | 419-290-9 |  | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
608-036-00-4 | 3-(2-{}{4-[2-(4-cyanophenyl)vinyl]phenyl}}vinyl)benzonitrile | 419-060-8 | 79026-02-1 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
608-037-00-X | reaction mass of: (E)-2,12-tridecadiennitrile; (E)-3,12-tridecadiennitrile; (Z)-3,12-tridecadiennitrile | 422-190-8 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
608-038-00-5 | 2,2,4-trimethyl-4-phenyl-butane-nitrile | 422-580-8 | 75490-39-0 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
608-039-00-0 | 2-phenylhexanenitrile | 423-460-8 | 3508-98-3 | Acute Tox. 4 Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  | Oral: ATE = 500 mg/kg bw |  | CLP00/ATP14
608-040-00-6 | 4,4'-dithiobis(5-amino-1-(2,6-dichloro-4-(trifluoromethyl)phenyl)-1H-pyrazole-3-carbonitrile) | 423-490-1 | 130755-46-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
608-041-00-1 | 4'-((2-butyl-4-oxo-1,3-diazaspiro[4.4]non-1-ene-3-yl)methyl)(1,1'-biphenyl)-2-carbonitrile | 423-500-4 | 138401-24-8 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
608-042-00-7 | (S)-2,2-diphenyl-2-(3-pyrrolidinyl)acetonitrile hydrobromide | 421-810-4 | 194602-27-2 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H317 H411 |  |  |  | ATP01
608-043-00-2 | 3-(cis-3-hexenyloxy)propanenitril | 415-220-6 | 142653-61-0 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H302 H400 H410 | GHS06 GHS09 Dgr | H331 H302 H410 |  |  |  | CLP00
608-044-00-8 | 2-cyclohexylidene-2-phenylacetonitrile | 423-740-1 | 10461-98-0 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | ATP01
608-046-00-9 | 5-(4-chloro-2-nitro-phenylazo)-1,2-dihydro-6-hydroxy-1,4-dimethyl-2-oxo-pyridine-3-carbonitrile | 425-310-7 | 77889-90-8 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
608-047-00-4 | 2-piperidin-1-yl-benzonitrile | 427-330-1 | 72752-52-4 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
608-048-00-X | 1-(3-cyclopentyloxy-4-methoxyphenyl)-4-oxo-cyclohexanecarbonitrile | 427-450-4 | 152630-47-2 | Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Chronic 2 | H302 H373 ** H317 H411 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H317 H411 |  |  |  | ATP01
608-049-00-5 | 2-(4-(4-(butyl-(1-methylhexyl)amino)phenyl)-3-cyano-5-oxo-1,5-dihydropyrrol-2-ylidene)propandinitrile | 429-180-2 | 157362-53-3 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | ATP01
608-050-00-0 | reaction mass of: 5-(2-cyano-4-nitrophenylazo)-2-(2-(2-hydroxyethoxy)ethylamino)-4-methyl-6-phenylaminonicotinonitrile; 5-(2-cyano-4-nitrophenylazo)-6-(2-(2-hydroxyethoxy)ethylamino)-4-methyl-2-phenylaminonicotinonitrile | 429-760-5 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
608-051-00-6 | (R)-4-(4-dimethylamino-1-(4-fluorophenyl)-1-hydroxybutyl)-3-(hydroxymethyl)benzonitrile | 430-760-2 | 219861-18-4 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | ATP01
608-052-00-1 | (S)-4-(4-dimethylamino-1-(4-fluorophenyl)-1-hydroxybutyl)-3-(hydroxymethyl)benzonitrile | 430-770-7 | 128173-52-4 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | ATP01
608-053-00-7 | (R,S)-4-(4-dimethylamino-1-(4-fluorophenyl)-1-hydroxybutyl)-3-(hydroxymethyl)benzonitrile | 430-780-1 | 103146-25-4 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | ATP01
608-054-00-2 | (R,S)-4-(4-dimethylamino-1-(4-fluorophenyl)-1-hydroxybutyl)-3-(hydroxymethyl)benzonitrile hemisulfate | 430-790-6 | - | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H317 H411 |  |  |  | ATP01
608-055-00-8 | fipronil (ISO); (±)-5-amino-1-(2,6-dichloro-α,α,α-trifluoro-para-tolyl)-4-trifluoromethylsulfinyl-pyrazole-3-carbonitrile | 424-610-5 | 120068-37-3 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H372 * H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H372 * H410 |  | M=1000 M=10000 |  | ATP01/ATP10
608-056-00-3 | N-methyl-N-cyanomethylmorpholiniummethylsulfate | 429-340-1 | - | Acute Tox. 4 * Eye Dam. 1 | H302 H318 | GHS05 GHS07 Dgr | H302 H318 |  |  |  | ATP01
608-057-00-9 | 4-(cyanomethyl)-4-methylmorpholin-4-ium hydrogen sulfate | 431-200-1 | 208538-34-5 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 | H302 H318 H317 | GHS05 GHS07 Dgr | H302 H318 H317 |  |  |  | ATP01/ATP01corr
608-058-00-4 | esfenvalerate (ISO); (S)-α-cyano- 3-phenoxybenzyl-(S)- 2-(4-chlorophenyl)- 3-methylbutyrate |  | 66230-04-4 | Acute Tox. 3 Acute Tox. 3 STOT SE 1 STOT RE 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H370 (nervous system) H373 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H301 H317 H370 (nervous system) H373 H410 |  | Inhalation: ATE = 0.53 mg/L (dusts/mists) Oral: ATE = 88.5 mg/kg bw M=10000 M=10000 |  | CLP00/ATP17
608-059-00-X | 5-amino-1-(2,6-dichloro-4-(trifluoromethyl)phenyl)-1H-pyrazole-3-carbonitrile | 421-240-6 | 120068-79-3 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
608-060-00-5 | 5-methyl-2-[(2-nitrophenyl)amino]-3-thiophenecarbonitrile | 421-300-1 | 138564-59-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
608-062-00-6 | 2-fluoro-4-hydroxybenzonitrile | 422-810-7 | 82380-18-5 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | ATP01
608-063-00-1 | (S)-α-hydroxy-3-phenoxy-benzeneacetonitrile | 441-070-6 | 61826-76-4 | Acute Tox. 3 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H318 H317 H400 H410 | GHS06 GHS05 GHS09 Dgr | H301 H318 H317 H410 |  |  |  | ATP01
608-064-00-7 | cyanomethyltrimethylammoniummethylsulfate | 433-720-2 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
608-065-00-2 | salts of bromoxynil with the exception of those specified elsewhere in this Annex |  |  | Repr. 2 Acute Tox. 2 * Acute Tox. 3 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H330 H301 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H361d *** H330 H301 H317 H410 |  | M=10 | A | CLP00
608-066-00-8 | salts of ioxynil with the exception of those specified elsewhere in this Annex |  |  | Repr. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H331 H301 H312 H373 ** H319 H400 H410 | GHS06 GHS08 GHS09 Dgr | H361d *** H331 H301 H312 H373 ** H319 H410 |  | M=10 | A | CLP00
608-067-00-3 | 3,7-dimethylocta-2,6-dienenitrile | 225-918-0 | 5146-66-7 | Muta. 1B | H340 | GHS08 Dgr | H340 |  |  |  | ATP09
608-068-00-9 | flutianil (ISO); (2Z)-{[2-fluoro-5-(trifluoro­methyl)phenyl]thio}[3-(2-methoxyphenyl)-1,3-thiazolidin-2-ylidene]acetonitrile |  | 958647-10-4 | Aquatic Chronic 1 | H410 | GHS09 Wng | H410 |  | M=100 |  | ATP13
608-069-00-4 | fludioxonil (ISO); 4-(2,2-difluoro-1,3-benzodioxol-4-yl)-1H-pyrrole-3-carbonitrile |  | 131341-86-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1 M=10 |  | ATP14
609-001-00-6 | 1-nitropropane | 203-544-9 | 108-03-2 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H226 H332 H312 H302 | GHS02 GHS07 Wng | H226 H332 H312 H302 |  | * |  | CLP00
609-002-00-1 | 2-nitropropane | 201-209-1 | 79-46-9 | Flam. Liq. 3 Carc. 1B Acute Tox. 4 * Acute Tox. 4 * | H226 H350 H332 H302 | GHS02 GHS08 GHS07 Dgr | H226 H350 H332 H302 |  |  |  | CLP00
609-003-00-7 | nitrobenzene | 202-716-0 | 98-95-3 | Carc. 2 Repr. 1B Acute Tox. 3 Acute Tox. 3 Acute Tox. 3 STOT RE 1 Aquatic Chronic 3 | H351 H360F H331 H311 H301 H372 (blood) H412 | GHS06 GHS08 Dgr | H331 H311 H301 H351 H360F H372 H412 |  |  |  | CLP00/ATP05
609-004-00-2 | dinitrobenzene [1] 1,4-dinitrobenzene [2] 1,3-dinitrobenzene [3] 1,2-dinitrobenzene [4] | 246-673-6 [1] 202-833-7 [2] 202-776-8 [3] 208-431-8 [4] | 25154-54-5 [1] 100-25-4 [2] 99-65-0 [3] 528-29-0 [4] | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H310 H300 H373 ** H410 |  |  |  | CLP00
609-005-00-8 | 1,3,5-trinitrobenzene | 202-752-7 | 99-35-4 | Expl. 1.1 Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H201 H310 H330 H300 H373 ** H400 H410 | GHS01 GHS06 GHS08 GHS09 Dgr | H201 H330 H310 H300 H373 ** H410 |  |  |  | CLP00/ATP01
609-006-00-3 | 4-nitrotoluene | 202-808-0 | 99-99-0 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 2 | H331 H311 H301 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H411 |  |  |  | CLP00
609-007-00-9 | 2,4-dinitrotoluene [1] dinitrotoluene [2] | 204-450-0 [1] 246-836-1 [2] | 121-14-2 [1] 25321-14-6 [2] | Carc. 1B Muta. 2 Repr. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H341 H361f *** H331 H311 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H341 H361f *** H331 H311 H301 H373 ** H410 |  |  |  | CLP00/ATP01
609-008-00-4 | 2,4,6-trinitrotoluene; TNT | 204-289-6 | 118-96-7 | Expl. 1.1 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 2 | H201 H331 H311 H301 H373 ** H411 | GHS01 GHS06 GHS08 GHS09 Dgr | H201 H331 H311 H301 H373 ** H411 |  |  |  | CLP00
609-009-00-X | 2,4,6-trinitrophenol; picric acid | 201-865-9 | 88-89-1 | Expl. 1.1 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H201 H331 H311 H301 | GHS01 GHS06 Dgr | H201 H331 H311 H301 |  |  |  | CLP00/ATP01
609-010-00-5 | salts of picric acid |  |  | Unst. Expl. Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H200 H331 H311 H301 | GHS01 GHS06 Dgr | H200 H331 H311 H301 |  |  | T | CLP00
609-011-00-0 | 2,4,6-trinitroanisole |  | 606-35-9 | Expl. 1.1 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 2 | H201 H332 H312 H302 H411 | GHS01 GHS07 GHS09 Wng | H201 H332 H312 H302 H411 |  |  |  | CLP00
609-012-00-6 | 2,4,6-trinitro-m-cresol | 210-027-1 | 602-99-3 | Expl. 1.1 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H201 H332 H312 H302 | GHS01 GHS07 Wng | H201 H332 H312 H302 |  |  |  | CLP00
609-013-00-1 | 2,4,6-trinitro-m-xylene | 211-187-5 | 632-92-8 | Expl. 1.1 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * | H201 H332 H312 H302 H373 ** | GHS01 GHS08 GHS07 Wng | H201 H332 H312 H302 H373 ** |  |  |  | CLP00
609-015-00-2 | 4-nitrophenol; p-nitrophenol | 202-811-7 | 100-02-7 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * | H332 H312 H302 H373 ** | GHS08 GHS07 Wng | H332 H312 H302 H373 ** |  |  |  | CLP00
609-016-00-8 | dinitrophenol (reaction mass of isomers) [1] 2,4(or 2,6)-dinitrophenol [2] | 247-096-2 [1] 275-732-9 [2] | 25550-58-7 [1] 71629-74-8 [2] | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H410 |  |  |  | CLP00
609-018-00-9 | 2,4,6-trinitroresorcinol; styphnic acid | 201-436-6 | 82-71-3 | Expl. 1.1 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H201 H332 H312 H302 | GHS01 GHS07 Dgr | H201 H332 H312 H302 |  |  |  | CLP00/ATP01
609-019-00-4 | lead 2,4,6-trinitro-m-phenylene dioxide; lead 2,4,6-trinitroresorcinoxide; lead styphnate | 239-290-0 | 15245-44-0 | Repr. 1A Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 Unst. Expl. | H360Df H332 H302 H373 ** H400 H410 H200 | GHS01 GHS08 GHS07 GHS09 Dgr | H200 H360Df H332 H302 H373 ** H410 |  |  | 1 | CLP00
609-019-01-1 | lead 2,4,6-trinitro-m-phenylene dioxide; lead 2,4,6-trinitroresorcinoxide; lead styphnate (≥ 20 % phlegmatiser) | 239-290-0 | 15245-44-0 | Expl. 1.1 Repr. 1A Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H201 H360Df H332 H302 H373 ** H400 H410 | GHS01 GHS08 GHS07 GHS09 Dgr | H201 H360Df H332 H302 H373 ** H410 |  |  | 1 | CLP00
609-020-00-X | DNOC (ISO); 4,6-dinitro-o-cresol | 208-601-1 | 534-52-1 | Muta. 2 Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H341 H310 H330 H300 H315 H318 H317 H400 H410 | GHS06 GHS08 GHS05 GHS07 GHS09 Dgr | H341 H330 H310 H300 H315 H318 H317 H410 | EUH044 |  |  | CLP00
609-021-00-5 | sodium salt of DNOC; sodium 4,6-dinitro-o-cresolate [1] potassium salt of DNOC; potassium 4,6-dinitro-o-cresolate [2] | 219-007-7 [1] | 2312-76-7 [1] 5787-96-2 [2] | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H410 |  |  |  | CLP00
609-022-00-0 | ammonium salt of DNOC; ammonium 4,6-dinitro-o-tolyl oxide | 221-037-0 | 2980-64-5 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H330 H300 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H310 H300 H373 ** H410 |  |  |  | CLP00
609-023-00-6 | dinocap (ISO); (RS)-2,6-dinitro-4-octylphenyl crotonates and (RS)-2,4-dinitro-6-octylphenyl crotonates in which “octyl” is a reaction mass of 1-methylheptyl, 1-ethylhexyl and 1-propylpentyl groups | 254-408-0 | 39300-45-3 | Repr. 1B Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D *** H332 H302 H373 ** H315 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H360D *** H332 H302 H373 ** H315 H317 H410 |  | M=100 |  | CLP00/ATP01
609-024-00-1 | binapacryl (ISO); 2-sec-butyl-4,6-dinitrophenyl-3-methylcrotonate | 207-612-9 | 485-31-4 | Repr. 1B Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H360D *** H312 H302 H400 H410 | GHS08 GHS07 GHS09 Dgr | H360D *** H312 H302 H410 |  |  |  | CLP00
609-025-00-7 | dinoseb (ISO); 6-sec-butyl-2,4-dinitrophenol | 201-861-7 | 88-85-7 | Repr. 1B Acute Tox. 3 * Acute Tox. 3 * Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H360Df H311 H301 H319 H400 H410 | GHS06 GHS08 GHS09 Dgr | H360Df H311 H301 H319 H410 | EUH044 |  |  | CLP00
609-026-00-2 | salts and esters of dinoseb, with the exception of those specified elsewhere in this Annex |  |  | Repr. 1B Acute Tox. 3 * Acute Tox. 3 * Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H360Df H311 H301 H319 H400 H410 | GHS06 GHS08 GHS09 Dgr | H360Df H311 H301 H319 H410 | EUH044 |  | A | CLP00
609-027-00-8 | dinocton; reaction mass of isomers: methyl 2-octyl-4,6-dinitrophenyl carbonate, methyl 4-octyl-2,6-dinitrophenyl carbonate |  | 63919-26-6 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
609-028-00-3 | dinex (ISO); 2-cyclohexyl-4,6-dinitrophenol | 205-042-5 | 131-89-5 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H301 H410 |  |  |  | CLP00
609-029-00-9 | salts and esters of dinex |  |  | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H301 H410 |  |  | A | CLP00
609-030-00-4 | dinoterb (ISO); 2-tert-butyl-4,6-dinitrophenol | 215-813-8 | 1420-07-1 | Repr. 1B Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H360D *** H300 H311 H400 H410 | GHS06 GHS08 GHS09 Dgr | H360D *** H300 H311 H410 | EUH044 |  |  | CLP00
609-031-00-X | salts and esters of dinoterb |  |  | Repr. 1B Acute Tox. 2 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H360D *** H300 H311 H400 H410 | GHS06 GHS08 GHS09 Dgr | H360D *** H300 H311 H410 |  |  | A | CLP00
609-032-00-5 | bromofenoxim (ISO); 3,5-dibromo-4-hydroxybenzaldehyde-O-(2,4-dinitrophenyl)-oxime | 236-129-6 | 13181-17-4 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
609-033-00-0 | dinosam (ISO); 2-(1-methylbutyl)-4,6-dinitrophenol |  | 4097-36-3 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H301 H410 |  |  |  | CLP00
609-034-00-6 | salts and esters of dinosam |  |  | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H301 H410 |  |  | A | CLP00
609-035-00-1 | nitroethane | 201-188-9 | 79-24-3 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * | H226 H332 H302 | GHS02 GHS07 Wng | H226 H332 H302 |  | * |  | CLP00
609-036-00-7 | nitromethane | 200-876-6 | 75-52-5 | Flam. Liq. 3 Acute Tox. 4 * | H226 H302 | GHS02 GHS07 Wng | H226 H302 |  | * |  | CLP00
609-037-00-2 | 5-nitroacenaphthene | 210-025-0 | 602-87-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
609-038-00-8 | 2-nitronaphthalene | 209-474-5 | 581-89-5 | Carc. 1B Aquatic Chronic 2 | H350 H411 | GHS08 GHS09 Dgr | H350 H411 |  |  |  | CLP00
609-039-00-3 | 4-nitrobiphenyl | 202-204-7 | 92-93-3 | Carc. 1B Aquatic Chronic 2 | H350 H411 | GHS08 GHS09 Dgr | H350 H411 |  |  |  | CLP00
609-040-00-9 | nitrofen (ISO); 2,4-dichlorophenyl 4-nitrophenyl ether | 217-406-0 | 1836-75-5 | Carc. 1B Repr. 1B Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H360D *** H302 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H360D *** H302 H410 |  |  |  | CLP00
609-041-00-4 | 2,4-dinitrophenol | 200-087-7 | 51-28-5 | Acute Tox. 2 Acute Tox. 3 * Acute Tox. 3 STOT RE 1 Aquatic Acute 1 | H300 H331 H311 H372 H400 | GHS06 GHS08 GHS09 Dgr | H331 H311 H300 H372 H400 |  | Oral: ATE = 30 mg/kg bw Dermal: ATE = 300 mg/kg bw |  | CLP00/ATP15
609-042-00-X | pendimethalin (ISO); N-(1-ethylpropyl)-2,6-dinitro-3,4-xylidine | 254-938-2 | 40487-42-1 | Repr. 2 Aquatic Acute 1 Aquatic Chronic 1 | H361d H400 H410 | GHS08 GHS09 Wng | H361d H410 |  | M = 100 M = 10 |  | CLP00/ATP18
609-043-00-5 | quintozene (ISO); pentachloronitrobenzene | 201-435-0 | 82-68-8 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
609-044-00-0 | tecnazene (ISO); 1,2,4,5-tetrachloro-3-nitrobenzene | 204-178-2 | 117-18-0 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
609-045-00-6 | reaction mass of: 4,6-dinitro-2-(3-octyl)phenyl methyl carbonate and 4,6-dinitro-2-(4-octyl)phenyl methyl carbonate; dinocton-6 |  | 8069-76-9 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
609-046-00-1 | trifluralin (ISO) (containing < 0.5 ppm NPDA); α,α,α-trifluoro-2,6-dinitro-N,N-dipropyl-p-toluidine (containing < 0.5 ppm NPDA); 2,6-dinitro-N,N-dipropyl-4-trifluoromethylaniline (containing < 0.5 ppm NPDA); N,N-dipropyl-2,6-dinitro-4-trifluoromethylaniline (containing < 0.5 ppm NPDA) | 216-428-8 | 1582-09-8 | Carc. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H317 H410 |  | M=10 |  | CLP00/ATP01
609-047-00-7 | 2-nitroanisole | 202-052-1 | 91-23-6 | Carc. 1B Acute Tox. 4 * | H350 H302 | GHS08 GHS07 Dgr | H350 H302 |  |  |  | CLP00
609-048-00-2 | sodium 3-nitrobenzenesulphonate | 204-857-3 | 127-68-4 | Eye Irrit. 2 Skin Sens. 1 | H319 H317 | GHS07 Wng | H319 H317 |  |  |  | CLP00
609-049-00-8 | 2,6-dinitrotoluene | 210-106-0 | 606-20-2 | Carc. 1B Muta. 2 Repr. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 3 | H350 H341 H361f *** H331 H311 H301 H373 ** H412 | GHS06 GHS08 Dgr | H350 H341 H361f *** H331 H311 H301 H373 ** H412 |  |  |  | CLP00
609-050-00-3 | 2,3-dinitrotoluene | 210-013-5 | 602-01-7 | Carc. 1B Muta. 2 Repr. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H341 H361f *** H331 H311 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H341 H361f *** H331 H311 H301 H373 ** H410 |  |  |  | CLP00
609-051-00-9 | 3,4-dinitrotoluene | 210-222-1 | 610-39-9 | Carc. 1B Muta. 2 Repr. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 2 | H350 H341 H361f *** H331 H311 H301 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H350 H341 H361f *** H331 H311 H301 H373 ** H411 |  |  |  | CLP00
609-052-00-4 | 3,5-dinitrotoluene | 210-566-2 | 618-85-9 | Carc. 1B Muta. 2 Repr. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 3 | H350 H341 H361f *** H331 H311 H301 H373 ** H412 | GHS06 GHS08 Dgr | H350 H341 H361f *** H331 H311 H301 H373 ** H412 |  |  |  | CLP00
609-053-00-X | hydrazine-trinitromethane | 414-850-9 |  | Expl. 1.1 **** Self-react. A Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Skin Sens. 1 | H201 H240 H350 H331 H301 H317 | GHS01 GHS06 GHS08 Dgr | H201 H240 H350 H331 H301 H317 |  |  |  | CLP00
609-054-00-5 | 2,3-dinitrophenol [1] 2,5-dinitrophenol [2] 2,6-dinitrophenol [3] 3,4-dinitrophenol [4] salts of dinitrophenol [5] | 200-628-7 [1] 206-348-1 [2] 209-357-9 [3] 209-415-3 [4] | 66-56-8 [1] 329-71-5 [2] 573-56-8 [3] 577-71-9 [4] | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 2 | H331 H311 H301 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H411 |  |  |  | CLP00
609-055-00-0 | 2,5-dinitrotoluene | 210-581-4 | 619-15-8 | Carc. 1B Muta. 2 Repr. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 2 | H350 H341 H361f *** H331 H311 H301 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H350 H341 H361f *** H331 H311 H301 H373 ** H411 |  |  |  | CLP00
609-056-00-6 | 2,2-dibromo-2-nitroethanol | 412-380-9 | 69094-18-4 | Expl. 1.1 Carc. 2 Acute Tox. 4 * STOT RE 2 * Skin Corr. 1A Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H201 H351 H302 H373 ** H314 H317 H400 H410 | GHS01 GHS08 GHS05 GHS07 GHS09 Dgr | H201 H351 H302 H373 ** H314 H317 H410 |  | * STOT SE 3; H335: C ≥ 1 % | T | CLP00
609-057-00-1 | 3-chloro-2,4-difluoronitrobenzene | 411-980-8 | 3847-58-3 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H314 H317 H410 |  |  |  | CLP00
609-058-00-7 | 2-nitro-2-phenyl-1,3-propanediol | 410-360-4 | 5428-02-4 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Skin Sens. 1 Aquatic Chronic 2 | H312 H302 H372 ** H317 H411 | GHS08 GHS07 GHS09 Dgr | H372 ** H312 H302 H317 H411 | EUH070 |  |  | CLP00
609-059-00-2 | 2-chloro-6-(ethylamino)-4-nitrophenol | 411-440-1 | 131657-78-8 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | CLP00
609-060-00-8 | 4-[(3-hydroxypropyl)amino]-3-nitrophenol | 406-305-9 | 92952-81-3 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
609-061-00-3 | (E,Z)-4-chlorophenyl(cyclopropyl)ketone O-(4-nitrophenylmethyl)oxime | 406-100-4 | 94097-88-8 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
609-062-00-9 | 2-bromo-2-nitropropanol | 407-030-7 | 24403-04-1 | Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H311 H302 H373 ** H314 H317 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H311 H302 H373 ** H314 H317 H410 |  |  |  | CLP00
609-063-00-4 | 2-[(4-chloro-2-nitrophenyl)amino]ethanol | 413-280-8 | 59320-13-7 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
609-064-00-X | mesotrione (ISO); 2-[4-(methylsulfonyl)-2-nitrobenzoyl]-1,3-cyclohexanedione |  | 104206-82-8 | Repr. 2 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H361d H373 (eyes, nervous system) H400 H410 | GHS08 GHS09 Wng | H361d H373 (eyes, nervous system) H410 |  | M=10 M=10 |  | CLP00/ATP15
609-065-00-5 | 2-nitrotoluene | 201-853-3 | 88-72-2 | Carc. 1B Muta. 1B Repr. 2 Acute Tox. 4 * Aquatic Chronic 2 | H350 H340 H361f *** H302 H411 | GHS08 GHS07 GHS09 Dgr | H350 H340 H361f *** H302 H411 |  |  |  | CLP00
609-067-00-6 | sodium and potassium 4-(3-aminopropylamino)-2,6-bis[3-(4-methoxy-2-sulfophenylazo)-4-hydroxy-2-sulfo-7-naphthylamino]-1,3,5-triazine | 416-280-6 | 156769-97-0 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
609-068-00-1 | musk xylene; 5-tert-butyl-2,4,6-trinitro-m-xylene | 201-329-4 | 81-15-2 | Expl. 1.1 Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H201 H351 H400 H410 | GHS01 GHS08 GHS09 Wng | H201 H351 H410 |  |  | T | CLP00
609-069-00-7 | musk ketone; 3,5-dinitro-2,6-dimethyl-4-tert-butylacetophenone; 4'-tert-butyl-2',6'-dimethyl-3',5'-dinitroacetophenone | 201-328-9 | 81-14-1 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  |  |  | ATP01
609-070-00-2 | 1,4-dichloro-2-(1,1,2,3,3,3-hexafluoropropoxy)-5-nitrobenzene | 415-580-4 | 130841-23-5 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
609-071-00-8 | reaction mass of: 2-methylsulfanyl-4,6-bis-(2-hydroxy-4-methoxy-phenyl)-1,3,5-triazine; 2-(4,6-bis-methylsulfanyl-1,3,5-triazin-2-yl)-5-methoxy-phenol | 423-520-3 | 156137-33-6 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
609-072-00-3 | 4-mesyl-2-nitrotoluene | 430-550-0 | 1671-49-4 | Repr. 2 Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 3 | H361f *** H302 H317 H412 | GHS08 GHS07 Wng | H302 H317 H361f *** H412 |  |  |  | ATP01/ATP01corr
609-073-00-9 | lithium potassium sodium N,N''-bis{6-[7-[4-(4-chloro-1,3,5-triazin-2-yl)amino-4-(2-ureidophenylazo)]naphthalene-1,3,6-trisulfonato]}-N'-(2-aminoethyl)piperazine | 427-850-9 | - | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
609-074-00-4 | 1,4-dichloro-2-nitrobenzene | 201-923-3 | 89-61-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | ATP22
610-001-00-3 | trichloronitromethane; chloropicrin | 200-930-9 | 76-06-2 | Acute Tox. 2 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H330 H302 H335 H315 H319 | GHS06 Dgr | H330 H302 H319 H335 H315 |  |  |  | CLP00
610-002-00-9 | 1,1-dichloro-1-nitroethane | 209-854-0 | 594-72-9 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H331 H311 H301 | GHS06 Dgr | H331 H311 H301 |  |  |  | CLP00
610-003-00-4 | chlorodinitrobenzene |  |  | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H410 |  |  | C | CLP00
610-004-00-X | 2-chloro-1,3,5-trinitrobenzene | 201-864-3 | 88-88-0 | Expl. 1.1 Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H201 H310 H330 H300 H400 H410 | GHS01 GHS06 GHS09 Dgr | H201 H330 H310 H300 H410 |  |  |  | CLP00
610-005-00-5 | 1-chloro-4-nitrobenzene | 202-809-6 | 100-00-5 | Carc. 2 Muta. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 2 | H351 H341 H331 H311 H301 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H351 H341 H331 H311 H301 H373 ** H411 |  |  |  | CLP00
610-006-00-0 | chloronitroanilines with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Chronic 2 | H310 H330 H300 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H330 H310 H300 H373 ** H411 |  |  | A C | CLP00
610-007-00-6 | 1-chloro-1-nitropropane | 209-990-0 | 600-25-9 | Acute Tox. 4 * Acute Tox. 4 * | H332 H302 | GHS07 Wng | H332 H302 |  | * |  | CLP00
610-008-00-1 | 2,6-dichloro-4-nitroanisole | 403-350-6 | 17742-69-7 | Acute Tox. 3 * Aquatic Chronic 2 | H301 H411 | GHS06 GHS09 Dgr | H301 H411 |  |  |  | CLP00
610-009-00-7 | 2-chloro-4-nitroaniline | 204-502-2 | 121-87-9 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
610-010-00-2 | 2-bromo-1-(2-furyl)-2-nitroethylene | 406-110-9 | 35950-52-8 | Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H314 H317 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H302 H373 ** H314 H317 H410 |  |  |  | CLP00
611-001-00-6 | azobenzene | 203-102-5 | 103-33-3 | Carc. 1B Muta. 2 Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H341 H332 H302 H373 ** H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H341 H332 H302 H373 ** H410 |  |  |  | CLP00
611-002-00-1 | azoxybenzene | 207-802-1 | 495-48-7 | Acute Tox. 4 * Acute Tox. 4 * | H332 H302 | GHS07 Wng | H332 H302 |  |  |  | CLP00
611-003-00-7 | fenaminosulf (ISO); sodium 4-dimethylaminobenzenediazosulphonate | 205-419-4 | 140-56-7 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Chronic 3 | H301 H312 H412 | GHS06 Dgr | H301 H312 H412 |  |  |  | CLP00
611-004-00-2 | methyl-ONN-azoxymethyl acetate; methyl azoxy methyl acetate | 209-765-7 | 592-62-1 | Carc. 1B Repr. 1B | H350 H360D *** | GHS08 Dgr | H350 H360D *** |  |  |  | CLP00
611-005-00-8 | disodium {}{5-[(4'-((2,6-hydroxy-3-((2-hydroxy-5-sulphophenyl)azo)phenyl)azo)(1,1'-biphenyl)-4-yl)azo]salicylato(4-)}}cuprate(2-); CI Direct Brown 95 | 240-221-1 | 16071-86-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
611-006-00-3 | 4-o-tolylazo-o-toluidine; 4-amino-2',3-dimethylazobenzene; fast garnet GBC base; AAT; o-aminoazotoluene | 202-591-2 | 97-56-3 | Carc. 1B Skin Sens. 1 | H350 H317 | GHS08 Dgr | H350 H317 |  |  |  | CLP00
611-007-00-9 | tricyclazole (ISO); 5-methyl-1,2,4-triazolo(3,4-b)benzo-1,3-thiazole | 255-559-5 | 41814-78-2 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
611-008-00-4 | 4-aminoazobenzene; 4-phenylazoaniline | 200-453-6 | 60-09-3 | Carc. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H400 H410 | GHS08 GHS09 Dgr | H350 H410 |  |  |  | CLP00
611-009-00-X | sodium (1-(5-(4-(4-anilino-3-sulphophenylazo)-2-methyl-5-methylsulphonamidophenylazo)-4-hydroxy-2-oxido-3-(phenylazo)phenylazo)-5-nitro-4-sulphonato-2-naphtholato)iron(II) | 401-220-3 |  | Acute Tox. 4 * Aquatic Chronic 3 | H332 H412 | GHS07 Wng | H332 H412 |  |  |  | CLP00
611-010-00-5 | 2'-(2-cyano-4,6-dinitrophenylazo)-5'-(N,N-dipropylamino)propionanilide | 403-010-7 | 106359-94-8 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
611-011-00-0 | N,N,N',N'-tetramethyl-3,3'-(propylenebis(iminocarbonyl-4,1-phenylenazo(1,6-dihydro-2-hydroxy-4-methyl-6-oxopyridine-3,1-diyl)))di(propylammonium) dilactate | 403-340-1 |  | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
611-012-00-6 | reaction mass of 2,2-iminodiethanol 6-methyl-2-(4-(2,4,6-triaminopyrimidin-5-ylazo)phenyl)benzothiazole-7-sulfonate and 2-methylaminoethanol 6-methyl-2-(4-(2,4,6-triaminopyrimidin-5-ylazo)phenyl)benzothiazole-7-sulfonate and N,N-diethylpropane-1,3-diamine 6-methyl-2-(4-(2,4,6-triaminopyrimidin-5-ylazo)phenyl)benzothiazole-7-sulfonate | 403-410-1 | 114565-65-0 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-013-00-1 | trilithium-1-hydroxy-7-(3-sulfonatoanilino)-2-(3-methyl-4-(2-methoxy-4-(3-sulfonatophenylazo)phenylazo)phenylazo)naphthalene-3-sulfonate | 403-650-7 | 117409-78-6 | Expl. 1.3 **** Aquatic Chronic 2 | H203 H411 | GHS01 GHS09 Dgr | H203 H411 |  |  |  | CLP00
611-014-00-7 | (tetrasodium 1-(4-(3-acetamido-4-(4'-nitro-2,2'-disulfonatostilben-4-ylazo)anilino)-6-(2,5-disulfonatoanilino)-1,3,5-triazin-2-yl)-3-carboxypyridinium) hydroxide | 404-250-5 | 115099-55-3 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-015-00-2 | tetrasodium 4-amino-5-hydroxy-6-(4-(2-(2-(sulfonatooxy)ethylsulfonyl)ethylcarbamoyl)phenylazo)-3-(4-(2-(sulfonatooxy)ethylsulfonyl)phenylazo)naphthalene-2,7-disulfonate | 404-320-5 | 116889-78-2 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-016-00-8 | reaction mass of 1,1'-((dihydroxyphenylene)bis(azo-3,1-phenylenazo(1-(3-dimethylaminopropyl)-1,2-dihydro-6-hydroxy-4-methyl-2-oxopyridine-5,3-diyl)))dipyridinium dichloride dihydrochloride, mixed isomers and 1-(1-(3-dimethylaminopropyl)-5-(3-((4-(1-(3-dimethylaminopropyl)-1,6-dihydro-2-hydroxy-4-methyl-6-oxo-5-pyridinio-3-pyridylazo)phenylazo)-2,4(or2,6 or3,5)-dihydroxyphenylazo)phenylazo)-1,2-dihydro-6-hydroxy-4-methyl-2-oxo-3-pyridyl)pyridinium dichloride | 404-540-1 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-017-00-3 | 2-(4-(diethylaminopropylcarbamoyl)phenylazo)-3-oxo-N-(2,3-dihydro-2-oxobenzimidazol-5-yl)butyramide | 404-910-2 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
611-018-00-9 | tetraammonium 5-(4-(7-amino-1-hydroxy-3-sulfonato-2-naphthylazo)-6-sulfonato-1-naphthylazo)isophthalate | 405-130-5 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-019-00-4 | tetralithium 6-amino-4-hydroxy-3-(7-sulfonato-4-(4-sulfonatophenylazo)-1-naphthylazo)naphthalene-2,7-disulfonate | 405-150-4 | 106028-58-4 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-020-00-X | tetrakis(tetramethylammonium) 6-amino-4-hydroxy-3-(7-sulfonato-4-(4-sulfonatophenylazo)-1-naphthylazo)naphthalene-2,7-disulfonate | 405-170-3 | 116340-05-7 | Acute Tox. 3 * Skin Sens. 1 Aquatic Chronic 3 | H301 H317 H412 | GHS06 Dgr | H301 H317 H412 |  |  |  | CLP00
611-021-00-5 | 2-(4-(4-cyano-3-methylisothiazol-5-ylazo)-N-ethyl-3-methylanilino)ethyl acetate | 405-480-9 |  | Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Aquatic Chronic 4 | H302 H373 ** H315 H413 | GHS08 GHS07 Wng | H302 H373 ** H315 H413 |  |  |  | CLP00
611-022-00-0 | 4-dimethylaminobenzenediazonium 3-carboxy-4-hydroxybenzenesulfonate | 404-980-4 |  | Self-react. C Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H242 H331 H301 H312 H373 ** H318 H317 H400 H410 | GHS02 GHS06 GHS08 GHS05 GHS09 Dgr | H242 H331 H301 H312 H373 ** H318 H317 H410 |  |  | T | CLP00
611-023-00-6 | disodium 7-(4,6-dichloro-1,3,5-triazin-2-ylamino)-4-hydroxy-3-(4-(2-(sulfonatooxy)ethylsulfonyl)phenylazo) naphthalene-2-sulfonate | 404-600-7 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-024-00-1 | Benzidine based azo dyes; 4,4'-diarylazobiphenyl dyes, with the exception of those specified elsewhere in this Annex |  |  | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | A | CLP00
611-025-00-7 | disodium 4-amino-3-[[4'-[(2,4-diaminophenyl)azo][1,1'-biphenyl]-4-yl]azo]-5-hydroxy-6-(phenylazo)naphtalene-2,7-disulphonate; C.I. Direct Black 38 | 217-710-3 | 1937-37-7 | Carc. 1B Repr. 2 | H350 H361d *** | GHS08 Dgr | H350 H361d *** |  |  |  | CLP00
611-026-00-2 | tetrasodium 3,3'-[[1,1'-biphenyl]-4,4'-diylbis(azo)]bis[5-amino-4-hydroxynaphthalene-2,7-disulphonate]; C.I. Direct Blue 6 | 220-012-1 | 2602-46-2 | Carc. 1B Repr. 2 | H350 H361d *** | GHS08 Dgr | H350 H361d *** |  |  |  | CLP00
611-027-00-8 | disodium 3,3'-[[1,1'-biphenyl]-4,4'-diylbis(azo)]bis(4-aminonaphthalene-1-sulphonate); C.I. Direct Red 28 | 209-358-4 | 573-58-0 | Carc. 1B Repr. 2 | H350 H361d *** | GHS08 Dgr | H350 H361d *** |  |  |  | CLP00
611-028-00-3 | C,C'-azodi(formamide) | 204-650-8 | 123-77-3 | Resp. Sens. 1 | H334 | GHS08 Dgr | H334 |  |  | G | CLP00/ATP01corr
611-029-00-9 | o-dianisidine based azo dyes; 4,4'-diarylazo-3,3'-dimethoxybiphenyl dyes with the exception of those mentioned elsewhere in this Annex |  |  | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | A | CLP00
611-030-00-4 | o-tolidine based dyes; 4,4'-diarylazo-3,3'-dimethylbiphenyl dyes, with the exception of those mentioned elsewhere in this Annex |  |  | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | A | CLP00
611-031-00-X | 4,4'-(4-iminocyclohexa-2,5-dienylidenemethylene)dianiline hydrochloride; C.I. Basic Red 9 | 209-321-2 | 569-61-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
611-032-00-5 | 1,4,5,8-tetraaminoanthraquinone; C.I. Disperse Blue 1 | 219-603-7 | 2475-45-8 | Carc. 1B Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H350 H315 H318 H317 | GHS08 GHS05 GHS07 Dgr | H350 H315 H318 H317 |  |  |  | CLP00
611-033-00-0 | hexasodium [4,4''-azoxybis(2,2'-disulfonatostilbene-4,4'-diylazo)]-bis[5'-sulfonatobenzene-2,2'- diolato-O(2),O(2),N(1)]-copper(II) | 400-020-3 | 82027-60-9 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
611-034-00-6 | N-(5-(bis(2-methoxyethyl)amino)-2-((5-nitro-2,1-benzisothiazol-3-yl)azo)phenylacetamide | 402-430-8 | 105076-77-5 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
611-035-00-1 | tetralithium 6-amino-4-hydroxy-3-[7-sulfonato-4-(5-sulfonato-2-naphthylazo)-1-naphthylazo]naphthalene-2,7-disulfonate | 403-660-1 | 107246-80-0 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00/ATP01
611-036-00-7 | 2-(4-(5,6(or 6,7)-dichloro-1,3-benzothiazol-2-ylazo)-N-methyl-m-toluidino)ethyl acetate | 405-440-0 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-037-00-2 | 3(or 5)-(4-(N-benzyl-N-ethylamino)-2-methylphenylazo)-1,4-dimethyl-1,2,4-triazolium methylsulphate | 406-055-0 | 124584-00-5 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H317 H411 |  |  |  | CLP00
611-038-00-8 | trisodium 1-hydroxynaphthalene-2-azo-4'(5',5''-dimethylbiphenyl)-4''-azo(4''-phenylsulfonyloxybenzene)- 2',2'',4-trisulfonate | 406-820-9 |  | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
611-039-00-3 | 7-[((4,6-dichloro-1,3,5-triazin-2-yl)amino)-4-hydroxy-3-(4-((2-sulfoxy)ethyl)sulfonyl)phenylazo]naphthalene-2-sulfonic acid | 407-050-6 | 117715-57-8 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-040-00-9 | 3-(5-acetylamino-4-(4-[4,6-bis(3-diethylaminopropylamino)-1,3,5-triazin-2-ylamino]phenylazo)-2-(2-methoxyethoxy)phenylazo)-6-amino-4-hydroxy-2-naphthalenesulfonic acid | 407-670-7 | 115099-58-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
611-041-00-4 | 2-[[4[[4,6-bis[[3-(diethylamino)propyl]amino]-1,3,5-triazine-2-yl]amino]phenyl]azo]-N-(2,3-dihydro-2-oxo-1H-benzimidazol-5-yl)-3-oxobutanamide | 407-680-1 | 98809-11-1 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H318 H317 H411 |  |  |  | CLP00
611-042-00-X | trisodium 5-amino-3-[5-(2-bromoacryloylamino)-2-sulfonatophenylazo]-4-hydroxy-6-(4-vinylsulfonylphenylazo)naphthalene-2,7-disulfonate | 411-770-6 | 136213-71-3 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
611-043-00-5 | reaction mass of: trisodium N(1')-N(2):N(1''')-N(2'')-η-6-[2-amino-4-(or 6)-hydroxy-(or 4-amino-2-hydroxy)phenylazo]-6''-(1-carbaniloyl-2-hydroxyprop-1-enylazo)-5',5'''-disulfamoyl-3,3''-disulfonatobis(naphthalene-2,1'-azobenzene-1,2'-diolato-O(1),O(2'))-chromate; trisodium N(1')-N(2):N(1''')-N(2'')-η-6,6''-bis(1-carbaniloyl-2-hydroxyprop-1-enylazo)-5',5'''-disulfamoyl-3,3''-disulfonatobis(naphthalene-2,1'azobenzene-1,2'-diolato-O(1),O(2'))-chromate; trisodium N(1')-N(2):N(1''')-N(2'')-η-6,6''-bis[2-amino-4-(or 6)-hydroxy-(or 4-amino-2-hydroxy)phenylazo]5',5'''-disulfamoyl-3,3''-disulfonatobis(naphthalene-2,1'azobenzene-1,2'-diolato-O(1),O(2'))-chromate (2:1:1) | 402-850-1 |  | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
611-044-00-0 | reaction mass of: tert-alkyl(C12-C14)ammonium bis[1-[(2-hydroxy-5-nitrophenyl)azo]-2-naphthalenolato(2-)]-chromate(1-); tert-alkyl(C12-C14)ammonium bis[1-[(2-hydroxy-4-nitrophenyl)azo]-2-naphthalenolato(2-)]-chromate(1-); tert-alkyl(C12-C14)ammonium bis[1-[[5-(1,1-dimethylpropyl)-2-hydroxy-3-nitrophenyl]azo]-2-naphthalenolato(2-)]-chromate(1-); tert-alkyl(C12-C14)ammonium [[1-[(2-hydroxy-5-nitrophenyl)azo]-2-naphthalenolato(2-)]-[1-[(2-hydroxy-5-nitrophenyl)azo]-2-naphthalenolato(2-)]]-chromate(1-); tert-alkyl(C12-C14)ammonium [[1-[[5-(1,1-dimethylpropyl)-2-hydroxy-3-nitrophenyl]azo]-2-naphthalenolato(2-)]-[1-[(2-hydroxy-5-nitrophenyl)azo]-2-naphthalenolato(2-)]]-chromate(1-); tert-alkyl(C12-C14)ammonium ((1-(4(or 5)-nitro-2-oxidophenylazo)-2-naphtholato)(1-(3-nitro-2-oxido-5-pentylphenylazo)-2-naphtholato))chromate(1-) | 403-720-7 | 117527-94-3 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
611-045-00-6 | 2-[4-[N-(4-acetoxybutyl)-N-ethyl]amino-2-methylphenylazo]-3-acetyl-5-nitrothiophene | 404-830-8 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
611-046-00-1 | 4,4'-diamino-2-methylazobenzene | 407-590-2 | 43151-99-1 | Acute Tox. 3 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H373 ** H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H301 H373 ** H317 H410 |  |  |  | CLP00
611-047-00-7 | reaction mass of: 2-[[4-[N-ethyl-N-(2-acetoxyethyl)amino]phenyl]azo]-5,6-dichlorobenzothiazole; 2-[[4-[N-ethyl-N-(2-acetoxyethyl)amino]phenyl]azo]-6,7-dichlorobenzothiazole (1:1) | 407-890-3 | 111381-11-4 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
611-048-00-2 | reaction mass of: 2-[[4-[bis(2-acetoxyethyl)amino]phenyl]azo]-5,6-dichlorobenzothiazole; 2-[[4-[bis(2-acetoxyethyl)amino]phenyl]azo]-6,7-dichlorobenzothiazole (1:1) | 407-900-6 | 111381-12-5 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
611-049-00-8 | reaction mass of 7-[4-(3-diethylaminopropylamino)-6-(3-diethylammoniopropylamino)-1,3,5-triazin-2-ylamino]-4-hydroxy-3-(4-phenylazophenylazo)-naphthalene-2-sulfonate, acetic acid, lactic acid (2:1:1) | 408-000-6 | 118658-98-3 | STOT RE 2 * Skin Sens. 1 Aquatic Chronic 3 | H373 ** H317 H412 | GHS08 Wng | H373 ** H317 H412 |  |  |  | CLP00
611-050-00-3 | reaction mass of: pentasodium 7-amino-3-[[4-[[4-[[4-[[4-[(6-amino-1-hydroxy-3-sulfonato-2-naphthyl)azo]-7-sulfonato-1-naphthyl]azo]phenyl]amino]-3-sulfonatophenyl]azo]-6-sulfonato-1-naphthyl]azo]-4-hydroxynaphthalen-2-sulfonate; pentasodium 7-amino-8-[4-[4-[4-[4-(2-amino-5-hydroxy-7-sulfonato-naphthalen-1-ylazo)-7-sulfonatonaphthalen-1-ylazo]-phenylamino]-3-sulfonato-phenylazo]-6-sulfonato-naphthalen-1-ylazo]-4-hydroxy-naphthalene-2-sulfonate; pentasodium 7-amino-8-[4-[4-[4-[4-(6-amino-1-hydroxy-3-sulfonato-naphthalen-1-ylazo)-7-sulfonatonaphthalen-1-ylazo]-phenylamino]-3-sulfonato-phenylazo]-6-sulfonato-naphthalen-1-ylazo]-4-hydroxy-naphthalene-2-sulfonate; tetrasodium 7-amino-4-hydroxy-3-[4-[4-[4-(4-hydroxy-7-sulfonato-naphthalen-1-ylazo)-2-sulfonato-phenylamino]phenylazo]-6-sulfonato-naphthalen-1-ylazo]naphthalene-2-sulfonate; tetrasodium 7-amino-4-hydroxy-3-[4-[4-[4-(4-amino-7-sulfonato-naphthalen-1-ylazo)-2-sulfonato-phenylamino]phenylazo]-6-sulfonato-naphthalen-1-ylazo]naphthalene-2-sulfonate | 415-350-3 | - | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
611-051-00-9 | 2-(4-(N-ethyl-N-(2-hydroxy)ethyl)amino-2-methylphenyl)azo-6-methoxy-3-methyl-benzothiazolium chloride | 411-110-7 | 136213-74-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
611-052-00-4 | monosodium aqua-[5-[[2,4-dihydroxy-5-[(2-hydroxy-3,5-dinitrophenyl)azo]phenyl]azo]-2-naphthalensulfonate], iron complex | 400-720-9 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
611-053-00-X | 2,2'-azobis[2-methylpropionamidine] dihydrochloride | 221-070-0 | 2997-92-4 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  |  |  | CLP00
611-055-00-0 | C.I. Disperse Yellow 3; N-[4-[(2-hydroxy-5-methylphenyl)azo]phenyl]acetamide | 220-600-8 | 2832-40-8 | Carc. 2 Skin Sens. 1 | H351 H317 | GHS08 GHS07 Wng | H351 H317 |  |  |  | CLP00
611-056-00-6 | C.I. Solvent Yellow 14; 1-phenylazo-2-naphthol | 212-668-2 | 842-07-9 | Carc. 2 Muta. 2 Skin Sens. 1 Aquatic Chronic 4 | H351 H341 H317 H413 | GHS08 GHS07 Wng | H351 H341 H317 H413 |  |  |  | CLP00
611-057-00-1 | 6-hydroxy-1-(3-isopropoxypropyl)-4-methyl-2-oxo-5-[4-(phenylazo)phenylazo]-1,2-dihydro-3-pyridinecarbonitrile | 400-340-3 | 85136-74-9 | Carc. 1B Aquatic Chronic 4 | H350 H413 | GHS08 Wng | H350 H413 |  |  |  | CLP00
611-058-00-7 | (6-(4-hydroxy-3-(2-methoxyphenylazo)-2-sulfonato-7-naphthylamino)-1,3,5-triazin-2,4-diyl)bis[(amino-1-methylethyl)ammonium] formate | 402-060-7 | 108225-03-2 | Carc. 1B Eye Dam. 1 Aquatic Chronic 2 | H350 H318 H411 | GHS08 GHS05 GHS09 Dgr | H350 H318 H411 |  |  |  | CLP00
611-059-00-2 | octasodium 2-(6-(4-chloro-6-(3-(N-methyl-N-(4-chloro-6-(3,5-disulfonato-2-naphthylazo)-1-hydroxy-6-naphthylamino)-1,3,5-triazin-2-yl)aminomethyl)phenylamino)-1,3,5-triazin-2-ylamino)-3,5-disulfonato-1-hydroxy-2-naphthylazo)naphthalene-1,5-disulfonate | 412-960-1 | 148878-21-1 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | CLP00
611-060-00-8 | reaction mass of: sodium 5-[8-[4-[4-[4-[7-(3,5-dicarboxylatophenylazo)-8-hydroxy-3,6-disulfonatonaphthalen-1-ylamino]-6-hydroxy-1,3,5-triazin-2-yl]-2,5-dimethylpiperazin-1-yl]-6-hydroxy-1,3,5-triazin-2-ylamino]-1-hydroxy-3,6-disulfonatonaphthalen-2-ylazo]-isophthalate; ammonium 5-[8-[4-[4-[4-[7-(3,5-dicarboxylatophenylazo)-8-hydroxy-3,6-disulfonatonaphthalen-1-ylamino]-6-hydroxy-1,3,5-triazin-2-yl]-2,5-dimethylpiperazin-1-yl]-6-hydroxy-1,3,5-triazin-2-ylamino]-1-hydroxy-3,6-disulfonatonaphthalen-2-ylazo]-isophthalate; 5-[8-[4-[4-[4-[7-(3,5-dicarboxylatophenylazo)-8-hydroxy-3,6-disulfonatonaphthalen-1-ylamino]-6-hydroxy-1,3,5-triazin-2-yl]-2,5-dimethylpiperazin-1-yl]-6-hydroxy-1,3,5-triazin-2-ylamino]-1-hydroxy-3,6-disulfonaphthalen-2-ylazo]-isophthalic acid | 413-180-4 | 187285-15-0 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
611-061-00-3 | disodium 5-[5-[4-(5-chloro-2,6-difluoropyrimidin-4-ylamino)benzamido]-2-sulfonatophenylazo]-1-ethyl-6-hydroxy-4-methyl-2-oxo-3-pyridylmethylsulfonate | 412-530-3 |  | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
611-062-00-9 | octasodium 2-(8-(4-chloro-6-(3-((4-chloro-6-(3,6-disulfonato-2-(1,5-disulfonatonaphthalen-2-ylazo)-1-hydroxynaphthalen-8-ylamino)-1,3,5-triazin-2-yl)aminomethyl)phenylamino)-1,3,5-triazin-2-ylamino)-3,6-disulfonato-1-hydroxynaphthalen-2-ylazo)naphthalene-1,5-disulfonate | 413-550-5 |  | Skin Irrit. 2 Eye Dam. 1 | H315 H318 | GHS05 Dgr | H315 H318 |  |  |  | CLP00
611-063-00-4 | trisodium [4'-(8-acetylamino-3,6-disulfonato-2-naphthylazo)-4''-(6-benzoylamino-3-sulfonato-2-naphthylazo)-biphenyl-1,3',3'',1'''-tetraolato-O,O',O'',O''']copper(II) | 413-590-3 | 164058-22-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
611-064-00-X | 4-(3,4-dichlorophenylazo)-2,6-di-sec-butyl-phenol | 410-600-8 | 124719-26-2 | STOT RE 2 * Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H315 H400 H410 | GHS08 GHS07 GHS09 Wng | H373 ** H315 H410 |  |  |  | CLP00
611-065-00-5 | 4-(4-nitrophenylazo)-2,6-di-sec-butyl-phenol | 410-610-2 | 111850-24-9 | STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H315 H319 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H373 ** H319 H315 H317 H410 |  |  |  | CLP00
611-066-00-0 | tetrasodium 5-[4-chloro-6-(N-ethyl-anilino)-1,3,5-triazin-2-ylamino]-4-hydroxy-3-(1,5-disulfonatonaphthalen-2-ylazo)-naphthalene-2,7-disulfonate | 411-540-5 | 130201-57-9 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H318 H317 H411 |  |  |  | CLP00
611-067-00-6 | reaction mass of: bis(tris(2-(2-hydroxy(1-methyl)ethoxy)ethyl)ammonium) 7-anilino-4-hydroxy-3-(2-methoxy-5-methyl-4-(4-sulfonatophenylazo)phenylazo)naphthalene-2-sulfonate; bis(tris(2-(2-hydroxy(2-methyl)ethoxy)ethyl)ammonium) 7-anilino-4-hydroxy-3-(2-methoxy-5-methyl-4-(4-sulfonatophenylazo)phenylazo)naphthalene-2-sulfonate | 406-910-8 | - | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00/ATP01
611-068-00-1 | tetrasodium 4-amino-3,6-bis(5-[4-chloro-6-(2-hydroxyethylamino)-1,3,5-triazin-2-ylamino]-2-sulfonatophenylazo)-5-hydroxynaphthalene-2,7-disulfonate | 400-690-7 | 85665-98-1 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
611-069-00-7 | N,N-di-[poly(oxyethylene)-co-poly(oxypropylene)]-4-[(3,5-dicyano-4-methyl-2-thienyl)azo)]-3-methylaniline | 413-380-1 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
611-070-00-2 | reaction mass of: disodium (6-(4-anisidino)-3-sulfonato-2-(3,5-dinitro-2-oxidophenylazo)-1-naphtholato)(1-(5-chloro-2-oxidophenylazo)-2-naphtholato)chromate(1-); trisodium bis(5-(4-anisidino)-3-sulfonato-2-(3,5-dinitro-2-oxidophenylazo)-1-naphtholato)chromate(1-) | 405-665-4 |  | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
611-071-00-8 | tris(tetramethylammonium) 5-hydroxy-1-(4-sulphonatophenyl)-4-(4-sulphonatophenylazo)pyrazole-3-carboxylate | 406-073-9 | 131013-81-5 | Acute Tox. 3 * Aquatic Chronic 3 | H301 H412 | GHS06 Dgr | H301 H412 |  |  |  | CLP00
611-072-00-3 | 2,4-bis[2,2'-[2-(N,N-dimethylamino)ethyloxycarbonyl]phenylazo]-1,3-dihydroxybenzene, dihydrochloride | 407-010-8 | 118208-02-9 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | CLP00
611-073-00-9 | dimethyl 3,3'-(N-(4-(4-bromo-2,6-dicyanophenylazo)-3-hydroxyphenyl)imino)dipropionate | 407-310-9 | 122630-55-1 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
611-074-00-4 | reaction mass of: sodium/potassium (3-(4-(5-(5-chloro-2,6-difluoropyrimidin-4-ylamino)-2-methoxy-3-sulfonatophenylazo)-2-oxidophenylazo)-2,5,7-trisulfonato-4-naphtholato)copper(II); sodium/potassium (3-(4-(5-(5-chloro-4,6-difluoropyrimidin-2-ylamino)-2-methoxy-3-sulfonatophenylazo)-2-oxidophenylazo)-2,5,7-trisulfonato-4-naphtholato)copper(II) | 407-100-7 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-075-00-X | reaction mass of: tris(3,5,5-trimethylhexylammonium) 4-amino-3-(4-(4-(2-amino-4-hydroxyphenylazo)anilino)-3-sulfonatophenylazo)-5,6-dihydro-5-oxo-6-phenylhydrazononaphthalene-2,7-disulfonate; tris(3,5,5-trimethylhexylammonium) 4-amino-3-(4-(4-(4-amino-2-hydroxyphenylazo)anilino)-3-sulfonatophenylazo)-5,6-dihydro-5-oxo-6-phenylhydrazononaphthalene-2,7-disulfonate (2:1) | 406-000-0 |  | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
611-076-00-5 | 3-(2,6-dichloro-4-nitrophenylazo)-1-methyl-2-phenylindole | 406-280-4 | 117584-16-4 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
611-077-00-0 | dilithium disodium (5,5'-diamino-(μ-4,4'-dihydroxy-1:2-κ-2,O4,O4',-3,3'-[3,3'-dihydroxy-1:2-κ-2-O3,O3'-biphenyl-4,4'-ylenebisazo-1:2-(N3,N4-η:N3',N4'-η)]-dinaphthalene-2,7-disulfonato(8)))dicuprate(2-) | 407-230-4 | 126637-70-5 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  |  |  | CLP00
611-078-00-6 | (2,2'-(3,3'-dioxidobiphenyl-4,4'-diyldiazo)bis(6-(4-(3-(diethylamino)propylamino)-6-(3-(diethylammonio)propylamino)-1,3,5-triazin-2-ylamino)-3-sulfonato-1-naphtholato))dicopper(II) acetate lactate | 407-240-9 | 159604-94-1 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
611-079-00-1 | disodium 7-[4-chloro-6-(N-ethyl-o-toluidino)-1,3,5-triazin-2-ylamino]-4-hydroxy-3-(4-methoxy-2-sulfonatophenylazo)-2-naphthalenesulfonate | 410-390-8 | 147703-64-8 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
611-080-00-7 | sodium 3-(2-acetamido-4-(4-(2-hydroxybutoxy)phenylazo)phenylazo)benzenesulfonate | 410-150-2 | 147703-65-9 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-081-00-2 | tetrasodium [7-(2,5-dihydroxy-KO2-7-sulfonato-6-[4-(2,5,6-trichloro-pyrimidin-4-ylamino)phenylazo]-(N1,N7-N)-1-naphthylazo)-8-hydroxy-KO8-naphthalene-1,3,5-trisulfonato(6-)]cuprate(II) | 411-470-5 | 141048-13-7 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
611-082-00-8 | reaction mass of: pentasodium bis(1-(3(or 5)-(4-anilino-3-sulfonatophenylazo)-4-hydroxy-2-oxidophenylazo)-6-nitro-4-sulfonato-2-naphtholato)ferrate(1-); pentasodium [(1-(3-(4-anilino-3-sulfonatophenylazo)-4-hydroxy-2-oxidophenylazo)-6-nitro-4-sulfonato-2-naphtholato)-(5-(4-anilino-3-sulfonatophenylazo)-4-hydroxy-2-oxidophenylazo)-6-nitro-4-sulfonato-2-naphtholato]ferrate(1-) | 407-570-3 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
611-083-00-3 | reaction mass of: 2-[N-ethyl-4-[(5,6-dichlorobenzothiazol-2-yl)azo]-m-toludino]ethyl acetate; 2-[N-ethyl-4-[(6,7-dichlorobenzothiazol-2-yl)azo]-m-toludino]ethyl acetate (1:1) | 411-560-4 |  | STOT RE 1 Skin Sens. 1 Aquatic Chronic 2 | H372 ** H317 H411 | GHS08 GHS07 GHS09 Dgr | H372 ** H317 H411 |  |  |  | CLP00
611-085-00-4 | reaction mass of: 3-cyano-5-(2-cyano-4-nitro-phenylazo)-2-(2-hydroxy-ethylamino)-4-methyl-6-[3-(2-phenoxyethoxy)propylamino]pyridine; 3-cyano-5-(2-cyano-4-nitro-phenylazo)-6-(2-hydroxy-ethylamino)-4-methyl-2-[3-(2-phenoxyethoxy)propylamino]pyridine; 3-cyano-5-(2-cyano-4-nitro-phenylazo)-2-amino-4-methyl-6-[3-(3-hydroxypropoxy)propylamino]pyridine; 3-cyano-5-(2-cyano-4-nitro-phenylazo)-6-amino-4-methyl-2-[3-(3-methoxypropoxy)propylamino]pyridine | 411-880-4 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
611-086-00-X | monolithium 5-[[2,4-dihydroxy-5-[(2-hydroxy-3,5-dinitrophenyl)azo]phenyl]azo]-2-naphthalenesulfonate], iron complex, monohydrate | 411-360-7 |  | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
611-087-00-5 | reaction mass of: 3-((5-cyano-1,6-dihydro-1,4-dimethyl-2-hydroxyl-6-oxo-3-pyridinyl)azo)-benzoyloxy-2-phenoxyethane; 3-((5-cyano-1,6-dihydro-1,4-dimethyl-2-hydroxy-6-oxo-3-pyridinyl)azo)-benzoyloxy-2-ethyloxy-2-(ethylphenol) | 411-710-9 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
611-088-00-0 | reaction mass of: trilithium 4-amino-3-((4-((4-((2-amino-4-hydroxyphenyl)azo)phenyl)amino)-3-sulfophenyl)azo)-5-hydroxy-6-(phenylazo)naphthalene-2,7-disulfonate; trilithium 4-amino-3-((4-((4-((4-amino-2-hydroxyphenyl)azo)phenyl)amino)-3-sulfophenyl)azo)-5-hydroxy-6-(phenylazo)naphthalene-2,7-disulfonate | 411-890-9 |  | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | CLP00
611-089-00-6 | 2-((4-(ethyl-(2-hydroxyethyl)amino)-2-methylphenyl)azo)-6-methoxy-3-methyl-benzothiazolium methylsulfate | 411-100-2 | 136213-73-5 | STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H373 ** H317 H410 |  |  |  | CLP00
611-090-00-1 | 2,5-dibutoxy-4-(morpholin-4-yl)benzenediazonium 4-methylbenzenesulfonate | 413-290-2 | 93672-52-7 | Self-react. C Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H242 H302 H318 H317 H412 | GHS02 GHS05 GHS07 Dgr | H242 H302 H318 H317 H412 |  |  | T | CLP00
611-091-00-7 | sodium (1.0-1.95)/lithium (0.05-1) 5-((5-((5-chloro-6-fluoro-pyrimidin-4-yl)amino)-2-sulfonatophenyl)azo)-1,2-dihydro-6-hydroxy-1,4-dimethyl-2-oxo-3-pyridinemethylsulfonate | 413-470-0 | 134595-59-8 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-092-00-2 | tert-(dodecyl/tetradecyl)-ammonium bis(3-(4-((5-(1,1-dimethyl-propyl)-2-hydroxy-3-nitrophenyl)azo)-3-methyl-5-hydroxy-(1H)pyrazol-1-yl)benzenesulfonamidato)chromate | 413-210-6 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
611-093-00-8 | sodium 2-(4-(4-fluoro-6-(2-sulfo-ethylamino)-[1,3,5]triazin-2-ylamino)-2-ureido-phenylazo)-5-(4-sulfophenylazo)benzene-1-sulfonate | 410-770-3 | 146177-84-6 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-094-00-3 | reaction mass of: 2-[2-acetylamino-4-[N,N-bis[2-ethoxy-carbonyloxy)ethyl]amino]phenylazo]-5,6-dichloro-1,3-benzothiazole; 2-[2-acetylamino-4-[N,N-bis[2-ethoxy-carbonyloxy)ethyl]amino]phenylazo]-6,7-dichloro-1,3-benzotriazole (1:1) | 411-600-0 | 143145-93-1 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
611-095-00-9 | hexasodium 1,1'-[(1-amino-8-hydroxy-3,6-disulfonate-2,7-naphthalenediyl)bis(azo(4-sulfonate-1,3-phenyl)imino[6-[(4-chloro-3-sulfonatophenyl)amino]-1,3,5-triazin-2,4-diyl]]]bis[3-carboxypyridinium] dihydroxide | 412-240-7 | 89797-03-5 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
611-096-00-4 | methyl N-[3-acetylamino)-4-(2-cyano-4-nitrophenylazo)phenyl]-N-[(1-methoxy)acetyl]glycinate | 413-040-2 | 149850-30-6 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-097-00-X | reaction mass of iron complexes of: 1,3-dihydroxy-4-[(5-phenylaminosulfonyl)-2-hydroxyphenylazo]-n-(5-amino-sulfonyl-2-hydroxyphenylazo)benzene and: 1,3-dihydroxy-4-[(5-phenylaminosulfonyl)-2-hydroxyphenylazo]-n-[4-(4-nitro-2-sulfophenylamino)phenylazo]benzene (n=2,5,6) | 414-150-3 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
611-098-00-5 | tetrakis(tetramethylammonium)3,3'-(6-(2-hydroxyethylamino)1,3,5-triazine-2,4-diylbisimino(2-methyl-4,1-phenyleneazo))bisnaphthalene-1,5-disulfonate | 405-950-3 | 131013-83-7 | Acute Tox. 3 * Aquatic Chronic 3 | H301 H412 | GHS06 Dgr | H301 H412 |  |  |  | CLP00
611-099-00-0 | (methylenebis(4,1-phenylenazo(1-(3-(dimethylamino)propyl)-1,2-dihydro-6-hydroxy-4-methyl-2-oxopyridine-5,3-diyl)))-1,1'-dipyridinium dichloride dihydrochloride | 401-500-5 | 118658-99-4 | Carc. 1B Aquatic Chronic 2 | H350 H411 | GHS08 GHS09 Dgr | H350 H411 |  |  |  | CLP00
611-100-00-4 | potassium sodium 3,3'-(3(or4)-methyl-1,2-phenylenebis(imino(6-chloro)-1,3,5-triazine-4,2-diylimino(2-acetamido-5-methoxy)-4,1-phenylenazo)dinaphthalene-1,5-disulfonate | 403-810-6 | 140876-13-7 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
611-101-00-X | 2'-(4-chloro-3-cyano-5-formyl-2-thienyl)azo-5'-diethylaminoacetanilide | 405-200-5 | 104366-25-8 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-102-00-5 | reaction product of: C.I. Leuco Sulfur Black 1 and reaction mass of: disodium-4-{4-[8-amino-1-hydroxy-7-(4-sulfamoylphenylazo)-3,6-disulfonato-2-naphthylazo]phenylsulfonylamino}benzendiazoniumchlorid; disodium-4-{4-[2,6-dihydroxy-3-(8-hydroxy-3,6-disulfonato-1-naphthylazo)phenylazo]phenylsulfonylamino}benzen-diazoniumchlorid | 424-500-7 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
611-103-00-0 | trisodium (1-(3-carboxylato-2-oxido-5-sulfonatophenylazo)-5-hydroxy-7-sulfonatonaphthalen-2-amido)nickel(II) | 407-110-1 |  | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H318 H317 H411 |  |  |  | CLP00
611-104-00-6 | reaction mass of: trisodium (2,4(or 2,6 or 4,6)-bis(3,5-dinitro-2-oxidophenylazo)-5-hydroxyphenolato)(2(or 4or 6)-(3,5-dinitro-2-oxidophenylazo)-5-hydroxy-4(or 2or 6)-(4-(4-nitro-2-sulfonatoanilino)phenylazo)phenolato)ferrate(1-); trisodium bis(2,4(or 2,6 or 4,6)-bis(3,5-dinitro-2-oxidophenylazo)-5-hydroxyphenolato)ferrate(1-); trisodium (2,4(or 2,6 or 4,6)-bis(3,5-dinitro-2-oxidophenylazo)-5-hydroxyphenolato)(2(or 4 or 6)-(3,5-dinitro-2-oxidophenylazo)-5-hydroxy-4(or 2 or 6)-(4-nitro-2-sulfonatophenylazo)phenolato)ferrate(1-); trisodium (2,4(or 2,6 or 4,6)-bis(3,5-dinitro-2-oxidophenylazo)-5-hydroxyphenolato)(2(or 4 or 6)-(3,5-dinitro-2-oxidophenylazo)-5-hydroxy-4(or 2 or 6)-(3-sulfonatophenylazo)phenolato)ferrate(1-); disodium 3,3'-(2,4-dihydroxy-1,3(or 1,5 or 3,5)-phenylenediazo)dibenzenesulfonate | 406-870-1 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
611-105-00-1 | sodium 4-(4-chloro-6-(N-ethylanilino)-1,3,5-triazin-2-ylamino)-2-(1-(2-chlorophenyl)-5-hydroxy-3-methyl-1H-pyrazol-4-ylazo)benzenesulfonate | 407-800-2 | 136213-75-7 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
611-106-00-7 | hexasodium 4,4'-dihydroxy-3,3'-bis[2-sulfonato-4-(4-sulfonatophenylazo)phenylazo]-7,7'[p-phenylenebis[imino(6-chloro-1,3,5-triazine-4,2-diyl)imino]]dinaphthalene-2-sulfonate | 410-180-6 | 157627-99-1 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
611-107-00-2 | potassium sodium 4-(4-chloro-6-(3,6-disulfonato-7-(5,8-disulfonato-naphthalen-2-ylazo)-8-hydroxy-naphthalen-1-ylamino)-1,3,5-triazin-2-ylamino)-5-hydroxy-6-(4-(2-sulfatoethanesulfonyl)-phenylazo)-naphthalene-1,7-disulfonate | 412-490-7 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-108-00-8 | disodium 5-((4-((4-chloro-3-sulfonatophenyl)azo)-1-naphthyl)azo)-8-(phenylamino)-1-naphthalenesulfonate | 413-600-6 | 6527-62-4 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
611-109-00-3 | Reaction products of: copper(II) sulfate and tetrasodium 2,4-bis[6-(2-methoxy-5-sulfonatophenylazo)-5-hydroxy-7-sulfonato-2-naphthylamino]-6-(2-hydroxyethylamino)-1,3,5-triazine (2:1) | 407-710-3 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
611-110-00-9 | tetra-sodium/lithium 4,4'-bis-(8-amino-3,6-disulfonato-1-naphthol-2-ylazo)-3-methylazobenzene | 408-210-8 | 124605-82-9 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
611-111-00-4 | disodium 2-[[4-(2-chloroethylsulfonyl)phenyl]-[(2-hydroxy-5-sulfo-3-[3-[2-(2-(sulfooxy)ethylsulfonyl)ethylazo]-4-sulfobenzoato(3-)cuprate(1-) | 414-230-8 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-112-00-X | tetrasodium 4-hydroxy-5-[4-[3-(2-sulfatoethanesulfonyl)phenylamino]-6-morpholin-4-yl-1,3,5-triazin-2-ylamino]-3-(1-sulfonatonaphthalen-2-ylazo)naphthalene-2,7-disulfonate | 413-070-6 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-113-00-5 | lithium sodium (2-(((5-((2,5-dichlorophenyl)azo)-2-hydroxyphenyl)methylene)amino)benzoato(2-))(2-((4,5-dihydro-3-methyl-5-oxo-1-phenyl-1H-pyrazol-4-yl)azo)-5-sulfobenzoato(3-)) chromate(2-) | 414-280-0 | 149626-00-6 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
611-114-00-0 | lithium sodium (4-((5-chloro-2-hydroxyphenyl)azo)-2,4-dihydro-5-methyl-3H-pyrazol-3-onato(2-))(3-((4,5-dihydro-3-methyl-1-(4-methylphenyl)-5-oxo-1H-pyrazol-4-yl)azo)-4-hydroxy-5-nitrobenzenesulfonato(3-)) chromate(2-) | 414-250-7 | 149564-66-9 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | CLP00
611-115-00-6 | trilithium bis(4-((4-(diethylamino)-2-hydroxyphenyl)azo)-3-hydroxy-1-naphthalenesulfonato(3-))chromate(3-) | 414-290-5 | 149564-65-8 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
611-116-00-1 | reaction mass of: trisodium 5-{}{4-chloro-6-[2-(2,6-dichloro-5-cyanopyrimidin-4-ylamino)-propylamino]-1,3,5-triazin-2-ylamino}}-4-hydroxy-3-(1-sulfonatonaphthalene-2-ylazo)-naphthalene-2,7-disulfonate; trisodium 5-{}{4-chloro-6-[2-(2,6-dichloro-5-cyanopyrimidin-4-ylamino)-1-methyl-ethylamino]-1,3,5-triazin-2-ylamino}}-4-hydroxy-3-(1-sulfonatonaphthalene-2-ylazo)-naphthalene-2,7-disulfonate; trisodium 5-{}{4-chloro-6-[2-(4,6-dichloro-5-cyanopyrimidin-2-ylamino)-propylamino]-1,3,5-triazin-2-ylamino}}-4-hydroxy-3-(1-sulfonatonaphthalen-2-ylazo)-naphthalene-2,7-disulfonate; trisodium 5-{}{4-chloro-6-[2-(4,6-dichloro-5-cyanopyrimidin-2-ylamino)-1-methyl-ethylamino]-1,3,5-triazin-2-ylamino}}-4-hydroxy-3-(1-sulfonatonaphthalen-2-ylazo)-naphthalene-2,7-disulfonate | 414-620-8 |  | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
611-117-00-7 | 1,3-bis{}{6-fluoro-4-[1,5-disulfo-4-(3-aminocarbonyl-1-ethyl-6-hydroxy-4-methyl-pyrid-2-on-5-ylazo)-phenyl-2-ylamino]-1,3,5-triazin-2-ylamino}}propane lithium-, sodium salt | 415-100-3 | 149850-29-3 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-118-00-2 | sodium 1,2-bis[4-[4-{}{4-(4-sulfophenylazo)-2-sulfophenylazo}}-2-ureido-phenyl-amino]-6-fluoro-1,3,5-triazin-2-ylamino]-propane, sodium salt | 413-990-8 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
611-119-00-8 | tetrasodium 4-[4-chloro-6-(4-methyl-2-sulfophenylamino)-1,3,5-triazin-2-ylamino]-6-(4,5-dimethyl-2-sulfophenylazo)-5-hydroxynaphthalene-2,7-disulfonate | 415-400-4 | 148878-22-2 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
611-120-00-3 | 5-{}{4-[5-amino-2-[4-(2-sulfoxyethylsulfonyl)phenylazo]-4-sulfo-phenylamino]-6-chloro-1,3,5-triazin-2-ylamino}}-4-hydroxy-3-(1-sulfo-naphthalen-2-ylazo)-naphthalene-2,7-disulfonicacid sodium salt | 418-340-7 | 157707-94-3 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
611-121-00-9 | Main component 6 (isomer): asym. 1:2 Cr(III)-complex of: A: 3-hydroxy-4-(2-hydroxy-naphthalene-1-ylazo)naphthalene-1-sulfonic acid, Na-salt and B: 1-[2-hydroxy-5-(4-methoxy-phenylazo)phenylazo]naphthalene-2-ol; Main component 8 (isomer): asym. 1:2 Cr-complex of: A: 3-hydroxy-4-(2-hydroxy-naphthalene-1-ylazo)-naphthalene-1-sulfonic acid, Na-salt and B: 1-[2-hydroxy-5-(4-methoxy-phenylazo)-phenylazo]-naphthalene-2-ol | 417-280-9 | 30785-74-1 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | CLP00
611-122-00-4 | hexasodium (di[N-(3-(4-[5-(5-amino-3-methyl-1-phenylpyrazol-4-yl-azo)-2,4-disulfo-anilino]-6-chloro-1,3,5-triazin-2-ylamino)phenyl)-sulfamoyl](di-sulfo)-phthalocyaninato)nickel | 417-250-5 | 151436-99-6 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
611-123-00-X | 3-(2,4-bis(4-((5-(4,6-bis(2-aminopropylamino)-1,3,5-triazin-2-ylamino)-4-hydroxy-2,7-disulfonaphthalen-3-yl)azo)phenylamino)-1,3,5-triazin-6-ylamino)propyldiethylammonium lactate | 424-310-4 | 178452-66-9 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
611-124-00-5 | reaction mass of: pentasodium 5-amino-3-(5-{}{4-chloro-6-[4-(2-sulfoxyethoxysulfonato)phenylamino]-1,3,5-triazin-2-ylamino}}-2-sulfonatophenylazo)-6-[5-(2,3-dibromopropionylamino)-2-sulfonatophenylazo]-4-hydroxynaphthalene-2,7-disulfonate; pentasodium 5-amino-6-[5-(2-bromoacryloylamino)-2-sulfonatophenylazo]-3-(5-{}{4-chloro-6-[4-(2-sulfoxyethoxysulfonato)phenylamino]-1,3,5-triazin-2-ylamino}}-2-sulfonatophenylazo)-4-hydroxynaphthalene-2,7-disulfonate; tetrasodium 5-amino-3-[5-{}{4-chloro-6-[4-(vinylsulfonyl)phenylamino]-1,3,5-triazin-2-ylamino}}-2-sulfonatophenylazo]-6-[5-(2,3-dibromopropionylamino)-2-sulfonatophenylazo]-4-hydroxynaphthalene-2,7-disulfonate | 424-320-9 |  | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
611-125-00-0 | reaction mass of: Disodium 6-[3-carboxy-4,5-dihydro-5-oxo-4-sulfonatophenyl)pyrazolin-4-yl-azo]-3-[2-oxido-4-(ethensulfonyl)-5-methoxyphenylazo]-4-oxidonaphthalene-2-sulfonate copper (II) complex; Disodium 6-[3-carboxy-4,5-dihydro-5-oxo-4-sulfonatophenyl)pyrazolin-4-yl-azo]-3-[2-oxido-4-(2-hydroxyethylsulfonyl)-5-methoxyphenylazo]-4-oxidonaphthalene-2-sulfonate copper (II) complex | 423-940-7 |  | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
611-126-00-6 | 2,6-bis-(2-(4-(4-amino-phenylamino)-phenylazo)-1,3-dimethyl-3H-imidazolium)-4-dimethylamino-1,3,5-triazine, dichloride | 424-120-1 | 174514-06-8 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | CLP00
611-127-00-1 | pentasodium 4-amino-6-(5-(4-(2-ethyl-phenylamino)-6-(2-sulfatoethanesulfonyl)-1,3,5-triazin-2-ylamino)-2-sulfonatophenylazo)-5-hydroxy-3-(4-(2-sulfatoethanesulfonyl)phenylazo)naphthalene-2,7-disulfonate | 423-790-2 |  | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  | G | CLP00
611-128-00-7 | N,N'-bis{}{6-chloro-4-[6-(4-vinylsulfonylphenylazo)-2,7-disulfonicacid-5-hydroxynapht-4-ylamino]-1,3,5-triazin-2-yl}}-N-(2-hydroxyethyl)ethane-1,2-diamine, sodium salt | 419-500-9 | 171599-85-2 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
611-129-00-2 | reaction mass of: 5-[(4-[(7-amino-1-hydroxy-3-sulfo-2-naphthyl)azo]-2,5-diethoxyphenyl)azo]-2-[(3-phosphonophenyl)azo]benzoic acid; 5-[(4-[(7-amino-1-hydroxy-3-sulfo-2-naphthyl)azo]-2,5-diethoxyphenyl)azo]-3-[(3-phosphonophenyl)azo]benzoic acid | 418-230-9 | 163879-69-4 | Expl. 1.3 **** Repr. 2 STOT RE 2 * Skin Sens. 1 Aquatic Chronic 2 | H203 H361f *** H373 ** H317 H411 | GHS01 GHS08 GHS07 GHS09 Dgr | H203 H361f *** H373 ** H317 H411 |  |  |  | CLP00
611-130-00-8 | tetra-ammonium 2-[6-[7-(2-carboxylato-phenylazo)-8-hydroxy-3,6-disulfonato-1-naphthylamino]-4-hydroxy-1,3,5-triazin-2-ylamino]benzoate | 418-520-5 | 183130-96-3 | Eye Irrit. 2 Aquatic Chronic 3 | H319 H412 | GHS07 Wng | H319 H412 |  |  |  | CLP00/ATP01
611-131-00-3 | 2-[2-hydroxy-3-(2-chlorophenyl)carbamoyl-1-naphthylazo]-7-[2-hydroxy-3-(3-methylphenyl)carbamoyl-1-naphthylazo]fluoren-9-one | 420-580-2 | 151798-26-4 | Repr. 1B Aquatic Chronic 4 | H360D *** H413 | GHS08 Dgr | H360D *** H413 |  |  |  | CLP00
611-132-00-9 | pentasodium bis{}{7-[4-(1-butyl-5-cyano-1,2-dihydro-2-hydroxy-4-methyl-6-oxo-3-pyridylazo)phenylsulfonylamino]-5'-nitro-3,3'-disulfonatonaphthalene-2-azobenzene-1,2'-diolato}} chromate (III) | 419-210-2 | 178452-71-6 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
611-133-00-4 | Product by process iron complex of azo dyestuffs obtained by coupling a mixture of diazotized 2-amino-1-hydroxybenzene-4-sulfanilide and 2-amino-1-hydroxybenzene-4-sulfonamide with resorcin, the obtained mixture being subsequently submitted to a second coupling reaction with a mixture of diazotized 3-aminobenzene-1-sulfonic acid (metanilic acid) and 4'-amino-4-nitro-1,1'-diphenylamine-2-sulfonic acid and metallization with ferric chloride, sodium salt | 419-260-5 |  | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
611-134-00-X | trisodium 2-{}{α[2-hydroxy-3-[4-chloro-6-[4-(2,3-dibromopropionylamino)-2-sulfonatophenylamino]-1,3,5-triazin-2-ylamino]-5-sulfonatophenylazo]-benzylidenehydrazino}}-4-sulfonatobenzoate, copper complex | 423-770-3 |  | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | CLP00
611-135-00-5 | Reaction product of: 2-[[4-amino-2-ureidophenylazo]-5-[(2-(sulfooxy)ethyl)sulfonyl]]benzenesulfonic acid with 2,4,6-trifluoropyrimidine and partial hydrolysis to the corresponding vinylsulfonyl derivative,mixed potassium/sodium salt | 424-250-9 |  | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | CLP00
611-136-00-0 | 2-{}{4-(2-ammoniopropylamino)-6-[4-hydroxy-3-(5-methyl-2-methoxy-4-sulfamoylphenylazo)-2-sulfonatonaphth-7-ylamino]-1,3,5-triazin-2-ylamino}}-2-aminopropyl formate | 424-260-3 |  | Repr. 2 Eye Dam. 1 Aquatic Chronic 2 | H361f *** H318 H411 | GHS05 GHS08 GHS09 Dgr | H361f *** H318 H411 |  |  |  | CLP00
611-137-00-6 | 6-tert-butyl-7-chloro-3-tridecyl-7,7a-dihydro-1H-pyrazolo[5,1-c]-1,2,4-triazole | 419-870-1 | 159038-16-1 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
611-138-00-1 | 2-(4-aminophenyl)-6-tert-butyl-1H-pyrazolo[1,5-b][1,2,4]triazole | 415-910-7 | 152828-25-6 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
611-139-00-7 | reaction product of: C.I. Leuco Sulfur Black 1 with (3-chloro-2-hydroxypropyl)trimethylammonium chloride | 424-510-1 | - | Eye Dam. 1 Aquatic Chronic 2 | H318 H411 | GHS05 GHS09 Dgr | H318 H411 |  |  |  | ATP01
611-140-00-2 | azafenidin (ISO); 2-(2,4-dichloro-5-prop-2-ynyloxyphenyl)-5,6,7,8-tetrahydro-1,2,4-triazolo[4,3-a]pyridin-3(2H)-one |  | 68049-83-2 | Repr. 1B STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H360Df H373 ** H400 H410 | GHS08 GHS09 Dgr | H360Df H373 ** H410 |  | M=1000 |  | CLP00
611-141-00-8 | 5-(4-[4-[4-(3,5-dicarboxy-phenyl-azo)phenylamino]-6-morpholin-4-yl-1,3,5-triazin-2-ylamino]phenylazo)isophthalic acid, mixed monosodium and diammonium salt | 414-410-6 | - | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
611-142-00-3 | product-by-process definition polyazodyestuff obtained by coupling 4-[4-(1-amino-8-hydroxy-3,6-disulfo-2-naphthylazo)phenylsulfonylamino]benzenediazonium with reaction mass of 4-carboxybenzenediazonium and diphenylamine-3-sulfo-4,4'-bisdiazonium, and further coupling of the obtained compounds with reaction mass of naphth-2-ol and 3-aminophenol, sodium salts; sodium chloride | 425-740-5 | - | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
611-143-00-9 | reaction mass of: trisodium 2-(2-[α-(2-carboxylato-κ-O-4-sulfonatophenylazo)benzylidene]hydrazino-κ-N')-6-(2,6-difluoropyrimidin-4-ylamino)-4-sulfonatophenolatocuprate (II); trisodium 2-(2-[α-(2-carboxylato-κ-O-4-sulfonatophenylazo)benzylidene]hydrazino-κ-N')-6-(4,6-difluoropyrimidin-2-ylamino)-4-sulfonatophenolatocuprate (II) | 428-260-4 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
611-144-00-4 | reaction mass of: 7-amino-3,8-bis-[4-(2-sulfoxyethylsulfonyl)phenylazo]-4-hydroxynaphthalene-2-sulfonic acid, Na/K salt; 7-amino-3-[4-(2-sulfoxyethylsulfonyl)phenylazo]-4-hydroxy-8-[4-(2-sulfoxyethylsulfonyl)-2-sulfophenylazo]naphthalene-2-sulfonic acid, Na/K salt; 7-amino-8-[4-(2-sulfoxyethylsulfonyl)-phenylazo]-4-hydroxy-3-[4-(2-sulfoxyethylsulfonyl)-2-sulfophenylazo]naphthalene-2-sulfonic acid, Na/K salt; 7-amino-3,8-bis-[4-(2-sulfoxyethylsulfonyl)-2-sulfophenylazo]-4-hydroxynaphthalene-2-sulfonic acid, Na/K salt | 429-070-4 | 214362-06-8 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
611-145-00-X | reaction mass of: tetrasodium 3-(1,5-disulfonatonaphthalene-2-ylazo)-4-hydroxy-7-{4-chloro-6-[4-(2-sulfoxyethylsulfonyl)phenylamino]-1,3,5-triazine-2-ylamino}naphthalene-2-sulfonate; 3-(2,5-disulfophenylazo)-4-hydroxy-7-{4-chloro-6-[4-(2-sulfoxyethylsulfonyl)phenylamino]-1,3,5-triazine-2-ylamino}naphthalene-2-sulfonic acid, sodium salt | 429-440-5 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
611-146-00-5 | reaction mass of: pentasodium 3-(4-(4-(7-(2,4-diamino-5-sulfonato-3-(4-sulfonatophenylazo)phenylazo)-1-hydroxy-3-sulfonatonaphthalen-2-ylazo)-2-sulfonatophenylamino)phenylazo)-4-hydroxy-6-(2-oxo-1-phenylcarbamoylpropylazo)naphthalene-2-sulfonate; pentasodium 6-((2,4-diamino-5-sulfonatophenyl)azo)-3-((4-((4-((7-((2,4-diamino-5-sulfonatophenyl)azo)-1-hydroxy-3-sulfonatonaphthalen-2-yl)azo)phenyl)amino)-2-sulfonatophenyl)azo)-4-hydroxynaphthalene-2-sulfonate; pentasodium 6-((2,4-diamino-5-sulfonato-3-((4-sulfonatophenyl)azo)phenyl)azo)-3-((4-((4-((1,7-dihydroxy-3-sulfonatonaphthalen-2-yl)azo)-2-sulfonatophenyl)amino)phenyl)azo)-4-hydroxynaphthalene-2-sulfonate; hexasodium 6-((2,4-diamino-5-sulfonatophenyl)azo)-3-((4-((4-((7-((2,4-diamino-5-sulfonato-3-((4-sulfonatophenyl)azo)phenyl)azo)-1-hydroxy-3-sulfonatonaphthalen-2-yl)azo)-2-sulfonatophenyl)amino)phenyl)azo)-4-hydroxynaphthalene-2-sulfonate | 430-070-1 | - | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
611-147-00-0 | sodium, potassium, lithium 5-amino-3,6-bis(5-(4-chloro-6-(methyl-(2-methylaminoacetyl)amino)-1,3,5-triazin-2-ylamino)-2-sulfonatophenylazo)-4-hydroxynaphthalene-2,7-disulfonate | 430-090-0 | 205764-96-1 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
611-148-00-6 | reaction mass of: 2-(3-(2,6-dichloro-4-nitrophenylazo)carbazol-9-yl)ethanol; 2-(2-(3-(2,6-dichloro-4-nitro-phenylazo)-carbazol-9-yl)-ethoxy)ethanol; 3-(2,6-dichloro-4-nitrophenylazo)carbazol | 429-590-1 | - | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | ATP01
611-149-00-1 | 2-(2-chloroacetoxy)ethyl 3-((4-(2,5-dichloro-4-fluorosulfonylphenylazo)-3-methylphenyl)ethylamino)propionate | 427-570-7 | 193486-83-8 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
611-150-00-7 | tetralithium 2-[6-[7-[2-(carboxylato)phenylazo]-8-hydroxy-3,6-disulfonato-1-naphthylamino]-4-hydroxy-1,3,5-triazine-2-ylamino]benzoate | 440-460-3 | - | Eye Irrit. 2 Aquatic Chronic 3 | H319 H412 | GHS07 Wng | H319 H412 |  |  |  | ATP01
611-151-00-2 | chrysoidine; 4-(phenylazo)benzene-1,3-diamine | 207-803-7 | 495-54-5 | Muta. 2 Acute Tox. 4 * Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H341 H302 H315 H400 H410 | GHS08 GHS07 GHS09 Wng | H341 H302 H315 H410 |  |  |  | ATP01
611-152-00-8 | chrysoidine monohydrochloride; 4-phenylazophenylene-1,3-diamine monohydrochloride [1] chrysoidine monoacetate; 4-(phenylazo)benzene-1,3-diamine monoacetate [2] chrysoidine acetate; 4-(phenylazo)benzene-1,3-diamine acetate [3] chrysoidine-p-dodecylbenzenesulfonate; dodecylbenzenesulfonic acid, compound with 4-(phenylazo)benzene-1,3-diamine (1:1) [4] chrysoidine dihydrochloride; 4-(phenylazo)benzene-1,3-diamine dihydrochloride [5] chrysoidine sulfate; bis[4-(phenylazo)benzene-1,3-diamine] sulfate [6] | 208-545-8 [1] 278-290-5 [2] 279-116-0 [3] 264-409-8 [4] 281-549-5 [5] 282-432-1 [6] | 532-82-1 [1] 75660-25-2 [2] 79234-33-6 [3] 63681-54-9 [4] 83968-67-6 [5] 84196-22-5 [6] | Muta. 2 Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H341 H302 H315 H318 H400 H410 | GHS05 GHS08 GHS07 GHS09 Dgr | H341 H302 H315 H318 H410 |  |  |  | ATP01
611-153-00-3 | chrysoidine C10-14-alkyl derivatives; benzenesulfonic acid, mono-C10-14-alkyl derivatives, compounds with 4-(phenylazo)-1,3-benzenediamine [1] chrysoidine compound with dibutylnaphthalene sulfonic acid; dibutylnaphthalenesulfonic acid, compound with 4-(phenylazo)benzene-1,3-diamine (1:1) [2] | 286-946-7 [1] 304-236-8 [2] | 85407-90-5 [1] 94247-67-3 [2] | Muta. 2 Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 | H341 H302 H315 H318 | GHS05 GHS08 GHS07 Dgr | H341 H302 H315 H318 |  |  |  | ATP01
611-154-00-9 | trisodium 5-benzamido-4-hydroxy-3-(4-methyl-2-sulfonatophenylazo)naphthalene-2,7-disulfonate | 403-670-6 | 92408-46-3 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
611-155-00-4 | 4,4'-oxybis(benzenesulfonylazide) | 431-850-4 | 7456-68-0 | Expl. 1.1 **** STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H201 H373 ** H400 H410 | GHS01 GHS08 GHS09 Dgr | H201 H373 ** H410 |  |  |  | ATP01
611-156-00-X | triammonium 4-[4-[7-(4-carboxylatoanilino)-1-hydroxy-3-sulfonato-2-naphthylazo]-2,5-dimethoxyphenylazo]benzoate | 432-270-4 | 221354-37-6 | Repr. 2 STOT RE 2 * Aquatic Chronic 2 | H361f *** H373 ** H411 | GHS08 GHS09 Wng | H361f *** H373 ** H411 |  |  |  | ATP01
611-157-00-5 | benzenesulfonic acid, 3,3'-(methylenebis((dihydroxyphenylene)azo))bis-, potassium sodium salt; potassium sodium 3-[(E)-(6-{3,4-dihydroxy-2-[(Z)-(3-sulfonatophenyl)diazenyl]benzyl}-2,3-dihydroxyphenyl)diazenyl]benzenesulfonate | 432-590-4 | 243869-48-9 | Eye Irrit. 2 Aquatic Chronic 3 | H319 H412 | GHS07 Wng | H319 H412 |  |  |  | ATP01
611-158-00-0 | reaction product of: 2,3,4,2',3',4'-hexahydroxy-5,5'-diacethyl-diphenylmethane and 6-diazo-5,6-dihydro-5-oxo-1-naphthalenesulfonylchloride and 3-diazo-3,4-dihydro-6-methoxy-4-oxo-1-naphthalenesulfonylchloride | 421-520-8 | - | **** Aquatic Chronic 4 | **** H413 | **** | **** H413 |  |  |  | ATP01
611-160-00-1 | reaction mass of: 1,1,1-tris(phenyl-4'-(3''-diazo-3'',4''-dihydro-4''-oxo-naphthalene-1''-sulfonato)ethane; 1,1,1-tris(phenyl-4'-(6''-diazo-5'',6''-dihydro-5''-oxo-naphthalene-1''-sulfonato)ethane; reaction product of 1,1,1-tris(p-hydroxyphenyl)ethane with 6-diazo-5,6-dihydro-5-oxo-1-naphthylsulfonylchloride and 3-diazo-3,4-dihydro-4-oxo-1-naphthylsulfonylchloride (2:1); reaction product of 1,1,1-tris(p-hydroxyphenyl)ethane with 6-diazo-5,6-dihydro-5-oxo-1-naphthylsulfonylchloride and 3-diazo-3,4-dihydro-4-oxo-1-naphthylsulfonylchloride (1:2) | 422-760-6 | - | **** Aquatic Chronic 4 | **** H413 | **** | **** H413 |  |  |  | ATP01
611-161-00-7 | trisodium [1,2'-(2-(8-amino-3,5-disulfonatonaphthalene)azo)-(4'-nitrobenzene)diolato-O,O,N][(Z)-2,2-((phenylcarbamoylprop-1'-enyl)azo)-5-sulfamoylbenzene)diolato-O,O,N]chromate(III) | 423-100-1 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
611-162-00-2 | 2,4-bis(((2-(dimethylammonio)ethyloxy)carbonyl)phen-2-ylazo)benzene-1,3-diolbis(methanesulfonate) | 429-600-4 | - | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | ATP01
611-163-00-8 | 2,4-bis(((2-(dimethylammonio)ethyloxy)carbonyl)phen-2-ylazo)benzene-1,3-diol sulfate | 429-610-9 | - | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 2 | H302 H318 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H411 |  |  |  | ATP01
611-164-00-3 | reaction mass of: 2,2’-dimethyl-2,2’-azobutanenitrile; 2-methylpentanenitrile-2-azo-2’-(2’-methylpropanenitrile); 2,2’-dimethyl-2,2’-azoheptanenitrile; 2-methylheptanenitrile-2-azo-2’-(2’-methylpropanenitrile); 2-methylheptanenitrile-2-azo-2’-(2’-methylbutanenitrile) | 429-710-2 |  | Self-react. D Acute Tox. 4 * Aquatic Chronic 2 | H242 H302 H411 | GHS02 GHS07 GHS09 Dgr | H242 H302 H411 |  |  |  | ATP01/ATP01corr
611-165-00-9 | reaction mass of: tetrasodium 4-amino-6-(5-(2,6-difluoropyrimidin-4-ylamino)-2-sulfonatophenylazo)-5-hydroxy-3-(4-(sulfatoethylsulfonyl)phenylazo)naphthalene-2,7-disulfonate; tetrasodium 4-amino-6-(5-(4,6-difluoropyrimidin-2-ylamino)-2-sulfonatophenylazo)-5-hydroxy-3-(4-(2-sulfatoethylsulfonyl)phenylazo)naphthalene-2,7-disulfonate | 431-830-5 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
611-166-00-4 | reaction mass of: pentasodium 4-amino-5-hydroxy-3-{(E)-4-[2-(sulfonatooxy)ethylsulfonyl]phenylazo}-6-{(E)-2-sulfonato-4-[2-(sulfonatooxy)ethylsulfonyl]phenylazo}naphthalene-2,7-disulfonate; tetrasodium 4-amino-5-hydroxy-3-{(E)-4-[2-(sulfonatooxy)ethylsulfonyl]phenylazo}-6-[(E)-2-sulfonato-4-(vinylsulfonyl)phenylazo]naphthalene-2,7-disulfonate; tetrasodium 4-amino-5-hydroxy-6-{(E)-2-sulfonato-4-[2-(sulfonatooxy)ethylsulfonyl]phenylazo}-3-[(E)-4-(vinylsulfonyl)phenylazo]naphthalene-2,7-disulfonate | 432-100-9 | - | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
611-167-00-X | sodium bis[tris(2-hydroxyethyl)ammonium][6-anilino-4'-(4,8-disulfonato-2-naphthylazo)-5'-methyl-3-sulfonatonaphthalene-2-azobenzene-1,2'-diolato]cuprate(II) | 435-240-9 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
611-168-00-5 | reaction mass of: 3-[[4-chloro-6-[[7-[(1,5-disulfo-2-naphthalenyl)azo]-8-hydroxy-3,6-disulfo-1-naphthalenyl]amino]-1,3,5-triazin-2-yl]amino]-5-[[4-chloro-6-[[8-hydroxy-3,6-disulfo-7-[(2-sulfophenyl)azo]-1-naphthalenyl]amino]-1,3,5-triazin-2-yl]amino]benzoic acid; 3,5-bis[[4-chloro-6-[[7-[(1,5-disulfo-2-naphthalenyl)azo]-8-hydroxy-3,6-disulfo-1-naphthalenyl]amino]-1,3,5-triazin-2-yl]amino]benzoic acid | 435-440-6 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
611-169-00-0 | sodium 5-(2-carboxyphenylazo)-6-hydroxynaphthalene-2-sulfonate | 435-800-2 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
611-170-00-6 | reaction mass of: trisodium 2-((1-(2-hydroxy-κ-O-5-(2-sulfonatoethansulfonyl)phenylazo-κ-N2)-1-phenylmethyl)azo-κ-N1)-4-sulfonatobenzoate(5-)-κ-O)cuprate(II); disodium 2-((1-(5-ethenesulfonyl-2-hydroxy-κ-O-phenylazo-κ-N2)-1-phenylmethyl)azo-κ-N1)-4-sulfonatobenzoate-κ-O-(5-))cuprate(II) | 435-880-9 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
611-171-00-1 | reaction mass of: trisodium 3-(5-(2,6-difluoropyrimidin-4-ylamino)-2-sulfonatophenylazo)-5-(4-fluoro-6-morpholin-4-yl-1,3,5-triazin-2-ylamino)-4-hydroxy-2,7-naphthalenedisulfonate; trisodium 3-(5-(4,6-difluoropyrimidin-2-ylamino)-2-sulfonatophenylazo)-5-(4-fluoro-6-morpholin-4-yl-1,3,5-triazin-2-ylamino)-4-hydroxy-2,7-naphthalenedisulfonate | 436-890-6 | - | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
611-172-00-7 | reaction mass of: triammonium 6-amino-3-((2,5-diethoxy-4-(3-phosphonophenyl)azo)phenyl)azo-4-hydroxy-2-naphthalenesulfonate; diammonium 3-((4-((7-amino-1-hydroxy-3-sulfo-naphthalen-2-yl)azo)-2,5-diethoxyphenyl)azo)benzoate | 438-310-7 | - | Self-react. C **** Repr. 2 Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 3 | H242 H361f *** H302 H373 ** H412 | GHS02 GHS08 GHS07 Dgr | H242 H361f *** H302 H373 ** H412 |  |  |  | ATP01
611-173-00-2 | reaction mass of: 3-[3-carbamoyl-5-(5-{4-chloro-6-[4-(2-sulfonatooxyethylsulfonyl)anilino]-1,3,5-triazin-2-ylamino}-2-sulfonatophenylazo)-1,2-dihydro-6-hydroxy-4-methyl-2-oxo-1-pyridyl]propanoic acid, trisodium salt; 3-[3-carbamoyl-5-(5-{4-chloro-6-[4-(vinylsulfonyl)anilino]-1,3,5-triazin-2-ylamino}-2-sulfonatophenylazo)-1,2-dihydro-6-hydroxy-4-methyl-2-oxo-1-pyridyl]propanoic acid, disodium salt | 440-510-4 | - | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
611-174-00-8 | reaction mass of: 3-[5-(4-ethenesulfonylbutyrylamino)-2-sulfophenylazo]-5-{4-chloro-[6-(4-(3-amino-5-hydroxy-2,7-disulfonaphthalene-4-ylazo)-3-sulfophenylamino]-1,3,5-triazin-2-ylamino}-4-hydroxynaphthalene-2,7-disulfonic acid, sodium salt; 3-[5-(4-(2-chloroethanesulfonyl)butyrylamino)-2-sulfophenylazo]-5-{4-chloro-[6-(4-(3-amino-5-hydroxy-2,7-disulfonaphthalene-4-ylazo)-3-sulfophenylamino]-1,3,5-triazin-2-ylamino}-4-hydroxynaphthalene-2,7-disulfonic acid, sodium salt | 442-290-5 | 457624-86-1 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
611-175-00-3 | reaction mass of: trisodium 5-{4-chloro-6-[N-ethyl-(3-(2-sulfonatooxy)ethylsulfonyl)anilino]-1,3,5-triazin-2-ylamino}-4-hydroxy-3-[4-(vinylsulfonyl)phenylazo]naphthalene-2,7-disulfonate; trisodium 5-{4-chloro-6-[N-ethyl-3-(vinylsulfonyl)anilino]-1,3,5-triazin-2-ylamino}-4-hydroxy-3-[4-(2-(sulfonatooxy)ethylsulfonyl)phenylazo]naphthalene-2,7-disulfonate; disodium 5-{4-chloro-6-[N-ethyl-3-(vinylsulfonyl)anilino]-1,3,5-triazin-2-ylamino}-4-hydroxy-3-[(4-vinylsulfonyl)phenylazo]naphthalene-2,7-disulfonate; tetrasodium 5-{4-chloro-6-[N-ethyl-3-(2-(sulfonatooxy)ethylsulfonyl)anilino]-1,3,5-triazin-2-ylamino}-3-[4-(2-(sulfonatooxy)ethylsulfonyl)phenylazo]-4-hydroxynaphthalene-2,7-disulfonate | 444-050-5 | - | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
611-176-00-9 | 2,6-bis(2,3,4-trihydroxybenzyl)-p-cresol ester with 6-diazo-5,6-dihydro-5-oxo-1-naphthalenesulfonate | 444-250-2 | - | Self-react. C **** Aquatic Chronic 2 | H242 H411 | GHS02 GHS09 Dgr | H242 H411 |  |  |  | ATP01
611-177-00-4 | reaction mass of: pentasodium bis[6-anilino-3,5'-disulfonatonaphthalene-2-azobenzene-1,2'-diolato]cobaltate(III); tetrasodium [6-anilino-3,5'-disulfonatonaphthalene-2-azobenzene-1,2'-diolato][6-anilino-5'-sulfamoyl-3-sulfonatonaphthalene-2-azobenzene-1,2'-diolato]cobaltate(III); trisodium bis[6-anilino-5'-sulfamoyl-3-sulfonatonaphthalene-2-azobenzene-1,2'-diolato]cobaltate(III) | 444-290-0 | 508202-43-5 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | ATP01
611-178-00-X | reaction mass of: pentasodium 4-amino-5-hydroxy-3-{(E)-4-[2-(sulfonatooxy)ethylsulfonyl]phenylazo}-6-{(E)-2-sulfonato-4-[2-(sulfonatooxy)ethylsulfonyl]phenylazo}naphthalene-2,7-disulfonate; tetrasodium 4-amino-5-hydroxy-3-{(E)-4-[2-(sulfonatooxy)ethylsulfonyl]phenylazo}-6-[(E)-2-sulfonato-4-(vinylsulfonyl)phenylazo]naphthalene-2,7-disulfonate; tetrasodium 4-amino-5-hydroxy-6-[(E)-2-sulfonato-4-[2-(sulfonatooxy)ethylsulfonyl]phenylazo}-3-[(E)-4-(vinylsulfonyl)phenylazo]naphthalene-2,7-disulfonate; trisodium 4-amino-5-hydroxy-3-[(E)-4-(vinylsulfonyl)phenylazo]-6-[(E)-2-sulfonato-4-(vinylsulfonyl)phenylazo]naphthalene-2,7-disulfonate; trisodium 4-amino-5-hydroxy-3-[(2-hydroxyethylsulfonyl)-phenylazo]-6-[(E)-2-sulfonato-4-(vinylsulfonyl)phenylazo]naphthalene-2,7-disulfonate; trisodium 4-amino-5-hydroxy-3-[(E)-4-(vinylsulfonyl)phenylazo]-6-[-2-sulfonato-4-(2-hydroxyethylsulfonyl)phenylazo]naphthalene-2,7-disulfonate | 445-280-9 | - | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | ATP01
611-179-00-5 | reaction mass of: pentasodium 2-[[8-[[4-chloro-6-[[4-(2-sulfonato ethylsulfonyl)]phenyl]amino]-1,3,5-triazin-2-yl]amino-1- hydroxy-3,6-disulfonato-2-naphthalenyl]azo]naphthalene-1,5-disulfonate; 2-[[8-[[4-chloro-6-[[4-[[2-ethenyl]sulfonyl]phenyl]amino]-1,3,5-triazin-2-yl]amino]-1-hydroxy-3,6-disulfonato-2-naphthalenyl]azo]naphthalene-1,5-disulfonate | 450-010-8 | - | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
611-180-00-0 | iron, complexes with diazotised 4-aminobenzenesulfonamide,diazotised 3-aminobenzenesulfonic acid, diazotised 3-amino-4-hydroxybenzenesulfonamide,diazotised 3-amino-4-hydroxy-N-phenylbenzenesulfonamide, diazotised 5-amino-2-(phenylamino)benzenesulfonic acid and resorcinol, sodium salts | 417-850-7 | - | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
611-181-00-6 | potassium (oxido-N­NO-azoxy)cyclohexane; cyclohexylhydroxydia­zene 1-oxide, potassium salt; [K-HDO] |  | 66603-10-9 | Flam. Sol. 1 Acute Tox. 3 STOT RE 2 Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 2 | H228 H301 H373 (liver) H315 H318 H411 | GHS02 GHS06 GHS08 GHS05 GHS09 Dgr | H228 H301 H315 H318 H373 (liver) H411 |  | Oral: ATE = 136 mg/kg bw |  | ATP15
611-182-00-1 | 2-[ethyl[3-methyl-4-[(5-nitrothiazol-2-yl)azo]phenyl]amino]ethanol | 271-183-4 | 68516-81-4 | Skin Sens. 1A | H317 | GHS07 Wng | H317 |  | Skin Sens. 1A; H317: C ≥ 0,001 % |  | ATP22
612-001-00-9 | mono-methylamine [1] di-methylamine [2] tri-methylamine [3] | 200-820-0 [1] 204-697-4 [2] 200-875-0 [3] | 74-89-5 [1] 124-40-3 [2] 75-50-3 [3] | Flam. Gas 1 Press. Gas Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 | H220 H332 H335 H315 H318 | GHS02 GHS04 GHS05 GHS07 Dgr | H220 H332 H335 H315 H318 |  | * Skin Irrit. 2; H315: C ≥ 5 % Eye Dam. 1; H318: C ≥ 5 % Eye Irrit. 2; H319: 0,5 % ≤ C < 5 % STOT SE 3; H335: C ≥ 5 % | 5 U | CLP00
612-001-01-6 | mono-methylamine ... % [1] di-methylamine ... % [2] tri-methylamine ... % [3] | 200-820-0 [1] 204-697-4 [2] 200-875-0 [3] | 74-89-5 [1] 124-40-3 [2] 75-50-3 [3] | Flam. Liq. 1 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H224 H332 H302 H314 | GHS02 GHS05 GHS07 Dgr | H224 H332 H302 H314 |  | * STOT SE 3; H335: C ≥ 5 % | B | CLP00
612-002-00-4 | ethylamine | 200-834-7 | 75-04-7 | Flam. Gas 1 Press. Gas STOT SE 3 Eye Irrit. 2 | H220 H335 H319 | GHS02 GHS04 GHS07 Dgr | H220 H319 H335 |  |  | U | CLP00
612-003-00-X | diethylamine | 203-716-3 | 109-89-7 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A | H225 H332 H312 H302 H314 | GHS02 GHS05 GHS07 Dgr | H225 H332 H312 H302 H314 |  | STOT SE 3; H335: C ≥ 1 % |  | CLP00
612-004-00-5 | triethylamine | 204-469-4 | 121-44-8 | Flam. Liq. 2 Acute Tox. 3 Acute Tox. 3 Acute Tox. 3 Skin Corr. 1A Eye Dam. 1 | H225 H331 H311 H301 H314 H318 | GHS02 GHS06 GHS05 Dgr | H225 H331 H311 H301 H314 |  | inhalation: ATE = 7.2 mg/L(vapours) oral: ATE = 100 mg/kg bw dermal: ATE = 300 mg/kg bw STOT SE 3; H335: C >= 1 % |  | CLP00/ATP21
612-005-00-0 | butylamine | 203-699-2 | 109-73-9 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A | H225 H332 H312 H302 H314 | GHS02 GHS05 GHS07 Dgr | H225 H332 H312 H302 H314 |  | STOT SE 3; H335: C ≥ 1 % |  | CLP00
612-006-00-6 | ethylenediamine; 1,2-diaminoethane | 203-468-6 | 107-15-3 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1 | H226 H312 H302 H314 H334 H317 | GHS02 GHS08 GHS05 GHS07 Dgr | H226 H312 H302 H314 H334 H317 |  |  |  | CLP00
612-007-00-1 | 2-aminopropane; isopropylamine | 200-860-9 | 75-31-0 | Flam. Liq. 1 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H224 H335 H315 H319 | GHS02 GHS07 Dgr | H224 H319 H335 H315 |  |  |  | CLP00
612-008-00-7 | aniline | 200-539-3 | 62-53-3 | Carc. 2 Muta. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 | H351 H341 H331 H311 H301 H372 ** H318 H317 H400 | GHS06 GHS08 GHS05 GHS09 Dgr | H351 H341 H331 H311 H301 H372 ** H318 H317 H400 |  | * STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,2 % ≤ C < 1 % |  | CLP00
612-009-00-2 | salts of aniline |  |  | Carc. 2 Muta. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 | H351 H341 H331 H311 H301 H372 ** H318 H317 H400 | GHS06 GHS08 GHS05 GHS09 Dgr | H351 H341 H331 H311 H301 H372 ** H318 H317 H400 |  | * STOT RE 1; H372: C ≥ 1 % STOT RE 2; H373: 0,2 % ≤ C < 1 % | A | CLP00
612-010-00-8 | chloroanilines, with exception of those specified elsewhere in this Annex |  |  | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H410 |  |  | C | CLP00
612-011-00-3 | 4-nitrosoaniline | 211-535-6 | 659-49-4 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H332 H312 H302 | GHS07 Wng | H332 H312 H302 |  |  |  | CLP00
612-012-00-9 | o-nitroaniline [1] m-nitroaniline [2] p-nitroaniline [3] | 201-855-4 [1] 202-729-1 [2] 202-810-1 [3] | 88-74-4 [1] 99-09-2 [2] 100-01-6 [3] | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 3 | H331 H311 H301 H373 ** H412 | GHS06 GHS08 Dgr | H331 H311 H301 H373 ** H412 |  |  | C | CLP00
612-013-00-4 | 3-aminobenzene sulphonic acid; metanilic acid | 204-473-6 | 121-47-1 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H332 H312 H302 | GHS07 Wng | H332 H312 H302 |  |  |  | CLP00
612-014-00-X | sulphanilic acid; 4-aminobenzenesulphonic acid | 204-482-5 | 121-57-3 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H315 H319 H317 | GHS07 Wng | H319 H315 H317 |  |  |  | CLP00
612-015-00-5 | N-methylaniline | 202-870-9 | 100-61-8 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H410 |  |  |  | CLP00
612-016-00-0 | N,N-dimethylaniline | 204-493-5 | 121-69-7 | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Chronic 2 | H351 H331 H311 H301 H411 | GHS06 GHS08 GHS09 Dgr | H351 H331 H311 H301 H411 |  |  |  | CLP00
612-017-00-6 | N-methyl-N-2,4,6-tetranitroaniline; tetryl | 207-531-9 | 479-45-8 | Expl. 1.1 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 | H201 H331 H311 H301 H373 ** | GHS01 GHS06 GHS08 Dgr | H201 H331 H311 H301 H373 ** |  |  |  | CLP00/ATP01
612-018-00-1 | bis(2,4,6-trinitrophenyl)amine; hexyl | 205-037-8 | 131-73-7 | Expl. 1.1 Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 Aquatic Chronic 2 | H201 H310 H330 H300 H373 ** H411 | GHS01 GHS06 GHS08 GHS09 Dgr | H201 H330 H310 H300 H373 ** H411 |  |  |  | CLP00/ATP01
612-019-00-7 | dipicrylamine, ammonium salt | 220-639-0 | 2844-92-0 | Expl. 1.1 Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 Aquatic Chronic 2 | H201 H310 H330 H300 H373 ** H411 | GHS01 GHS06 GHS08 GHS09 Dgr | H201 H330 H310 H300 H373 ** H411 |  |  |  | CLP00/ATP01
612-020-00-2 | 1-naphthylamine | 205-138-7 | 134-32-7 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
612-022-00-3 | 2-naphthylamine | 202-080-4 | 91-59-8 | Carc. 1A Acute Tox. 4 * Aquatic Chronic 2 | H350 H302 H411 | GHS08 GHS07 GHS09 Dgr | H350 H302 H411 |  | Carc. 1A; H350: C ≥ 0,01 % |  | CLP00
612-023-00-9 | phenylhydrazine [1] phenylhydrazinium chloride [2] phenylhydrazine hydrochloride [3] phenylhydrazinium sulphate (2:1) [4] | 202-873-5 [1] 200-444-7 [2] 248-259-0 [3] 257-622-2 [4] | 100-63-0 [1] 59-88-1 [2] 27140-08-5 [3] 52033-74-6 [4] | Carc. 1B Muta. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 | H350 H341 H331 H311 H301 H372 ** H315 H319 H317 H400 | GHS06 GHS08 GHS09 Dgr | H350 H341 H331 H311 H301 H372 ** H319 H315 H317 H400 |  |  |  | CLP00
612-024-00-4 | m-toluidine; 3-aminotoluene | 203-583-1 | 108-44-1 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 | H331 H311 H301 H373 ** H400 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H400 |  |  |  | CLP00
612-025-00-X | nitrotoluidines, with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 2 | H331 H311 H301 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H411 |  |  | C | CLP00
612-026-00-5 | diphenylamine | 204-539-4 | 122-39-4 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H410 |  |  |  | CLP00
612-027-00-0 | xylidines with the exception of those specified elsewhere in this Annex; dimethyl anilines with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 2 | H331 H311 H301 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H411 |  |  | C | CLP00
612-028-00-6 | p-phenylenediamine | 203-404-7 | 106-50-3 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H319 H317 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H301 H319 H317 H410 |  |  |  | CLP00
612-029-00-1 | benzene-1,4-diamine dihydrochloride; p-phenylenediamine dihydrochloride | 210-834-9 | 624-18-0 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H319 H317 H400 H410 | GHS06 GHS09 Dgr | H331 H311 H301 H319 H317 H410 |  |  |  | CLP00
612-030-00-7 | 2-methyl-p-phenylenediamine sulphate [1] 2-methyl-p-phenylenediamine sulphate [2] | 210-431-8 [1] 228-871-4 [2] | 615-50-9 [1] 6369-59-1 [2] | Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H301 H332 H312 H317 H411 | GHS06 GHS09 Dgr | H301 H332 H312 H317 H411 |  |  |  | CLP00
612-031-00-2 | N,N-dimethylbenzene-1,3-diamine [1] 4-amino-N,N-dimethylaniline; 3-amino-N,N'-dimethylaniline [2] | 220-623-3 [1] 202-807-5 [2] | 2836-04-6 [1] 99-98-9 [2] | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H331 H311 H301 | GHS06 Dgr | H331 H311 H301 |  |  | C | CLP00
612-032-00-8 | N,N,N',N'-tetramethyl-p-phenylenediamine | 202-831-6 | 100-22-1 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H332 H312 H302 | GHS07 Wng | H332 H312 H302 |  |  |  | CLP00
612-033-00-3 | 2-aminophenol | 202-431-1 | 95-55-6 | Muta. 2 Acute Tox. 4 * Acute Tox. 4 * | H341 H332 H302 | GHS08 GHS07 Wng | H341 H332 H302 |  |  |  | CLP00
612-034-00-9 | 2-amino-4,6-dinitrophenol; picramic acid | 202-544-6 | 96-91-3 | Expl. 1.1 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 3 | H201 H332 H312 H302 H412 | GHS01 GHS07 Dgr | H201 H332 H312 H302 H412 |  |  |  | CLP00/ATP01
612-034-01-6 | 2-amino-4,6-dinitrophenol; picramic acid; [≥ 20 % water] | 202-544-6 | 96-91-3 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 3 | H332 H312 H302 H412 | GHS07 Wng | H332 H312 H302 H412 |  |  | G | CLP00
612-035-00-4 | 2-methoxyaniline; o-anisidine | 201-963-1 | 90-04-0 | Carc. 1B Muta. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * | H350 H341 H331 H311 H301 | GHS06 GHS08 Dgr | H350 H341 H331 H311 H301 |  |  |  | CLP00
612-036-00-X | 3,3'-dimethoxybenzidine; o-dianisidine | 204-355-4 | 119-90-4 | Carc. 1B Acute Tox. 4 * | H350 H302 | GHS08 GHS07 Dgr | H350 H302 |  |  |  | CLP00
612-037-00-5 | salts of 3,3'-dimethoxybenzidine; salts of o-dianisidine |  |  | Carc. 1B Acute Tox. 4 * | H350 H302 | GHS08 GHS07 Dgr | H350 H302 |  |  | A | CLP00
612-038-00-0 | 2-nitro-p-anisidine; 4-methoxy-2-nitroaniline | 202-547-2 | 96-96-8 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Chronic 3 | H310 H330 H300 H373 ** H412 | GHS06 GHS08 Dgr | H330 H310 H300 H373 ** H412 |  |  |  | CLP00
612-039-00-6 | 2-ethoxyaniline; o-phenetidine | 202-356-4 | 94-70-2 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * | H331 H311 H301 H373 ** | GHS06 GHS08 Dgr | H331 H311 H301 H373 ** |  |  |  | CLP00
612-040-00-1 | 2,4-dinitroaniline | 202-553-5 | 97-02-9 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Chronic 2 | H310 H330 H300 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H330 H310 H300 H373 ** H411 |  |  |  | CLP00
612-041-00-7 | 4,4'-bi-o-toluidine | 204-358-0 | 119-93-7 | Carc. 1B Acute Tox. 4 * Aquatic Chronic 2 | H350 H302 H411 | GHS08 GHS07 GHS09 Dgr | H350 H302 H411 |  |  |  | CLP00
612-042-00-2 | benzidine; 1,1'-biphenyl-4,4'-diamine; 4,4'-diaminobiphenyl; biphenyl-4,4'-ylenediamine | 202-199-1 | 92-87-5 | Carc. 1A Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H302 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H302 H410 |  | Carc. 1A; H350: C ≥ 0,01 % |  | CLP00
612-043-00-8 | N,N'-dimethylbenzidine |  | 2810-74-4 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H332 H312 H302 | GHS07 Wng | H332 H312 H302 |  |  |  | CLP00
612-044-00-3 | N,N'-diacetylbenzidine | 210-338-2 | 613-35-4 | Carc. 1B Muta. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H350 H341 H332 H312 H302 | GHS08 GHS07 Dgr | H350 H341 H332 H312 H302 |  |  |  | CLP00/ATP01
612-046-00-4 | allylamine | 203-463-9 | 107-11-9 | Flam. Liq. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Chronic 2 | H225 H331 H311 H301 H411 | GHS02 GHS06 GHS09 Dgr | H225 H331 H311 H301 H411 |  |  |  | CLP00
612-047-00-X | benzylamine | 202-854-1 | 100-46-9 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H312 H302 H314 | GHS05 GHS07 Dgr | H312 H302 H314 |  |  |  | CLP00
612-048-00-5 | dipropylamine | 205-565-9 | 142-84-7 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A | H225 H332 H312 H302 H314 | GHS02 GHS05 GHS07 Dgr | H225 H332 H312 H302 H314 |  | STOT SE 3; H335: C ≥ 1 % |  | CLP00
612-049-00-0 | di-n-butylamine | 203-921-8 | 111-92-2 | Flam. Liq. 3 Acute Tox. 2 Acute Tox. 3 Acute Tox. 3 Skin Corr. 1B Eye Dam. 1 | H226 H330 H311 H301 H314 H318 | GHS02 GHS06 GHS05 Dgr | H226 H330 H311 H301 H314 | EUH071 | inhalation: ATE = 1.2 mg/L(vapours) oral: ATE = 220 mg/kg bw dermal: ATE = 300 mg/kg bw |  | CLP00/ATP21
612-050-00-6 | cyclohexylamine | 203-629-0 | 108-91-8 | Flam. Liq. 3 Repr. 2 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H226 H361f *** H312 H302 H314 | GHS02 GHS05 GHS08 GHS07 Dgr | H226 H361f *** H312 H302 H314 |  |  |  | CLP00/ATP01corr
612-051-00-1 | 4,4'-diaminodiphenylmethane; 4,4'-methylenedianiline | 202-974-4 | 101-77-9 | Carc. 1B Muta. 2 STOT SE 1 STOT RE 2 * Skin Sens. 1 Aquatic Chronic 2 | H350 H341 H370 ** H373 ** H317 H411 | GHS08 GHS07 GHS09 Dgr | H350 H341 H370 ** H373 ** H317 H411 |  |  |  | CLP00
612-052-00-7 | (S)-sec-butylamine; (S)-2-aminobutane [1] (R)-sec-butylamine; (R)-2-aminobutane [2] sec-butylamine; 2-aminobutane [3] | 208-164-7 [1] 236-232-6 [2] 237-732-7 [3] | 513-49-5 [1] 13250-12-9 [2] 13952-84-6 [3] | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A Aquatic Acute 1 | H225 H332 H302 H314 H400 | GHS02 GHS05 GHS07 GHS09 Dgr | H225 H332 H302 H314 H400 |  |  | C | CLP00
612-053-00-2 | N-ethylaniline | 203-135-5 | 103-69-5 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * | H331 H311 H301 H373 ** | GHS06 GHS08 Dgr | H331 H311 H301 H373 ** |  |  |  | CLP00
612-054-00-8 | N,N-diethylaniline | 202-088-8 | 91-66-7 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 2 | H331 H311 H301 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H411 |  | * |  | CLP00
612-055-00-3 | N-methyl-o-toluidine [1] N-methyl-m-toluidine [2] N-methyl-p-toluidine [3] | 210-260-9 [1] 211-795-0 [2] 210-769-6 [3] | 611-21-2 [1] 696-44-6 [2] 623-08-5 [3] | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 3 | H331 H311 H301 H373 ** H412 | GHS06 GHS08 Dgr | H331 H311 H301 H373 ** H412 |  |  | C | CLP00
612-056-00-9 | N,N-dimethyl-m-toluidine;[1] N,N-dimethyl-o-toluidine[2] | 204-495-6[1] 210-199-8[2] | 121-72-2[1] 609-72-3[2] | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 3 | H331 H311 H301 H373 ** H412 | GHS06 GHS08 Dgr | H331 H311 H301 H373 ** H412 |  | * | C | CLP00/ATP21
612-057-00-4 | piperazine; [solid] | 203-808-3 | 110-85-0 | Repr. 2 Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1 | H361fd H314 H334 H317 | GHS05 GHS08 Dgr | H361fd H314 H334 H317 |  |  |  | CLP00/ATP01
612-057-01-1 | piperazine; [liquid] | 203-808-3 | 110-85-0 | Repr. 2 Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1 | H361fd H314 H334 H317 | GHS05 GHS08 Dgr | H361fd H314 H334 H317 |  |  |  | ATP01
612-058-00-X | 2,2'-iminodiethylamine; diethylenetriamine | 203-865-4 | 111-40-0 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 | H312 H302 H314 H317 | GHS05 GHS07 Dgr | H312 H302 H314 H317 |  |  |  | CLP00
612-059-00-5 | 3,6-diazaoctanethylenediamin; triethylenetetramine | 203-950-6 | 112-24-3 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 3 | H312 H314 H317 H412 | GHS05 GHS07 Dgr | H312 H314 H317 H412 |  |  |  | CLP00
612-060-00-0 | 3,6,9-triazaundecamethylenediamine; tetraethylenepentamine | 203-986-2 | 112-57-2 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H312 H302 H314 H317 H411 | GHS05 GHS07 GHS09 Dgr | H312 H302 H314 H317 H411 |  |  |  | CLP00
612-061-00-6 | 3-aminopropyldimethylamine; N,N-dimethyl-1,3-diaminopropane | 203-680-9 | 109-55-7 | Flam. Liq. 3 Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 | H226 H302 H314 H317 | GHS02 GHS05 GHS07 Dgr | H226 H302 H314 H317 |  |  |  | CLP00
612-062-00-1 | 3-aminopropyldiethylamine; N,N-diethyl-1,3-diaminopropane | 203-236-4 | 104-78-9 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 | H226 H312 H302 H314 H317 | GHS02 GHS05 GHS07 Dgr | H226 H312 H302 H314 H317 |  |  |  | CLP00
612-063-00-7 | 3,3'-iminodi(propylamine); dipropylenetriamine | 200-261-2 | 56-18-8 | Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1A Skin Sens. 1 | H330 H311 H302 H314 H317 | GHS06 GHS05 Dgr | H330 H311 H302 H314 H317 |  |  |  | CLP00
612-064-00-2 | 3,6,9,12-tetra-azatetradecamethylenediamine; pentacthylenehexamine | 223-775-9 | 4067-16-7 | Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H314 H317 H410 |  |  |  | CLP00
612-065-00-8 | polyethlyenepolyamines with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H312 H302 H314 H317 H410 |  |  |  | CLP00
612-066-00-3 | dicyclohexylamine | 202-980-7 | 101-83-7 | Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H314 H410 |  |  |  | CLP00
612-067-00-9 | 3-aminomethyl-3,5,5-trimethylcyclohexylamine | 220-666-8 | 2855-13-2 | Acute Tox. 4 Skin Corr. 1B Eye Dam. 1 Skin Sens. 1A | H302 H314 H318 H317 | GHS05 GHS07 Dgr | H302 H314 H317 |  | Oral: ATE = 1030 mg/kg bw Skin Sens. 1A; H317: C ≥ 0,001 % |  | CLP00/ATP17
612-068-00-4 | 3,3'-dichlorobenzidine; 3,3'-dichlorobiphenyl-4,4'-ylenediamine | 202-109-0 | 91-94-1 | Carc. 1B Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H312 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H312 H317 H410 |  |  |  | CLP00
612-069-00-X | salts of 3,3'-dichlorobenzidine; salts of 3,3'-dichlorobiphenyl-4,4'-ylenediamine |  |  | Carc. 1B Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H312 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H312 H317 H410 |  |  | A | CLP00
612-070-00-5 | salts of benzidine [ [1] salts of benzidine [ [2] salts of benzidine [ [3] salts of benzidine [ [4] | 208-519-6 [1] 208-520-1 [2] 244-236-4 [3] 252-984-8 [4] | 531-85-1 [1] 531-86-2 [2] 21136-70-9 [3] 36341-27-2 [4] | Carc. 1A Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H302 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H302 H410 |  |  | A | CLP00
612-071-00-0 | salts of 2-naphthylamine [1] salts of 2-naphthylamine [2] | 209-030-0 [1] 210-313-6 [2] | 553-00-4 [1] 612-52-2 [2] | Carc. 1A Acute Tox. 4 * Aquatic Chronic 2 | H350 H302 H411 | GHS08 GHS07 GHS09 Dgr | H350 H302 H411 |  |  | A | CLP00
612-072-00-6 | biphenyl-4-ylamine; xenylamine; 4-aminobiphenyl | 202-177-1 | 92-67-1 | Carc. 1A Acute Tox. 4 * | H350 H302 | GHS08 GHS07 Dgr | H350 H302 |  |  |  | CLP00
612-073-00-1 | salts of biphenyl-4-ylamine; salts of xenylamine; salts of 4-aminobiphenyl |  |  | Carc. 1A Acute Tox. 4 * | H350 H302 | GHS08 GHS07 Dgr | H350 H302 |  |  | A | CLP00
612-074-00-7 | benzyldimethylamine | 203-149-1 | 103-83-3 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 3 | H226 H332 H312 H302 H314 H412 | GHS02 GHS05 GHS07 Dgr | H226 H332 H312 H302 H314 H412 |  |  |  | CLP00
612-075-00-2 | 2-aminoethyldimethylamine; 2-dimethylaminoethylamine | 203-541-2 | 108-00-9 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A | H225 H312 H302 H314 | GHS02 GHS05 GHS07 Dgr | H225 H312 H302 H314 |  |  |  | CLP00
612-076-00-8 | ethyldimethylamine | 209-940-8 | 598-56-1 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H225 H332 H302 H314 | GHS02 GHS05 GHS07 Dgr | H225 H332 H302 H314 |  |  |  | CLP00/ATP01
612-077-00-3 | dimethylnitrosoamine; N-nitrosodimethylamine | 200-549-8 | 62-75-9 | Carc. 1B Acute Tox. 2 * Acute Tox. 3 * STOT RE 1 Aquatic Chronic 2 | H350 H330 H301 H372 ** H411 | GHS06 GHS08 GHS09 Dgr | H350 H330 H301 H372 ** H411 |  | Carc. 1B; H350: C ≥ 0,001 % |  | CLP00
612-078-00-9 | 2,2'-dichloro-4,4'-methylenedianiline; 4,4'-methylene bis(2-chloroaniline) | 202-918-9 | 101-14-4 | Carc. 1B Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H302 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H302 H410 |  |  |  | CLP00
612-079-00-4 | salts of 2,2'-dichloro-4,4'-methylenedianiline; salts of 4,4'-methylenebis(2-chloroaniline) |  |  | Carc. 1B Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H302 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H302 H410 |  |  | A | CLP00
612-080-00-X | 4-amino-N,N-diethylaniline; N,N-diethyl-p-phenylendiamine | 202-214-1 | 93-05-0 | Acute Tox. 3 * Skin Corr. 1B | H301 H314 | GHS06 GHS05 Dgr | H301 H314 |  |  |  | CLP00
612-081-00-5 | salts of 4,4'-bi-o-toluidine; salts of 3,3'-dimethylbenzidine; salts of o-tolidine [1] salts of 4,4'-bi-o-toluidine; salts of 3,3'-dimethylbenzidine; salts of o-tolidine [2] salts of 4,4'-bi-o-toluidine; salts of 3,3'-dimethylbenzidine; salts of o-tolidine [3] | 210-322-5 [1] 265-294-7 [2] 277-985-0 [3] | 612-82-8 [1] 64969-36-4 [2] 74753-18-7 [3] | Carc. 1B Acute Tox. 4 * Aquatic Chronic 2 | H350 H302 H411 | GHS08 GHS07 GHS09 Dgr | H350 H302 H411 |  |  | A | CLP00
612-082-00-0 | thiourea; thiocarbamide | 200-543-5 | 62-56-6 | Carc. 2 Repr. 2 Acute Tox. 4 * Aquatic Chronic 2 | H351 H361d *** H302 H411 | GHS08 GHS07 GHS09 Wng | H351 H361d *** H302 H411 |  |  |  | CLP00
612-083-00-6 | 1-methyl-3-nitro-1-nitrosoguanidine | 200-730-1 | 70-25-7 | Carc. 1B Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H350 H332 H315 H319 H411 | GHS08 GHS07 GHS09 Dgr | H350 H332 H319 H315 H411 |  | Carc. 1B; H350: C ≥ 0,01 % |  | CLP00/ATP01
612-084-00-1 | dapsone; 4,4'-diamino diphenyl sulfone | 201-248-4 | 80-08-0 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
612-085-00-7 | 4,4'-methylenedi-o-toluidine | 212-658-8 | 838-88-0 | Carc. 1B Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H302 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H350 H302 H317 H410 |  |  |  | CLP00
612-086-00-2 | amitraz (ISO); N,N-bis(2,4-xylyliminomethyl) methylamine | 251-375-4 | 33089-61-1 | Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H317 H410 |  | M=10 |  | CLP00
612-087-00-8 | guazatine (ISO) |  | 108173-90-6 | Acute Tox. 2 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H312 H302 H335 H315 H318 H400 H410 | GHS06 GHS05 GHS09 Dgr | H330 H312 H302 H335 H315 H318 H410 |  |  |  | CLP00
612-088-00-3 | simazine (ISO); 6-chloro-N,N'-diethyl-1,3,5-triazine-2,4-diamine | 204-535-2 | 122-34-9 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  |  |  | CLP00
612-089-00-9 | 1,5-naphthylenediamine | 218-817-8 | 2243-62-1 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  |  |  | CLP00
612-090-00-4 | 2,2'-(nitrosoimino)bisethanol | 214-237-4 | 1116-54-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
612-091-00-X | o-toluidine; 2-aminotoluene | 202-429-0 | 95-53-4 | Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Eye Irrit. 2 Aquatic Acute 1 | H350 H331 H301 H319 H400 | GHS06 GHS08 GHS09 Dgr | H350 H331 H301 H319 H400 |  |  |  | CLP00
612-092-00-5 | N,N'-(2,2-dimethylpropylidene)hexamethylenediamine | 401-660-6 | 1000-78-8 | Skin Irrit. 2 Skin Sens. 1 | H315 H317 | GHS07 Wng | H315 H317 |  |  |  | CLP00
612-093-00-0 | 3,5-dichloro-4-(1,1,2,2-tetrafluoroethoxy)aniline | 401-790-3 | 104147-32-2 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
612-094-00-6 | 4-(2-chloro-4-trifluoromethyl)phenoxy-2-fluoroaniline hydrochloride | 402-190-4 | 113674-95-6 | Acute Tox. 4 * STOT RE 1 STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H372 ** H373 ** H318 H317 H400 H410 | GHS05 GHS08 GHS07 GHS09 Dgr | H372 ** H302 H373 ** H318 H317 H410 |  |  |  | CLP00/ATP01
612-095-00-1 | benzyl-2-hydroxydodecyldimethylammonium benzoate | 402-610-6 | 113694-52-3 | Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H400 H410 | GHS05 GHS07 GHS09 Dgr | H314 H302 H410 |  |  |  | CLP00
612-096-00-7 | 4,4'-carbonimidoylbis[N,N-dimethylaniline] | 207-762-5 | 492-80-8 | Carc. 2 Acute Tox. 4 * Eye Irrit. 2 Aquatic Chronic 2 | H351 H302 H319 H411 | GHS08 GHS07 GHS09 Wng | H351 H302 H319 H411 |  |  |  | CLP00
612-097-00-2 | salts of 4,4'-carbonimidoylbis[N,N-dimethylaniline] |  |  | Carc. 2 Acute Tox. 4 * Eye Irrit. 2 Aquatic Chronic 2 | H351 H302 H319 H411 | GHS08 GHS07 GHS09 Wng | H351 H302 H319 H411 |  |  | A | CLP00
612-098-00-8 | nitrosodipropylamine | 210-698-0 | 621-64-7 | Carc. 1B Acute Tox. 4 * Aquatic Chronic 2 | H350 H302 H411 | GHS08 GHS07 GHS09 Dgr | H350 H302 H411 |  | Carc. 1B; H350: C ≥ 0,001 % |  | CLP00/ATP01
612-099-00-3 | 4-methyl-m-phenylenediamine; 2,4-toluenediamine | 202-453-1 | 95-80-7 | Carc. 1B Muta. 2 Repr. 2 Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Chronic 2 | H350 H341 H361f *** H301 H312 H373 ** H317 H411 | GHS06 GHS08 GHS09 Dgr | H350 H341 H361f *** H301 H312 H373 ** H317 H411 |  |  |  | CLP00/ATP01
612-100-00-7 | propylenediamine | 201-155-9 | 78-90-0 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A | H226 H312 H302 H314 | GHS02 GHS05 GHS07 Dgr | H226 H312 H302 H314 |  |  |  | CLP00
612-101-00-2 | methenamine; hexamethylenetetramine | 202-905-8 | 100-97-0 | Flam. Sol. 2 Skin Sens. 1 | H228 H317 | GHS02 GHS07 Wng | H228 H317 |  |  |  | CLP00/ATP01
612-102-00-8 | N,N-bis(3-aminopropyl)methylamine | 203-336-8 | 105-83-9 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1B | H331 H311 H302 H314 | GHS06 GHS05 Dgr | H331 H311 H302 H314 |  |  |  | CLP00
612-103-00-3 | N,N,N',N'-tetramethylethylenediamine | 203-744-6 | 110-18-9 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H225 H332 H302 H314 | GHS02 GHS05 GHS07 Dgr | H225 H332 H302 H314 |  |  |  | CLP00
612-104-00-9 | hexamethylenediamine | 204-679-6 | 124-09-4 | Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Corr. 1B | H312 H302 H335 H314 | GHS05 GHS07 Dgr | H312 H302 H335 H314 |  |  |  | CLP00
612-105-00-4 | 2-piperazin-1-ylethylamine | 205-411-0 | 140-31-8 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 3 | H312 H302 H314 H317 H412 | GHS05 GHS07 Dgr | H312 H302 H314 H317 H412 |  |  |  | CLP00
612-106-00-X | 2,6-diethylaniline | 209-445-7 | 579-66-8 | Acute Tox. 4 * | H302 |  | H302 |  |  |  | CLP00
612-107-00-5 | 1-phenylethylamine [1] Dl-α-methylbenzylamine [2] | 202-706-6 [1] 210-545-8 [2] | 98-84-0 [1] 618-36-0 [2] | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H312 H302 H314 | GHS05 GHS07 Dgr | H312 H302 H314 |  |  |  | CLP00
612-108-00-0 | 3-aminopropyltriethoxysilane | 213-048-4 | 919-30-2 | Acute Tox. 4 * Skin Corr. 1B | H302 H314 | GHS05 GHS07 Dgr | H302 H314 |  |  |  | CLP00
612-109-00-6 | bis(2-dimethylaminoethyl)(methyl)amine | 221-201-1 | 3030-47-5 | Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1B | H311 H302 H314 | GHS06 GHS05 Dgr | H311 H302 H314 |  |  |  | CLP00
612-110-00-1 | 2,2'-dimethyl-4,4'-methylenebis(cyclohexylamine) | 229-962-1 | 6864-37-5 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1A Aquatic Chronic 2 | H331 H311 H302 H314 H411 | GHS06 GHS05 GHS09 Dgr | H331 H311 H302 H314 H411 |  |  |  | CLP00
612-111-00-7 | 2-methyl-m-phenylenediamine; 2,6-toluenediamine | 212-513-9 | 823-40-5 | Muta. 2 Acute Tox. 4 * Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H341 H312 H302 H317 H411 | GHS08 GHS07 GHS09 Wng | H341 H312 H302 H317 H411 |  |  |  | CLP00
612-112-00-2 | p-anisidine; 4-methoxyaniline | 203-254-2 | 104-94-9 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * STOT RE 2 * Aquatic Acute 1 | H310 H330 H300 H373 ** H400 | GHS06 GHS08 GHS09 Dgr | H330 H310 H300 H373 ** H400 |  |  |  | CLP00
612-113-00-8 | 6-methyl-2,4-bis(methylthio)phenylene-1,3-diamine | 403-240-8 | 106264-79-3 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
612-114-00-3 | R,R-2-hydroxy-5-(1-hydroxy-2-(4-phenylbut-2-ylamino)ethyl)benzamide hydrogen 2,3-bis(benzoyloxy)succinate | 404-390-7 |  | Flam. Sol. 1 Skin Sens. 1 Aquatic Chronic 3 | H228 H317 H412 | GHS02 GHS07 Wng | H228 H317 H412 |  |  |  | CLP00
612-115-00-9 | dimethyldioctadecylammonium hydrogen sulfate | 404-050-8 | 123312-54-9 | Eye Irrit. 2 Aquatic Chronic 4 | H319 H413 | GHS07 Wng | H319 H413 |  |  |  | CLP00
612-116-00-4 | C8-18alkylbis(2-hydroxyethyl)ammonium bis(2-ethylhexyl)phosphate | 404-690-8 | 68132-19-4 | Acute Tox. 3 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H314 H317 H400 H410 | GHS06 GHS05 GHS09 Dgr | H331 H314 H317 H410 |  |  |  | CLP00
612-117-00-X | C12-14-tert-alkylamine, methylphosphonic acid salt | 404-750-3 | 119415-07-5 | Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 2 | H302 H314 H411 | GHS05 GHS07 GHS09 Dgr | H302 H314 H411 |  |  |  | CLP00
612-118-00-5 | A reaction mass of: (1,3-dioxo-2H-benz(de)isoquinolin-2-ylpropyl)hexadecyldimethylammonium 4-toluenesulfonate; (1,3-dioxo-2H-benz(de)isoquinolin-2-ylpropyl)hexadecyldimethylammonium bromide | 405-080-4 |  | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | CLP00
612-119-00-0 | benzyldimethyloctadecylammonium 3-nitrobenzenesulfonate | 405-330-2 |  | Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H400 H410 | GHS05 GHS09 Dgr | H315 H318 H410 |  |  |  | CLP00
612-120-00-6 | aclonifen (ISO); 2-chloro-6-nitro-3-phenoxyaniline | 277-704-1 | 74070-46-5 | Carc. 2 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 1 | H351 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H317 H351 H410 |  | M=100 M=10 |  | CLP00/ATP05
612-121-00-1 | amines, polyethylenepoly-; HEPA | 268-626-9 | 68131-73-7 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H312 H302 H314 H317 H410 |  |  |  | CLP00
612-122-00-7 | hydroxylamine ....% [> 55 % in aqueous solution] | 232-259-2 | 7803-49-8 | Unst. Expl. Met. Corr. 1 Carc. 2 Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 | H200 H290 H351 H312 H302 H335 H373 ** H315 H318 H317 H400 | GHS01 GHS05 GHS08 GHS07 GHS09 Dgr | H200 H290 H351 H312 H302 H373 ** H335 H315 H318 H317 H400 |  |  | B | CLP00/ATP01
612-122-01-4 | hydroxylamine ...% [≤ 55% in aqueous solution] | 232-259-2 | 7803-49-8 | Met. Corr. 1 Carc. 2 Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 | H290 H351 H312 H302 H335 H373 ** H315 H318 H317 H400 | GHS05 GHS08 GHS07 GHS09 Dgr | H290 H351 H312 H302 H373 ** H335 H315 H318 H317 H400 |  |  | B | ATP01
612-123-00-2 | hydroxylammonium chloride; hydroxylamine hydrochloride [1] bis(hydroxylammonium) sulfate; hydroxylamine sulfate (2:1) [2] | 226-798-2 [1] 233-118-8 [2] | 5470-11-1 [1] 10039-54-0 [2] | Met. Corr. 1 Carc. 2 Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 | H290 H351 H312 H302 H373 ** H315 H319 H317 H400 | GHS05 GHS08 GHS07 GHS09 Wng | H290 H351 H312 H302 H373 ** H319 H315 H317 H400 |  |  |  | CLP00/ATP01
612-124-00-8 | N,N,N-trimethylanilinium chloride | 205-319-0 | 138-24-9 | Acute Tox. 3 * Acute Tox. 3 * | H311 H301 | GHS06 Dgr | H311 H301 |  |  |  | CLP00
612-125-00-3 | 2-methyl-p-phenylenediamine; 2,5-toluenediamine | 202-442-1 | 95-70-5 | Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H301 H332 H312 H317 H411 | GHS06 GHS09 Dgr | H301 H332 H312 H317 H411 |  |  |  | CLP00
612-126-00-9 | toluene-2,4-diammonium sulphate; 4-methyl-m-phenylenediamine sulfate | 265-697-8 | 65321-67-7 | Carc. 1B Acute Tox. 3 * Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H350 H301 H312 H319 H317 H411 | GHS06 GHS08 GHS09 Dgr | H350 H301 H312 H319 H317 H411 |  |  |  | CLP00
612-127-00-4 | 3-aminophenol | 209-711-2 | 591-27-5 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 2 | H332 H302 H411 | GHS07 GHS09 Wng | H332 H302 H411 |  |  |  | CLP00
612-128-00-X | 4-aminophenol | 204-616-2 | 123-30-8 | Muta. 2 Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H341 H332 H302 H400 H410 | GHS08 GHS07 GHS09 Wng | H341 H332 H302 H410 |  |  |  | CLP00
612-129-00-5 | diisopropylamine | 203-558-5 | 108-18-9 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H225 H332 H302 H314 | GHS02 GHS05 GHS07 Dgr | H225 H332 H302 H314 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
612-130-00-0 | 2,6-diamino-3,5-diethyltoluene; 4,6-diethyl-2-methyl-1,3-benzenediamine [1] 2,4-diamino-3,5-diethyltoluene; 2,4-diethyl-6-methyl-1,3-benzenediamine [2] diethylmethylbenzenediamine [3] | 218-255-3 [1] 218-256-9 [2] 270-877-4 [3] | 2095-01-4 [1] 2095-02-5 [2] 68479-98-1 [3] | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H373 ** H319 H400 H410 | GHS08 GHS07 GHS09 Wng | H312 H302 H373 ** H319 H410 |  |  | C | CLP00
612-131-00-6 | didecyldimethylammonium chloride | 230-525-2 | 7173-51-5 | Acute Tox. 4 * Skin Corr. 1B | H302 H314 | GHS05 GHS07 Dgr | H302 H314 |  |  |  | CLP00
612-132-00-1 | N,N'-diphenyl-p-phenylenediamine; N,N'-diphenyl-1,4-benzenediamine | 200-806-4 | 74-31-7 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
612-133-00-7 | (4-ammonio-m-tolyl)ethyl(2-hydroxyethyl)ammonium sulphate; 4-(N-ethyl-N-2-hydroxyethyl)-2-methylphenylenediamine sulphate | 247-162-0 | 25646-77-9 | Acute Tox. 3 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H373 ** H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H301 H373 ** H317 H410 |  |  |  | CLP00
612-134-00-2 | N-(2-(4-amino-N-ethyl-m-toluidino)ethyl)methanesulphonamide sesquisulphate; 4-(N-ethyl-N-2-methanesulphonylaminoethyl)-2-methylphenylenediamine sesquisulphate monohydrate | 247-161-5 | 25646-71-3 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
612-135-00-8 | N-2-naphthylaniline; N-phenyl-2-naphthylamine | 205-223-9 | 135-88-6 | Carc. 2 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H351 H315 H319 H317 H411 | GHS08 GHS07 GHS09 Wng | H351 H319 H315 H317 H411 |  |  |  | CLP00
612-136-00-3 | N-isopropyl-N'-phenyl-p-phenylenediamine | 202-969-7 | 101-72-4 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  | Skin Sens. 1; H317: C ≥ 0,1 % |  | CLP00
612-137-00-9 | 4-chloroaniline | 203-401-0 | 106-47-8 | Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H331 H311 H301 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H331 H311 H301 H317 H410 |  |  |  | CLP00
612-138-00-4 | furalaxyl (ISO); methyl N-(2,6-dimethylphenyl)-N-(2-furylcarbonyl)-Dl-alaninate | 260-875-1 | 57646-30-7 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
612-139-00-X | mefenacet (ISO); 2-(benzothiazol-2-yloxy)-N-methyl-N-phenylacetamide | 277-328-8 | 73250-68-7 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
612-140-00-5 | quaternary ammonium compounds, benzyl-C8-18-alkyldimethyl, chlorides | 264-151-6 | 63449-41-2 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 | H312 H302 H314 H400 | GHS05 GHS07 GHS09 Dgr | H312 H302 H314 H400 |  |  |  | CLP00
612-141-00-0 | 4,4'-methylenebis(2-ethylaniline); 4,4'-methylenebis(2-ethylbenzeneamine) | 243-420-1 | 19900-65-3 | Carc. 2 Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H302 H410 |  |  |  | CLP00
612-142-00-6 | biphenyl-2-ylamine | 201-990-9 | 90-41-5 | Carc. 2 Acute Tox. 4 * Aquatic Chronic 3 | H351 H302 H412 | GHS08 GHS07 Wng | H351 H302 H412 |  |  |  | CLP00
612-143-00-1 | N5,N5-diethyltoluene-2,5-diamine monohydrochloride; 4-diethylamino-2-methylaniline monohydrochloride | 218-130-3 | 2051-79-8 | Acute Tox. 3 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H319 H317 H400 H410 | GHS06 GHS09 Dgr | H301 H319 H317 H410 |  |  |  | CLP00
612-144-00-7 | flumetralin (ISO); N-(2-chloro-6-fluorobenzyl)-N-ethyl-α,α,α-trifluoro-2,6-dinitro-p-toluidine |  | 62924-70-3 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H319 H317 H400 H410 | GHS07 GHS09 Wng | H319 H315 H317 H410 |  |  |  | CLP00
612-145-00-2 | o-phenylenediamine | 202-430-6 | 95-54-5 | Carc. 2 Muta. 2 Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H341 H301 H332 H312 H319 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H341 H301 H332 H312 H319 H317 H410 |  |  |  | CLP00
612-146-00-8 | o-phenylenediamine dihydrochloride | 210-418-7 | 615-28-1 | Carc. 2 Muta. 2 Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H341 H301 H332 H312 H319 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H351 H341 H301 H332 H312 H319 H317 H410 |  |  |  | CLP00
612-147-00-3 | m-phenylenediamine | 203-584-7 | 108-45-2 | Muta. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H341 H331 H311 H301 H319 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H341 H331 H311 H301 H319 H317 H410 |  |  |  | CLP00
612-148-00-9 | m-phenylenediamine dihydrochloride | 208-790-0 | 541-69-5 | Muta. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H341 H331 H311 H301 H319 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H341 H331 H311 H301 H319 H317 H410 |  |  |  | CLP00
612-149-00-4 | 1,3-diphenylguanidine | 203-002-1 | 102-06-7 | Repr. 2 Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H361f *** H302 H335 H315 H319 H411 | GHS08 GHS07 GHS09 Wng | H361f *** H302 H319 H335 H315 H411 |  |  |  | CLP00
612-150-00-X | spiroxamine (ISO); 8-tert-butyl-1,4-dioxaspiro[4.5]decan-2-ylmethyl(ethyl)(propyl)amine |  | 118134-30-8 | Repr. 2 Acute Tox. 4 Acute Tox. 4 Acute Tox. 4 STOT RE 2 Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d H332 H312 H302 H373 (eye) H315 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H332 H312 H302 H315 H317 H361d H373 (eye) H410 |  | M=100 M=100 |  | CLP00/ATP10
612-151-00-5 | methyl-phenylene diamine; diaminotoluene; [technical product – reaction mass of 4-methyl-m-phenylene diamine (EC No 202-453-1) and 2-methyl-m-phenylene diamine (EC No 212-513-9)] | - | - | Carc. 1B Muta. 2 Repr. 2 Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H350 H341 H361f *** H301 H312 H373 ** H319 H317 H411 | GHS06 GHS08 GHS09 Dgr | H350 H341 H361f *** H301 H312 H373 ** H319 H317 H411 |  |  |  | CLP00/ATP01
612-152-00-0 | N,N-diethyl-N',N'-dimethylpropan-1,3-diyl-diamine | 406-610-7 | 62478-82-4 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1A Aquatic Chronic 3 | H226 H332 H302 H373 ** H314 H412 | GHS02 GHS08 GHS05 GHS07 Dgr | H226 H332 H302 H373 ** H314 H412 |  |  |  | CLP00
612-153-00-6 | 4-[N-ethyl-N-(2-hydroxyethyl)amino]-1-(2-hydroxyethyl)amino-2-nitrobenzene, monohydrochloride | 407-020-2 | 132885-85-9 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 3 | H302 H317 H412 | GHS07 Wng | H302 H317 H412 |  |  |  | CLP00
612-154-00-1 | 6'-(isobutylethylamino)-3'-methyl-2'-phenylamino-spiro[isobenzo-2-oxofuran-7,9'-[9H]-xanthene] | 410-890-6 | 95235-29-3 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
612-155-00-7 | 2'-anilino-6'-((3-ethoxypropyl)ethylamino)-3'-methylspiro(isobenzo-3-oxofuran)-1-(1H)-9'-xanthene | 411-730-8 | 93071-94-4 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
612-156-00-2 | reaction mass of: trihexadecylmethylammonium chloride; dihexadecyldimethylammonium chloride | 405-620-9 |  | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | CLP00
612-157-00-8 | (Z)-1-benzo[b]thien-2-ylethanone oxime hydrochloride | 410-780-8 |  | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H373 ** H318 H317 H411 | GHS08 GHS05 GHS07 GHS09 Dgr | H302 H373 ** H318 H317 H411 |  |  |  | CLP00
612-158-00-3 | reaction mass of: bis(5-dodecyl-2-hydroxybenzald-oximate) copper (II) C12-alkyl group is branched; 4-dodecylsalicylaldoxime | 410-820-4 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
612-159-00-9 | Reaction products of: trimethylhexamethylene diamine (a mixture of 2,2,4-trimethyl-1,6-hexanediamine and 2,4,4-trimethyl-1,6-hexanediamine, EINECS listed), Epoxide 8 (mono[(C10-C16-alkyloxy)methyl]oxirane derivatives) and p-toluene-sulfonic acid | 410-880-1 |  | Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H314 H410 |  |  |  | CLP00
612-160-00-4 | p-toluidine; 4-aminotoluene [1] toluidinium chloride [2] toluidine sulphate (1:1) [3] | 203-403-1 [1] 208-740-8 [2] 208-741-3 [3] | 106-49-0 [1] 540-23-8 [2] 540-25-0 [3] | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 | H351 H331 H311 H301 H319 H317 H400 | GHS06 GHS08 GHS09 Dgr | H351 H331 H311 H301 H319 H317 H400 |  |  |  | CLP00
612-161-00-X | 2,6-xylidine; 2,6-dimethylaniline | 201-758-7 | 87-62-7 | Carc. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Aquatic Chronic 2 | H351 H332 H312 H302 H335 H315 H411 | GHS08 GHS07 GHS09 Wng | H351 H332 H312 H302 H335 H315 H411 |  |  |  | CLP00
612-162-00-5 | dimethyldioctadecylammonium chloride; DODMAC | 203-508-2 | 107-64-2 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | CLP00
612-163-00-0 | metalaxyl-M (ISO); mefenoxam; (R)-2-[(2,6-dimethylphenyl)-methoxyacetylamino]propionic acid methyl ester |  | 70630-17-0 | Acute Tox. 4 * Eye Dam. 1 | H302 H318 | GHS05 GHS07 Dgr | H302 H318 |  |  |  | CLP00
612-164-00-6 | 2-butyl-2-ethyl-1,5-diaminopentane | 412-700-7 | 137605-95-9 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 3 | H312 H302 H373 ** H314 H317 H412 | GHS08 GHS05 GHS07 Dgr | H312 H302 H373 ** H314 H317 H412 |  |  |  | CLP00
612-165-00-1 | N,N'-diphenyl-N,N'-bis(3-methylphenyl)-(1,1'-diphenyl)-4,4'-diamine | 413-810-8 | 65181-78-4 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
612-166-00-7 | reaction mass of: cis-(5-ammonium-1,3,3-trimethyl)-cyclohexanemethylammonium phosphate (1:1); trans-(5-ammonium-1,3,3-trimethyl)-cyclohexanemethylammonium phosphate (1:1) | 411-830-1 | 114765-88-7 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | CLP00
612-167-00-2 | 5-acetyl-3-amino-10,11-dihydro-5H-dibenz[b,f]azepine-hydrochloride | 410-490-1 |  | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H373 ** H318 H317 H411 | GHS08 GHS05 GHS07 GHS09 Dgr | H302 H373 ** H318 H317 H411 |  |  |  | CLP00
612-168-00-8 | 3,5-dichloro-2,6-difluoropyrdine-4-amine | 220-630-1 | 2840-00-8 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 2 | H312 H302 H411 | GHS07 GHS09 Wng | H312 H302 H411 |  |  |  | CLP00
612-169-00-3 | bis(N-methyl-N-phenylhydrazine)sulfate | 423-170-1 | 618-26-8 | Flam. Liq. 2 Acute Tox. 4 * STOT RE 1 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H225 H302 H372 ** H318 H317 H400 H410 | GHS02 GHS05 GHS08 GHS07 GHS09 Dgr | H225 H372 ** H302 H318 H317 H410 |  |  |  | ATP01
612-170-00-9 | 4-chlorophenyl cyclopropyl ketone O-(4-aminobenzyl)oxime | 405-260-2 |  | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
612-171-00-4 | N,N,N',N'-tetraglycidyl-4,4'-diamino-3,3'-diethyldiphenylmethane | 410-060-3 | 130728-76-6 | Muta. 2 Skin Sens. 1 Aquatic Chronic 2 | H341 H317 H411 | GHS08 GHS09 Wng | H341 H317 H411 |  |  |  | CLP00
612-172-00-X | 4,4'-methylenebis(N,N'-dimethylcyclohexanamine | 412-840-9 | 13474-64-1 | Acute Tox. 4 * STOT RE 2 * Skin Corr. 1A Aquatic Chronic 3 | H302 H373 ** H314 H412 | GHS08 GHS05 GHS07 Dgr | H302 H373 ** H314 H412 |  |  |  | CLP00
612-173-00-5 | lithium 1-amino-4-(4-tert-butylanilino)anthraquinone-2-sulfonate | 411-140-0 | 125328-86-1 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H318 H317 H411 |  |  |  | CLP00
612-174-00-0 | 4,4-dimethoxybutylamine | 407-690-6 | 19060-15-2 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 3 | H302 H314 H317 H412 | GHS05 GHS07 Dgr | H302 H314 H317 H412 |  |  |  | CLP00
612-175-00-6 | 2-(O-aminooxy)ethylamine dihydrochloride | 412-310-7 | 37866-45-8 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
612-176-00-1 | Polymer of 1,3-dibromopropane and N,N-diethyl-N',N'-dimethyl-1,3-propanediamine | 410-570-6 | 143747-73-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
612-177-00-7 | 2-naphthylamino-6-sulfomethylamide | 412-120-4 | 104295-55-8 | STOT RE 2 * Skin Sens. 1 Aquatic Chronic 2 | H373 ** H317 H411 | GHS08 GHS09 Wng | H373 ** H317 H411 |  |  |  | CLP00
612-178-00-2 | 1,4,7,10-tetraazacyclododecane disulfate | 412-080-8 | 112193-77-8 | Acute Tox. 4 * STOT SE 3 Eye Dam. 1 Aquatic Chronic 3 | H302 H335 H318 H412 | GHS05 GHS07 Dgr | H302 H335 H318 H412 |  |  |  | CLP00
612-179-00-8 | 1-(2-propenyl)pyridinium chloride | 412-740-5 | 25965-81-5 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  |  |  | CLP00
612-180-00-3 | 3-aminobenzylamine | 412-230-2 | 4403-70-7 | Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 2 | H302 H314 H411 | GHS05 GHS07 GHS09 Dgr | H302 H314 H411 |  |  |  | CLP00
612-181-00-9 | 2-phenylthioaniline | 413-030-8 | 1134-94-7 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
612-182-00-4 | 1-ethyl-1-methylmorpholinium bromide | 418-250-8 | 65756-41-4 | Muta. 2 | H341 | GHS08 Wng | H341 |  |  |  | CLP00
612-183-00-X | 1-ethyl-1-methylpyrrolidinium bromide | 418-200-5 | 69227-51-6 | Muta. 2 | H341 | GHS08 Wng | H341 |  |  |  | CLP00
612-184-00-5 | 6'-(dibutylamino)-3'-methyl-2'-(phenylamino)spiro[isobenzofuran-1(3H),9-(9H)-xanthen]-3-one | 403-830-5 | 89331-94-2 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
612-185-00-0 | 1-[3-[4-((heptadecafluorononyl)oxy)-benzamido]propyl]-N,N,N-trimethylammonium iodide | 407-400-8 | 59493-72-0 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | CLP00
612-186-00-6 | bis(N-(7-hydroxy-8-methyl-5-phenylphenazin-3-ylidene)dimethylammonium) sulfate | 406-770-8 | 149057-64-7 | STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H318 H317 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H373 ** H318 H317 H410 |  |  |  | CLP00
612-187-00-1 | 2,3,4-trifluoroaniline | 407-170-9 | 3862-73-5 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 2 | H312 H302 H373 ** H315 H318 H411 | GHS08 GHS05 GHS07 GHS09 Dgr | H312 H302 H373 ** H315 H318 H411 |  |  |  | CLP00
612-188-00-7 | 4,4'-(9H-fluoren-9-ylidene)bis(2-chloroaniline) | 407-560-9 | 107934-68-9 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
612-189-00-2 | 4-amino-2-(aminomethyl)phenol dihydrochloride | 412-510-4 | 135043-64-0 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
612-190-00-8 | 4,4'-methylenebis(2-isopropyl-6-methylaniline) | 415-150-6 | 16298-38-7 | STOT RE 2 * Aquatic Chronic 2 | H373 ** H411 | GHS08 GHS09 Wng | H373 ** H411 |  |  |  | CLP00
612-191-00-3 | Polymer of allylamine hydrochloride | 415-050-2 | 71550-12-4 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  |  |  | CLP00
612-192-00-9 | 2-isopropyl-4-(N-methyl)aminomethylthiazole | 414-800-6 | 154212-60-9 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 2 | H312 H302 H315 H318 H411 | GHS05 GHS07 GHS09 Dgr | H312 H302 H315 H318 H411 |  |  |  | CLP00
612-193-00-4 | 3-methylaminomethylphenylamine | 414-570-7 | 18759-96-1 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H312 H302 H314 H317 H410 |  |  |  | CLP00
612-194-00-X | 2-hydroxy-3-[(2-hydroxyethyl)-[2-(1-oxotetradecyl)amino]ethyl]amino]-N,N,N-trimethyl-1-propanammonium chloride | 414-670-0 | 141890-30-4 | Acute Tox. 4 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H318 H410 |  |  |  | CLP00
612-195-00-5 | bis[tributyl 4-(methylbenzyl)ammonium] 1,5-naphthalenedisulfonate | 415-210-1 | 160236-81-7 | Acute Tox. 4 * Acute Tox. 4 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H332 H302 H318 H410 |  |  |  | CLP00
612-196-00-0 | 4-chloro-o-toluidine [1] 4-chloro-o-toluidine hydrochloride [2] | 202-441-6 [1] 221-627-8 [2] | 95-69-2 [1] 3165-93-3 [2] | Carc. 1B Muta. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H350 H341 H331 H311 H301 H400 H410 | GHS06 GHS08 GHS09 Dgr | H350 H341 H331 H311 H301 H410 |  |  |  | CLP00
612-197-00-6 | 2,4,5-trimethylaniline [1] 2,4,5-trimethylaniline hydrochloride [2] | 205-282-0 [1] | 137-17-7 [1] 21436-97-5 [2] | Carc. 1B Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Chronic 2 | H350 H331 H311 H301 H411 | GHS06 GHS08 GHS09 Dgr | H350 H331 H311 H301 H411 |  |  |  | CLP00
612-198-00-1 | 4,4'-thiodianiline and its salts | 205-370-9 | 139-65-1 | Carc. 1B Acute Tox. 4 * Aquatic Chronic 2 | H350 H302 H411 | GHS08 GHS07 GHS09 Dgr | H350 H302 H411 |  |  |  | CLP00
612-199-00-7 | 4,4'-oxydianiline and its salts; p-aminophenyl ether | 202-977-0 | 101-80-4 | Carc. 1B Muta. 1B Repr. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Chronic 2 | H350 H340 H361f *** H331 H311 H301 H411 | GHS06 GHS08 GHS09 Dgr | H350 H340 H361f *** H331 H311 H301 H411 |  |  |  | CLP00
612-200-00-0 | 2,4-diaminoanisole; 4-methoxy-m-phenylenediamine [1] 2,4-diaminoanisole sulphate [2] | 210-406-1 [1] 254-323-9 [2] | 615-05-4 [1] 39156-41-7 [2] | Carc. 1B Muta. 2 Acute Tox. 4 * Aquatic Chronic 2 | H350 H341 H302 H411 | GHS08 GHS07 GHS09 Dgr | H350 H341 H302 H411 |  |  |  | CLP00
612-201-00-6 | N,N,N',N'-tetramethyl-4,4'-methylendianiline | 202-959-2 | 101-61-1 | Carc. 1B Aquatic Acute 1 Aquatic Chronic 1 | H350 H400 H410 | GHS08 GHS09 Dgr | H350 H410 |  |  |  | CLP00
612-202-00-1 | 3,4-dichloroaniline | 202-448-4 | 95-76-1 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H318 H317 H400 H410 | GHS06 GHS05 GHS09 Dgr | H331 H311 H301 H318 H317 H410 |  |  |  | CLP00
612-203-00-7 | C8-10 alkyl dimethyl hydroxyethyl ammoniumchloride (chain < C8: <3%, chain = C8: 15%-70%, chain = C10: 30%-85%, chain > C10: <3%) | 417-360-3 | - | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 | H312 H302 H315 | GHS07 Wng | H312 H302 H315 |  |  |  | ATP01
612-204-00-2 | C.I. Basic Violet 3; 4-[4,4'-bis(dimethylamino) benzhydrylidene]cyclohexa-2,5-dien-1-ylidene]dimethylammonium chloride | 208-953-6 | 548-62-9 | Carc. 2 Acute Tox. 4 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H318 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H351 H302 H318 H410 |  |  |  | CLP00
612-205-00-8 | C.I. Basic Violet 3 with ≥ 0.1 % of Michler's ketone (EC no. 202-027-5) | 208-953-6 | 548-62-9 | Carc. 1B Acute Tox. 4 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H302 H318 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H350 H302 H318 H410 |  |  |  | CLP00
612-206-00-3 | famoxadone (ISO); 3-anilino-5-methyl-5-(4-phenoxyphenyl)-1,3-oxazolidine-2,4-dione |  | 131807-57-3 | STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H400 H410 | GHS08 GHS09 Wng | H373 ** H410 |  |  |  | CLP00
612-207-00-9 | 4-ethoxyaniline; p-phenetidine | 205-855-5 | 156-43-4 | Muta. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 | H341 H332 H312 H302 H319 H317 | GHS08 GHS07 Wng | H341 H332 H312 H302 H319 H317 |  |  |  | CLP00
612-208-00-4 | N-methylbenzene-1,2-diammonium hydrogen phosphate | 424-460-0 | - | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | ATP01
612-209-00-X | 6-methoxy-m-toluidine; p-cresidine | 204-419-1 | 120-71-8 | Carc. 1B Acute Tox. 4 * | H350 H302 | GHS08 GHS07 Dgr | H350 H302 |  |  |  | CLP00
612-210-00-5 | 5-nitro-o-toluidine [1] 5-nitro-o-toluidine hydrochloride [2] | 202-765-8 [1] 256-960-8 [2] | 99-55-8 [1] 51085-52-0 [2] | Carc. 2 Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * Aquatic Chronic 3 | H351 H331 H311 H301 H412 | GHS06 GHS08 Dgr | H351 H331 H311 H301 H412 |  |  |  | CLP00
612-211-00-0 | N-[(benzotriazole-1-yl)methyl)]-4-carboxybenzenesulfonamide | 416-470-9 | 170292-97-4 | Eye Irrit. 2 Aquatic Chronic 2 | H319 H411 | GHS07 GHS09 Wng | H319 H411 |  |  |  | CLP00
612-212-00-6 | 2,6-dichloro-4-trifluoromethylaniline | 416-430-0 | 24279-39-8 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H315 H317 H400 H410 | GHS07 GHS09 Wng | H332 H302 H315 H317 H410 |  |  |  | CLP00
612-213-00-1 | isobutylidene-(2-(2-isopropyl-4,4-dimethyloxazolidine-3-yl)-1,1-dimethylethyl)amine | 419-850-2 | 148348-13-4 | Skin Corr. 1B Aquatic Chronic 3 | H314 H412 | GHS05 Dgr | H314 H412 |  |  |  | CLP00
612-214-00-7 | 4-(2,2-diphenylethenyl)-N,N-di-phenylbenzenamine | 421-390-2 | 89114-90-9 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
612-215-00-2 | 3-chloro-2-(isopropylthio)aniline | 421-700-6 | 179104-32-6 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
612-216-00-8 | 1-amino-1-cyanamino-2,2-dicyanoethylene, sodium salt | 425-870-2 | 19450-38-5 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
612-217-00-3 | 1-methoxy-2-propylamine | 422-550-4 | 37143-54-7 | Flam. Liq. 2 Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 3 | H225 H302 H314 H412 | GHS02 GHS05 GHS07 Dgr | H225 H314 H302 H412 |  |  |  | CLP00
612-219-00-4 | (2-hydroxy-3-(3,4-dimethyl-9-oxo-10-thiaanthracen-2-yloxy)propyl)trimethylammonium chloride | 402-200-7 | - | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
612-220-00-X | N-nitro-N-(3-methyl-3,6-dihydro-2H-1,3,5-oxadiazin-4-yl)amine | 431-060-1 | 153719-38-1 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 3 | H302 H317 H412 | GHS07 Wng | H302 H317 H412 |  |  |  | ATP01
612-221-00-5 | 2-amino-4-(trifluoromethyl)benzenethiol hydrochloride | 429-560-8 | 4274-38-8 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 | H332 H312 H302 H373 ** H314 H317 H400 | GHS05 GHS08 GHS07 GHS09 Dgr | H314 H332 H312 H302 H373 ** H317 H400 |  |  |  | ATP01
612-222-00-0 | cis-1-(3-(4-fluorophenoxy)propyl)-3-methoxy-4-piperidinamine | 425-080-8 | 104860-26-6 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H373 ** H318 H400 H410 | GHS05 GHS08 GHS07 GHS09 Dgr | H312 H302 H373 ** H318 H410 |  |  |  | ATP01
612-223-00-6 | N-benzyl-N-ethyl-(4-(5-nitro-benzo[c]isothiazol-3-ylazo)phenyl)amine | 425-300-2 | 186450-73-7 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
612-224-00-1 | N2,N4,N6-tris{4-[(1,4-dimethylpentyl)amino]phenyl}-1,3,5-triazine-2,4,6-triamine | 426-150-0 | 121246-28-4 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | ATP01
612-225-00-7 | 1,4,7,10-tetraazacyclododecane | 425-450-9 | 294-90-6 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H314 H400 H410 | GHS05 GHS07 GHS09 Dgr | H314 H312 H302 H410 |  |  |  | ATP01
612-226-00-2 | 3-(2'-phenoxyethoxy)propylamine | 427-870-8 | 6903-18-0 | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H302 H315 H318 H412 | GHS05 GHS07 Dgr | H302 H315 H318 H412 |  |  |  | ATP01
612-227-00-8 | benzyl-N-(2-(2-methoxyphenoxy)ethyl)amine hydrochloride | 428-290-8 | 120606-08-8 | Acute Tox. 4 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H318 H410 |  |  |  | ATP01
612-228-00-3 | reaction mass of: N-(3-(trimethoxysilyl)propyl)ethylenediamine; N-benzyl-N-(3-(trimethoxysilyl)propyl)ethylenediamine; N-benzyl-N'-[3-(trimethoxysilyl)propyl]ethylenediamine; N,N'-bis-benzyl-N'-[3-(trimethoxysilyl)propyl]ethylenediamine; N,N,N'-tris-benzyl-N'-[3-(trimethoxysilyl)propyl]ethylenediamine; N,N-bis-benzyl-N'-[3-(trimethoxysilyl)propyl]ethylenediamine | 414-340-6 | - | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H226 H332 H312 H302 H371 H318 H317 H412 | GHS02 GHS05 GHS08 GHS07 Dgr | H226 H332 H312 H302 H371 H318 H317 H412 |  |  |  | ATP01
612-229-00-9 | mepanipyrim; 4-methyl-N-phenyl-6-(1-propynyl)-2-pyrimidinamine | - | 110235-47-7 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  |  |  | ATP01
612-230-00-4 | N,N-bis(cocoyl-2-oxypropyl)-N,N-dibutylammonium bromide | 431-530-4 | - | Skin Corr. 1A Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H314 H317 H410 |  |  |  | ATP01
612-231-00-X | 3-((C12-18)-acylamino)-N-(2-((2-hydroxyethyl)amino)-2-oxoethyl)-N,N-dimethyl-1-propanaminium chloride | 427-370-1 | 164288-56-6 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | ATP01
612-232-00-5 | reaction mass of: triisopropanolamine salt of 1-amino-4-(3-propionamidoanilino)anthraquinone-2-sulfonic acid; triisopropanolamine salt of 1-amino-4-[3,4-dimethyl-5-(2-hydroxyethylaminosulfonyl)anilino]anthraquinone-2-sulfonic acid | 430-410-9 | 186148-38-9 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
612-237-00-2 | hydroxylammonium hydrogensulfate; hydroxylamine sulfate(1:1) [1] hydroxylamine phosphate [2] hydroxylamine dihydrogenphosphate [3] hydroxylamine 4-methylbenzenesulfonate [4] | 233-154-4 [1] 244-077-0 [2] 242-818-2 [3] 258-872-5 [4] | 10046-00-1 [1] 20845-01-6 [2] 19098-16-9 [3] 53933-48-5 [4] | Expl. 1.1 Carc. 2 Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 | H201 H351 H312 H302 H373 ** H315 H319 H317 H400 | GHS01 GHS08 GHS07 GHS09 Dgr | H201 H302 H312 H315 H319 H317 H351 H373 ** H400 |  |  | T | ATP01/ATP01corr
612-238-00-8 | (3-chloro-2-hydroxypropyl) trimethylammonium chloride ...% | 222-048-3 | 3327-22-8 | Carc. 2 Aquatic Chronic 3 | H351 H412 | GHS08 Wng | H351 H412 |  |  | B | ATP01
612-239-00-3 | biphenyl-3,3',4,4'-tetrayltetraamine; diaminobenzidine | 202-110-6 | 91-95-2 | Carc. 1B Muta. 2 | H350 H341 | GHS08 Dgr | H350 H341 |  |  |  | ATP01
612-240-00-9 | pyrimethanil (ISO); N-(4,6-dimethylpyrimidin-2-yl)aniline | - | 53112-28-0 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
612-241-00-4 | piperazine hydrochloride [1] piperazine dihydrochloride [2] piperazine phosphate [3] | 228-042-7 [1] 205-551-2 [2] 217-775-8 [3] | 6094-40-2 [1] 142-64-3 [2] 1951-97-9 [3] | Repr. 2 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 Skin Sens. 1 Aquatic Chronic 3 | H361fd H315 H319 H334 H317 H412 | GHS08 Dgr | H361fd H319 H315 H334 H317 H412 |  |  |  | ATP01
612-242-00-X | cyprodinil (ISO); 4-cyclopropyl-6-methyl-N-phenylpyrimidin-2-amine | - | 121552-61-2 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  | M=10 |  | ATP01
612-243-00-5 | (1S-cis)-4-(3,4-dichlorophenyl)-1,2,3,4-tetrahydro-N-methyl-1-naphthalenamine 2-hydroxy-2-phenylacetate | 420-560-3 | 79617-97-3 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  | M=10 |  | ATP01
612-244-00-0 | 3-(piperazin-1-yl)-benzo[d]isothiazole hydrochloride | 421-310-6 | 87691-88-1 | Repr. 2 Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361f *** H302 H319 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H361f *** H302 H319 H317 H410 |  |  |  | ATP01
612-245-00-6 | 2-ethylphenylhydrazine hydrochloride | 421-460-2 | 19398-06-2 | Carc. 2 Acute Tox. 4 * STOT RE 1 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H372 ** H318 H317 H400 H410 | GHS05 GHS08 GHS07 GHS09 Dgr | H351 H372 ** H302 H318 H317 H410 |  | M=10 |  | ATP01
612-246-00-1 | (2-chloroethyl)(3-hydroxypropyl)ammonium chloride | 429-740-6 | 40722-80-3 | Carc. 1B Muta. 1B STOT RE 2 * Skin Sens. 1 Aquatic Chronic 3 | H350 H340 H373 ** H317 H412 | GHS08 GHS07 Dgr | H350 H340 H373 ** H317 H412 |  |  |  | ATP01
612-247-00-7 | N-[3-(1,1-dimethylethyl)-1H-pyrazol-5-yl]-N'-hydroxy-4-nitrobenzenecarboximidamide | 423-530-8 | 152828-23-4 | Acute Tox. 4 * STOT RE 1 Aquatic Chronic 3 | H302 H372 ** H412 | GHS08 GHS07 Dgr | H372 ** H302 H412 |  |  |  | ATP01
612-248-00-2 | reaction product of diphenylamine, phenothiazine, and alkenes, branched (C8-10, C9-rich) | 439-540-0 | - | Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 4 | H315 H317 H413 | GHS07 Wng | H315 H317 H413 |  |  |  | ATP01
612-249-00-8 | 4-[(3-chlorophenyl)(1H-imidazol-1-yl)methyl]-1,2-benzenediamine dihydrochloride | 425-030-5 | 159939-85-2 | Repr. 2 Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H361f *** H302 H314 H317 H411 | GHS05 GHS08 GHS07 GHS09 Dgr | H361f *** H302 H314 H317 H411 |  |  |  | ATP01
612-250-00-3 | chloro-N,N-dimethylformiminium chloride | 425-970-6 | 3724-43-4 | Repr. 1B Acute Tox. 4 * Skin Corr. 1A | H360D *** H302 H314 | GHS05 GHS08 GHS07 Dgr | H360D *** H302 H314 | EUH014 |  |  | ATP01
612-251-00-9 | cis-1-(3-chloroallyl)-3,5,7-triaza-1-azoniaadamantane chloride | 426-020-3 | 51229-78-8 | Flam. Sol. 2 Repr. 2 Acute Tox. 4 * Skin Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H228 H361d *** H302 H315 H317 H411 | GHS02 GHS08 GHS07 GHS09 Wng | H228 H361d *** H302 H315 H317 H411 |  |  |  | ATP01
612-252-00-4 | imidacloprid (ISO); (E)-1-(6-chloro- 3-pyridylmethyl)-N-nitroimidazolidin-2-ylideneamine;(2E)-1-[(6-chloropyridin- 3-yl) methyl]-N-nitroimidazolidin-2-imine | 428-040-8 | 138261-41-3 | Acute Tox. 3 Aquatic Acute 1 Aquatic Chronic 1 | H301 H400 H410 | GHS06 GHS09 Dgr | H301 H410 |  | Oral: ATE = 131 mg/kg bw M=100 M=1000 |  | ATP01/ATP17
612-253-00-X | 7-methoxy-6-(3-morpholin-4-yl-propoxy)-3H-quinazolin-4-one; [containing < 0.5 % formamide (EC No 200-842-0)] | 429-400-7 | 199327-61-2 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
612-253-01-7 | 7-methoxy-6-(3-morpholin-4-yl-propoxy)-3H-quinazolin-4-one; [containing ≥ 0.5 % formamide (EC No 200-842-0) ] | 429-400-7 | 199327-61-2 | Repr. 1B Aquatic Chronic 3 | H360D *** H412 | GHS08 Dgr | H360D *** H412 |  |  |  | ATP01
612-254-00-5 | reaction products of diisopropanolamine with formaldehyde (1:4) | 432-440-8 | 220444-73-5 | Carc. 2 Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H351 H302 H314 H317 H411 | GHS05 GHS08 GHS07 GHS09 Dgr | H351 H302 H314 H317 H411 |  |  |  | ATP01
612-255-00-0 | 1-(3-methoxypropyl)-4-piperidinamine | 431-950-8 | 179474-79-4 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B Aquatic Chronic 3 | H312 H302 H314 H412 | GHS05 GHS07 Dgr | H312 H302 H314 H412 |  |  |  | ATP01
612-256-00-6 | benzyl(S)-2-[(2'-cyanobiphenyl-4-ylmethyl)pentanoylamino]-3-methylbutyrate | 427-470-3 | 137864-22-3 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  |  |  | ATP01
612-257-00-1 | tripropylammonium dihydrogenphosphate | 433-700-3 | 35687-90-2 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | ATP01
612-259-00-2 | N-ethyl-3-trimethoxysilyl-2-methyl-propanamine | 437-720-3 | 227085-51-0 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
612-261-00-3 | 3,5-dichloro-2-fluoro-4-(1,1,2,3,3,3-hexafluoropropoxy)aniline | 441-190-9 | 121451-05-6 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  | M=10 |  | ATP01
612-265-00-5 | bis(2-hydroxyethyl)-(2-hydroxypropyl)ammonium acetate | 444-360-0 | 191617-13-7 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
612-266-00-0 | 3-chloro-4-(3-fluorobenzyloxy)aniline | 445-590-4 | 202197-26-0 | Muta. 2 Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H341 H302 H373 ** H400 H410 | GHS08 GHS07 GHS09 Wng | H341 H302 H373 ** H410 |  |  |  | ATP01
612-267-00-6 | bis(hydrogenated tallow C16-18-alkyl)hydroxylamine | 418-370-0 | - | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
612-269-00-7 | reaction mass of: 1-[di(4-octylphenyl)aminomethyl]-5-methyl-1H-benzotriazole; 1-[di(4-octylphenyl)aminomethyl]-4-methyl-1H-benzotriazole; reaction mass of: N-[(5-methyl-1H-benzotriazol-1-yl)methyl]-4-octyl-N-(4-octylphenyl)aniline; N-[(4-methyl-1H-benzotriazol-1-yl)methyl]-4-octyl-N-(4-octylphenyl)aniline | 420-720-2 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
612-270-00-2 | (S)-azetidine-2-carboxylic acid 4-cyanobenzylamide hydrochloride | 433-010-2 | - | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 3 | H302 H317 H412 | GHS07 Wng | H302 H317 H412 |  |  |  | ATP01
612-271-00-8 | reaction mass of: ethyl 2-((4-(5,6-dichlorobenzothiazol-2-ylazo)phenyl)ethylamino)benzoate; ethyl 2-((4-(6,7-dichlorobenzothiazol-2-ylazo)phenyl)ethylamino)benzoate | 434-970-5 | 160987-57-5 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
612-272-00-3 | ammonium (η-6-2-(2-(1,2-dicarboxylatoethylamino)ethylamino)butane-1,4-dioato(4-))iron(3+) monohydrate | 435-210-5 | - | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
612-273-00-9 | alkyl(rapeseed oil), bis(2-hydroxyethyl)ammonium fluoride | 435-650-8 | - | Acute Tox. 4 * Skin Corr. 1A Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H314 H410 |  |  |  | ATP01
612-274-00-4 | (R,S)-1-[2-amino-1(4-methoxyphenyl)ethyl]cyclohexanol acetate | 445-750-3 | - | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H302 H318 H317 H412 | GHS05 GHS07 Dgr | H302 H318 H317 H412 |  |  |  | ATP01
612-275-00-X | fatty acids, C18-unsatd., dimers, reaction products with 1-piperazineethanamine and tall oil | 447-880-6 | 206565-89-1 | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H315 H318 H317 H410 |  | M=10 |  | ATP01
612-276-00-5 | 1-amino-4-[(4-amino-2-sulfofenyl)amino]-9,10-dihydro-9,10-dioxo-2-anthracenesulfonic acid, disodium salt, reaction products with 2-[[3-[(4,6-dichloro-1,3,5-triazin-2-yl)ethylamino]phenyl]sulfonyl]ethyl hydrogen sulfate, sodium salts | 451-430-4 | 500717-36-2 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | ATP01
612-277-00-0 | reaction mass of: 4-amino-3-(4-ethenesulfonyl-2-sulfonatophenylazo)-5-hydroxy-6-(5-{4-chloro-6-[4-(2-sulfonatooxyethanesulfonyl)phenylamino]-1,3,5-triazin-2-ylamino}-2-sulfonatophenylazo)naphthalene-2,7-disulfonate potassium/sodium; 4-amino-5-hydroxy-6-(5-{4-chloro-6-[4-(2-sulfonatooxyethanesulfonyl)phenylamino]-1,3,5-triazin-2-ylamino}-2-sulfonatophenylazo)-3-(2-sulfonato-4-(2-sulfonatooxyethanesulfonyl)phenylazo)naphthalene-2,7-disulfonate potassium/sodium | 451-440-9 | 586372-44-3 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
612-278-00-6 | ethidium bromide; 3,8-diamino-1-ethyl-6-phenylphenantridinium bromide | 214-984-6 | 1239-45-8 | Muta. 2 Acute Tox. 2 * Acute Tox. 4 * | H341 H330 H302 | GHS06 GHS08 Dgr | H341 H330 H302 |  |  |  | ATP01
612-279-00-1 | (R,S)-2-amino-3,3-dimethylbutane amide | 447-860-7 | 144177-62-8 | Repr. 2 STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H361f *** H373 ** H315 H319 H317 | GHS08 GHS07 Wng | H361f *** H373 ** H319 H315 H317 |  |  |  | ATP01
612-280-00-7 | 3-amino-9-ethyl carbazole; 9-ethylcarbazol-3-ylamine | 205-057-7 | 132-32-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | ATP01
612-281-00-2 | leucomalachite green; N,N,N',N'-tetramethyl-4,4'-benzylidenedianiline | 204-961-9 | 129-73-7 | Carc. 2 Muta. 2 | H351 H341 | GHS08 Wng | H351 H341 |  |  |  | ATP03
612-282-00-8 | octadecylamine | 204-695-3 | 124-30-1 | Asp. Tox. 1 STOT RE 2 Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H304 H373 (gastro-intestinal tract, liver, immune system) H315 H318 H400 H410 | GHS05 GHS08 GHS09 Dgr | H315 H318 H373 (gastro-intestinal tract, liver, immune system) H304 H410 |  | M=10 M=10 |  | ATP05
612-283-00-3 | (Z)-octadec-9-enylamine | 204-015-5 | 112-90-3 | Acute Tox. 4 Asp. Tox. 1 STOT SE 3 STOT RE 2 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H302 H304 H335 H373 (gastro-intestinal tract, liver, immune system) H314 H400 H410 | GHS05 GHS07 GHS08 GHS09 Dgr | H302 H314 H335 H373 (gastro-intestinal tract, liver, immune system) H304 H410 |  | M=10 M=10 |  | ATP05
612-284-00-9 | amines, hydrogenated tallow alkyl | 262-976-6 | 61788-45-2 | Asp. Tox. 1 STOT RE 2 Skin Irrit. 2 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H304 H373 (gastro-intestinal tract, liver, immune system) H315 H318 H400 H410 | GHS08 GHS05 GHS09 Dgr | H315 H318 H373 (gastro-intestinal tract, liver, immune system) H304 H410 |  | M=10 M=10 |  | ATP05
612-285-00-4 | amines, coco alkyl | 262-977-1 | 61788-46-3 | Acute Tox. 4 Asp. Tox. 1 STOT SE 3 STOT RE 2 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H302 H304 H335 H373 H314 H400 H410 | GHS05 GHS07 GHS08 GHS09 Dgr | H302 H314 H335 H373 (gastro-intestinal tract, liver, immune system) H304 H410 |  | M=10 M=10 |  | ATP05
612-286-00-X | amines, tallow alkyl | 263-125-1 | 61790-33-8 | Acute Tox. 4 Asp. Tox. 1 STOT RE 2 Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H302 H304 H373 (gastro-intestinal tract, liver, immune system) H314 H400 H410 | GHS05 GHS07 GHS08 GHS09 Dgr | H302 H314 H373 (gastro-intestinal tract, liver, immune system) H304 H410 |  | M=10 M=10 |  | ATP05
612-287-00-5 | fluazinam (ISO); 3-chloro-N-[3-chloro-2,6-dinitro-4-(trifluoromethyl)phenyl]-5-(trifluoromethyl)pyridin-2-amine |  | 79622-59-6 | Repr. 2 Acute Tox. 4 Eye Dam. 1 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 1 | H361d H332 H318 H317 H400 H410 | GHS08 GHS07 GHS05 GHS09 Dgr | H332 H318 H317 H361d H410 |  | M=10 M=10 |  | ATP06
612-288-00-0 | bupirimate (ISO); 5-butyl-2-ethylamino-6-methylpyrimidin-4-yl dimethylsulphamate | 255-391-2 | 41483-43-6 | Carc. 2 Skin Sens. 1B Aquatic Chronic 1 | H351 H317 H410 | GHS08 GHS07 GHS09 Wng | H317 H351 H410 |  | M=1 |  | ATP09
612-289-00-6 | triflumizole (ISO); (1E)-N-[4-chloro-2-(trifluoromethyl)phenyl]-1-(1H-imidazol-1-yl)-2-propoxyethanimine |  | 68694-11-1 | Repr. 1B Acute Tox. 4 STOT RE 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D H302 H373 (liver) H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H302 H317 H360D H373 (liver) H410 |  | M=1 M=1 |  | ATP09
612-290-00-1 | reaction products of paraformaldehyde and 2-hydroxypropylamine (ratio 3:2); [formaldehyde released from 3,3′-methylenebis[5-methyloxazolidine]; [formaldehyde released from oxazolidin]; [MBO] |  |  | Carc. 1B Muta. 2 Acute Tox. 3 Acute Tox. 4 Acute Tox. 4 STOT RE 2 Skin Corr. 1B Eye Dam. 1 Skin Sens. 1A Aquatic Chronic 2 | H350 H341 H311 H332 H302 H373 (gastrointestinal tract, respiratory tract) H314 H318 H317 H411 | GHS08 GHS06 GHS05 GHS09 Dgr | H332 H311 H302 H314 H317 H341 H350 H373 (gastrointestinal tract, respiratory tract) H411 | EUH071 |  | 8 9 | ATP10
612-291-00-7 | reaction products of paraformaldehyde with 2-hydroxypropylamine (ratio 1:1); [formaldehyde released from α,α,α-trimethyl-1,3,5-triazine-1,3,5(2H,4H,6H)-triethanol]; [HPT] |  |  | Carc. 1B Muta. 2 Acute Tox. 4 Acute Tox. 4 STOT RE 2 Skin Corr. 1C Eye Dam. 1 Skin Sens. 1A Aquatic Chronic 2 | H350 H341 H332 H302 H373 (gastrointestinal tract, respiratory tract) H314 H318 H317 H411 | GHS08 GHS07 GHS05 GHS09 Dgr | H332 H302 H314 H317 H341 H350 H373 (gastrointestinal tract, respiratory tract) H411 | EUH071 |  | 8 9 | ATP10
612-292-00-2 | methylhydrazine | 200-471-4 | 60-34-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | ATP10
612-293-00-8 | reaction mass of 1-[2-(2-aminobutoxy)ethoxy]but-2-ylamine and 1-({[2-(2-aminobutoxy)ethoxy]methyl}propoxy)but-2-ylamine | 447-920-2 |  | Repr. 2 Acute Tox. 4 Skin Corr. 1B Eye Dam. 1 | H361f H302 H314 H318 | GHS08 GHS07 GHS05 Dgr | H302 H314 H361f | EUH071 |  |  | ATP13
612-294-00-3 | mecetronium etilsul­fate; N-ethyl-N,N-dimethyl­hexadecan-1-aminium ethyl sulfate; mecetronium ethyl sulphate; [MES] | 221-106-5 | 3006-10-8 | Skin Corr. 1 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H314 H318 H400 H410 | GHS05 GHS09 Dgr | H314 H410 | EUH071 | M=100 M=1000 |  | ATP15
612-295-00-9 | benfluralin (ISO); N-butyl-N-ethyl-α,α,α-trifluoro-2,6-dinitro-p-toluidine | 217-465-2 | 1861-40-1 | Carc. 2 Repr. 2 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H361d H315 H319 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H361d H315 H319 H317 H410 |  | M = 10 M = 10 |  | ATP21
612-296-00-4 | N,N-dimethyl-p-toluidine | 202-805-4 | 99-97-8 | Carc. 1B Acute Tox. 3 Acute Tox. 4 STOT RE 2 Aquatic Chronic 3 | H350 H301 H332 H373 (blood system, respiratory tract) H412 | GHS06 GHS08 Dgr | H350 H332 H301 H373 (blood system, respiratory tract) H412 |  | inhalation: ATE = 1.4 mg/L(dusts or mists) oral: ATE = 140 mg/kg bw |  | ATP21
612-297-00-X | 1-phenylethan-1-one (1-phenylethylidene)hydrazone | 211-979-0 | 729-43-1 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP21
612-298-00-5 | 1,4-Benzenediamine, N,N’-mixed Ph and tolyl derivs.; | 273-227-8 | 68953-84-4 | Repr. 1B Skin Sens. 1 | H360FD H317 | GHS08 GHS07 Dgr | H360FD H317 |  |  |  | ATP21
612-299-00-0 | fenpropidin (ISO); (R,S)-1-[3-(4-tert-butylphenyl)-2-methylpropyl]piperidine | - | 67306-00-7 | Repr. 2 Acute Tox. 4 Acute Tox. 4 STOT SE 3 STOT SE 3 STOT RE 2 Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361d H332 H302 H335 H336 H373 (nervous system, eyes, lungs) H315 H318 H317 H400 H410 | GHS08 GHS07 GHS05 GHS09 Dgr | H361d H332 H302 H335 H336 H373 (nervous system, eyes, lungs) H315 H318 H317 H410 |  | Inhalation: ATE = 1,2 mg/L (dusts or mists) oral: ATE = 1 330 mg/kg bw M = 1000 M = 10000 |  | ATP22
613-001-00-1 | ethyleneimine; aziridine | 205-793-9 | 151-56-4 | Flam. Liq. 2 Carc. 1B Muta. 1B Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Skin Corr. 1B Aquatic Chronic 2 | H225 H350 H340 H310 H330 H300 H314 H411 | GHS02 GHS06 GHS08 GHS05 GHS09 Dgr | H225 H350 H340 H330 H310 H300 H314 H411 |  |  | D | CLP00
613-002-00-7 | pyridine | 203-809-9 | 110-86-1 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H225 H332 H312 H302 | GHS02 GHS07 Dgr | H225 H332 H312 H302 |  | * |  | CLP00
613-003-00-2 | 1,2,3,4-tetranitrocarbazole | - | 6202-15-9 | Expl. 1.1 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * | H201 H332 H312 H302 | GHS01 GHS07 Dgr | H201 H332 H312 H302 |  |  |  | CLP00/ATP01
613-004-00-8 | crimidine (ISO); 2-chloro-6-methylpyrimidin-4-yldimethylamine | 208-622-6 | 535-89-7 | Acute Tox. 2 * | H300 | GHS06 Dgr | H300 |  |  |  | CLP00
613-007-00-4 | desmetryne (ISO); 6-isopropylamino-2-methylamino-4-methylthio-1,3,5-triazine | 213-800-1 | 1014-69-3 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H400 H410 | GHS07 GHS09 Wng | H312 H302 H410 |  |  |  | CLP00
613-008-00-X | dazomet (ISO); tetrahydro-3,5-dimethyl-1,3,5-thiadiazine-2-thione | 208-576-7 | 533-74-4 | Acute Tox. 4 * Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H410 |  |  |  | CLP00
613-009-00-5 | 2,4,6-trichloro-1,3,5-triazine; cyanuric chloride | 203-614-9 | 108-77-0 | Acute Tox. 2 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 | H330 H302 H314 H317 | GHS06 GHS05 Dgr | H330 H302 H314 H317 | EUH014 | STOT SE 3; H335: C ≥ 5 % |  | CLP00
613-010-00-0 | ametryn (ISO); N-ethyl-N'-isopropyl-6-(methylthio)-1,3,5-triazine-2,4-diamine | 212-634-7 | 834-12-8 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  | M=100 |  | CLP00/ATP01corr
613-011-00-6 | amitrole (ISO); 1,2,4-triazol-3-ylamine | 200-521-5 | 61-82-5 | Repr. 2 STOT RE 2 * Aquatic Chronic 2 | H361d *** H373 ** H411 | GHS08 GHS09 Wng | H361d *** H373 ** H411 |  |  |  | CLP00
613-012-00-1 | bentazone (ISO); 3-isopropyl-2,1,3-benzothiadiazine-4-one-2,2-dioxide | 246-585-8 | 25057-89-0 | Repr. 2 Acute Tox. 4 Eye Irrit. 2 Skin Sens. 1 | H361d H302 H319 H317 | GHS08 GHS07 Wng | H361d H302 H319 H317 |  | oral: ATE = 1 600 mg/kg bw |  | CLP00/ATP18
613-013-00-7 | cyanazine (ISO); 2-(4-chloro-6-ethylamino-1,3,5-triazine-2-ylamino)-2-methylpropionitrile | 244-544-9 | 21725-46-2 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
613-014-00-2 | ethoxyquin (ISO); 6-ethoxy-1,2-dihydro-2,2,4-trimethylquinoline | 202-075-7 | 91-53-2 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
613-015-00-8 | fenazaflor (ISO); phenyl 5,6-dichloro-2-trifluoromethylbenzimidazole-1-carboxylate | 238-134-9 | 14255-88-0 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H400 H410 | GHS07 GHS09 Wng | H312 H302 H410 |  |  |  | CLP00
613-016-00-3 | fuberidazole (ISO); 2-(2-furyl)-1H-benzimidazole | 223-404-0 | 3878-19-1 | Carc. 2 Acute Tox. 4 STOT RE 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H373 (heart) H317 H400 H410 | GHS07 GHS08 GHS09 Wng | H351 H302 H373 (heart) H317 H410 |  | M=1 |  | CLP00/ATP03
613-017-00-9 | bis (8-hydroxyquinolinium) sulphate | 205-137-1 | 134-31-6 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
613-018-00-4 | morfamquat (ISO); 1,1'-bis(3,5-dimethylmorpholinocarbonylmethyl)-4,4'-bipyridilium ion |  | 7411-47-4 | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 3 | H302 H335 H315 H319 H412 | GHS07 Wng | H302 H319 H335 H315 H412 |  |  |  | CLP00
613-019-00-X | thioquinox (ISO); 2-thio-1,3-dithiolo(4,5,b)quinoxaline | 202-272-8 | 93-75-4 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
613-020-00-5 | tridemorph (ISO); 2,6-dimethyl-4-tridecylmorpholine | 246-347-3 | 24602-86-6 | Repr. 1B Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H360D *** H332 H302 H315 H400 H410 | GHS08 GHS07 GHS09 Dgr | H360D *** H332 H302 H315 H410 |  |  |  | CLP00
613-021-00-0 | dithianon (ISO); 5,10-dihydro-5,10-dioxonaphtho(2,3-b)(1,4)dithiazine-2,3-dicarbonitrile | 222-098-6 | 3347-22-6 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
613-022-00-6 | pyrethrins including cinerins, with the exception of those specified elsewhere in this Annex |  |  | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  |  | A | CLP00
613-023-00-1 | 2-methyl-4-oxo-3-(penta-2,4-dienyl)cyclopent-2-enyl [1R-[1α[S*(Z)],3β]]-chrysanthemate; pyrethrin I | 204-455-8 | 121-21-1 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  |  |  | CLP00
613-024-00-7 | 2-methyl-4-oxo-3-(penta-2,4-dienyl)cyclopent-2-enyl[1R-[1α[S*(Z)](3β)]]-3-(3-methoxy-2-methyl-3-oxoprop-1-enyl)-2,2-dimethylcyclopropanecarboxylate; pyrethrin II | 204-462-6 | 121-29-9 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  |  |  | CLP00
613-025-00-2 | cinerin I; 3-(but-2-enyl)-2-methyl-4-oxocyclopent-2-enyl 2,2-dimethyl-3-(2-methylprop-1-enyl)cyclopropanecarboxylate | 246-948-0 | 25402-06-6 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
613-026-00-8 | cinerin II; 3-(but-2-enyl)-2-methyl-4-oxocyclopent-2-enyl 2,2-dimethyl-3-(3-methoxy-2-methyl-3-oxoprop-1-enyl)cyclopropanecarboxylate | 204-454-2 | 121-20-0 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
613-027-00-3 | piperidine | 203-813-0 | 110-89-4 | Flam. Liq. 2 Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B | H225 H331 H311 H314 | GHS02 GHS06 GHS05 Dgr | H225 H331 H311 H314 |  | * |  | CLP00
613-028-00-9 | morpholine | 203-815-1 | 110-91-8 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H226 H332 H312 H302 H314 | GHS02 GHS05 GHS07 Dgr | H226 H332 H312 H302 H314 |  |  |  | CLP00
613-029-00-4 | dichloro-1,3,5-triazinetrione; dichloroisocyanuric acid | 220-487-5 | 2782-57-2 | Ox. Sol. 2 Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H272 H302 H335 H319 H400 H410 | GHS03 GHS07 GHS09 Dgr | H272 H302 H319 H335 H410 | EUH031 |  | T | CLP00
613-030-00-X | troclosene potassium [1] troclosene sodium [2] | 218-828-8 [1] 220-767-7 [2] | 2244-21-5 [1] 2893-78-9 [2] | Ox. Sol. 2 Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H272 H302 H335 H319 H400 H410 | GHS03 GHS07 GHS09 Dgr | H272 H302 H319 H335 H410 | EUH031 | * STOT SE 3; H335: C ≥ 10 % EUH031: C ≥ 10 % | G | CLP00/ATP01
613-030-01-7 | troclosene sodium, dihydrate | 220-767-7 | 51580-86-0 | Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H335 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H335 H410 | EUH031 |  |  | CLP00
613-031-00-5 | symclosene; trichloroisocyanuric acid; trichloro-1,3,5-triazinetrion | 201-782-8 | 87-90-1 | Ox. Sol. 2 Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H272 H302 H335 H319 H400 H410 | GHS03 GHS07 GHS09 Dgr | H272 H302 H319 H335 H410 | EUH031 |  |  | CLP00
613-032-00-0 | methyl-2,3,5,6-tetrachloro-4-pyridylsulphone; 2,3,5,6-tetrachloro-4-(methylsulphonyl)pyridine | 236-035-5 | 13108-52-6 | Acute Tox. 4 * Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 | H312 H302 H319 H317 | GHS07 Wng | H312 H302 H319 H317 |  |  |  | CLP00
613-033-00-6 | 2-methylaziridine; propyleneimine | 200-878-7 | 75-55-8 | Flam. Liq. 2 Carc. 1B Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Eye Dam. 1 Aquatic Chronic 2 | H225 H350 H310 H330 H300 H318 H411 | GHS02 GHS06 GHS08 GHS05 GHS09 Dgr | H225 H350 H330 H310 H300 H318 H411 |  | Carc. 1B; H350: C ≥ 0,01 % |  | CLP00
613-034-00-1 | 1,2-dimethylimidazole | 217-101-2 | 1739-84-0 | Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 | H302 H315 H318 | GHS05 GHS07 Dgr | H302 H315 H318 |  |  |  | CLP00
613-035-00-7 | 1-methylimidazole | 210-484-7 | 616-47-7 | Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1B | H312 H302 H314 | GHS05 GHS07 Dgr | H312 H302 H314 |  |  |  | CLP00
613-036-00-2 | 2-methylpyridine; 2-picoline | 203-643-7 | 109-06-8 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Eye Irrit. 2 | H226 H332 H312 H302 H335 H319 | GHS02 GHS07 Wng | H226 H332 H312 H302 H319 H335 |  |  |  | CLP00
613-037-00-8 | 4-methylpyridine; 4-picoline | 203-626-4 | 108-89-4 | Flam. Liq. 3 Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H226 H311 H332 H302 H335 H315 H319 | GHS02 GHS06 Dgr | H226 H311 H332 H302 H319 H335 H315 |  |  |  | CLP00
613-038-00-3 | 6-phenyl-1,3,5-triazine-2,4-diyldiamine; 6-phenyl-1,3,5-triazine-2,4-diamine; benzoguanamine | 202-095-6 | 91-76-9 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
613-039-00-9 | ethylene thiourea; imidazolidine-2-thione; 2-imidazoline-2-thiol | 202-506-9 | 96-45-7 | Repr. 1B Acute Tox. 4 * | H360D *** H302 | GHS08 GHS07 Dgr | H360D *** H302 |  |  |  | CLP00
613-040-00-4 | azaconazole (ISO); 1-{}{[2-(2,4-dichlorophenyl)-1,3-dioxolan-2-yl]methyl}}-1H-1,2.4-triazole | 262-102-3 | 60207-31-0 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
613-041-00-X | morpholine-4-carbonyl chloride | 239-213-0 | 15159-40-7 | Carc. 2 Skin Irrit. 2 Eye Irrit. 2 | H351 H315 H319 | GHS08 Wng | H351 H319 H315 | EUH014 |  |  | CLP00
613-042-00-5 | imazalil (ISO); 1-[2-(allyloxy)-2-(2,4-dichlorophenyl)ethyl]-1H-imidazole | 252-615-0 | 35554-44-0 | Carc. 2 Acute Tox. 3 Acute Tox. 4 Eye Dam. 1 Aquatic Chronic 1 | H351 H301 H332 H318 H410 | GHS08 GHS06 GHS05 GHS09 Dgr | H301 H332 H318 H351 H410 |  | M=10 |  | CLP00/ATP07
613-043-00-0 | imazalil sulphate (ISO) powder; 1- [2-(allyloxy)ethyl-2-(2,4-dichlorophenyl)]-1H-imidazolium hydrogen sulphate [1] (±)-1- [2-(allyloxy)ethyl-2-(2,4-dichlorophenyl)]-1H-imidazolium hydrogen sulphate [2] | 261-351-5 [1] 281-291-3 [2] | 58594-72-2 [1] 83918-57-4 [2] | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
613-043-01-8 | imazalil sulphate (ISO), aqueous solution; 1- [2-(allyloxy)ethyl-2-(2,4-dichlorophenyl)]-1H-imidazolium hydrogen sulphate [1] (±)-1- [2-(allyloxy)ethyl-2-(2,4-dichlorophenyl)]-1H-imidazolium hydrogen sulphate [2] | 261-351-5 [1] 281-291-3 [2] | 58594-72-2 [1] 83918-57-4 [2] | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H317 H400 H410 | GHS05 GHS07 GHS09 Wng | H302 H314 H317 H410 |  | Skin Corr. 1B; H314: C ≥ 50 % Skin Irrit. 2; H315: 30 % ≤ C < 50 % Eye Dam. 1; H318: 15 % ≤ C < 50 % Eye Irrit. 2; H319: 5 % ≤ C < 15 % |  | CLP00
613-044-00-6 | captan (ISO); 1,2,3,6-tetrahydro-N-(trichloromethylthio)phthalimide | 205-087-0 | 133-06-2 | Carc. 2 Acute Tox. 3 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 | H351 H331 H318 H317 H400 | GHS06 GHS05 GHS08 GHS09 Dgr | H351 H331 H318 H317 H400 |  | M=10 |  | CLP00/ATP01
613-045-00-1 | folpet (ISO); N-(trichloromethylthio)phthalimide | 205-088-6 | 133-07-3 | Carc. 2 Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 | H351 H332 H319 H317 H400 | GHS08 GHS07 GHS09 Wng | H351 H332 H319 H317 H400 |  | M=10 |  | CLP00/ATP01
613-046-00-7 | captafol (ISO); 1,2,3,6-tetrahydro-N-(1,1,2,2-tetrachloroethylthio)phthalimide | 219-363-3 | 2425-06-1 | Carc. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H350 H317 H400 H410 | GHS08 GHS09 Dgr | H350 H317 H410 |  |  |  | CLP00
613-047-00-2 | 1-dimethylcarbamoyl-5-methylpyrazol-3-yl dimethylcarbamate; dimetilan (ISO) | 211-420-0 | 644-64-4 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H312 H400 H410 | GHS06 GHS09 Dgr | H301 H312 H410 |  |  |  | CLP00
613-048-00-8 | carbendazim (ISO); methyl benzimidazol-2-ylcarbamate | 234-232-0 | 10605-21-7 | Muta. 1B Repr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H340 H360FD H317 H400 H410 | GHS07 GHS08 GHS09 Dgr | H317 H340 H360FD H410 |  | M=10 M=10 |  | CLP00/ATP17
613-049-00-3 | benomyl (ISO); methyl 1-(butylcarbamoyl)benzimidazol-2-ylcarbamate | 241-775-7 | 17804-35-2 | Muta. 1B Repr. 1B STOT SE 3 Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H340 H360FD H335 H315 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H340 H360FD H335 H315 H317 H410 |  | M=10 |  | CLP00
613-050-00-9 | carbadox (INN); methyl 3-(quinoxalin-2-ylmethylene)carbazate 1,4-dioxide; 2-(methoxycarbonylhydrazonomethyl)quinoxaline 1,4-dioxide | 229-879-0 | 6804-07-5 | Flam. Sol. 1 Carc. 1B Acute Tox. 4 * | H228 H350 H302 | GHS02 GHS08 GHS07 Dgr | H228 H350 H302 |  |  | T | CLP00
613-051-00-4 | molinate (ISO); S-ethyl 1-perhydroazepinecarbothioate; S-ethyl perhydroazepine-1-carbothioate | 218-661-0 | 2212-67-1 | Carc. 2 Repr. 2 Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H361f *** H332 H302 H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H361f *** H332 H302 H373 ** H317 H410 |  | M=100 |  | CLP00
613-052-00-X | trifenmorph (ISO); 4-tritylmorpholine | 215-812-2 | 1420-06-0 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
613-053-00-5 | anilazine (ISO); 2-chloro-N-(4,6-dichloro-1,3,5-triazin-2-yl)aniline | 202-910-5 | 101-05-3 | Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H315 H319 H400 H410 | GHS07 GHS09 Wng | H319 H315 H410 |  |  |  | CLP00
613-054-00-0 | thiabendazole (ISO); 2-(thiazol-4-yl)benzimidazole | 205-725-8 | 148-79-8 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1 M=1 |  | CLP00/ATP14
613-056-00-1 | 1,2-dimethyl-3,5-diphenylpyrazolium methylsulphate; difenzoquat methyl sulfate | 256-152-5 | 43222-48-6 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS09 Wng | H302 H410 |  |  |  | CLP00
613-057-00-7 | dodemorph (ISO); 4-cyclododecyl-2,6-dimethylmorpholine | 216-474-9 | 1593-77-7 | Repr. 2 STOT RE 2 Skin Corr. 1C Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 1 | H361d H373 (liver) H314 H317 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H314 H317 H361d H373 (liver) H410 | EUH071 | M=1 M=1 |  | CLP00/ATP07
613-058-00-2 | permethrin (ISO); m-phenoxybenzyl 3-(2,2-dichlorovinyl)-2,2-dimethylcyclopropanecarboxylate | 258-067-9 | 52645-53-1 | Acute Tox. 4 * Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H317 H400 H410 | GHS07 GHS09 Wng | H332 H302 H317 H410 |  | M=1000 |  | CLP00
613-059-00-8 | profluralin (ISO); N-(cyclopropylmethyl)-α,α,α-trifluoro-2,6-dinitro-N-propyl-p-toluidine | 247-656-6 | 26399-36-0 | Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H319 H400 H410 | GHS07 GHS09 Wng | H319 H410 |  |  |  | CLP00
613-060-00-3 | resmethrin (ISO); 5-benzyl-3-furylmethyl (±)-cis-trans-chrysanthemate | 233-940-7 | 10453-86-8 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  | M=1000 |  | CLP00/ATP01
613-061-00-9 | 6-(1α,5aβ,8aβ,9-pentahydroxy-7β-isopropyl-2β,5β,8β-trimethylperhydro-8bα,9-epoxy-5,8-ethanocyclopenta[1,2-b]indenyl) pyrrole-2-carboxylate; ryania | 239-732-2 | 15662-33-6 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H400 H410 | GHS07 GHS09 Wng | H312 H302 H410 |  |  |  | CLP00
613-062-00-4 | sabadilla (ISO); veratrine |  | 8051-02-3 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H335 H315 H319 | GHS07 Wng | H319 H335 H315 |  |  |  | CLP00
613-063-00-X | secbumeton (ISO); 2-sec-butylamino-4-ethylamino-6-methoxy-1,3,5-triazine | 247-554-1 | 26259-45-0 | Acute Tox. 4 * Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H410 |  |  |  | CLP00
613-064-00-5 | 5-(3,6,9-trioxa-2-undecyloxy)benzo(d)-1,3-dioxolane; sesamex |  | 51-14-9 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
613-065-00-0 | simetryn (ISO); 2,4-bis(ethylamino)-6-methylthio-1,3,5-triazine | 213-801-7 | 1014-70-6 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
613-066-00-6 | terbumeton (ISO); 2-tert-butylamino-4-ethylamino-6-methoxy-1,3,5-triazine | 251-637-8 | 33693-04-8 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
613-067-00-1 | propazine (ISO); 2-chloro-4,6-bis(isopropylamino)-1,3,5-triazine | 205-359-9 | 139-40-2 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  |  |  | CLP00
613-068-00-7 | atrazine (ISO); 2-chloro-4-ethylamine-6-isopropylamine-1,3,5-triazine | 217-617-8 | 1912-24-9 | STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H317 H400 H410 | GHS08 GHS09 Wng | H373 ** H317 H410 |  |  |  | CLP00
613-069-00-2 | ε-caprolactam | 203-313-2 | 105-60-2 | Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 | H332 H302 H335 H315 H319 | GHS07 Wng | H332 H302 H319 H335 H315 |  |  |  | CLP00
613-070-00-8 | propylenethiourea |  | 2122-19-2 | Repr. 2 Acute Tox. 4 * Aquatic Chronic 3 | H361d *** H302 H412 | GHS08 GHS07 Wng | H361d *** H302 H412 |  |  |  | CLP00
613-071-00-3 | 2-fluoro-5-trifluoromethylpyridine | 400-290-2 | 69045-82-5 | Flam. Liq. 3 Skin Sens. 1 Aquatic Chronic 3 | H226 H317 H412 | GHS02 GHS07 Wng | H226 H317 H412 |  |  |  | CLP00
613-072-00-9 | N,N-bis(2-ethylhexyl)-((1,2,4-triazol-1-yl)methyl)amine | 401-280-0 | 91273-04-0 | Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H314 H317 H411 | GHS05 GHS07 GHS09 Dgr | H314 H317 H411 |  |  |  | CLP00
613-073-00-4 | N,N-dimethyl-2-(3-(4-chlorophenyl)-4,5-dihydropyrazol-1-ylphenylsulphonyl)ethylamine | 401-410-6 | 10357-99-0 | STOT RE 2 * Skin Sens. 1 Aquatic Chronic 2 | H373 ** H317 H411 | GHS08 GHS09 Wng | H373 ** H317 H411 |  |  |  | CLP00
613-074-00-X | 3-(3-methylpent-3-yl)isoxazol-5-ylamine | 401-460-9 | 82560-06-3 | Acute Tox. 3 * Acute Tox. 3 * Eye Dam. 1 Aquatic Chronic 3 | H331 H301 H318 H412 | GHS06 GHS05 Dgr | H331 H301 H318 H412 |  |  |  | CLP00
613-075-00-5 | 1,3-dichloro-5-ethyl-5-methylimidazolidine-2,4-dione | 401-570-7 | 89415-87-2 | Ox. Sol. 1 **** Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 | H271 H331 H302 H314 H317 H400 | GHS03 GHS06 GHS05 GHS09 Dgr | H271 H331 H314 H302 H317 H400 |  |  |  | CLP00
613-076-00-0 | 3-chloro-5-trifluoromethyl-2-pyridylamine | 401-670-0 | 79456-26-1 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
613-077-00-6 | reaction mass of 5-heptyl-1,2,4-triazol-3-ylamine and 5-nonyl-1,2,4-triazol-3-ylamine | 401-940-8 |  | Acute Tox. 4 * Eye Irrit. 2 Aquatic Chronic 2 | H302 H319 H411 | GHS07 GHS09 Wng | H302 H319 H411 |  |  |  | CLP00
613-078-00-1 | N,N,N,N-tetrakis(4,6-bis(butyl-(N-methyl-2,2,6,6-tetramethylpiperidin-4-yl)amino)triazin-2-yl)-4,7-diazadecane-1,10-diamine | 401-990-0 | 106990-43-6 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
613-079-00-7 | 4-(1(or 4 or 5 or 6)-methyl-8,9,10-trinorborn-5-en-2-yl)pyridine, reaction mass of isomers | 402-520-7 |  | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H315 H317 H400 H410 | GHS07 GHS09 Wng | H312 H302 H315 H317 H410 |  |  |  | CLP00
613-080-00-2 | 3-(bis(2-ethylhexyl)aminomethyl)benzothiazole-2(3H)-thione | 402-540-6 | 105254-85-1 | Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H314 H317 H410 |  |  |  | CLP00
613-081-00-8 | 1-butyl-2-methylpyridinium bromide | 402-680-8 | 26576-84-1 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
613-082-00-3 | 2-methyl-1-pentylpyridinium bromide | 402-690-2 |  | Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 3 | H312 H302 H412 | GHS07 Wng | H312 H302 H412 |  |  |  | CLP00
613-083-00-9 | 2-(4-(3-(4-chlorophenyl)-2-pyrazolin-1-yl)phenylsulfonyl)ethyldimethylammonium formate | 402-120-2 |  | STOT RE 2 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H314 H317 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H314 H373 ** H317 H410 |  |  |  | CLP00
613-084-00-4 | 2-(4-(3-(4-chlorophenyl)-4,5-dihydropyrazolyl)phenylsulphonyl)ethyldimethylammonium hydrogen phosphonate | 402-490-5 | 106359-93-7 | Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H319 H400 H410 | GHS07 GHS09 Wng | H319 H410 |  |  |  | CLP00
613-085-00-X | reaction mass of 1,1'-(methylenebis(4,1-phenylene))dipyrrole-2,5-dione and N-(4-(4-(2,5-dioxopyrrol-1-yl)benzyl)phenyl)acetamide and 1-(4-(4-(5-oxo-2H-2-furylidenamino)benzyl)phenyl)pyrrole-2,5-dione | 401-970-1 |  | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
613-086-00-5 | caffeine | 200-362-1 | 58-08-2 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
613-087-00-0 | tetrahydrothiophene | 203-728-9 | 110-01-0 | Flam. Liq. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 3 | H225 H332 H312 H302 H315 H319 H412 | GHS02 GHS07 Dgr | H225 H332 H312 H302 H319 H315 H412 |  |  |  | CLP00
613-088-00-6 | 1,2-benzisothiazol-3(2H)-one; 1,2-benzisothiazolin-3-one | 220-120-9 | 2634-33-5 | Acute Tox. 2 Acute Tox. 4 Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 1 | H330 H302 H315 H318 H317 H400 H410 | GHS06 GHS05 GHS09 Dgr | H330 H302 H315 H318 H317 H410 |  | inhalation: ATE = 0.21 mg/L(dusts or mists) oral: ATE = 450 mg/kg bw Skin Sens. 1A; H317: C >= 0.036 % M = 1 M = 1 |  | CLP00/ATP21
613-089-00-1 | diquat dibromide [1] diquat dichloride [2] 6,7-dihydrodipyrido[1,2-α:2',1'-c]pyrazinediylium dihydroxide [3] | 201-579-4 [1] 223-714-6 [2] 301-467-6 [3] | 85-00-7 [1] 4032-26-2 [2] 94021-76-8 [3] | Acute Tox. 2 * Acute Tox. 4 * STOT SE 3 STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H302 H335 H372 ** H315 H319 H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H372 ** H302 H319 H335 H315 H317 H410 |  |  |  | CLP00
613-090-00-7 | paraquat dichloride; 1,1-dimethyl-4,4'-bipyridinium dichloride [1] paraquat dimethylsulfate; 1,1-dimethyl-4,4'-bipyridinium dimethyl sulphate [2] | 217-615-7 [1] 218-196-3 [2] | 1910-42-5 [1] 2074-50-2 [2] | Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H330 H311 H301 H335 H372 ** H315 H319 H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H311 H301 H372 ** H319 H335 H315 H410 |  |  |  | CLP00
613-091-00-2 | morfamquat dichloride [1] morfamquat sulfate [2] | 225-062-8 [1] | 4636-83-3 [1] 29873-36-7 [2] | Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 3 | H302 H335 H315 H319 H412 | GHS07 Wng | H302 H319 H335 H315 H412 |  |  |  | CLP00
613-092-00-8 | 1,10-phenanthroline | 200-629-2 | 66-71-7 | Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H400 H410 | GHS06 GHS09 Dgr | H301 H410 |  |  |  | CLP00
613-093-00-3 | hexasodium 6,13-dichloro-3,10-bis((4-(2,5-disulfonatoanilino)-6-fluoro-1,3,5-triazin-2-ylamino)prop-3-ylamino)-5,12-dioxa-7,14-diazapentacene-4,11-disulfonate | 400-050-7 | 85153-92-0 | Resp. Sens. 1 Skin Sens. 1 | H334 H317 | GHS08 Dgr | H334 H317 |  |  |  | CLP00
613-094-00-9 | 4-methoxy-N,6-dimethyl-1,3,5-triazin-2-ylamine | 401-360-5 | 5248-39-5 | Acute Tox. 4 * STOT RE 2 * | H302 H373 ** | GHS08 GHS07 Wng | H302 H373 ** |  |  |  | CLP00
613-095-00-4 | sodium 3-(2H-benzotriazol-2-yl)-5-sec-butyl-4-hydroxybenzenesulfonate | 403-080-9 | 92484-48-5 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
613-096-00-X | 2-amino-6-ethoxy-4-methylamino-1,3,5-triazine | 403-580-7 | 62096-63-3 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
613-097-00-5 | 7-amino-3-((5-carboxymethyl-4-methyl-1,3-thiazol-2-ylthio)methyl)-8-oxo-5-thia-1-azabicyclo(4.2.0)oct-2-ene-2-carboxylic acid | 403-690-5 | 111298-82-9 | Resp. Sens. 1 Skin Sens. 1 Aquatic Chronic 3 | H334 H317 H412 | GHS08 Dgr | H334 H317 H412 |  |  |  | CLP00
613-098-00-0 | N-(n-octyl)-2-pyrrolidone | 403-700-8 | 2687-94-7 | Skin Corr. 1B Aquatic Chronic 2 | H314 H411 | GHS05 GHS09 Dgr | H314 H411 |  |  |  | CLP00
613-099-00-6 | 1-dodecyl-2-pyrrolidone | 403-730-1 | 2687-96-9 | Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H314 H317 H410 |  |  |  | CLP00
613-100-00-X | 2,9-bis(3-(diethylamino)propylsulfamoyl)quino(2,3-b)acridine-7,14-dione | 404-230-6 |  | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
613-101-00-5 | N-tert-pentyl-2-benzothiazolesulfenamide | 404-380-2 | 110799-28-5 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
613-102-00-0 | dimethomorph (ISO); (E,Z)- 4-(3-(4-chlorophenyl)- 3-(3,4-dimethoxyphenyl) acryloyl)morpholine | 404-200-2 | 110488-70-5 | Repr. 1B Aquatic Chronic 2 | H360F H411 | GHS08 GHS09 Dgr | H360F H411 |  |  |  | CLP00/ATP17
613-103-00-6 | sodium 5-n-butylbenzotriazole | 404-450-2 | 118685-34-0 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H302 H314 H317 H411 | GHS05 GHS07 GHS09 Dgr | H302 H314 H317 H411 |  |  |  | CLP00
613-104-00-1 | 5-tert-butyl-3-isoxazolylamine hydrochloride | 404-840-2 |  | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Aquatic Chronic 3 | H302 H373 ** H318 H412 | GHS08 GHS05 GHS07 Dgr | H302 H373 ** H318 H412 |  |  |  | CLP00
613-105-00-7 | hexakis(tetramethylammonium) 4,4'-vinylenebis((3-sulfonato-4,1-phenylene)imino(6-morpholino-1,3,5-triazine-4,2-diyl)imino)bis(5-hydroxy-6-phenylazonaphthalene-2,7-disulfonate) | 405-160-9 | 124537-30-0 | Acute Tox. 3 * Skin Sens. 1 Aquatic Chronic 3 | H301 H317 H412 | GHS06 Dgr | H301 H317 H412 |  |  |  | CLP00
613-106-00-2 | tetrapotassium 2-(4-(5-(1-(2,5-disulfonatophenyl)-3-ethoxycarbonyl-5-hydroxypyrazol-4-yl)penta-2,4-dienylidene)-3-ethoxycarbonyl-5-oxo-2-pyrazolin-1-yl)benzene-1,4-disulfonate | 405-240-3 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
613-107-00-8 | hexasodium 2,2'-vinylenebis((3-sulfonato-4,1-phenylene)imino(6-(N-cyanoethyl-N-(2-hydroxypropyl)amino)-1,3,5-triazine-4,2-diyl)imino)dibenzene-1,4-disulfonate | 405-280-1 | 76508-02-6 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | CLP00
613-108-00-3 | benzothiazole-2-thiol | 205-736-8 | 149-30-4 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
613-109-00-9 | bis(piperidinothiocarbonyl) disulphide | 202-328-1 | 94-37-1 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H335 H315 H319 H317 | GHS07 Wng | H319 H335 H315 H317 |  |  |  | CLP00
613-110-00-4 | dimepiperate (ISO); S-(1-methyl-1-phenylethyl) piperidine-1-carbothioate | 262-784-2 | 61432-55-1 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
613-111-00-X | 1,2,4-triazole | 206-022-9 | 288-88-0 | Repr. 1B Acute Tox. 4 Eye Irrit. 2 | H360FD H302 H319 | GHS08 GHS07 Dgr | H302 H319 H360FD |  | Oral: ATE = 1320 mg/kg bw |  | CLP00/ATP17
613-112-00-5 | octhilinone (ISO); 2-octyl-2H-isothiazol-3-one; [OIT] | 247-761-7 | 26530-20-1 | Acute Tox. 2 Acute Tox. 3 Acute Tox. 3 Skin Corr. 1 Eye Dam. 1 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 1 | H330 H311 H301 H314 H318 H317 H400 H410 | GHS06 GHS05 GHS09 Dgr | H330 H311 H301 H314 H317 H410 | EUH071 | Inhalation: ATE = 0.27 mg/L (dusts/mists) Dermal: ATE = 311 mg/kg bw Oral: ATE = 125 mg/kg bw Skin Sens. 1A; : C ≥ ,0015 % M=100 M=100 |  | CLP00/ATP15
613-113-00-0 | 2-(morpholinothio)benzothiazole | 203-052-4 | 102-77-2 | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H315 H319 H317 H411 | GHS07 GHS09 Wng | H319 H315 H317 H411 |  |  |  | CLP00
613-114-00-6 | 2,2',2"-(hexahydro-1,3,5-triazine-1,3,5-triyl)triethanol; 1,3,5-tris(2-hydroxyethyl)hexahydro-1,3,5-triazine | 225-208-0 | 4719-04-4 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  | Skin Sens. 1; H317: C ≥ 0,1 % |  | CLP00
613-115-00-1 | hymexazol (ISO); 3-hydroxy-5-methylisoxazole | 233-000-6 | 10004-44-1 | Repr. 2 Acute Tox. 4 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H361d H302 H318 H317 H411 | GHS08 GHS07 GHS05 GHS09 Dgr | H302 H318 H317 H361d H411 |  | Oral: ATE = 1600 mg/kg bw |  | CLP00/ATP15
613-116-00-7 | tolylfluanid (ISO); dichloro-N- [(dimethylamino)sulphonyl]fluoro-N-(p-tolyl)methanesulphenamide; [containing ≥ 0.1 % (w/w) of particles with an aerodynamic diameter of below 50 µm] | 211-986-9 | 731-27-1 | Acute Tox. 2 * STOT SE 3 STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 | H330 H335 H372 ** H315 H319 H317 H400 | GHS06 GHS08 GHS09 Dgr | H330 H315 H319 H317 H335 H372 ** H400 |  | M=10 |  | CLP00/ATP01
613-116-01-4 | tolylfluanid (ISO); dichloro-N-[(dimethylamino)sulphonyl]fluoro-N-(p-tolyl)methanesulphenamide; [containing < 0.1% (w/w) of particles with an aerodynamic diameter of below 50 μm] | 211-986-9 | 731-27-1 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 | H335 H315 H319 H317 H400 | GHS07 GHS09 Wng | H319 H335 H315 H317 H400 |  | M=10 |  | ATP01
613-117-00-2 | diniconazole (ISO); (E)-β-[(2,4-dichlorophenyl)methylene]-α-(1,1-dimethylethyl)-1H-1,2,4-triazol-1-ethanol; (E)-(RS)-1-(2,4-dichlorophenyl)-4,4-dimethyl-2-(1H-1,2,4-triazol-1-yl)pent-1-en-3-ol |  | 76714-88-0 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
613-118-00-8 | flubenzimine (ISO); N-[3-phenyl-4,5-bis[(trifluoromethyl)imino]thiazolidin-2-ylidene]aniline | 253-703-1 | 37893-02-0 | Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H319 H400 H410 | GHS07 GHS09 Wng | H319 H410 |  |  |  | CLP00
613-119-00-3 | (benzothiazol-2-ylthio)methyl thiocyanate; TCMTB | 244-445-0 | 21564-17-0 | Acute Tox. 2 * Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H302 H315 H319 H317 H400 H410 | GHS06 GHS09 Dgr | H330 H302 H319 H315 H317 H410 |  |  |  | CLP00
613-120-00-9 | bioresmethrin (ISO); (5-benzyl-3-furyl)methyl (1R)-2,2-dimethyl-3-(2-methylprop-1-en-1-yl)cyclopropanecarboxylate | 249-014-0 | 28434-01-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1000 |  | CLP00/ATP01corr
613-121-00-4 | chlorsulfuron (ISO); 2-chloro-N-[[(4-methoxy-6-methyl-1,3,5-triazin-2-yl)amino]carbonyl]benzenesulphonamide | 265-268-5 | 64902-72-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1000 M=100 |  | CLP00/ATP09
613-122-00-X | diclobutrazole (ISO); (R*, R*)-(±)-β-[(2,4-dichlorophenyl)methyl]-α-(1,1-dimethylethyl)-1H-1,2,4-triazole-1-ethanol; (2RS, 3RS)-1-(2,4-dichlorophenyl)-4,4-dimethyl-2-(1H-1,2,4-triazol-1yl)pentan-3-ol |  | 75736-33-3 | Eye Irrit. 2 Aquatic Chronic 2 | H319 H411 | GHS07 GHS09 Wng | H319 H411 |  |  |  | CLP00
613-123-00-5 | 5,6-dihydro-3H-imidazo[2,1-c]-1,2,4-dithiazole-3-thione; etem | 251-684-4 | 33813-20-6 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
613-124-00-0 | fenpropimorph (ISO); cis-4-[3-(p-tert-butylphenyl)-2-methylpropyl]-2,6-dimethylmorpholine | 266-719-9 | 67564-91-4 | Repr. 2 Acute Tox. 4 * Skin Irrit. 2 Aquatic Chronic 2 | H361d *** H302 H315 H411 | GHS08 GHS07 GHS09 Wng | H361d *** H302 H315 H411 |  |  |  | CLP00
613-125-00-6 | hexythiazox (ISO); trans-5-(4-chlorophenyl)-N-cyclohexyl-4-methyl-2-oxo-3-thiazolidine-carboxamide |  | 78587-05-0 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1 M=1 |  | CLP00/ATP15
613-126-00-1 | imazapyr (ISO); 2-[4,5-dihydro-4-methyl-4-(1-methylethyl)-5-oxo-1H-imidazol-2-yl]-3-pyridine carboxylate |  | 81334-34-1 | Eye Irrit. 2 Aquatic Chronic 3 | H319 H412 | GHS07 Wng | H319 H412 |  |  |  | CLP00
613-127-00-7 | mepiquat chloride (ISO); 1,1-dimethylpiperidinium chloride | 246-147-6 | 24307-26-4 | Acute Tox. 3 Acute Tox. 4 Aquatic Chronic 3 | H301 H332 H412 | GHS06 Dgr | H332 H301 H412 |  | inhalation: ATE = 2.8 mg/L(dusts or mists) oral: ATE = 270 mg/kg bw |  | CLP00/ATP21
613-128-00-2 | prochloraz (ISO); N-propyl-N-[2-(2,4,6-trichlorophenoxy)ethyl]-1H-imidazole-1-carboxamide | 266-994-5 | 67747-09-5 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
613-129-00-8 | metamitron (ISO); 4-amino-3-methyl-6-phenyl-1,2,4-triazin-5-one | 255-349-3 | 41394-05-2 | Acute Tox. 4 * Aquatic Acute 1 | H302 H400 | GHS07 GHS09 Wng | H302 H400 |  |  |  | CLP00
613-131-00-9 | pyroquilon (ISO); 1,2,5,6-tetrahydropyrrolo[3,2,1-ij]quinolin-4-one |  | 57369-32-1 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
613-132-00-4 | hexazinone (ISO); 3-cyclohexyl-6-dimethylamino-1-methyl-1,2,3,4-tetrahydro-1,3,5-triazine-2,4-dione | 257-074-4 | 51235-04-2 | Acute Tox. 4 * Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H319 H400 H410 | GHS07 GHS09 Wng | H302 H319 H410 |  |  |  | CLP00
613-133-00-X | etridiazole (ISO); 5-ethoxy-3-trichloromethyl-1,2,4-thiadiazole | 219-991-8 | 2593-15-9 | Carc. 2 Acute Tox. 4 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H317 H351 H410 |  | M=1 M=1 |  | CLP00/ATP07
613-134-00-5 | myclobutanil (ISO); 2-(4-chlorophenyl)-2-(1H-1,2,4-triazol-1-ylmethyl)hexanenitrile |  | 88671-89-0 | Repr. 2 Acute Tox. 4 * Eye Irrit. 2 Aquatic Chronic 2 | H361d *** H302 H319 H411 | GHS08 GHS07 GHS09 Wng | H361d *** H302 H319 H411 |  |  |  | CLP00
613-135-00-0 | di(benzothiazol-2-yl) disulphide | 204-424-9 | 120-78-5 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 | EUH031 |  |  | CLP00
613-136-00-6 | N-cyclohexylbenzothiazole-2-sulphenamide | 202-411-2 | 95-33-0 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
613-137-00-1 | methabenzthiazuron (ISO); 1-(1,3-benzothiazol-2-yl)1,3-dimethylurea | 242-505-0 | 18691-97-9 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
613-138-00-7 | quinoxyfen (ISO); 5,7-dichloro-4-(4-fluorophenoxy)quinoline |  | 124495-18-7 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
613-139-00-2 | metsulfuron-methyl (ISO); methyl 2-{[(4-methoxy-6-methyl-1,3,5-triazin-2-yl)carbamoyl]sulfamoyl}benzoate |  | 74223-64-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1000 |  | CLP00/ATP01corr
613-140-00-8 | cycloheximide (ISO); 4-{}{(2R)-2-[(1S,3S,5S)-3,5-dimethyl-2-oxocyclohexyl]-2-hydroxyethyl}}piperidine-2,6-dione | 200-636-0 | 66-81-9 | Muta. 2 Repr. 1B Acute Tox. 2 * Aquatic Chronic 2 | H341 H360D *** H300 H411 | GHS06 GHS08 GHS09 Dgr | H341 H360D *** H300 H411 |  |  |  | CLP00
613-141-00-3 | 1,4-diamino-2-(2-butyltetrazol-5-yl)-3-cyanoanthraquinone | 401-470-3 | 93686-63-6 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
613-142-00-9 | trans-N-methyl-2-styryl-[4'-aminomethine-(1-acetyl-1-(2-methoxyphenyl)acetamido)]pyridinium acetate | 405-860-4 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
613-143-00-4 | 1-(3-phenylpropyl)-2-methylpyridinium bromide | 405-930-4 | 10551-42-5 | Acute Tox. 4 * Eye Irrit. 2 Aquatic Chronic 3 | H302 H319 H412 | GHS07 Wng | H302 H319 H412 |  |  |  | CLP00
613-144-00-X | Reaction products of: poly(vinyl acetate), partially hydrolyzed, with (E)-2-(4-formylstyryl)-3,4-dimethylthiazoliummethyl sulfate | 406-460-2 | 125139-08-4 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
613-145-00-5 | (S)-3-benzyloxycarbonyl-1,2,3,4-tetrahydro-isoquinolinium 4-methylbenzenesulfonate | 406-960-0 | 77497-97-3 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
613-146-00-0 | N-ethyl-N-methylpiperidinium iodide | 407-780-5 | 4186-71-4 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
613-147-00-6 | 4-[2-(1-methyl-2-(4-morpholinyl)ethoxy)ethyl]morpholine | 407-940-4 | 111681-72-2 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
613-148-00-1 | tetrasodium 1,2-bis(4-fluoro-6-[5-(1-amino-2-sulfonatoanthrachinon-4-ylamino)-2,4,6-trimethyl-3-sulfonatophenylamino]-1,3,5-triazin-2-ylamino)ethane | 411-240-4 | 143683-23-2 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
613-149-00-7 | pyridaben (ISO); 2-tert-butyl-5-(4-tert-butylbenzylthio)-4-chloropyridazin-3(2H)-one | 405-700-3 | 96489-71-3 | Acute Tox. 3 Acute Tox. 3 Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H301 H410 |  | M=1000 M=1000 |  | CLP00/ATP07
613-150-00-2 | 2,2'-[3,3'-(piperazine-1,4-diyl)dipropyl]bis(1H-benzimidazo[2,1-b]benzo[l,m,n][3,8]phenanthroline-1,3,6-trione | 406-295-6 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
613-151-00-8 | 1-(3-mesyloxy-5-trityloxymethyl-2-D-threofuryl)thymine | 406-360-9 | 104218-44-2 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
613-152-00-3 | phenyl N-(4,6-dimethoxypyrimidin-2-yl)carbamate | 406-600-2 | 89392-03-0 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
613-153-00-9 | 2,3,5-trichloropyridine | 407-270-2 | 16063-70-0 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
613-154-00-4 | 2-amino-4-chloro-6-methoxypyrimidine | 410-050-9 | 5734-64-5 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
613-155-00-X | 5-chloro-2,3-difluoropyridine | 410-090-7 | 89402-43-7 | Flam. Liq. 3 Acute Tox. 4 * Aquatic Chronic 3 | H226 H302 H412 | GHS02 GHS07 Wng | H226 H302 H412 |  |  |  | CLP00
613-156-00-5 | 2-butyl-4-chloro-5-formylimidazole | 410-260-0 | 83857-96-9 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
613-157-00-0 | 2,4-diamino-5-methoxymethylpyrimidine | 410-330-0 | 54236-98-5 | Acute Tox. 4 * STOT RE 2 * Eye Irrit. 2 | H302 H373 ** H319 | GHS08 GHS07 Wng | H302 H373 ** H319 |  |  |  | CLP00
613-158-00-6 | 2,3-dichloro-5-trifluoromethyl-pyridine | 410-340-5 | 69045-84-7 | Acute Tox. 4 * Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H332 H302 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H332 H302 H318 H317 H411 |  |  |  | CLP00
613-159-00-1 | fenazaquin (ISO); 4-[2-[4-(1,1-dimethylethyl)phenyl]-ethoxy]quinazoline | 410-580-0 | 120928-09-8 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H332 H400 H410 | GHS06 GHS09 Dgr | H301 H332 H410 |  |  |  | CLP00
613-160-00-7 | (1S)-2-methyl-2,5-diazobicyclo[2.2.1]heptane dihydrobromide | 411-000-9 | 125224-62-6 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
613-161-00-2 | (2,4-diaminopteridin-6-yl)methanol hydrobromide | 430-620-0 | 76145-91-0 | STOT RE 2 * Skin Sens. 1 Aquatic Chronic 3 | H373 ** H317 H412 | GHS08 GHS07 Wng | H317 H373 ** H412 |  |  |  | ATP01/ATP01corr
613-162-00-8 | (6R-trans)-1-((7-ammonio-2-carboxylato-8-oxo-5-thia-1-azabicyclo-[4.2.0]oct-2-en-3-yl)methyl)pyridinium iodide | 423-260-0 | 100988-63-4 | Muta. 2 Skin Sens. 1 Aquatic Chronic 2 | H341 H317 H411 | GHS08 GHS07 GHS09 Wng | H341 H317 H411 |  |  |  | ATP01
613-163-00-3 | azimsulfuron (ISO); 1-(4,6-dimethoxypyrimidin-2-yl)-3-[1-methyl-4-(2-methyl-2H-tetrazol-5-yl)pyrazol-5-ylsulfonyl]urea | - | 120162-55-2 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1000 |  | CLP00/ATP01
613-164-00-9 | flufenacet (ISO); N-(4-fluorophenyl)-N-isopropyl-2-(5-trifluoromethyl-[1,3,4]thiadiazol-2-yloxy)acetamide | - | 142459-58-3 | Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H317 H410 |  | M=100 |  | CLP00/ATP01
613-165-00-4 | flupyrsulfuron-methyl-sodium (ISO); methyl 2-[[(4,6-dimethoxypyrimidin-2-ylcarbamoyl)sulfamoyl]-6-trifluoromethyl]nicotinate, monosodium salt | - | 144740-54-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=100 |  | CLP00/ATP01
613-166-00-X | flumioxazin (ISO); N-(7-fluoro-3,4-dihydro- 3-oxo-4-prop- 2-ynyl-2H-1,4-benzoxazin- 6-yl)cyclohex- 1-ene-1,2-dicarboximide |  | 103361-09-7 | Repr. 2 Aquatic Acute 1 Aquatic Chronic 1 | H361d H400 H410 | GHS08 GHS09 Wng | H361d H410 |  | M=1000 M=1000 |  | CLP00/ATP17
613-167-00-5 | reaction mass of 5-chloro-2-methyl-2H-isothiazol-3-one and 2-methyl-2H-isothiazol-3-one (3:1) |  | 55965-84-9 | Acute Tox. 2 Acute Tox. 2 Acute Tox. 3 Skin Corr. 1C Eye Dam. 1 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 1 | H330 H310 H301 H314 H318 H317 H400 H410 | GHS06 GHS05 GHS09 Dgr | H330 H310 H301 H314 H317 H410 | EUH071 | Skin Corr. 1C; : C ≥ ,6 % Skin Irrit. 2; H315: ,06 % ≤ C < ,6 % Eye Dam. 1; : C ≥ ,6 % Eye Irrit. 2; H319: ,06 % ≤ C < ,6 % Skin Sens. 1A; : C ≥ ,0015 % M=100 M=100 | B | CLP00/ATP13
613-168-00-0 | 1-vinyl-2-pyrrolidone | 201-800-4 | 88-12-0 | Carc. 2 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * STOT SE 3 STOT RE 2 * Eye Dam. 1 | H351 H332 H312 H302 H335 H373 ** H318 | GHS06 GHS05 GHS09 Dgr | H351 H332 H312 H302 H373 ** H335 H318 |  |  | D | CLP00
613-169-00-6 | 9-vinylcarbazole | 216-055-0 | 1484-13-5 | Muta. 2 Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H341 H312 H302 H315 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H341 H312 H302 H315 H317 H410 |  | M=100 |  | CLP00/ATP01
613-170-00-1 | 2,2-ethylmethylthiazolidine | 404-500-3 | 694-64-4 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H317 H411 |  |  |  | CLP00
613-171-00-7 | hexaconazole (ISO); (RS)-2-(2,4-dichlorophenyl)-1-(1H-1,2,4-triazol-1-yl)hexan-2-ol | 413-050-7 | 79983-71-4 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | CLP00
613-172-00-2 | 5-chloro-1,3-dihydro-2H-indol-2-one | 412-200-9 | 17630-75-0 | Repr. 2 Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 3 | H361f *** H302 H317 H412 | GHS08 GHS07 Wng | H361f *** H302 H317 H412 |  |  |  | CLP00
613-173-00-8 | fluquinconazole (ISO); 3-(2,4-dichlorophenyl)-6-fluoro-2-(1H-1,2,4-triazol-1-yl)quinazolin-4-(3H)-one | 411-960-9 | 136426-54-5 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 4 * STOT RE 1 Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H312 H372 ** H315 H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H301 H372 ** H312 H315 H410 |  |  |  | CLP00
613-174-00-3 | tetraconazole (ISO); (±) 2-(2,4-dichlorophenyl)-3-(1H-1,2,4-triazol-1-yl)propyl-1,1,2,2-tetrafluoroethylether | 407-760-6 | 112281-77-3 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 2 | H332 H302 H411 | GHS07 GHS09 Wng | H332 H302 H411 |  |  |  | CLP00/ATP01
613-175-00-9 | epoxiconazole (ISO); (2RS,3SR)-3-(2-chlorophenyl)-2-(4-fluorophenyl)-[( 1H-1,2,4-triazol-1-yl)methyl]oxirane | 406-850-2 | 133855-98-8 | Carc. 2 Repr. 1B Aquatic Chronic 2 | H351 H360Df H411 | GHS08 GHS09 Dgr | H351 H360Df H411 |  |  |  | CLP00/ATP05
613-176-00-4 | 2-methyl-2-azabicyclo[2.2.1]heptane | 404-810-9 | 4524-95-2 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B | H226 H312 H302 H373 ** H314 | GHS02 GHS08 GHS05 GHS07 Dgr | H226 H312 H302 H373 ** H314 |  |  |  | CLP00
613-177-00-X | 8-amino-7-methylquinoline | 412-760-4 | 5470-82-6 | Acute Tox. 4 * Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H312 H302 H317 H411 | GHS07 GHS09 Wng | H312 H302 H317 H411 |  |  |  | CLP00
613-178-00-5 | 4-ethyl-2-methyl-2-isopentyl-1,3-oxazolidine | 410-470-2 | 137796-06-6 | Skin Corr. 1B Skin Sens. 1 | H314 H317 | GHS05 GHS07 Dgr | H314 H317 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
613-179-00-0 | lithium 3-oxo-1,2(2H)-benzisothiazol-2-ide | 411-690-1 | 111337-53-2 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H302 H314 H317 H411 | GHS05 GHS07 Dgr | H302 H314 H317 H411 |  |  |  | CLP00
613-180-00-6 | N-(1,1-dimethylethyl)bis(2-benzothiazolesulfen)amide | 407-430-1 | 3741-80-8 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
613-181-00-1 | 5,5-dimethyl-perhydro-pyrimidin-2-one α-(4-trifluoromethylstyryl)-α-(4-trifluoromethyl)cinnamylidenehydrazone | 405-090-9 | 67485-29-4 | Acute Tox. 4 * STOT RE 1 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H372 ** H319 H400 H410 | GHS08 GHS07 GHS09 Dgr | H372 ** H302 H319 H410 |  |  |  | CLP00
613-182-00-7 | 1-(1-naphthylmethyl)quinolinium chloride | 406-220-7 | 65322-65-8 | Carc. 2 Muta. 2 Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H351 H341 H302 H315 H318 H412 | GHS08 GHS05 GHS07 Dgr | H351 H341 H302 H315 H318 H412 |  |  |  | CLP00
613-183-00-2 | reaction mass of: 5-(N-methylperfluorooctylsulfonamido)methyl-3-octadecyl-1,3-oxazolidin-2-one; 5-(N-methylperfluoroheptylsulfonamido)methyl-3-octadecyl-1,3-oxazolidin-2-one | 413-640-4 |  | STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H400 H410 | GHS08 GHS09 Wng | H373 ** H410 |  |  |  | CLP00
613-184-00-8 | nitrilotriethyleneammoniopropane-2-ol 2-ethylhexanoate | 413-670-8 |  | Eye Irrit. 2 Skin Sens. 1 | H319 H317 | GHS07 Wng | H319 H317 |  |  |  | CLP00
613-185-00-3 | 2,3,5,6-tetrahydro-2-methyl-2H-cyclopenta[d]-1,2-thiazol-3-one | 407-630-9 | 82633-79-2 | Acute Tox. 3 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H318 H317 H400 H410 | GHS06 GHS05 GHS09 Dgr | H301 H318 H317 H410 |  |  |  | CLP00
613-186-00-9 | (2R,3R)-3-((R)-1-(tert-butyldimethylsiloxy)ethyl)-4-oxoazetidin-2-yl acetate | 408-050-9 | 76855-69-1 | Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H319 H317 H411 | GHS07 GHS09 Wng | H319 H317 H411 |  |  |  | CLP00
613-187-00-4 | 5-(2-amino-5-cyano-6-[2-(2-hydroxyethoxy)ethylamino]-4-methylpyridin-3-ylazo)-3-methyl-2,4-dicarbonitrilethiophene | 410-530-8 | - | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
613-188-00-X | 1-(3-(4-fluorophenoxy)propyl)-3-methoxy-4-piperidinone | 411-500-7 | 116256-11-2 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H317 H411 |  |  |  | CLP00
613-189-00-5 | 1,4,7,10-tetrakis(p-toluensulfonyl)-1,4,7,10-tetraazacyclododecane | 414-030-0 | 52667-88-6 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
613-190-00-0 | disodium 1-amino-4-(2-(5-chloro-6-fluoro-pyrimidin-4-ylamino-methyl)-4-methyl-6-sulfo-phenylamino)-9,10-dioxo-9,10-dihydro-anthracene-2-sulfonate | 414-040-5 | 149530-93-8 | Acute Tox. 4 * Skin Sens. 1 | H302 H317 | GHS07 Wng | H302 H317 |  |  |  | CLP00
613-191-00-6 | 3-ethyl-2-methyl-2-(3-methylbutyl)-1,3-oxazolidine | 421-150-7 | 143860-04-2 | Repr. 1B Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H360F *** H314 H400 H410 | GHS08 GHS05 GHS09 Dgr | H360F *** H314 H410 |  |  |  | CLP00
613-192-00-1 | 3-benzyl-exo-6-nitro-2,4-dioxo-3-aza-cis-bicyclo[3.1.0]hexane | 426-750-2 | 151860-15-0 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
613-193-00-7 | pentakis[3-(dimethylammonio)propylsulfamoyl]-[(6-hydroxy-4,4,8,8-tetramethyl-4,8-diazoniaundecane-1,11-diyldisulfamoyl)di[phthalocyaninecopper(II)]] heptalactate | 414-930-3 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
613-194-00-2 | 6,13-dichloro-3,10-bis{}{2-[4-fluoro-6-(2-sulfophenylamino)-1,3,5-triazin-2-ylamino]propylamino}}benzo[5,6][1,4]oxazino[2,3-.b.]phenoxazine-4,11-disulphonic acid, lithium-, sodium salt | 418-000-8 | 163062-28-0 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
613-195-00-8 | 2,2-(1,4-phenylene)bis((4H-3,1-benzoxazine-4-one) | 418-280-1 | 18600-59-4 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
613-196-00-3 | 5-[[4-chloro-6-[[2-[[4-fluoro-6-[[5-hydroxy-6-[(4-methoxy-2-sulfophenyl)azo]-7-sulfo-2-naphthalenyl]amino]-1,3,5-triazin-2-yl]amino]-1-methylethyl]amino]-1,3,5-triazin-2-yl]amino]-3-[[4-(ethenylsulfonyl)phenyl]azo]-4-hydroxy-naphtalene-2,7-disulfonic acid, sodium salt | 418-380-5 | 168113-78-8 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
613-197-00-9 | reaction mass of: 2,4,6-tri(butylcarbamoyl)-1,3,5-triazine; 2,4,6-tri(methylcarbamoyl)-1,3,5-triazine; [(2-butyl-4,6-dimethyl)tricarbamoyl]-1,3,5-triazine; [(2,4-dibutyl-6-methyl)tricarbamoyl]-1,3,5-triazine | 420-390-1 | 187547-46-2 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
613-198-00-4 | 2-amino-4-dimethylamino-6-trifluoroethoxy-1,3,5-triazine | 415-500-8 | 145963-84-4 | Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 3 | H302 H373 ** H412 | GHS08 GHS07 Wng | H302 H373 ** H412 |  |  |  | ATP01
613-199-00-X | reaction mass of: 1,3,5-tris(3-aminomethylphenyl)-1,3,5-(1H,3H,5H)-triazine-2,4,6-trione; reaction mass of oligomers of 3,5-bis(3-aminomethylphenyl)-1-poly[3,5-bis(3-aminomethylphenyl)-2,4,6-trioxo-1,3,5-(1H,3H,5H)-triazin-1-yl]-1,3,5-(1H,3H,5H)-triazine-2,4,6-trione | 421-550-1 |  | Carc. 1B Repr. 1B Skin Sens. 1 Aquatic Chronic 3 | H350 H360D *** H317 H412 | GHS08 Dgr | H350 H360D *** H317 H412 |  |  |  | CLP00
613-200-00-3 | Reaction product of: copper, (29H,31H-phthalocyaninato(2-)-N29,N30,N31,N32)-, chlorosulfuric acid and 3-(2-sulfooxyethylsulfonyl)aniline, sodium salts | 420-980-7 |  | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
613-201-00-9 | (R)-5-bromo-3-(1-methyl-2-pyrrolidinyl methyl)-1H-indole | 422-390-5 | 143322-57-0 | Repr. 2 Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361f *** H332 H302 H372 ** H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H361f *** H372 ** H332 H302 H317 H410 | EUH070 |  |  | CLP00
613-202-00-4 | pymetrozine (ISO); (E)-4,5-dihydro-6-methyl-4-(3-pyridylmethyleneamino)-1,2,4-triazin-3(2H)-one |  | 123312-89-0 | Carc. 2 Repr. 2 Aquatic Chronic 1 | H351 H361fd H410 | GHS08 GHS09 Wng | H351 H361fd H410 |  | M=1 |  | CLP00/ATP15
613-203-00-X | pyraflufen-ethyl (ISO); 2-chloro-5-(4-chloro-5-difluoromethoxy-1-methylpyrazol-3-yl)-4-fluorophenoxyacetic acid ethyl ester [1] pyraflufen (ISO); 2-chloro-5-(4-chloro-5-difluoromethoxy-1-methylpyrazol-3-yl)-4-fluorophenoxyacetic acid [2] |  | 129630-19-9 [1] 129630-17-7 [2] | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1000 |  | CLP00/ATP01
613-204-00-5 | oxadiargyl (ISO); 3-[2,4-dichloro-5-(2-propynyloxy)phenyl]-5-(1,1-dimethylethyl)-1,3,4-oxadiazol-2(3H)-one | 254-637-6 | 39807-15-3 | Repr. 2 STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H361d *** H373 ** H400 H410 | GHS08 GHS09 Wng | H361d *** H373 ** H410 |  | M=1000 |  | CLP00/ATP01corr
613-205-00-0 | propiconazole (ISO); (2RS,4RS;2RS,4SR)-1-{[2-(2,4-dichlorophenyl)-4-propyl-1,3-dioxolan-2-yl]methyl}-1H-1,2,4-triazole | 262-104-4 | 60207-90-1 | Repr. 1B Acute Tox. 4 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D H302 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H302 H317 H360D H410 |  | M=1 M=1 |  | CLP00/ATP13
613-206-00-6 | fenamidone (ISO); (S)-5-methyl-2-methylthio-5-phenyl-3-phenylamino-3,5-dihydroimidazol-4-one |  | 161326-34-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
613-208-00-7 | imazamox (ISO); (RS)-2-(4-isopropyl-4-methyl-5-oxo-2-imidazolin-2-yl)-5-methoxymethylnicotinic acid |  | 114311-32-9 | Repr. 2 Aquatic Acute 1 Aquatic Chronic 1 | H361d H400 H410 | GHS08 GHS09 Wng | H361d H410 |  | M=10 M=10 |  | CLP00/ATP17
613-209-00-2 | cis-1-(3-chloropropyl)-2,6-dimethyl-piperidin hydrochloride | 417-430-3 | 63645-17-0 | Acute Tox. 3 * STOT RE 2 * Skin Sens. 1 Aquatic Chronic 2 | H301 H373 ** H317 H411 | GHS06 GHS08 GHS09 Dgr | H301 H373 ** H317 H411 |  |  |  | CLP00
613-210-00-8 | 2-(3-chloropropyl)-2,5,5-trimethyl-1,3-dioxane | 417-650-1 | 88128-57-8 | STOT RE 2 * Aquatic Chronic 3 | H373 ** H412 | GHS08 Wng | H373 ** H412 |  |  |  | CLP00
613-211-00-3 | N-methyl-4-(p-formylstyryl)pyridinium methylsulfate | 418-240-3 | 74401-04-0 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
613-212-00-9 | 4-[4-(2-ethylhexyloxy)phenyl](1,4-thiazinane-1,1-dioxide) | 418-320-8 | 133467-41-1 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
613-213-00-4 | cis-1-benzoyl-4-[(4-methylsulfonyl)oxy]-L-proline | 416-040-0 | 120807-02-5 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
613-214-00-X | N,N-di-n-butyl-2-(1,2-dihydro-3-hydroxy-6-isopropyl-2-quinolylidene)-1,3-dioxoindan-5-carboxamide | 416-260-7 | 147613-95-4 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
613-215-00-5 | 2-chloromethyl-3,4-dimethoxypyridinium chloride | 416-440-5 | 72830-09-2 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H312 H302 H373 ** H315 H318 H317 H411 | GHS08 GHS05 GHS07 GHS09 Dgr | H312 H302 H373 ** H315 H318 H317 H411 |  |  |  | CLP00
613-216-00-0 | 6-tert-butyl-7-(6-diethylamino-2-methyl-3-pyridylimino)-3-(3-methylphenyl)pyrazolo[3,2-c][1,2,4]triazole | 416-490-8 | 162208-01-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
613-217-00-6 | 4-[3-(3,5-di-tert-butyl-4-hydroxyphenyl)propionyloxy]-1-[2-[3-(3,5-di-tert-butyl-4-hydrophenyl)propionyloxy]ethyl]-2,2,6,6-tetramethylpiperidine | 416-770-1 | 73754-27-5 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
613-218-00-1 | 6-hydroxyindole | 417-020-4 | 2380-86-1 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H302 H318 H317 H411 |  |  |  | CLP00
613-219-00-7 | 7a-ethyl-3,5-bis(1-methylethyl)-2,3,4,5-tetrahydrooxazolo[3,4-c]-2,3,4,5-tetrahydrooxazole | 417-140-7 | 79185-77-6 | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
613-220-00-2 | trans-(4S,6S)-5,6-dihydro-6-methyl-4H-thieno[2,3-b]thiopyran-4-ol, 7,7-dioxide | 417-290-3 | 147086-81-5 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
613-221-00-8 | 2-chloro-5-methyl-pyridine | 418-050-0 | 18368-64-4 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Aquatic Chronic 3 | H312 H302 H315 H412 | GHS07 Wng | H312 H302 H315 H412 |  |  |  | CLP00
613-222-00-3 | 4-(1-oxo-2-propenyl)-morpholine | 418-140-1 | 5117-12-4 | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 | H302 H373 ** H318 H317 | GHS08 GHS05 GHS07 Dgr | H302 H373 ** H318 H317 |  |  |  | CLP00
613-223-00-9 | N-isopropyl-3-(4-fluorophenyl)-1H-indole | 418-790-4 | 93957-49-4 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
613-224-00-4 | 2,5-dimercaptomethyl-1,4-dithiane | 419-770-8 | 136122-15-1 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H314 H317 H410 |  |  |  | CLP00
613-225-00-X | reaction mass of:[2-(anthraquinon-1-ylamino)-6-[(5-benzoylamino)-anthraquinone-1-ylamino]-4-phenyl]-1,3,5-triazine; 2,6-bis-[(5-benzoylamino)-anthraquinon-1-ylamino]-4-phenyl-1,3,5-triazine. | 421-290-9 |  | STOT RE 2 * Aquatic Chronic 4 | H373 ** H413 | GHS08 Wng | H373 ** H413 |  |  |  | CLP00
613-226-00-5 | 1-(2-(ethyl(4-(4-(4-(4-(ethyl(2-pyridinoethyl)amino)-2-methylphenylazo)benzoylamino)-phenylazo)-3-methylphenyl)amino)ethyl)-pyridinium dichloride | 420-950-3 | 163831-67-2 | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | CLP00
613-227-00-0 | (±)-[(R*,R*) and (R*,S*)]-6-fluoro-3,4-dihydro-2-oxiranyl-2H-1-benzopyran | 419-600-2 | 99199-90-3 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
613-228-00-6 | (±)-(R*,S*)-6-fluoro-3,4-dihydro-2-oxiranyl-2H-1-benzopyran | 419-630-6 | 793669-26-8 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
613-229-00-1 | 1-acetyl-4-(3-dodecyl-2,5-dioxo-1-pyrrolidinyl)-2,2,6,6-tetramethylpiperidine | 411-930-5 | 106917-31-1 | Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H317 H400 H410 | GHS07 GHS09 Wng | H315 H317 H410 |  |  |  | ATP01
613-230-00-7 | florasulam (ISO); 2',6',8-trifluoro-5-methoxy-5-triazolo[1,5-c]; pyrimidine-2-sulfonanilide |  | 145701-23-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
613-231-00-2 | 2,6-diamino-3-((pyridine-3-yl)azo)pyridine | 421-430-9 | 28365-08-4 | Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 2 | H302 H373 ** H411 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H411 |  |  |  | ATP01
613-232-00-8 | 3-(benzo[b]thien-2-yl)-5,6-dihydro-1,4,2-oxathiazine-4-oxide | 431-030-6 | 163269-30-5 | Acute Tox. 3 * STOT RE 2 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H373 ** H318 H400 H410 | GHS06 GHS05 GHS08 GHS09 Dgr | H331 H373 ** H318 H410 |  |  |  | ATP01
613-233-00-3 | 4,4'-(oxy-(bismethylene))-bis-1,3-dioxolane | 423-230-7 | 56552-15-9 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
613-234-00-9 | imidazo[1,2-b]pyridazin hydrochloride | 431-510-5 | 18087-70-2 | Acute Tox. 4 * Eye Irrit. 2 | H302 H319 | GHS07 Wng | H302 H319 |  |  |  | ATP01
613-235-00-4 | 2,3-dihydro-2,2-dimethyl-1H-perimidine | 424-060-6 | 6364-17-6 | Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H317 H410 |  |  |  | ATP01
613-236-00-X | 2-chloro-3-trifluoromethylpyridine | 424-520-6 | 65753-47-1 | Acute Tox. 3 * Acute Tox. 3 * STOT RE 1 Skin Corr. 1B Aquatic Chronic 3 | H311 H301 H372 ** H314 H412 | GHS06 GHS05 GHS08 Dgr | H311 H301 H372 ** H314 H412 |  |  |  | ATP01
613-237-00-5 | 6-tert-butyl-3-(3-dodecylsulfonyl)propyl-7H-1,2,4-triazolo[3.4b][1,3,4]thiadiazine | 424-950-4 | 133949-92-5 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
613-238-00-0 | sodium 2-[[4-[(4,6-dichloro-1,3,5-triazin-2-yl)amino]phenyl]sulfonyl]ethyl sulfate | 430-890-1 | 81992-66-7 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | ATP01
613-239-00-6 | 2-[3-(methylamino)propyl]-1H-benzimidazole | 425-760-4 | 64137-52-6 | Eye Dam. 1 Aquatic Chronic 3 | H318 H412 | GHS05 Dgr | H318 H412 |  |  |  | ATP01
613-241-00-7 | 3-(2H-tetrazol-5-yl)pyridine | 426-810-8 | 3250-74-6 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
613-242-00-2 | reaction products of 3,10-bis((2-aminopropyl)amino)-6,13-dichloro-4,11-triphenodioxazinedisulfonic acid with 2-amino-1,4-benzenedisulfonic acid, 2-((4-aminophenyl)sulfonyl)ethyl hydrogen sulfate and 2,4,6-trifluoro-1,3,5-triazine, sodium salts | 426-860-0 | 191877-09-5 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
613-243-00-8 | 4,4'-(1,6-hexamethylenebis(formylimino))bis(2,2,6,6-tetramethyl-1-oxylpiperidine) | 427-350-0 | 182235-14-9 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
613-244-00-3 | 5,7-dichloro-4-hydroxyquinoline | 427-420-0 | 21873-52-9 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
613-245-00-9 | 2-fluoro-6-trifluoromethylpyridine | 428-100-3 | 94239-04-0 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 3 | H226 H332 H302 H412 | GHS02 GHS07 Wng | H226 H332 H302 H412 |  |  |  | ATP01
613-246-00-4 | 2-hydroxymethyl-3-methyl-4-(2,2,2-trifluoroethoxy)pyridine | 428-200-7 | 103577-66-8 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
613-247-00-X | 3-(2-methoxy-4-methoxycarboxybenzyl)-5-nitroindole | 428-910-7 | 107786-36-7 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
613-248-00-5 | 3,4-dimethyl-1H-pyrazole | 429-130-1 | 2820-37-3 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | ATP01
613-249-00-0 | 1-(2-hydroxyethyl)-1H-pyrazol-4,5-diyldiammoniumsulfate | 429-300-3 | 155601-30-2 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H318 H317 H411 |  |  |  | ATP01
613-250-00-6 | reaction mass of: carbonato-bis-N-ethyl-2-isopropyl-1,3-oxazolidine; methyl carbonato-N-ethyl-2-isopropyl-1,3-oxazolidine; 2-isopropyl-N-hydroxyethyl 1,3-oxazolidine | 429-990-6 | - | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H318 H317 H412 | GHS05 GHS07 Dgr | H318 H317 H412 |  |  |  | ATP01
613-251-00-1 | (R)-3-[(1-methylpyrrolidin-2-yl)methyl]-5-[2-(phenylsulfonyl)ethenyl]-1H-indole | 430-560-5 | 180637-89-2 | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 | H302 H373 ** H318 H317 | GHS05 GHS08 GHS07 Dgr | H302 H373 ** H318 H317 |  |  |  | ATP01
613-253-00-2 | 2,2-dialkyl-4-hydroxymethyl-1,3-dioxolane; reaction products with ethylene oxide (alkyl is C1-12 and the sum to C13, average degree of ethoxylation is 3.5) | 430-580-4 | - | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 | EUH019 |  |  | ATP01
613-254-00-8 | forchlorfenuron (ISO); 1-(2-chloro-4-pyridyl)-3-phenylurea | - | 68157-60-8 | Carc. 2 Aquatic Chronic 2 | H351 H411 | GHS08 GHS09 Wng | H351 H411 |  |  |  | ATP01
613-255-00-3 | reaction mass of isomers of: sodium [(2-hydroxyethylsulfamoyl){[2-(2-piperazin-1-ylethylamino)ethylsulfamoyl][2-(4-aminoethylpiperazine-1-yl)ethylsulfamoyl](sulfamoyl)}(sulfonatophthalocyaninato)]copper(II) | 424-270-8 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
613-256-00-9 | 3'5'-anhydro thymidine | 425-810-5 | 38313-48-3 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
613-257-00-4 | 2-phthalimidoethyl N-[4-(2-cyano-4-nitrophenylazo)phenyl]-N-methyl-β-alaninate | 426-400-9 | 170222-39-6 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
613-258-00-X | reaction mass of: 4-chloro-7-methylbenzotriazole sodium salt; 4-chloro-5-methylbenzotriazole sodium salt; 5-chloro-4-methylbenzotriazole sodium salt | 427-730-6 | 202420-04-0 | Skin Corr. 1B Aquatic Chronic 3 | H314 H412 | GHS05 Dgr | H314 H412 |  |  |  | ATP01
613-259-00-5 | imiprothrin (ISO); reaction mass of: [2,4-dioxo-(2-propyn-1-yl)imidazolidin-3-yl]methyl(1R)-cis-chrysanthemate; [2,4-dioxo-(2-propyn-1-yl)imidazolidin-3-yl]methyl(1R)-trans-chrysanthemate | 428-790-6 | 72963-72-5 | Carc. 2 Acute Tox. 4 Acute Tox. 4 STOT SE 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H332 H302 H371 (nervous system; oral, inhalation) H400 | GHS08 GHS07 GHS09 Wng | H332 H302 H351 H371 (nervous system; oral) (inhalation) H410 |  | Inhalation: ATE = 1.4 mg/L (dusts/mists) Oral: ATE = 550 mg/kg bw M=10 M=10 |  | ATP01/ATP15
613-260-00-0 | (±)-4-(3-chlorophenyl)-6-[(4-chlorophenyl)hydroxy(1-methyl-1H-imidazol-5-yl)methyl]-1-methyl-2(1H)-quinolin | 430-730-9 | - | Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H400 H410 | GHS05 GHS09 Dgr | H318 H410 |  |  |  | ATP01
613-261-00-6 | pyrazole-1-carboxamidine monohydrochloride | 429-520-1 | 4023-02-3 | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H302 H373 ** H318 H317 H412 | GHS05 GHS08 GHS07 Dgr | H302 H373 ** H318 H317 H412 |  |  |  | ATP01
613-262-00-1 | disodium (E)-1,2-bis-(4-(4-methylamino-6-(4-methylcarbamoylphenylamino)-1,3,5-triazin-2-ylamino)phenyl-2-sulfonato)ethene | 427-310-2 | 180850-95-7 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
613-263-00-7 | monosodium 3-cyano-5-fluoro-6-hydroxypyridine-2-olate | 429-570-2 | - | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
613-266-00-3 | 2-chloro-5-chloromethylthiazole | 429-830-5 | 105827-91-6 | Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H311 H302 H314 H317 H411 | GHS06 GHS05 GHS09 Dgr | H311 H314 H302 H317 H411 |  |  |  | ATP01
613-267-00-9 | thiamethoxam (ISO); 3-(2-chloro-thiazol-5-ylmethyl)-5-methyl[1,3,5]oxadiazinan-4-ylidene-N-nitroamine | 428-650-4 | 153719-23-4 | Repr. 2 Acute Tox. 4 Aquatic Acute 1 Aquatic Chronic 1 | H361fd H302 H400 H410 | GHS07 GHS08 GHS09 Wng | H302 H361fd H410 |  | Oral: ATE = 780 mg/kg bw M=10 M=10 |  | ATP01/ATP17
613-268-00-4 | (4aS-cis-)-6-benzyl-octahydropyrrolo[3.4-b]pyridine | 425-930-8 | 151213-39-7 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Aquatic Chronic 2 | H332 H302 H373 ** H314 H411 | GHS05 GHS08 GHS07 GHS09 Dgr | H314 H332 H302 H373 ** H411 |  |  |  | ATP01
613-269-00-X | 2-thiazolidinylidenecyanamide | 427-720-1 | 26364-65-8 | Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 3 | H302 H373 ** H412 | GHS08 GHS07 Wng | H302 H373 ** H412 |  |  |  | ATP01
613-270-00-5 | 5-amino-N-(2,6-dichloro-3-methylphenyl)-1H-1,2,4-triazole-3-sulfonamide | 428-150-6 | 113171-13-4 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
613-271-00-0 | tritosulfuron (ISO) (containing ≤ 0,02% AMTT); 1-[4-methoxy-6-(trifluoromethyl)-1,3,5-triazin-2-yl]-3-[2-(trifluoromethyl)benzenesulfonyl]urea (containing ≤ 0,02% AMTT) | - | 142469-14-5 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  | M=10 |  | ATP01
613-272-00-6 | pyraclostrobin (ISO); methyl N-{2-[1-(4-chlorophenyl)-1H-pyrazol-3-yloxymethyl]phenyl}(N-methoxy)carbamate | - | 175013-18-0 | Repr. 2 Acute Tox. 3 Acute Tox. 4 STOT SE 3 STOT RE 2 Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H361d H331 H302 H335 H373 (liver, gastrointestinal tract, nasal cavity) H315 H400 H410 | GHS08 GHS06 GHS09 Dgr | H361d H331 H302 H335 H373 (liver, gastrointestinal tract, nasal cavity) H315 H410 |  | inhalation: ATE = 0,58 mg/L (dusts or mists) oral: ATE = 450 mg/kg bw M = 100 M = 100 |  | ATP01/ATP22
613-273-00-1 | tetrahydro-3-methyl-5-((2-phenylthio)thiazol-5-ylmethyl)-[4H]-1,3,5-oxadiazinan-4-ylidene-N-nitroamine | 427-600-9 | 192439-46-6 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
613-274-00-7 | 2,6-dichloro-1-fluoropyridiniumtetrafluoroborate | 427-400-1 | 140623-89-8 | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H314 H302 H317 H410 |  |  |  | ATP01
613-275-00-2 | 3-(2-chloroethyl)-6,7,8,9-tetra-hydro-2-methyl-4H-pyrido[1,2-a] pyrimidin-4-one monohydrochloride | 424-530-0 | 93076-03-0 | Acute Tox. 3 * STOT SE 2 STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H301 H371 ** H373 ** H318 H317 H411 | GHS06 GHS05 GHS08 GHS09 Dgr | H301 H318 H317 H371 ** H373 ** H411 |  |  |  | ATP01/ATP01corr
613-276-00-8 | 1-(2-chlorophenyl)-1,2-dihydro-5H-tetrazol-5-one | 426-110-2 | 98377-35-6 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
613-277-00-3 | (4-(6-diethylamino-2-methylpyridin-3-yl)imino-4,5-dihydro-3-methyl-1-(4-methylphenyl)-1H-pyrazol-5-one | 427-070-9 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
613-278-00-9 | (3-aminophenyl)pyridin-3-ylmethanone | 428-230-0 | 79568-06-2 | STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H400 H410 | GHS08 GHS09 Wng | H373 ** H410 |  |  |  | ATP01
613-279-00-4 | 2-ethyl-2,3-dihydro-2-methyl-1H-perimidine | 424-380-6 | 43057-68-7 | Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H410 |  |  |  | ATP01
613-280-00-X | tetrahydro-1,3-dimethyl-1H-pyrimidin-2-one; dimethyl propyleneurea | 230-625-6 | 7226-23-5 | Repr. 2 Acute Tox. 4 * Eye Dam. 1 | H361f *** H302 H318 | GHS05 GHS08 GHS07 Dgr | H361f *** H302 H318 |  |  |  | ATP01
613-281-00-5 | quinoline | 202-051-6 | 91-22-5 | Carc. 1B Muta. 2 Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H350 H341 H312 H302 H315 H319 H411 | GHS08 GHS07 GHS09 Dgr | H350 H341 H312 H302 H319 H315 H411 |  |  |  | ATP01
613-282-00-0 | triticonazole (ISO); (RS)-(E)-5-(4-chlorobenzylidene)-2,2-dimethyl-1-(1H-1,2,4-triazol-1-methyl)cyclopentanol |  | 138182-18-0 | Repr. 2 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H361f H373 H400 H410 | GHS08 GHS09 Wng | H361f H373 H410 |  | M=1 M=1 |  | ATP01/ATP17
613-283-00-6 | ketoconazole; 1-[4-[4-[[(2SR,4RS)-2-(2,4-dichlorophenyl)-2-(imidazol-1-ylmethyl)-1,3-dioxolan-4-yl]methoxy]phenyl]piperazin-1-yl]ethanone | 265-667-4 | 65277-42-1 | Repr. 1B Acute Tox. 3 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H360F *** H301 H373 ** H400 H410 | GHS06 GHS08 GHS09 Dgr | H360F *** H301 H373 ** H410 |  |  |  | ATP01
613-284-00-1 | metconazole (ISO); (1RS,5RS;1RS,5SR)-5-(4-chlorobenzyl)-2,2-dimethyl-1-(1H-1,2,4-triazol-1-ylmethyl)cyclopentanol | - | 125116-23-6 | Repr. 2 Acute Tox. 4 * Aquatic Chronic 2 | H361d *** H302 H411 | GHS08 GHS07 GHS09 Wng | H361d *** H302 H411 |  |  |  | ATP01
613-285-00-7 | 1-hydroxybenzotriazole, anhydrous [1] 1-hydroxybenzotriazole, monohydrated [2] | 219-989-7 [1] 219-989-7 [2] | 2592-95-2 [1] 123333-53-9 [2] | Expl. 1.3 | H203 | GHS01 Dgr | H203 |  |  |  | ATP01
613-286-00-2 | potassium 1-methyl-3-morpholinocarbonyl-4-[3-(1-methyl-3-morpholinocarbonyl-5-oxo-2-pyrazolin-4-ylidene)-1-propenyl]pyrazole-5-olate; [containing < 0.5 % N,N-dimethylformamide (EC no 200-679-5)] | 418-260-2 | 183196-57-8 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
613-286-01-X | potassium 1-methyl-3-morpholinocarbonyl-4-[3-(1-methyl-3-morpholinocarbonyl-5-oxo-2-pyrazolin-4-ylidene)-1-propenyl]pyrazole-5-olate; [containing ≥ 0.5 % N,N-dimethylformamide (EC No 200-679-5)] | 418-260-2 | 183196-57-8 | Repr. 1B Skin Sens. 1 | H360D *** H317 | GHS08 GHS07 Dgr | H360D *** H317 |  |  |  | ATP01
613-287-00-8 | 1-(3-iodo-4-aminobenzyl)-1H-1,2,4-triazole | 419-540-7 | 160194-26-3 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | ATP01
613-288-00-3 | 1,3-bis(dimethylcarbamoyl)-imidazolium chloride | 420-930-4 | 135756-61-5 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | ATP01
613-289-00-9 | 3-(4-chloro-2-fluoro-5-methylphenyl)-1-methyl-5-(trifluoromethyl)-1H-pyrazole | 432-020-4 | 142623-48-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
613-290-00-4 | 4-hydroxy-7-(2-aminoethyl)-1,3-benzothiazol-2(3H)-one hydrochloride | 432-470-1 | 189012-93-9 | Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H318 H317 H410 |  |  |  | ATP01
613-291-00-X | 2,4-dihydro-4-(4-(4-(4-hydroxyphenyl)-1-piperazinyl)phenyl)-2-(1-methylpropyl)-3H-1,2,4-triazol-3-one | 434-820-9 | 106461-41-0 | STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H400 H410 | GHS08 GHS09 Wng | H373 ** H410 |  |  |  | ATP01
613-292-00-5 | N,N',N''-tris(2-methyl-2,3-epoxypropyl)-perhydro-2,4,6-oxo-1,3,5-triazine | 435-010-8 | 26157-73-3 | Muta. 2 Aquatic Chronic 3 | H341 H412 | GHS08 Wng | H341 H412 |  |  |  | ATP01
613-293-00-0 | 2-(4-tert-butylphenyl)-6-cyano-5-[bis(ethoxycarbonylmethyl)carbamoyloxy]-1H-pyrrolo[1,2-b][1,2,4] triazole-7-carboxylic acid 2,6-di-tert-butyl-4-methylcyclohexylester | 448-050-6 | 444065-11-6 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
613-294-00-6 | 2-hexyldecanoic acid [4-(6-tert-butyl-7-chloro-1H-pyrazolo[1,5-b][1,2,4]triazol-2-yl)phenylcarbamoyl]methylester | 448-260-8 | 379268-96-9 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
613-295-00-1 | 11-amino-3-chloro-6,11-dihydro-5,5-dioxo-6-methyl-dibenzo[c,f][1,2]thiazepine hydrochloride | 448-720-8 | 363138-44-7 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | ATP01
613-296-00-7 | pentapotassium 2-(4-(5-[1-(2,5-disulfonatophenyl)-4,5-dihydro-3-methylcarbamoyl-5-oxopyrazol-4-ylidene]-3-methyl-1,3-pentadienyl)-3-methylcarbamoyl-5-oxidopyrazol-1-yl)benzene-1,4-disulfonate | 418-270-7 | - | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
613-297-00-2 | 5-(2-bromophenyl)-2-tert-butyl-2H-tetrazole | 420-820-6 | - | Flam. Liq. 3 Acute Tox. 4 * Aquatic Chronic 2 | H226 H302 H411 | GHS02 GHS07 GHS09 Wng | H226 H302 H411 |  |  |  | ATP01
613-298-00-8 | bis-(6-hydroxy-4-methyl-5-(3-methylimidazolium-1-yl)-3-(4-phenylazo)-1H-pyridin-2-one)ethylene dilactate | 421-560-6 | - | STOT RE 2 * Eye Dam. 1 Aquatic Chronic 2 | H373 ** H318 H411 | GHS05 GHS08 GHS09 Dgr | H373 ** H318 H411 |  |  |  | ATP01
613-299-00-3 | main component 1 (isomer 1): 2-{6-fluoro-4-[3-(2,5-disulfo-phenylazo)-4-hydroxy-2-sulfonapht-7-ylamino]-1,3,5-triazin-2-ylamino}-3-{6-fluoro-4-[3-(1,5-disulfonaphth-2-ylazo)-4-hydroxy-2-sulfonaphth-7-ylamino]-1,3,5-triazin-2-ylamino}-propane sodium salt; main component 1 (isomer 2): 2-{6-fluoro-4-[3-(2,5-disulfo-phenylazo)-4-hydroxy-2-sulfonaphth-7-ylamino]-1,3,5-triazin-2-ylamino}-3-{6-fluoro-4-[3-(2,5-disulfo-phenylazo)-4-hydroxy-2-sulfonaphth-7-ylamino]-1,3,5-triazin-2-ylamino}-propane sodium salt; main component 2: 2,3-bis-{6-fluoro-4-[3-(2,5-disulfo-phenylazo)-4-hydroxy-2-sulfonaphth-7-ylamino]-1,3,5-triazin-2-ylamino}-propane sodium salt; main component 3: 2,3-bis-{6-fluoro-4-[3-(1,5-disulfonaphth-2-ylazo)-4-hydroxy-2-sulfonaphth-7-ylamino]-1,3,5-triazin-2-ylamino}-propane sodium salt | 422-610-1 | - | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
613-300-00-7 | 1-imidazol-1-yl-octadecan-2-ol | 434-120-3 | - | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
613-301-00-2 | dimethyl-1-{[2-methoxy-5-(2-methyl-butoxycarbonyl)phenylcarbamoyl]-[2-octadecyl-1,1-dioxo-1,2,4-benzothiadiazin-3-yl]methyl} imidazole-4,5-dicarboxylate | 443-910-7 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
613-302-00-8 | disodium 2-(5-carbamoyl-1-ethyl-2-hydroxy-4-methyl-6-oxo-1,6-dihydro-pyridine-3-ylazo)-4-(4-fluoro-6-(4-(2-sulfonyloxy-ethylsulfonyl)-phenylamino)-1,3,5-triazine-2-ylamino)benzene sulfonate | 432-980-4 | 243858-60-8 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | ATP01
613-303-00-3 | 2-(1-methyl-2-(4-phenoxyphenoxy)ethoxy)pyridine | 429-800-1 | 95737-68-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
613-304-00-9 | 5,6-dihydroxy-2,3-dihydro-1H-indolium bromide | 421-170-6 | 138937-28-7 | Acute Tox. 4 * Eye Dam. 1 | H302 H318 | GHS05 GHS07 Dgr | H302 H318 |  |  |  | ATP01
613-305-00-4 | 2-(2-hydroxy-4-octyloxyphenyl)-2H-benzotriazole | 448-630-9 | 3147-77-1 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
613-306-00-X | (2,5-dioxopyrrolidin-1-yl)-9H-fluoren-9-ylmethyl carbonate | 433-520-5 | 82911-69-1 | Acute Tox. 4 * Skin Sens. 1 Aquatic Chronic 2 | H302 H317 H411 | GHS07 GHS09 Wng | H302 H317 H411 |  |  |  | ATP01
613-307-00-5 | clothianidin (ISO); (E)-1-(2-chloro–1,3-thiazol-5-ylmethyl)-3-methyl-2-nitroguanidine | 433-460-1 | 210880-92-5 | Repr. 2 Acute Tox. 4 STOT SE 1 Aquatic Acute 1 Aquatic Chronic 1 | H361f H302 H370 (nervous system) H400 H410 | GHS08 GHS07 GHS09 Dgr | H361f H302 H370 (nervous system) H410 |  | oral: ATE = 390 mg/kg bw M = 10 M = 100 |  | ATP01/ATP21
613-308-00-0 | 2-amino-5-methylthiazole | 423-800-5 | 7305-71-7 | Acute Tox. 4 * STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H410 |  |  |  | ATP01
613-309-00-6 | 1-methyl-3-phenyl-1-piperazine | 431-180-2 | 5271-27-2 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 3 | H312 H302 H315 H318 H412 | GHS05 GHS07 Dgr | H312 H302 H315 H318 H412 |  |  |  | ATP01
613-310-00-1 | (-)(3S,4R)-4-(4-fluorophenyl)-3-(3,4-methylenedioxy-phenoxymethyl)-N-benzylpiperidine hydrochloride | 432-360-3 | 105813-13-6 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | ATP01
613-311-00-7 | methyl-5-nitrophenyl-guanidine | 435-500-1 | 152460-07-6 | Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H302 H319 H317 H412 | GHS07 Wng | H302 H319 H317 H412 |  |  |  | ATP01
613-312-00-2 | 2-(4-methyl-2-phenyl-1-piperazinyl)benzenemethanol monohydrochloride | 420-200-5 | - | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H302 H318 H317 H412 | GHS05 GHS07 Dgr | H302 H318 H317 H412 |  |  |  | ATP01
613-313-00-8 | 2-(4-(4-(3-pyridinyl)-1H-imidazol-1-yl)butyl)-1H-isoindole-1,3(2H)-dione | 442-780-9 | 173838-67-0 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
613-314-00-3 | 4-decyloxazolidin-2-one; 4-decyl-1,3-oxazolidin-2-one | 443-770-7 | 7693-82-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
613-315-00-9 | tetrapotassium 4-[5-[3-carboxylato-4,5-dihydro-5-oxo-1-(4-sulfonatophenyl)pyrazol-4-ylidene]-3-(piperidinocarbonyl)penta-1,3-dienylidene]-5-hydroxy-1-(4-sulfonatophenyl)pyrazole-3-carboxylate | 430-390-1 | - | Acute Tox. 4 * Aquatic Chronic 3 | H332 H412 | GHS07 Wng | H332 H412 |  |  |  | ATP01
613-316-00-4 | trimethylopropane tri(3-aziridinylpropanoate); (TAZ) | 257-765-0 | 52234-82-9 | Muta. 2 Eye Dam. 1 Skin Sens. 1 | H341 H318 H317 | GHS05 GHS08 GHS07 Dgr | H341 H318 H317 |  |  |  | ATP01
613-317-00-X | penconazole (ISO); 1-[2-(2,4-dichlorophenyl)pentyl]-1H-1,2,4-triazole |  | 66246-88-6 | Repr. 2 Acute Tox. 4 Aquatic Acute 1 Aquatic Chronic 1 | H361d H302 H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H361d H410 |  | M=1 M=1 |  | ATP06
613-318-00-5 | fenpyrazamine (ISO); S-allyl 5-amino-2,3-dihydro-2-isopropyl-3-oxo-4-(o-tolyl)pyrazole-1-carbothioate; S-allyl 5-amino-2-isopropyl-4-(2-methylphenyl)-3-oxo-2,3-dihydropyrazole-1-carbothioate |  | 473798-59-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=10 M=1 |  | ATP06/ATP10
613-319-00-0 | imidazole | 206-019-2 | 288-32-4 | Repr. 1B Acute Tox. 4 Skin Corr. 1C | H360D H302 H314 | GHS08 GHS07 GHS05 Dgr | H302 H314 H360D |  |  |  | ATP07
613-320-00-6 | lenacil (ISO); 3-cyclohexyl-6,7-dihydro-1H-cyclopenta[d]pyrimidine-2,4(3H,5H)-dione | 218-499-0 | 2164-08-1 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  | M=10 M=10 |  | ATP07
613-321-00-1 | (RS)-4-[1-(2,3-dimethylphenyl)ethyl]-1H-imidazole; medetomidine |  | 86347-14-0 | Acute Tox. 2 Acute Tox. 2 STOT SE 3 STOT SE 1 STOT RE 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H300 H336 H370 (eye) H372 H400 H410 | GHS06 GHS08 GHS09 Dgr | H330 H300 H370 (eye) H336 H372 H410 |  | M=1 M=100 |  | ATP10
613-322-00-7 | triadimenol (ISO); (1RS,2RS;1RS,2SR)-1-(4-chlorophenoxy)-3,3-dimethyl-1-(1H-1,2,4-triazol-1-yl)butan-2-ol; α-tert-butyl-β-(4-chlorophenoxy)-1H-1,2,4-triazole-1-ethanol | 259-537-6 | 55219-65-3 | Repr. 1B Lact. Acute Tox. 4 Aquatic Chronic 2 | H360 H362 H302 H411 | GHS08 GHS07 GHS09 Dgr | H302 H360 H362 H411 |  |  |  | ATP10
613-323-00-2 | terbuthylazine (ISO); N-tert-butyl-6-chloro-N′-ethyl-1,3,5-triazine-2,4-diamine | 227-637-9 | 5915-41-3 | Acute Tox. 4 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 H400 H410 | GHS07 GHS08 GHS09 Wng | H302 H373 H410 |  | M=10 M=10 |  | ATP10
613-324-00-8 | quinolin-8-ol; 8-hydroxyquinoline | 205-711-1 | 148-24-3 | Repr. 1B Acute Tox. 3 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D H301 H318 H317 H400 H410 | GHS08 GHS06 GHS05 GHS09 Dgr | H301 H318 H317 H360D H410 |  | M=1 M=1 |  | ATP10
613-325-00-3 | thiacloprid (ISO); (Z)-3-(6-chloro-3-pyridylmethyl)-1,3-thiazolidin-2-ylidenecyanamide; {(2Z)-3-[(6-chloropyridin-3-yl)methyl]-1,3-thiazolidin-2-ylidene}cyanamide |  | 111988-49-9 | Carc. 2 Repr. 1B Acute Tox. 3 Acute Tox. 4 STOT SE 3 Aquatic Acute 1 Aquatic Chronic 1 | H351 H360FD H301 H332 H336 H400 H410 | GHS08 GHS06 GHS09 Dgr | H332 H301 H351 H360FD H336 H410 |  | M=100 M=100 |  | ATP10
613-326-00-9 | 2-methylisothiazol-3(2H)-one | 220-239-6 | 2682-20-4 | Acute Tox. 2 Acute Tox. 3 Acute Tox. 3 Skin Corr. 1B Eye Dam. 1 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 1 | H330 H311 H301 H314 H318 H317 H400 H410 | GHS05 GHS06 GHS09 Dgr | H330 H311 H301 H314 H317 H410 | EUH071 | Skin Sens. 1A; H317: C ≥ 0,0015 % M=10 M=1 |  | ATP13
613-327-00-4 | pyroxsulam (ISO); N-(5,7-dimethoxy[1,2,4]triazolo[1,5-a]pyrimidin-2-yl)-2-methoxy-4-(trifluoromethyl) pyridine-3-sulfonamide |  | 422556-08-9 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  | M=100 M=100 |  | ATP13
613-328-00-X | 1-vinylimidazole | 214-012-0 | 1072-63-5 | Repr. 1B | H360D | GHS08 Dgr | H360D |  | Repr. 1B; H360D: C ≥ 0,03 % |  | ATP13
613-329-00-5 | halosulfuron-methyl (ISO); methyl 3-chloro-5-{[(4,6-dimethoxypyrimidin-2-yl)carbamoyl]sulfamoyl}-1-methyl-1H-pyrazole-4-carboxylate |  | 100784-20-1 | Repr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H360D H400 H410 | GHS08 GHS09 Dgr | H360D H410 |  | M=1000 M=1000 |  | ATP14
613-330-00-0 | 2-methylimidazole | 211-765-7 | 693-98-1 | Repr. 1B | H360Df | GHS08 Dgr | H360Df |  |  |  | ATP14
613-331-00-6 | (2RS)-2-[4-(4-chloro­phenoxy)-2-(trifluoromethyl)phenyl]-1-(1H- 1,2,4-triazol-1-yl)propan-2-ol; mefentrifluconazole |  | 1417782-03-6 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  | M=1 M=1 |  | ATP15
613-332-00-1 | oxathiapiprolin (ISO); 1-(4-{4-[5-(2,6-di­fluorophenyl)-4,5-di­hydro-1,2-oxazol-3- yl]-1,3-thiazol-2-yl}piperidin-1-yl)-2-[5- methyl-3-(trifluoro­methyl)-1H-pyrazol-1-yl]ethanone |  | 1003318-67-9 | Aquatic Chronic 1 | H410 | GHS09 Wng | H410 |  | M=1 |  | ATP15
613-333-00-7 | pyrithione zinc; (T-4)- bis[1-(hydroxy-.kappa.O)pyridine-2(1H)- thionato-.kappa.S]zinc | 236-671-3 | 13463-41-7 | Repr. 1B Acute Tox. 2 Acute Tox. 3 STOT RE 1 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H360D H330 H301 H372 H318 H400 H410 | GHS08 GHS06 GHS05 GHS09 Dgr | H330 H301 H318 H360D H372 H410 |  | Inhalation: ATE = 0.14 mg/L (dusts/mists) Oral: ATE = 221 mg/kg bw M=1000 M=10 |  | ATP15
613-334-00-2 | flurochloridone (ISO); 3-chloro-4-(chloro­methyl)-1-[3-(trifluor­omethyl)phenyl]pyrrolidin-2-one | 262-661-3 | 61213-25-0 | Repr. 1B Acute Tox. 4 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H360FD H302 H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H302 H317 H360FD H410 |  | Oral: ATE = 500 mg/kg bw M=100 M=100 |  | ATP15
613-335-00-8 | 4,5-dichloro-2-octyl- 2H-isothiazol-3-one; [DCOIT] | 264-843-8 | 64359-81-5 | Acute Tox. 2 Acute Tox. 4 Skin Corr. 1 Eye Dam. 1 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 1 | H330 H302 H314 H318 H317 H400 H410 | GHS06 GHS05 GHS09 Dgr | H330 H302 H314 H317 H410 | EUH071 | Inhalation: ATE = 0.16 mg/L (dusts/mists) Oral: ATE = 567 mg/kg bw Skin Irrit. 2; H315: 0,025 % ≤ C < 5 % Eye Irrit. 2; H319: 0,025 % ≤ C < 3 % Skin Sens. 1A; H317: C ≥ 0,0015 % M=100 M=100 |  | ATP15
613-336-00-3 | 2-methyl-1,2-ben­zothiazol-3(2H)-one; [MBIT] |  | 2527-66-4 | Acute Tox. 3 Acute Tox. 4 Skin Corr. 1C Eye Dam. 1 Skin Sens. 1A Aquatic Acute 1 Aquatic Chronic 2 | H301 H312 H314 H318 H317 H400 H411 | GHS06 GHS05 GHS09 Dgr | H312 H301 H314 H317 H410 | EUH071 | Oral: ATE = 175 mg/kg bw Dermal: ATE = 1100 mg/kg bw Skin Sens. 1A; H317: C ≥ 0,0015 % M=1 |  | ATP15
613-337-00-9 | prothioconazole (ISO); 2-[2-(1-chlorocyclopropyl)-3-(2-chlorophenyl)-2-hydroxypropyl]-2,4-dihydro-3H-1,2,4-triazole-3-thione |  | 178928-70-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=10 M=1 |  | ATP17
613-338-00-4 | azamethiphos (ISO); S-[(6-chloro-2-oxooxazolo[4,5-b]pyridin-3(2H)-yl)methyl] O,O-dimethyl thiophosphate | 252-626-0 | 35575-96-3 | Carc. 2 Acute Tox. 3 Acute Tox. 4 STOT SE 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H331 H302 H370 (nervous system) H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H331 H302 H317 H351 H370 (nervous system) H410 |  | Inhalation: ATE = 0.5 mg/L (dusts/mists) Oral: ATE = 500 mg/kg bw M=1000 M=1000 |  | ATP17
613-339-00-X | 3-methylpyrazole | 215-925-7 | 1453-58-3 | Repr. 1B Acute Tox. 4 STOT RE 2 Skin Corr. 1 Eye Dam. 1 | H360D H302 H373 (lung) H314 H318 | GHS08 GHS07 GHS05 Dgr | H302 H314 H360D H373 (lung) |  | Oral: ATE = 500 mg/kg bw |  | ATP17
613-340-00-5 | clomazone (ISO); 2-(2-chlorobenzyl)-4,4-dimethyl-1,2-oxazolidin-3-one |  | 81777-89-1 | Acute Tox. 4 Acute Tox. 4 Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H400 H410 | GHS07 GHS09 Wng | H332 H302 H410 |  | Inhalation: ATE = 4.85 mg/L (dusts/mists) Oral: ATE = 768 mg/kg bw M=1 M=1 |  | ATP17
613-341-00-0 | clofentezine (ISO); 3,6-bis(o-chlorophenyl)-1,2,4,5-tetrazine | 277-728-2 | 74115-24-5 | Aquatic Chronic 1 | H410 | GHS09 Wng | H410 |  | M=1 |  | ATP18
613-342-00-6 | theophylline; 1,3-dimethyl-3,7-dihydro-1H-purine-2,6-dione | 200-385-7 | 58-55-9 | Repr. 1B | H360D | GHS08 Dgr | H360D |  |  |  | ATP18
613-343-00-1 | pyridalyl (ISO); 2,6-dichloro-4-(3,3-dichloroallyloxy)phenyl 3-[5-(trifluoromethyl)-2-pyridyloxy]propyl ether |  | 179101-81-6 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  | M=1000 M=100 |  | ATP18
613-344-00-7 | Pyridine-2-thiol 1-oxide, sodium salt; pyrithione sodium; sodium pyrithione | 223-296-5 240-062-8 | 3811-73-2 15922-78-8 | Acute Tox. 3 Acute Tox. 3 Acute Tox. 4 STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 2 | H331 H311 H302 H372 (nervous system) H315 H319 H317 H400 H411 | GHS06 GHS08 GHS09 Dgr | H331 H311 H302 H372 (nervous system) H315 H319 H317 H410 | EUH070 | inhalation: ATE = 0,5mg/L (dusts or mists) dermal: ATE = 790 mg/kg bw oral: ATE = 500 mg/kg bw M=100 |  | ATP18
613-345-00-2 | 1,3,5-triazine-2,4,6-triamine; melamine | 203-615-4 | 108-78-1 | Carc. 2 STOT RE 2 | H351 H373 (urinary tract) | GHS08 Wng | H351 H373 (urinary tract) |  |  |  | ATP18
613-346-00-8 | 4-nitrosomorpholine |  | 59-89-2 | Carc. 1B Muta. 2 STOT RE 1 | H350 H341 H372 (liver) | GHS08 Dgr | H350 H341 H372 (liver) |  | Carc. 1B; H350: C>= 0.001 % |  | ATP21
613-347-00-3 | difenoconazole (ISO); 1-({2-[2-chloro-4-(4-chlorophenoxy)phenyl]-4-methyl–1,3-dioxolan-2-yl}methyl)-1H-1,2,4-triazole; 3-chloro-4-[(2RS,4RS;2RS,4SR)-4-methyl-2-(1H-1,2,4-triazol-1-ylmethyl)-1,3-dioxolan-2-yl]phenyl 4-chlorophenyl ether |  | 119446-68-3 | Carc. 2 Acute Tox. 4 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H319 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H302 H319 H410 |  | oral: ATE = 1450 mg/kg bw M = 10 M = 10 |  | ATP21
613-348-00-9 | 9-[2-(ethoxycarbonyl)phenyl]–3,6-bis(ethylamino)–2,7-dimethylxanthylium chloride; Basic Red 1 | 213-584-9 | 989-38-8 | Acute Tox. 3 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H301 H318 H317 H400 H410 | GHS06 GHS05 GHS09 Dgr | H301 H318 H317 H410 |  | oral: ATE = 280 mg/kg bw M = 10 M = 1 |  | ATP21
613-349-00-4 | 4-methylimidazole | 212-497-3 | 822-36-6 | Carc. 1B Repr. 1B | H350 H360Fd | GHS08 Dgr | H350 H360Fd |  |  |  | ATP21
613-350-00-X | 1H-benzotriazole | 202-394-1 | 95-14-7 | Aquatic Chronic 2 | H411 | GHS09 Wng | H411 |  |  |  | ATP22
613-351-00-5 | methyl-1H-benzotriazole | 249-596-6 | 29385-43-1 | Aquatic Chronic 2 | H411 | GHS09 Wng | H411 |  |  |  | ATP22
614-001-00-4 | nicotine (ISO); 3-[(2S)-1-methylpyrrolidin-2-yl]pyridine | 200-193-3 | 54-11-5 | Acute Tox. 2 Acute Tox. 2 Acute Tox. 2 Aquatic Chronic 2 | H330 H310 H300 H411 | GHS06 GHS09 Dgr | H330 H310 H300 H411 |  | Inhalation: ATE = 0.19 mg/L (dusts/mists) Dermal: ATE = 70 mg/kg bw Oral: ATE = 5 mg/kg bw |  | CLP00/ATP13
614-002-00-X | salts of nicotine |  |  | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * Aquatic Chronic 2 | H310 H330 H300 H411 | GHS06 GHS09 Dgr | H330 H310 H300 H411 |  |  | A | CLP00
614-003-00-5 | strychnine | 200-319-7 | 57-24-9 | Acute Tox. 1 Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H310 H300 H400 H410 | GHS06 GHS09 Dgr | H310 H300 H410 |  |  |  | CLP00
614-004-00-0 | salts of strychnine |  |  | Acute Tox. 2 * Acute Tox. 2 * Aquatic Acute 1 Aquatic Chronic 1 | H330 H300 H400 H410 | GHS06 GHS09 Dgr | H330 H300 H410 |  |  | A | CLP00
614-005-00-6 | colchicine | 200-598-5 | 64-86-8 | Muta. 1B Acute Tox. 2 * | H340 H300 | GHS06 GHS08 Dgr | H340 H300 |  |  |  | CLP00/ATP01
614-006-00-1 | brucine; 2,3-dimethoxystrychnine | 206-614-7 | 357-57-3 | Acute Tox. 2 * Acute Tox. 2 * Aquatic Chronic 3 | H330 H300 H412 | GHS06 Dgr | H330 H300 H412 |  |  |  | CLP00
614-007-00-7 | brucine sulphate [1] brucine nitrate [2] Strychnidin-10-one, 2,3-dimethoxy-, mono[(R)-1-methylheptyl 1,2-benzenedicarboxylate] [3] Strychnidin-10-one, 2,3-dimethoxy-, compd. with (S)mono(1-methylheptyl)-1,2-benzenedicarboxylate (1:1) [4] | 225-432-9 [1] 227-317-9 [2] 269-439-5 [3] 269-710-8 [4] | 4845-99-2 [1] 5786-97-0 [2] 68239-26-9 [3] 68310-42-9 [4] | Acute Tox. 2 * Acute Tox. 2 * Aquatic Chronic 3 | H330 H300 H412 | GHS06 Dgr | H330 H300 H412 |  |  | A | CLP00
614-008-00-2 | aconitine | 206-121-7 | 302-27-2 | Acute Tox. 2 * Acute Tox. 2 * | H330 H300 | GHS06 Dgr | H330 H300 |  |  |  | CLP00
614-009-00-8 | salts of aconitine |  |  | Acute Tox. 2 * Acute Tox. 2 * | H330 H300 | GHS06 Dgr | H330 H300 |  |  | A | CLP00
614-010-00-3 | atropine | 200-104-8 | 51-55-8 | Acute Tox. 2 * Acute Tox. 2 * | H330 H300 | GHS06 Dgr | H330 H300 |  |  |  | CLP00
614-011-00-9 | salts of atropine |  |  | Acute Tox. 2 * Acute Tox. 2 * | H330 H300 | GHS06 Dgr | H330 H300 |  |  | A | CLP00
614-012-00-4 | hyoscyamine | 202-933-0 | 101-31-5 | Acute Tox. 2 * Acute Tox. 2 * | H330 H300 | GHS06 Dgr | H330 H300 |  |  |  | CLP00
614-013-00-X | salts of hyoscyamine |  |  | Acute Tox. 2 * Acute Tox. 2 * | H330 H300 | GHS06 Dgr | H330 H300 |  |  | A | CLP00
614-014-00-5 | hyoscine | 200-090-3 | 51-34-3 | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * | H310 H330 H300 | GHS06 Dgr | H330 H310 H300 |  |  |  | CLP00
614-015-00-0 | salts of hyoscine |  |  | Acute Tox. 1 Acute Tox. 2 * Acute Tox. 2 * | H310 H330 H300 | GHS06 Dgr | H330 H310 H300 |  |  | A | CLP00
614-016-00-6 | pilocarpine | 202-128-4 | 92-13-7 | Acute Tox. 2 * Acute Tox. 2 * | H330 H300 | GHS06 Dgr | H330 H300 |  |  |  | CLP00
614-017-00-1 | salts of pilocarpine |  |  | Acute Tox. 2 * Acute Tox. 2 * | H330 H300 | GHS06 Dgr | H330 H300 |  |  | A | CLP00
614-018-00-7 | papaverine | 200-397-2 | 58-74-2 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
614-019-00-2 | salts of papaverine |  |  | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  | A | CLP00
614-020-00-8 | physostigmine | 200-332-8 | 57-47-6 | Acute Tox. 2 * Acute Tox. 2 * | H330 H300 | GHS06 Dgr | H330 H300 |  |  |  | CLP00
614-021-00-3 | salts of physostigmine |  |  | Acute Tox. 2 * Acute Tox. 2 * | H330 H300 | GHS06 Dgr | H330 H300 |  |  | A | CLP00
614-022-00-9 | digitoxin | 200-760-5 | 71-63-6 | Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * | H331 H301 H373 ** | GHS06 GHS08 Dgr | H331 H301 H373 ** |  |  |  | CLP00
614-023-00-4 | ephedrine | 206-080-5 | 299-42-3 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
614-024-00-X | salts of ephedrine |  |  | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  | A | CLP00
614-025-00-5 | ouabain | 211-139-3 | 630-60-4 | Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * | H331 H301 H373 ** | GHS06 GHS08 Dgr | H331 H301 H373 ** |  |  |  | CLP00
614-026-00-0 | strophantin-K | 234-239-9 | 11005-63-3 | Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * | H331 H301 H373 ** | GHS06 GHS08 Dgr | H331 H301 H373 ** |  |  |  | CLP00
614-027-00-6 | bufa-4,20,22-trienolide, 6-(acetyloxy)-3-(β-D-glucopyranosyloxy)-8,14-dihydroxy-, (3β, 6β)-; red squill; scilliroside | 208-077-4 | 507-60-8 | Acute Tox. 2 * | H300 | GHS06 Dgr | H300 |  |  |  | CLP00
614-028-00-1 | reaction mass of: 2-ethylhexyl mono-D-glucopyranoside; 2-ethylhexyl di-D-glucopyranoside | 414-420-0 |  | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
614-029-00-7 | constitutional isomers of penta-O-allyl-β-D-fructofuranosyl-α-D-glucopyranoside; constitutional isomers of hexa-O-allyl-β-D-fructofuranosyl-α-D-glucopyranoside; constitutional isomers of hepta-O-allyl-β-D-fructofuransoyl-α-D-glucopyranoside | 419-640-0 | 68784-14-5 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
614-030-00-2 | emamectin benzoate (ISO); (4"R)-4"-deoxy-4"-(methylamino) avermectin B1 benzoate |  | 155569-91-8 | Acute Tox. 3 Acute Tox. 3 Acute Tox. 3 STOT SE 1 STOT RE 1 Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H311 H301 H370 (nervous system) H372 (nervous system) H318 H400 H410 | GHS06 GHS05 GHS08 GHS09 Dgr | H331 H311 H301 H318 H370 (nervous system) H372 (nervous system) H410 |  | Inhalation: ATE = 0.663 mg/L (dusts/mists) Dermal: ATE = 300 mg/kg bw Oral: ATE = 60 mg/kg bw STOT RE 1; H372: C ≥ 5 % STOT RE 2; H373: 0,5 % ≤ C < 5 % M=10000 M=10000 |  | ATP17
615-001-00-7 | methyl isocyanate | 210-866-3 | 624-83-9 | Flam. Liq. 2 Repr. 2 Acute Tox. 2 * Acute Tox. 3 * Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H225 H361d *** H330 H311 H301 H335 H315 H318 H334 H317 | GHS02 GHS06 GHS05 GHS08 Dgr | H225 H361d *** H330 H311 H301 H334 H317 H335 H315 H318 |  |  |  | CLP00/ATP01
615-002-00-2 | methyl isothiocyanate | 209-132-5 | 556-61-6 | Acute Tox. 3 * Acute Tox. 3 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H314 H317 H400 H410 | GHS06 GHS05 GHS09 Dgr | H331 H301 H314 H317 H410 |  |  |  | CLP00
615-003-00-8 | thiocyanic acid | 207-337-4 | 463-56-9 | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 3 | H332 H312 H302 H412 | GHS07 Wng | H332 H312 H302 H412 | EUH032 |  |  | CLP00
615-004-00-3 | salts of thiocyanic acid, with the exception of those specified elsewhere in this Annex | - | - | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 3 | H332 H312 H302 H412 | GHS07 Wng | H332 H312 H302 H412 | EUH032 |  | A | CLP00/ATP01
615-005-00-9 | 4,4'-methylenediphenyl diisocyanate; diphenylmethane-4,4'-diisocyanate [1] 2,2'-methylenediphenyl diisocyanate; diphenylmethane-2,2'-diisocyanate [2] o-(p-isocyanatobenzyl)phenyl isocyanate; diphenylmethane-2,4'-diisocyanate [3] methylenediphenyl diisocyanate [4] | 202-966-0 [1] 219-799-4 [2] 227-534-9 [3] 247-714-0 [4] | 101-68-8 [1] 2536-05-2 [2] 5873-54-1 [3] 26447-40-5 [4] | Carc. 2 Acute Tox. 4 * STOT SE 3 STOT RE 2 * Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 Skin Sens. 1 | H351 H332 H335 H373 ** H315 H319 H334 H317 | GHS08 GHS07 Dgr | H351 H332 H373 ** H319 H335 H315 H334 H317 |  | Eye Irrit. 2; H319: C ≥ 5 % Skin Irrit. 2; H315: C ≥ 5 % Resp. Sens. 1; H334: C ≥ 0,1 % STOT SE 3; H335: C ≥ 5 % | 2 C | CLP00/ATP01
615-006-00-4 | 2-methyl-m-phenylene diisocyanate; toluene-2,6-di-isocyanate [1] 4-methyl-m-phenylene diisocyanate; toluene-2,4-di-isocyanate [2] m-tolylidene diisocyanate; toluene-diisocyanate [3] | 202-039-0 [1] 209-544-5 [2] 247-722-4 [3] | 91-08-7 [1] 584-84-9 [2] 26471-62-5 [3] | Carc. 2 Acute Tox. 2 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 Skin Sens. 1 Aquatic Chronic 3 | H351 H330 H335 H315 H319 H334 H317 H412 | GHS06 GHS08 Dgr | H351 H330 H319 H335 H315 H334 H317 H412 |  | Resp. Sens. 1; H334: C ≥ 0,1 % | C | CLP00
615-008-00-5 | 3-isocyanatomethyl-3,5,5-trimethylcyclohexyl isocyanate; isophorone di-isocyanate | 223-861-6 | 4098-71-9 | Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 Skin Sens. 1 Aquatic Chronic 2 | H331 H335 H315 H319 H334 H317 H411 | GHS06 GHS08 GHS09 Dgr | H331 H319 H335 H315 H334 H317 H411 |  | * Resp. Sens. 1; H334: C ≥ 0,5 % Skin Sens.1; H317: C ≥ 0,5 % | 2 | CLP00
615-009-00-0 | 4,4'-methylenedi(cyclohexyl isocyanate); dicyclohexylmethane-4,4'-di-isocyanate | 225-863-2 | 5124-30-1 | Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 Skin Sens. 1 | H331 H335 H315 H319 H334 H317 | GHS06 GHS08 Dgr | H331 H319 H335 H315 H334 H317 |  | * Resp. Sens. 1; H334: C ≥ 0,5 % Skin Sens. 1; H317: C ≥ 0,5 % | 2 | CLP00
615-010-00-6 | 2,2,4-trimethylhexamethylene-1,6-di-isocyanate [1] 2,4,4-trimethylhexamethylene-1,6-di-isocyanate [2] | 241-001-8 [1] 239-714-4 [2] | 16938-22-0 [1] 15646-96-5 [2] | Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H331 H335 H315 H319 H334 | GHS06 GHS08 Dgr | H331 H319 H335 H315 H334 |  | * Resp. Sens. 1; H334: C ≥ 0,5 % Skin Sens. 1; H317: C ≥ 0,5 % | 2 C | CLP00
615-011-00-1 | hexamethylene-di-isocyanate | 212-485-8 | 822-06-0 | Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 Skin Sens. 1 | H331 H335 H315 H319 H334 H317 | GHS06 GHS08 Dgr | H331 H319 H335 H315 H334 H317 |  | * Resp. Sens. 1; H334: C ≥ 0,5 % Skin Sens. 1; H317: C ≥ 0,5 % | 2 | CLP00
615-012-00-7 | 4-isocyanatosulphonyltoluene; tosyl isocyanate | 223-810-8 | 4083-64-1 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H335 H315 H319 H334 | GHS08 GHS07 Dgr | H319 H335 H315 H334 | EUH014 | Eye Irrit.; H319: C ≥ 5 % STOT SE 3; H335: C ≥ 5 % Skin Irrit. 2; H315: C ≥ 5 % |  | CLP00
615-013-00-2 | cyanamide; carbamonitril | 206-992-3 | 420-04-2 | Carc. 2 Repr. 2 Acute Tox. 3 Acute Tox. 3 STOT RE 2 Skin Corr. 1 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H351 H361fd H311 H301 H373 (thyroid) H314 H318 H317 H412 | GHS08 GHS06 GHS05 Dgr | H311 H301 H314 H317 H351 H361fd H373 (thyroid) H412 |  |  |  | CLP00/ATP10
615-014-00-8 | tris(1-dodecyl-3-methyl-2-phenylbenzimidazolium)hexacyanoferrate |  | 7276-58-6 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
615-015-00-3 | 1,7,7-trimethylbicyclo(2,2,1)hept-2-yl thiocyanatoacetate; isobornyl thiocyanoacetate | 204-081-5 | 115-31-1 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
615-016-00-9 | potassium cyanate | 209-676-3 | 590-28-3 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
615-017-00-4 | calcium cyanamide | 205-861-8 | 156-62-7 | Acute Tox. 4 * STOT SE 3 Eye Dam. 1 | H302 H335 H318 | GHS05 GHS07 Dgr | H302 H335 H318 |  |  |  | CLP00
615-018-00-X | 2-(2-butoxyethoxy)ethyl thiocyanate | 203-985-7 | 112-56-1 | Flam. Liq. 3 Acute Tox. 3 * Acute Tox. 3 * | H226 H311 H301 | GHS02 GHS06 Dgr | H226 H311 H301 |  |  |  | CLP00
615-019-00-5 | dicyclohexylcarbodiimide | 208-704-1 | 538-75-0 | Acute Tox. 3 * Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 | H311 H302 H318 H317 | GHS06 GHS05 Dgr | H311 H302 H318 H317 |  |  |  | CLP00
615-020-00-0 | methylene dithiocyanate | 228-652-3 | 6317-18-6 | Acute Tox. 2 * Acute Tox. 3 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 | H330 H301 H314 H317 H400 | GHS06 GHS05 GHS09 Dgr | H330 H301 H314 H317 H400 |  |  |  | CLP00
615-021-00-6 | 1,3,5-tris(oxiranylmethyl)-1,3,5-triazine-2,4,6(1H,3H,5H)-trione; TGIC | 219-514-3 | 2451-62-9 | Muta. 1B Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H340 H331 H301 H373 ** H318 H317 H412 | GHS06 GHS08 GHS05 Dgr | H340 H331 H301 H373 ** H318 H317 H412 |  |  |  | CLP00
615-022-00-1 | methyl 3-isocyanatosulfonyl-2-thiophene-carboxylate | 410-550-7 | 79277-18-2 | STOT RE 2 * Resp. Sens. 1 Skin Sens. 1 | H373 ** H334 H317 | GHS08 Dgr | H373 ** H334 H317 | EUH014 |  |  | CLP00/ATP01
615-023-00-7 | 2-(isocyanatosulfonylmethyl)benzoic acid methyl ester; (alt.):methyl 2-(isocyanatosulfonylmethyl)benzoate | 410-900-9 | 83056-32-0 | Flam. Liq. 3 Muta. 2 Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Resp. Sens. 1 | H226 H341 H332 H373 ** H318 H334 | GHS02 GHS08 GHS05 GHS07 Dgr | H226 H341 H332 H373 ** H318 H334 | EUH014 |  |  | CLP00
615-024-00-2 | 2-phenylethylisocyanate | 413-080-0 | 1943-82-4 | Acute Tox. 3 * Acute Tox. 4 * Skin Corr. 1A Resp. Sens. 1 Skin Sens. 1 Aquatic Chronic 2 | H331 H302 H314 H334 H317 H411 | GHS06 GHS08 GHS05 GHS09 Dgr | H331 H302 H314 H334 H317 H411 |  |  |  | CLP00
615-025-00-8 | 4,4'-ethylidenediphenyl dicyanate | 405-740-1 | 47073-92-7 | Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H332 H302 H373 ** H318 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H332 H302 H373 ** H318 H410 |  |  |  | CLP00
615-026-00-3 | 4,4'-methylenebis(2,6-dimethylphenyl cyanate) | 405-790-4 | 101657-77-6 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
615-028-00-4 | ethyl 2-(isocyanatosulfonyl)benzoate | 410-220-2 | 77375-79-2 | Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Resp. Sens. 1 Skin Sens. 1 | H302 H373 ** H318 H334 H317 | GHS05 GHS08 GHS07 Dgr | H302 H373 ** H318 H334 H317 | EUH014 |  |  | CLP00/ATP01
615-029-00-X | 2,5-bis-isocyanatomethyl-bicyclo[2.2.1]heptane | 411-280-2 |  | Acute Tox. 2 * Acute Tox. 4 * Skin Corr. 1B Resp. Sens. 1 Skin Sens. 1 Aquatic Chronic 3 | H330 H302 H314 H334 H317 H412 | GHS06 GHS08 GHS05 Dgr | H330 H302 H314 H334 H317 H412 |  |  |  | CLP00
615-030-00-5 | alkali salts and alkali earth salts of thiocyanic acid, with the exception of those specified elsewhere in this Annex | - | - | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Chronic 3 | H332 H312 H302 H412 | GHS07 Wng | H332 H312 H302 H412 |  |  | A | CLP00/ATP01
615-031-00-0 | thallium thiocyanate | 222-571-7 | 3535-84-0 | Acute Tox. 2 * Acute Tox. 2 * Acute Tox. 4 * STOT RE 2 Aquatic Chronic 2 | H330 H300 H312 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H330 H300 H312 H373 ** H411 |  |  |  | CLP00/ATP01
615-032-00-6 | metal salts of thiocyanic acid, with the exception of those specified elsewhere in this Annex | - | - | Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H332 H312 H302 H400 H410 | GHS07 GHS09 Wng | H332 H312 H302 H410 |  |  | A | CLP00/ATP01
615-033-00-1 | reaction product of diphenylmethanediisocyanate, octylamine, oleylamine and cyclohexylamine (1:1.58:0.32:0.097) | 430-980-9 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
615-034-00-7 | reaction product of diphenylmethanediisocyanate, octylamine, 4-ethoxyaniline and ethylenediamine (1:0,37:1,53:0,05) | 430-750-8 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
615-035-00-2 | reaction product of diphenylmethanediisocyanate, octylamine and oleylamine (molar ratio 1:1.86:0.14) | 430-930-6 | 122886-55-9 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
615-036-00-8 | reaction product of diphenylmethanediisocyanate, toluenediisocyanate ( reaction mass of isomers: 65 % 2,4- and 35 % 2,6-diisocyanate), octylamine, oleylamine and 4-ethoxyaniline (molar ratio 4:1:7:1:2) | 430-940-0 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
615-037-00-3 | reaction product of diphenylmethanediisocyanate, toluenediisocyanate ( reaction mass of isomers: 65 % 2,4- and 35 % 2,6-diisocyanate), octylamine and oleylamine (molar ratio 4:1:9:1) | 430-950-5 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
615-038-00-9 | reaction product of toluenediisocyanate ( reaction mass of isomers: 65 % 2,4- and 35 % 2,6-diisocyanate) and aniline (molarratio 1:2) | 430-960-1 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
615-039-00-4 | reaction product of diphenylmethanediisocyanate, toluenediisocyanate ( reaction mass of isomers: 65 % 2,4- and 35 % 2,6-diisocyanate), octylamine, oleylamine and 4-ethoxyaniline (molar ratio 3.88:1:6.38:0.47:2.91) | 430-970-4 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
615-044-00-1 | 4-chlorophenylisocyanate | 203-176-9 | 104-12-1 | Acute Tox. 2 * Acute Tox. 4 * STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Resp. Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H330 H302 H335 H315 H318 H334 H400 H410 | GHS06 GHS05 GHS08 GHS09 Dgr | H330 H302 H335 H315 H318 H334 H410 |  |  |  | ATP01
615-045-00-7 | 4,4'-methylene bis(3-chloro-2,6-di-ethylphenylisocyanate) | 420-530-1 | - | Resp. Sens. 1 Skin Sens. 1 Aquatic Chronic 4 | H334 H317 H413 | GHS08 Dgr | H334 H317 H413 |  |  |  | ATP01
615-046-00-2 | 1,3-bis(1-isocyanato-1-methylethyl)benzene; [m-TMXDI] | 220-474-4 | 2778-42-9 | Resp. Sens. 1 Skin Sens. 1A | H334 H317 | GHS08 Dgr | H334 H317 |  |  |  | ATP18
615-047-00-8 | Bis(isocyanatomethyl)benzene; [m-XDI] | 222-852-4 | 3634-83-1 | Resp. Sens. 1 Skin Sens. 1A | H334 H317 | GHS08 Dgr | H334 H317 |  | Skin Sens. 1A; H317: C ≥ 0,001 % |  | ATP18
615-048-00-3 | 2,4,6-triisopropyl-m-phenylene diisocyanate | 218-485-4 | 2162-73-4 | Resp. Sens. 1 Skin Sens. 1 | H334 H317 | GHS08 Dgr | H334 H317 |  |  |  | ATP18
615-049-00-9 | 1,5-naphthylene diisocyanate [containing < 0.1 % (w/w) of particles with an aerodynamic diameter of below 50 µm] | 221-641-4 | 3173-72-6 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 Skin Sens. 1A Aquatic Chronic 3 | H335 H315 H319 H334 H317 H412 | GHS07 GHS08 Dgr | H335 H315 H319 H334 H317 H412 |  |  |  | ATP18
615-050-00-4 | 1,5-naphthylene diisocyanate [containing ≥ 0.1 % (w/w) of particles with an aerodynamic diameter of below 50 µm] | 221-641-4 | 3173-72-6 | Acute Tox. 2 STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 Skin Sens. 1A Aquatic Chronic 3 | H330 H335 H315 H319 H334 H317 H412 | GHS06 GHS08 Dgr | H330 H335 H315 H319 H334 H317 H412 |  | inhalation: ATE = 0,27 mg/L (dusts or mists) |  | ATP18
615-051-00-X | 3,3'-dimethylbiphenyl–4,4'-diyl diisocyanate | 202-112-7 | 91-97-4 | Carc. 2 Resp. Sens. 1 Skin Sens. 1A | H351 H334 H317 | GHS08 Dgr | H351 H334 H317 |  | Skin Sens. 1A; H317: C >= 0.001 % |  | ATP21
616-001-00-X | N,N-dimethylformamide; dimethyl formamide | 200-679-5 | 68-12-2 | Repr. 1B Acute Tox. 4 * Acute Tox. 4 * Eye Irrit. 2 | H360D *** H332 H312 H319 | GHS08 GHS07 Dgr | H360D *** H332 H312 H319 |  |  |  | CLP00
616-002-00-5 | 2-fluoroacetamide | 211-363-1 | 640-19-7 | Acute Tox. 2 * Acute Tox. 3 * | H300 H311 | GHS06 Dgr | H300 H311 |  |  |  | CLP00
616-003-00-0 | acrylamide; prop-2-enamide | 201-173-7 | 79-06-1 | Carc. 1B Muta. 1B Repr. 2 Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 1 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 | H350 H340 H361f *** H301 H332 H312 H372 ** H315 H319 H317 | GHS06 GHS08 Dgr | H350 H340 H361f *** H301 H372 ** H332 H312 H319 H315 H317 |  |  | D | CLP00
616-004-00-6 | allidochlor (ISO); N,N-diallylchloroacetamide | 202-270-7 | 93-71-0 | Acute Tox. 4 * Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H312 H302 H315 H319 H411 | GHS07 GHS09 Wng | H312 H302 H319 H315 H411 |  |  |  | CLP00
616-005-00-1 | chlorthiamid (ISO); 2,6-dichloro (thiobenzamide) | 217-637-7 | 1918-13-4 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
616-006-00-7 | dichlofluanid (ISO); N-[(dichlorofluoromethyl)thio]-N′,N′-dimethyl-N-phenylsulfamide | 214-118-7 | 1085-98-9 | Acute Tox. 4 Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 | H332 H319 H317 H400 | GHS07 GHS09 Wng | H332 H319 H317 H400 |  | M=10 |  | CLP00/ATP10
616-007-00-2 | diphenamid (ISO); N,N-dimethyl-2,2-diphenylacetamide | 213-482-4 | 957-51-7 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
616-008-00-8 | propachlor (ISO); 2-chloro-N-isopropylacetanilide; α-chloro-N-isopropylacetanilide | 217-638-2 | 1918-16-7 | Acute Tox. 4 * Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H319 H317 H400 H410 | GHS07 GHS09 Wng | H302 H319 H317 H410 |  |  |  | CLP00
616-009-00-3 | propanil (ISO); 3',4'-dichloropropionanilide | 211-914-6 | 709-98-8 | Acute Tox. 4 * Aquatic Acute 1 | H302 H400 | GHS07 GHS09 Wng | H302 H400 |  | M=10 |  | CLP00/ATP01
616-010-00-9 | tosylchloramide sodium | 204-854-7 | 127-65-1 | Acute Tox. 4 * Skin Corr. 1B Resp. Sens. 1 | H302 H314 H334 | GHS08 GHS05 GHS07 Dgr | H302 H314 H334 | EUH031 |  |  | CLP00
616-011-00-4 | N,N-dimethylacetamide | 204-826-4 | 127-19-5 | Repr. 1B Acute Tox. 4 * Acute Tox. 4 * | H360D *** H332 H312 | GHS08 GHS07 Dgr | H360D *** H332 H312 |  |  |  | CLP00/ATP09
616-012-00-X | N-(dichlorofluoromethylthio)phthalimide; N-(fluorodichloromethylthio)phthalimide | 211-952-3 | 719-96-0 | Skin Irrit. 2 | H315 | GHS07 Wng | H315 |  |  |  | CLP00
616-013-00-5 | butyraldehyde oxime | 203-792-8 | 110-69-0 | Acute Tox. 3 * Acute Tox. 4 * Eye Irrit. 2 | H311 H302 H319 | GHS06 Dgr | H311 H302 H319 |  |  |  | CLP00
616-014-00-0 | butanone oxime; ethyl methyl ketoxime; ethyl methyl ketone oxime | 202-496-6 | 96-29-7 | Carc. 1B Acute Tox. 3 Acute Tox. 4 STOT SE 3 STOT SE 1 STOT RE 2 Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 | H350 H301 H312 H336 H370 (upper respiratory tract) H373 (blood system) H315 H318 H317 | GHS08 GHS06 GHS05 Dgr | H312 H301 H315 H318 H317 H350 H370 (upper respiratory tract) H336 H373 (blood system) |  | Oral: ATE = 100 mg/kg bw Dermal: ATE = 1100 mg/kg bw |  | CLP00/ATP15
616-015-00-6 | alachlor (ISO); 2-chloro-2',6'-diethyl-N-(methoxymethyl)acetanilide | 240-110-8 | 15972-60-8 | Carc. 2 Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H302 H317 H410 |  | M=10 |  | CLP00
616-016-00-1 | 1-(3,4-dichlorophenylimino) thiosemicarbazide |  | 5836-73-7 | Acute Tox. 2 * | H300 | GHS06 Dgr | H300 |  |  |  | CLP00
616-017-00-7 | cartap hydrochloride | 239-309-2 | 15263-52-2 | Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H312 H302 H400 H410 | GHS07 GHS09 Wng | H312 H302 H410 |  |  |  | CLP00
616-018-00-2 | diethyltoluamide (ISO): N,N-diethyl-m-toluamide; [deet] | 205-149-7 | 134-62-3 | Acute Tox. 4 Skin Irrit. 2 Eye Irrit. 2 | H302 H315 H319 | GHS07 Wng | H302 H315 H319 |  | Oral: ATE = 1892 mg/kg bw |  | CLP00/ATP14
616-019-00-8 | perfluidone (ISO); 1,1,1-trifluoro-N-(4-phenylsulphonyl-o-tolyl)methanesulphonamide | 253-718-3 | 37924-13-3 | Acute Tox. 4 * Eye Irrit. 2 | H302 H319 | GHS07 Wng | H302 H319 |  |  |  | CLP00
616-020-00-3 | tebuthiuron (ISO); 1-(5-tert-butyl-1,3,4-thiadiazol-2-yl)-1,3-dimethylurea | 251-793-7 | 34014-18-1 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
616-021-00-9 | thiazafluron (ISO); 1,3-dimethyl-1-(5-trifluoromethyl-1,3,4-thiadiazol-2-yl)urea | 246-901-4 | 25366-23-8 | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
616-022-00-4 | acetamide | 200-473-5 | 60-35-5 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | CLP00
616-023-00-X | N-hexadecyl(or octadecyl)-N-hexadecyl(or octadecyl)benzamide | 401-980-6 |  | Skin Irrit. 2 Skin Sens. 1 | H315 H317 | GHS07 Wng | H315 H317 |  |  |  | CLP00
616-024-00-5 | 2-(4,4-dimethyl-2,5-dioxooxazolidin-1-yl)-2-chloro-5-(2-(2,4-di-tert-pentylphenoxy)butyramido)-4,4-dimethyl-3-oxovaleranilide | 402-260-4 | 54942-74-4 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-025-00-0 | valinamide | 402-840-7 | 20108-78-5 | Repr. 2 Eye Irrit. 2 Skin Sens. 1 | H361f *** H319 H317 | GHS08 Wng | H361f *** H319 H317 |  |  |  | CLP00
616-026-00-6 | thioacetamide | 200-541-4 | 62-55-5 | Carc. 1B Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 3 | H350 H302 H315 H319 H412 | GHS08 GHS07 Dgr | H350 H302 H319 H315 H412 |  |  |  | CLP00
616-027-00-1 | tris(2-(2-hydroxyethoxy)ethyl)ammonium 3-acetoacetamido-4-methoxybenzenesulfonate | 403-760-5 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
616-028-00-7 | N-(4-(3-(4-cyanophenyl)ureido)-3-hydroxyphenyl)-2-(2,4-di-tert-pentylphenoxy)octanamide | 403-790-9 | 108673-51-4 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
616-029-00-2 | N,N'-ethylenebis(vinylsulfonylacetamide) | 404-790-1 | 66710-66-5 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | CLP00
616-030-00-8 | ethidimuron (ISO); 1-(5-ethylsulphonyl-1,3,4-thiadiazol-2-yl)-1,3-dimethylurea | 250-010-6 | 30043-49-3 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
616-031-00-3 | dimethachlor (ISO); 2-chloro-N-(2,6-dimethylphenyl)-N-(2-methoxyethyl)acetamide | 256-625-6 | 50563-36-5 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
616-032-00-9 | diflufenican (ISO); N-(2,4-difluorophenyl)-2-[3-(trifluoromethyl)phenoxy]-3-pyridinecarboxamide; 2′,4′-difluoro-2-(α,α,α- trifluoro-m-tolyloxy) nicotinanilide; N-(2,4-difluorophenyl)-2-[3-(trifluoromethyl)phenoxy]-3-pyridinecarboxamide |  | 83164-33-4 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=10000 M=1000 |  | CLP00/ATP17
616-033-00-4 | cyprofuram (ISO); N-(3-chlorophenyl)-N-(tetrahydro-2-oxo-3-furyl)cyclopropanecarboxamide | 274-050-9 | 69581-33-5 | Acute Tox. 3 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H312 H400 H410 | GHS06 GHS09 Dgr | H301 H312 H410 |  |  |  | CLP00
616-034-00-X | pyracarbolid (ISO); 3,4-dihydro-6-methyl-2H-pyran-5-carboxanilide | 246-419-4 | 24691-76-7 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
616-035-00-5 | cymoxanil (ISO); 2-cyano-N-[(ethylamino)carbonyl]-2-(methoxyimino)acetamide;[1] (2E)-2-cyano-N-[(ethylamino)carbonyl]-2-(methoxyimino)acetamide;[2] | 261-043-0[1] -[2] | 57966-95-7[1] 166900-80-7[2] | Repr. 2 Acute Tox. 4 STOT RE 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361fd H302 H373 (blood ssytem, thymus, eyes) H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H361fd H302 H373 (blood ssytem, thymus, eyes) H317 H410 |  | oral: ATE = 360 mg/kg bw oral: ATE = 360 mg/kg bw M = 1 M = 1 |  | CLP00/ATP21
616-036-00-0 | 2-chloracetamide | 201-174-2 | 79-07-2 | Repr. 2 Acute Tox. 3 * Skin Sens. 1 | H361f *** H301 H317 | GHS06 GHS08 Dgr | H361f *** H301 H317 |  | Skin Sens. 1; H317: C ≥ 0,1 % |  | CLP00
616-037-00-6 | acetochlor (ISO); 2-chloro-N-(ethoxymethyl)-N-(2-ethyl-6-methylphenyl)acetamide | 251-899-3 | 34256-82-1 | Carc. 2 Repr. 2 Acute Tox. 4 STOT SE 3 STOT RE 2 Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H361f H332 H335 H373 (kidney) H315 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H332 H315 H317 H351 H361f H335 H373 (kidney) H410 |  | M=1000 M=100 |  | CLP00/ATP09
616-038-00-1 | (4-aminophenyl)-N-methylmethylensulfonamide hydrochloride | 406-010-5 | 88918-84-7 | Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H318 H317 H411 | GHS05 GHS07 GHS09 Dgr | H318 H317 H411 |  |  |  | CLP00
616-039-00-7 | 3',5'-dichloro-4'-ethyl-2'-hydroxypalmitanilide | 406-200-8 | 117827-06-2 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
616-040-00-2 | potassium N-(4-toluenesulfonyl)-4-toluenesulfonamide | 406-650-5 | 97888-41-0 | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
616-041-00-8 | 3',5'-dichloro-2-(2,4-di-tert-pentylphenoxy)-4'-ethyl-2'-hydroxyhexananilide | 406-840-8 | 101664-25-9 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-042-00-3 | N-(2-(6-ethyl-7-(4-methylphenoxy)-1H-pyrazolo[1,5-b][1,2,4]triazol-2-yl)propyl)-2-octadecyloxybenzamide | 407-070-5 | 142859-67-4 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
616-043-00-9 | isoxaben (ISO); N-[3-(1-ethyl-1-methylpropyl)-1,2-oxazol-5-yl]-2,6-dimethoxybenzamide | 407-190-8 | 82558-50-7 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-044-00-4 | N-(3,5-dichloro-4-ethyl-2-hydroxyphenyl)-2-(3-pentadecylphenoxy)-butanamide | 402-510-2 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
616-045-00-X | 2'-(4-chloro-3-cyano-5-formyl-2-thienylazo)-5'-diethylamino-2-methoxyacetanilide | 405-190-2 | 122371-93-1 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
616-046-00-5 | N-(2-(6-chloro-7-methylpyrazolo(1,5-b)-1,2,4-triazol-4-yl)propyl)-2-(2,4-di-tert-pentylphenoxy)octanamide | 406-390-2 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
616-047-00-0 | reaction mass of: 2,2',2'',2'''-(ethylenedinitrilotetrakis-N,N-di(C16)alkylacetamide; 2,2',2'',2'''-(ethylenedinitrilotetrakis-N,N-di(C18)alkylacetamide | 406-640-0 |  | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
616-048-00-6 | 3'-trifluoromethylisobutyranilide | 406-740-4 | 1939-27-1 | STOT RE 2 * Aquatic Chronic 2 | H373 ** H411 | GHS08 GHS09 Wng | H373 ** H411 |  |  |  | CLP00
616-049-00-1 | 2-(2,4-bis(1,1-dimethylethyl)phenoxy)-N-(3,5-dichloro-4-ethyl-2-hydroxyphenyl)-hexanamide | 408-150-2 | 99141-89-6 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-050-00-7 | lufenuron (ISO); N-[2,5-dichloro-4-(1,1,2,3,3,3-hexafluoropropoxy)-phenyl-aminocarbonyl]-2,6-difluorobenzamide | 410-690-9 | 103055-07-8 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
616-051-00-2 | reaction mass of: 2,4 -bis(N'-(4-methylphenyl)-ureido)-toluene; 2,6 -bis(N'-(4-methylphenyl)-ureido)-toluene | 411-070-0 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-052-00-8 | formamide | 200-842-0 | 75-12-7 | Repr. 1B | H360D *** | GHS08 Dgr | H360D *** |  |  |  | CLP00
616-053-00-3 | N-methylacetamide | 201-182-6 | 79-16-3 | Repr. 1B | H360D *** | GHS08 Dgr | H360D *** |  |  |  | CLP00
616-054-00-9 | iprodione (ISO); 3-(3,5-dichlorophenyl)-2,4-dioxo-N-isopropylimidazolidine-1-carboxamide | 253-178-9 | 36734-19-7 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  |  |  | CLP00
616-055-00-4 | propyzamide (ISO); 3,5-dichloro-N-(1,1-dimethylprop-2-ynyl)benzamide | 245-951-4 | 23950-58-5 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  |  |  | CLP00
616-056-00-X | N-methylformamide | 204-624-6 | 123-39-7 | Repr. 1B Acute Tox. 4 * | H360D *** H312 | GHS08 GHS07 Dgr | H360D *** H312 |  |  |  | CLP00
616-057-00-5 | reaction mass of: N-[3-hydroxy-2-(2-methylacryloylaminomethoxy)propoxymethyl]-2-methylacrylamide; N-[2,3-bis-(2-methylacryloylaminomethoxy)propoxymethyl]-2-methylacrylamide; methacrylamide; 2-methyl-N-(2-methylacryloylaminomethoxymethyl)-acrylamide; N-(2,3-dihydroxypropoxymethyl)-2-methylacrylamide | 412-790-8 |  | Carc. 1B Muta. 2 STOT RE 2 * | H350 H341 H373 ** | GHS08 Dgr | H350 H341 H373 ** |  |  |  | CLP00
616-058-00-0 | 1,3-bis(3-methyl-2,5-dioxo-1H-pyrrolinylmethyl)benzene | 412-570-1 | 119462-56-5 | STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H318 H317 H400 H410 | GHS08 GHS05 GHS07 GHS09 Dgr | H373 ** H318 H317 H410 |  |  |  | CLP00
616-059-00-6 | 4-((4-(diethylamino)-2-ethoxyphenyl)imino)-1,4-dihydro-1-oxo-N-propyl-2-naphthalenecarboxamide | 412-650-6 | 121487-83-0 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-060-00-1 | Condensation product of: 3-(7-carboxyhept-1-yl)-6-hexyl-4-cyclohexene-1,2-dicarboxylic acid with polyamines (primarily amino-ethyl-piperazine and triethylenetetramine) | 413-770-1 |  | Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H314 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H314 H317 H410 |  |  |  | CLP00
616-061-00-7 | N,N'-1,6-hexanediylbis(N-(2,2,6,6-tetramethyl-piperidin-4-yl)-formamide | 413-610-0 | 124172-53-8 | Eye Irrit. 2 Aquatic Chronic 3 | H319 H412 | GHS07 Wng | H319 H412 |  |  |  | CLP00
616-062-00-2 | N-[3-[(2-acetyloxy)ethyl](phenyl-methyl)amino]-4-methoxyphenylacetamide | 411-590-8 | 70693-57-1 | Skin Corr. 1B Aquatic Chronic 3 | H314 H412 | GHS05 Dgr | H314 H412 |  |  |  | CLP00
616-063-00-8 | 3-dodecyl-(1-(1,2,2,6,6-pentamethyl-4-piperidin)-yl)-2,5-pyrrolidindione | 411-920-0 | 106917-30-0 | Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1A Aquatic Acute 1 Aquatic Chronic 1 | H331 H302 H373 ** H314 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H331 H302 H373 ** H314 H410 |  |  |  | CLP00
616-064-00-3 | N-tert-butyl-3-methylpicolinamide | 406-720-5 | 32998-95-1 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
616-065-00-9 | 3'-(3-acetyl-4-hydroxyphenyl)-1,1-diethylurea | 411-970-3 | 79881-89-3 | Acute Tox. 4 * STOT RE 2 * | H302 H373 ** | GHS08 GHS07 Wng | H302 H373 ** |  |  |  | CLP00
616-066-00-4 | 5,6,12,13-tetrachloroanthra(2,1,9-def:6,5,10-d'e'f')diisoquinoline-1,3,8,10(2H,9H)-tetrone | 405-100-1 | 115662-06-1 | Repr. 2 | H361f *** | GHS08 Wng | H361f *** |  |  |  | CLP00
616-067-00-X | dodecyl 3-(2-(3-benzyl-4-ethoxy-2,5-dioxoimidazolidin-1-yl)-4,4-dimethyl-3-oxovaleramido)-4-chlorobenzoate | 407-300-4 | 92683-20-0 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-068-00-5 | potassium 4-(11-methacrylamidoundecanamido)benzenesulfonate | 406-500-9 | 174393-75-0 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
616-069-00-0 | 1-hydroxy-5-(2-methylpropyloxycarbonylamino)-N-(3-dodecyloxypropyl)-2-naphthoamide | 406-210-2 | 110560-22-0 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-070-00-6 | reaction mass of: 3,3'-dicyclohexyl-1,1'-methylenebis(4,1-phenylene)diurea; 3-cyclohexyl-1-(4-(4-(3-octadecylureido)benzyl)phenyl)urea; 3,3'-dioctadecyl-1,1'-methylenebis(4,1-phenylene)diurea | 406-530-2 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-071-00-1 | reaction mass of: bis(N-cyclohexyl-N'-phenyleneureido)methylene; bis(N-octadecyl-N'-phenyleneureido)methylene; bis(N-dicyclohexyl-N'-phenyleneureido)methylene (1:2:1) | 406-550-1 |  | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
616-072-00-7 | 1-(2-deoxy-5-O-trityl-β-D-threopentofuranosyl)thymine | 407-120-6 | 55612-11-8 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-073-00-2 | 4'-ethoxy-2-benzimidazoleanilide | 407-600-5 | 120187-29-3 | Muta. 2 Aquatic Chronic 4 | H341 H413 | GHS08 Wng | H341 H413 |  |  |  | CLP00
616-074-00-8 | N-butyl-2-(4-morpholinylcarbonyl)benzamide | 407-730-2 | 104958-67-0 | Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H319 H317 H412 | GHS07 Wng | H319 H317 H412 |  |  |  | CLP00
616-075-00-3 | D,L-(N,N-diethyl-2-hydroxy-2-phenylacetamide) | 408-120-9 | 65197-96-8 | Acute Tox. 4 * Eye Dam. 1 | H302 H318 | GHS05 GHS07 Dgr | H302 H318 |  |  |  | CLP00
616-076-00-9 | tebufenozide (ISO); N-tert-butyl-N'-(4-ethylbenzoyl)-3,5-dimethylbenzohydrazide | 412-850-3 | 112410-23-8 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
616-077-00-4 | reaction mass of: 2-(9-methyl-1,3,8,10-tetraoxo-2,3,9,10-tetrahydro-(1H,8H)-anthra[2,1,9-def: 6,5,10-d'e'f']diisoquinolin-2-ylethansulfonic acid; potassium 2-(9-methyl-1,3,8,10-tetraoxo-2,3,9,10-tetrahydro-(1H,8H)-anthra[2,1,9-def: 6,5,10-d'e'f']diisoquinolin-2-ylethansulfate | 411-310-4 |  | Eye Dam. 1 | H318 | GHS05 Dgr | H318 |  |  |  | CLP00
616-078-00-X | 2-[2,4-bis(1,1-dimethyl-ethyl)phenoxy]-N-(2-hydroxy-5-methyl-phenyl)hexanamide | 411-330-3 | 104541-33-5 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-079-00-5 | 1,6-hexanediyl-bis(2-(2-(1-ethylpentyl)-3-oxazolidinyl)ethyl)carbamate | 411-700-4 | 140921-24-0 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
616-080-00-0 | 4-(2-((3-ethyl-4-methyl-2-oxo-pyrrolin-1-yl)carboxamido)ethyl)benzenesulfonamide) | 411-850-0 | 119018-29-0 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
616-081-00-6 | 5-bromo-8-naphtholactam | 413-480-5 | 24856-00-6 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  |  |  | CLP00
616-082-00-1 | N-(5-chloro-3-((4-(diethylamino)-2-methylphenyl)imino-4-methyl-6-oxo-1,4-cyclohexadien-1-yl)benzamide | 413-200-1 | 129604-78-0 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
616-083-00-7 | [2-[(4-nitrophenyl)amino]ethyl]urea | 410-700-1 | 27080-42-8 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
616-084-00-2 | 2,4-bis[N'-(4-methylphenyl)ureido]toluene | 411-790-5 |  | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
616-085-00-8 | 3-(2,4-dichlorophenyl)-6-fluoro-quinazoline-2,4(1H,3H)-dione | 412-190-6 | 168900-02-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
616-086-00-3 | 2-acetylamino-6-chloro-4-[(4-diethylamino)2-methylphenyl-imino]-5-methyl-1-oxo-2,5-cyclohexadiene | 412-250-1 | 102387-48-4 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-087-00-9 | reaction mass of: 7,9,9-trimethyl-3,14-dioxa-4,13-dioxo-5,12-diazahexadecane-1,16-diyl-prop-2-enoate; 7,7,9-trimethyl-3,14-dioxa-4,13-dioxo-5,12-diazahexadecan-1,16-diyl-prop-2-enoate | 412-260-6 | 52658-19-2 | Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H319 H317 H411 | GHS07 GHS09 Wng | H319 H317 H411 |  |  |  | CLP00
616-088-00-4 | 2-aminosulfonyl-N,N-dimethylnicotinamide | 413-440-7 | 112006-75-4 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | CLP00
616-089-00-X | 5-(2,4-dioxo-1,2,3,4-tetrahydropyrimidine)-3-fluoro-2-hydroxymethyltetrahydrofuran | 415-360-8 | 41107-56-6 | Muta. 2 | H341 | GHS08 Wng | H341 |  |  |  | CLP00
616-090-00-5 | 1-(1,4-benzodioxan-2-ylcarbonyl)piperazine hydrochloride | 415-660-9 | 70918-74-0 | Acute Tox. 3 * Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Aquatic Chronic 2 | H331 H311 H301 H373 ** H411 | GHS06 GHS08 GHS09 Dgr | H331 H311 H301 H373 ** H411 |  |  |  | CLP00
616-091-00-0 | 1,3,5-tris-[(2S and 2R)-2,3-epoxypropyl]-1,3,5-triazine-2,4,6-(1H,3H,5H)-trione | 423-400-0 | 59653-74-6 | Muta. 1B Acute Tox. 3 * Acute Tox. 4 * STOT RE 2 * Eye Dam. 1 Skin Sens. 1 | H340 H331 H302 H373 ** H318 H317 | GHS06 GHS08 GHS05 Dgr | H340 H331 H302 H373 ** H318 H317 |  |  |  | CLP00
616-092-00-6 | Polymeric reaction product of bicyclo[2.2.1]hepta-2,5-diene, ethene, 1,4-hexadiene, 1-propene with N,N-di-2-propenylformamide | 404-035-6 |  | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | CLP00
616-093-00-1 | Reaction products of: aniline-terephthalaldehyde-o-toluidine condensate with maleic anhydride | 406-620-1 | 129217-90-9 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
616-094-00-7 | 3,3'-dicyclohexyl-1,1'-methylenebis(4,1-phenylene)diurea | 406-370-3 | 58890-25-8 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00/ATP10
616-095-00-2 | 3,3'-dioctadecyl-1,1'-methylenebis(4,1-phenylene)diurea | 406-690-3 | 43136-14-7 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-096-00-8 | N-(3-hexadecyloxy-2-hydroxyprop-1-yl)-N-(2-hydroxyethyl)palmitamide | 408-110-4 | 110483-07-3 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-097-00-3 | N,N'-1,4-phenylenebis(2-((2-methoxy-4-nitrophenyl)azo)-3-oxobutanamide | 411-840-6 | 83372-55-8 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-098-00-9 | 1-[4-chloro-3-((2,2,3,3,3-pentafluoropropoxy)methyl)phenyl]-5-phenyl-1H-1,2,4-triazole-3-carboxamide | 411-750-7 | 119126-15-7 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
616-099-00-4 | 2-[4-[(4-hydroxyphenyl)sulfonyl]phenoxy]-4,4-dimethyl-N-[5-[(methylsulfonyl)amino]-2-[4-(1,1,3,3-tetramethylbutyl)phenoxy]phenyl]-3-oxopentanamide | 414-170-2 | 135937-20-1 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-100-00-8 | 1,3-dimethyl-1,3-bis(trimethylsilyl)urea | 414-180-7 | 10218-17-4 | Acute Tox. 4 * Skin Irrit. 2 | H302 H315 | GHS07 Wng | H302 H315 |  |  |  | CLP00
616-101-00-3 | (S)-N-tert-butyl-1,2,3,4-tetrahydro-3-isoquinolinecarboxamide | 414-600-9 | 149182-72-9 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
616-102-00-9 | reaction mass of: α-[3-(3-mercaptopropanoxycarbonylamino)methylphenylaminocarbonyl]-ω-[3-(3-mercaptopropanoxycarbonylamino)methylphenylaminocarbonyloxy]-poly-(oxyethylene-co-oxypropylene); 1,2-(or 1,3-)bis[α-(3-mercaptopropanoxycarbonylamino)methylphenylaminocarbonyl)-ω-oxy-poly(oxyethylene-co-oxypropylene)]-3-(or 2-)propanol; 1,2,3-tris[α-(3-mercaptopropanoxycarbonyl-amino)methylphenylaminocarbonyl)-ω-oxy-poly-(oxyethylene-co-oxypropylene)]propane] | 415-870-0 |  | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
616-103-00-4 | (S,S)-trans-4-(acetylamino)-5,6-dihydro-6-methyl-7,7-dioxo-4H-thieno[2,3-b]thiopyran-2-sulfonamide | 415-030-3 | 120298-38-6 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  |  |  | CLP00
616-104-00-X | benalaxyl (ISO); methyl N-(2,6-dimethylphenyl)-N-(phenylacetyl)-DL-alaninate | 275-728-7 | 71626-11-4 | Acute Tox. 4 Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  | oral: ATE = 1000 mg/kg bw M = 1 M = 1 |  | CLP00/ATP21
616-105-00-5 | chlorotoluron (ISO); 3-(3-chloro-p-tolyl)-1,1-dimethylurea | 239-592-2 | 15545-48-9 | Carc. 2 Repr. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H361d *** H400 H410 | GHS08 GHS09 Wng | H351 H361d *** H410 |  |  |  | CLP00
616-106-00-0 | phenmedipham (ISO); methyl 3-(3-methylcarbaniloyloxy)carbanilate | 237-199-0 | 13684-63-4 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=10 M=10 |  | CLP00/ATP17
616-107-00-6 | cinidon ethyl (ISO); ethyl (Z)-2-chloro-3-[2-chloro-5-(cyclohex-1-ene-1,2-dicarboximido)phenyl]acrylate | - | 142891-20-1 | Carc. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H351 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H317 H410 |  |  |  | ATP01
616-108-00-1 | iodosulfuron-methyl-sodium; sodium ({}{[5-iodo-2-(methoxycarbonyl)phenyl]sulfonyl}}carbamoyl)(4-methoxy-6-methyl-1,3,5-triazin-2-yl)azanide |  | 144550-36-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
616-109-00-7 | sulfosulfuron (ISO); 1-(4,6-dimethoxypyrimidin-2-yl)-3-(2-ethylsulfonylimidazo[1,2-a]pyridin-3-yl)sulfonylurea |  | 141776-32-1 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
616-110-00-2 | cyclanilide (ISO); 1-(2,4-dichloroanilinocarbonyl)cyclopropanecarboxylic acid | 419-150-7 | 113136-77-9 | Acute Tox. 4 * Aquatic Chronic 2 | H302 H411 | GHS07 GHS09 Wng | H302 H411 |  |  |  | CLP00
616-111-00-8 | fenhexamid (ISO); N-(2,3-dichlor-4-hydroxyphenyl)-1-methylcyclohexancarboxamid | 422-530-5 | 126833-17-8 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
616-112-00-3 | oxasulfuron (ISO); oxetan-3-yl 2-[(4,6-dimethylpyrimidin-2-yl)-carbamoylsulfamoyl]benzoate |  | 144651-06-9 | STOT RE 2 * Aquatic Acute 1 Aquatic Chronic 1 | H373 ** H400 H410 | GHS08 GHS09 Wng | H373 ** H410 |  |  |  | CLP00
616-113-00-9 | desmedipham (ISO); ethyl 3-phenylcarbamoyloxyphenylcarbamate | 237-198-5 | 13684-56-5 | Repr. 2 Aquatic Acute 1 Aquatic Chronic 1 | H361d H400 H410 | GHS08 GHS09 Wng | H361d H410 |  | M=10 M=10 |  | CLP00/ATP17
616-114-00-4 | dodecanamide, N,N'-(9,9',10,10'-tetrahydro-9,9',10,10'-tetraoxo(1,1'-bianthracene)-4,4'-diyl)bis- | 418-010-2 | 136897-58-0 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-115-00-X | N-(3-acetyl-2-hydroxyphenyl)-4-(4-phenylbutoxy)benzamide | 416-150-9 | 136450-06-1 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-116-00-5 | N-(4-dimethylaminopyridinium)-3-methoxy-4-(1-methyl-5-nitroindol-3-ylmethyl)-N-(o-tolylsulfonyl)benzamidate | 416-790-9 | 143052-96-4 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-117-00-0 | N-[2-(3-acetyl-5-nitrothiophen-2-ylazo)-5-diethylaminophenyl]acetamide | 416-860-9 | 777891-21-1 | Repr. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H361f *** H317 H400 H410 | GHS08 GHS09 Wng | H361f *** H317 H410 |  |  |  | CLP00
616-118-00-6 | N-(2',6'-dimethylphenyl)-2-piperidinecarboxamide hydrochloride | 417-950-0 | 65797-42-4 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
616-119-00-1 | 2-(1-butyl-3,5-dioxo-2-phenyl-(1,2,4)-triazolidin-4-yl)-4,4-dimethyl-3-oxo-N-(2-methoxy-5-(2-(dodecyl-1-sulfonyl))propionylamino)-phenyl)-pentanamide | 418-060-5 | 118020-93-2 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-120-00-7 | reaction mass of: N-(3-dimethylamino-4-methyl-phenyl)-benzamide; N-(3-dimethylamino-2-methyl-phenyl)-benzamide; N-(3-dimethylamino-3-methyl-phenyl)-benzamide | 420-600-1 |  | STOT RE 2 * Aquatic Chronic 2 | H373 ** H411 | GHS08 GHS09 Wng | H373 ** H411 |  |  |  | CLP00
616-121-00-2 | 2,4-dihydroxy-N-(2-methoxyphenyl)benzamide | 419-090-1 | 129205-19-2 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | CLP00
616-122-00-8 | methyl neodecanamide | 414-460-9 | 105726-67-8 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | ATP01
616-123-00-3 | N-[3-[[4-(diethylamino)-2-methylphenyl]imino]-6-oxo-1,4-cyclohexadienyl]acetamide | 414-740-0 | 96141-86-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
616-124-00-9 | lithium bis(trifluoromethylsulfonyl)imide | 415-300-0 | 90076-65-6 | Acute Tox. 3 * Acute Tox. 3 * STOT RE 2 * Skin Corr. 1B Aquatic Chronic 3 | H311 H301 H373 ** H314 H412 | GHS06 GHS05 GHS08 Dgr | H311 H301 H373 ** H314 H412 |  |  |  | CLP00/ATP01
616-125-00-4 | 3-cyano-N-(1,1-dimethylethyl)androsta-3,5-diene-17-β-carboxamide | 415-730-9 | 151338-11-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
616-126-00-X | 1-methyl-4-nitro-3-propyl-1H-pyrazole-5-carboxamide | 423-960-6 | 139756-01-7 | Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 3 | H302 H373 ** H412 | GHS08 GHS07 Wng | H302 H373 ** H412 |  |  |  | ATP01
616-127-00-5 | reaction mass of N,N'-ethane-1,2-diylbis(decanamide) and 12-hydroxy-N-[2-[(1-oxodecyl)amino]ethyl]octadecanamide and N,N'-ethane-1,2-diylbis(12-hydroxyoctadecanamide) [1] reaction mass of N,N'-ethane-1,2-diylbis(decanamide) and 12-hydroxy-N-[2-[(1-oxodecyl)amino]ethyl]octadecanamide [2] | 430-050-2 [1] - [2] | - [1] - [2] | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  | M = 100 M = 10 |  | CLP00/ATP22
616-128-00-0 | N-(2-(1-allyl-4,5-dicyanoimidazol-2-ylazo)-5-(dipropylamino)phenyl)-acetamide | 417-530-7 | 123590-00-1 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-129-00-6 | N,N'-bis(2,2,6,6-tetramethyl-4-piperidyl)isophthalamide | 419-710-0 | 42774-15-2 | Acute Tox. 4 * Eye Irrit. 2 | H302 H319 | GHS07 Wng | H302 H319 |  |  |  | CLP00
616-130-00-1 | N-(3-(2-(4,4-dimethyl-2,5-dioxo-imidazolin-1-yl)-4,4-dimethyl-3-oxo-pentanoylamino)-4-methoxy-phenyl)-octadecanamide | 421-780-2 | 150919-56-5 | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | CLP00
616-131-00-7 | 1-aminocyclopentanecarboxamide | 422-950-9 | 17193-28-1 | Acute Tox. 4 * STOT RE 1 Eye Dam. 1 | H302 H372 ** H318 | GHS05 GHS08 GHS07 Dgr | H372 ** H302 H318 |  |  |  | ATP01
616-132-00-2 | N-[4-(4-cyano-2-furfurylidene-2,5-dihydro-5-oxo-3-furyl)phenyl]butane-1-sulfonamide | 423-250-6 | 130016-98-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
616-133-00-8 | N-cyclohexyl-S,S-dioxobenzo[b]tiophene-2-carboxamide | 423-990-1 | 149118-66-1 | Acute Tox. 4 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H318 H410 |  |  |  | CLP00
616-134-00-3 | 3,3'-bis(dioctyloxyphosphinothioylthio)-N,N'-oxybis(methylene)dipropionamide | 401-820-5 | 793710-14-2 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | CLP00
616-135-00-9 | (3S,4aS,8aS)-2-[(2R,3S)-3-amino-2-hydroxy-4-phenylbutyl]-N-tert-butyldecahydroisoquinoline-3-carboxamide | 430-230-0 | 136522-17-3 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | CLP00
616-136-00-4 | reaction product of cocoalkyldiethanolamides and cocoalkylmonoglycerides and molybdenumtrioxide (1.75-2.2: 0.75-1.0:0.1-1.1) | 430-380-7 | - | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP01
616-137-00-X | 4-dichloroacetyl-1-oxa-4-azaspiro[4.5]decane | 401-130-4 | 71526-07-3 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
616-138-00-5 | benzoic acid, N-tert-butyl-N'-(4-chlorobenzoyl)hydrazide | 431-600-4 | 112226-61-6 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
616-139-00-0 | (3S,4aS,8aS)-N-tert-butyldecahydro-3-isoquinolinecarboxamide | 420-380-5 | 136465-81-1 | Acute Tox. 4 * Eye Dam. 1 Aquatic Chronic 3 | H302 H318 H412 | GHS05 GHS07 Dgr | H302 H318 H412 |  |  |  | ATP01
616-140-00-6 | N,N''-(methylenedi-4,1-phenylene)bis[N'-(4-methylphenyl)urea] | 429-380-1 | 133336-92-2 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
616-141-00-1 | zoxamide (ISO); (RS)-3,5-dichloro-N-(3-chloro-1-ethyl-1-methyl-2-oxopropyl)-p-toluamide | - | 156052-68-5 | Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H317 H400 H410 | GHS07 GHS09 Wng | H317 H410 |  | M=10 |  | ATP01
616-142-00-7 | 1,3-Bis(vinylsulfonylacetamido)propane | 428-350-3 | 93629-90-4 | Muta. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H341 H318 H317 H412 | GHS08 GHS05 GHS07 Dgr | H341 H318 H317 H412 |  |  |  | CLP00
616-143-00-2 | N,N'-dihexadecyl-N,N'-bis(2-hydroxyethyl)propanediamide | 422-560-9 | 149591-38-8 | Repr. 2 Eye Irrit. 2 Aquatic Chronic 4 | H361f *** H319 H413 | GHS08 Wng | H361f *** H319 H413 |  |  |  | CLP00
616-144-00-8 | 3,4-dichloro-N-[5-chloro-4-[2-[4-dodecyloxyphenylsulfonyl]butyramido]-2-hydroxyphenyl]benzamide | 431-130-1 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-145-00-3 | pethoxamide (ISO); 2-chloro-N-(2-ethoxyethyl)-N-(2-methyl-1-phenylprop-1-enyl)acetamide | - | 106700-29-2 | Acute Tox. 4 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  | M=100 |  | ATP01
616-146-00-9 | N-(2-methoxy-5-octadecanoylaminophenyl)-2-(3-benzyl-2,5-dioxoimidazolidin-1-yl)-4,4-dimethyl-3-oxopentanoic acidamide | 431-330-7 | 142776-95-2 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-147-00-4 | 1-methyl-4-(2-methyl-2H-tetrazol-5-yl)-1H-pyrazole-5-sulfonamide | 424-160-1 | 139481-22-4 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | ATP01
616-148-00-X | N-[6,9-dihydro-9-[[2-hydroxy-1-(hydroxymethyl)ethoxy]methyl]-6-oxo-1H-purin-2-yl]acetamide | 424-550-1 | 84245-12-5 | Carc. 1B Muta. 1B Repr. 1B | H350 H340 H360FD | GHS08 Dgr | H350 H340 H360FD |  |  |  | ATP01
616-150-00-0 | (2R,3S)-N-(3-amino-2-hydroxy-4-phenylbutyl)-N-isobutyl-4-nitrobenzenesulfonamide hydrochloride | 425-260-6 | - | STOT RE 2 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H373 ** H318 H317 H411 | GHS05 GHS08 GHS07 GHS09 Dgr | H373 ** H318 H317 H411 |  |  |  | ATP01
616-151-00-6 | N-(2-amino-4,6-dichloropyrimidin-5-yl)formamide | 425-650-6 | 171887-03-9 | Acute Tox. 4 * Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 3 | H302 H318 H317 H412 | GHS05 GHS07 Dgr | H302 H318 H317 H412 |  |  |  | ATP01
616-152-00-1 | 4-(4-fluorophenyl)-2-(2-methyl-1-oxopropyl)-4-oxo-3,N-diphenylbutanamide | 425-850-3 | 125971-96-2 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-153-00-7 | 4-methyl-3-oxo-N-phenyl-2-(phenylmethylene)pentanamide | 425-860-8 | 125971-57-5 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
616-154-00-2 | 3,4-dichloro-N-[5-chloro-4-[2-[4-(hexadecyloxy)phenylsulfonyl]butyramido]-2-hydroxyphenyl]benzamide | 431-110-0 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-155-00-8 | N,N,N',N'-tetracyclohexyl-1,3-benzenedicarboxamide | 431-040-0 | 104560-40-9 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
616-156-00-3 | 6-(2-chloro-6-cyano-4-nitrophenylazo)-4-methoxy-3-[N-(methoxycarbonylmethyl)-N-(1-methoxycarbonylethyl)amino]acetanilide | 430-500-8 | 204277-61-2 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-157-00-9 | 3-amino-4-hydroxy-N-(3-isopropoxypropyl)benzenesulfonamide hydrochloride | 427-780-9 | 114565-70-7 | Acute Tox. 4 * Eye Dam. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H318 H400 H410 | GHS05 GHS07 GHS09 Dgr | H302 H318 H410 |  |  |  | ATP01/ATP01corr
616-158-00-4 | N-[4-cyano-3-trifluoromethylphenyl]methacrylamide | 427-880-2 | 90357-53-2 | STOT RE 2 * Aquatic Chronic 2 | H373 ** H411 | GHS08 GHS09 Wng | H373 ** H411 |  |  |  | ATP01
616-160-00-5 | 2,2'-azobis[N-(2-hydroxyethyl)-2-methylpropionamide] | 429-090-3 | 61551-69-7 | Skin Sens. 1 Aquatic Chronic 3 | H317 H412 | GHS07 Wng | H317 H412 |  |  |  | ATP01
616-161-00-0 | 2,4-dichloro-5-hydroxyacetanilide | 429-110-0 | 67669-19-6 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
616-162-00-6 | isostearic acid monoisopropanolamide | 431-540-9 | - | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | ATP01
616-163-00-1 | 4,4'-methylenebis[N-(4-chlorophenyl)-3-hydroxynaphthalene-2-carboxamide] | 430-350-3 | 192463-88-0 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-164-00-7 | dimoxystrobin (ISO); (2E)-2-{2-[(2,5-dimethylphenoxy)methyl]phenyl}-2-(methoxyimino)-N-methylacetamide; (E)-2-(methoxyimino)-N-methyl-2-[α-(2,5-xylyloxy)-o-tolyl]acetamide | - | 149961-52-4 | Carc. 2 Repr. 2 Acute Tox. 4 Aquatic Acute 1 Aquatic Chronic 1 | H351 H361d H332 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H361d H332 H410 |  | inhalation: ATE = 1,3 mg/L (dusts or mists) M = 100 M = 100 |  | ATP01/ATP18
616-165-00-2 | beflubutamid (ISO); (RS)-N-benzyl-2-(α,α,α,4-tetrafluoro-m-tolyoxy)butyramide | - | 113614-08-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=100 |  | ATP01
616-166-00-8 | cyazofamid (ISO); 4-chloro-2-cyano-N,N-dimethyl-5-p-tolylimidazole-1-sulfonamide | - | 120116-88-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=10 |  | ATP01
616-167-00-3 | N,N-dibutyl-(2,5-dihydro-5-thioxo-1H-tetrazol-1-yl)acetamide | 418-290-6 | 168612-06-4 | Eye Irrit. 2 Skin Sens. 1 | H319 H317 | GHS07 Wng | H319 H317 |  |  |  | ATP01
616-168-00-9 | 1-dimethylcarbamoyl-4-(2-sulfonatoethyl)pyridinium | 418-440-0 | 136997-71-2 | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | ATP01
616-169-00-4 | 4-[4-(2,2-dimethyl-propanamido)]phenylazo-3-(2-chloro-5-(2-(3-pentadecylphenoxy)butylamido)anilino)-1-(2,4,6-trichlorophenyl)-2-pyrazoline-5-one | 420-220-4 | 92771-56-7 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
616-170-00-X | (2R)-2-amino-2-phenylacetamide | 420-370-0 | 6485-67-2 | Eye Irrit. 2 Skin Sens. 1 | H319 H317 | GHS07 Wng | H319 H317 |  |  |  | ATP01
616-171-00-5 | 2-(para-chlorophenyl)glycineamide | 420-830-0 | 102333-75-5 | Eye Dam. 1 Skin Sens. 1 | H318 H317 | GHS05 GHS07 Dgr | H318 H317 |  |  |  | ATP01
616-172-00-0 | N-(2,2,6,6-tetramethyl-1-oxylpiperidin-4-yl)acetamide; (4-acetamido-2,2,6,6-tetramethyl-1-piperidinyl)oxidanyl | 423-840-3 | 14691-89-5 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | ATP01
616-174-00-1 | 2-butyl-1,3-diazaspiro[4.4]non-1-en-4-one hydrochloride | 424-560-4 | 151257-01-1 | Acute Tox. 4 * Eye Irrit. 2 | H302 H319 | GHS07 Wng | H302 H319 |  |  |  | ATP01
616-175-00-7 | 2-(2-hexyldecyloxy)benzamide | 431-230-3 | 202483-62-3 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-176-00-2 | 3-N,N-bis(methoxyethyl)aminoacetanilide | 432-530-7 | 24294-01-7 | Acute Tox. 4 * Aquatic Chronic 3 | H302 H412 | GHS07 Wng | H302 H412 |  |  |  | ATP01
616-177-00-8 | (3-(4-(2-(butyl-(4-methylphenylsulfonyl)amino)phenylthio)-5-oxo-1-(2,4,6-trichlorophenyl)-4,5-dihydro-1H-pyrazole-3-ylamino)-4-chlorophenyl)tetradecanamide; N-[3-({4-[(2-{butyl[(4-methylphenyl)sulfonyl]amino}phenyl)thio]-5-oxo-1-(2,4,6-trichlorophenyl)-4,5-dihydro-1H-pyrazol-3-yl}amino)-4-chlorophenyl]tetradecanamide | 432-970-1 | 217324-98-6 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-178-00-3 | N-(5-(bis(2-methoxyethyl)amino)-2-((2-cyano-4,6-dinitrophenyl)-azo)phenyl)acetamide | 434-500-9 | 52583-35-4 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-179-00-9 | 2-chloro-N-(4-methylphenyl)acetamide | 435-170-9 | 16634-82-5 | Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H318 H317 H410 |  |  |  | ATP01
616-180-00-4 | N,N-(dimethylamino)thioacetamide hydrochloride | 435-470-1 | 27366-72-9 | Repr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H360D *** H400 H410 | GHS08 GHS09 Dgr | H360D *** H410 |  |  |  | ATP01
616-181-00-X | 4'-methyldodecane-1-sulfonanilide | 435-490-9 | 17417-32-2 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
616-182-00-5 | N'-(1,3-dimethylbutylidene)-3-hydroxy-2-naphthohydrazide | 435-860-1 | 214417-91-1 | Skin Sens. 1 Aquatic Chronic 2 | H317 H411 | GHS07 GHS09 Wng | H317 H411 |  |  |  | ATP01
616-183-00-0 | N-dodecyl-4-methoxybenzamide | 442-340-6 | 1854-15-5 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-184-00-6 | 3-methyl-N-(5,8,13,14-tetrahydro-5,8,14-trioxonaphth[2,3-c]acridin-6-yl)benzamide | 442-560-2 | 105043-55-8 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-186-00-7 | N,N'-(2-chloro-1,4-phenylene)bis(3-oxobutaneamide) | 443-010-4 | 53641-10-4 | Aquatic Chronic 3 | H412 | - | H412 |  |  |  | ATP01
616-188-00-8 | 2-(5,5-dimethyl-2,4-dioxooxazolidin-3-yl)-4,4-dimethyl-3-oxo-N-(2-methoxy-5-octadecanoylaminophenyl)pentanoic acid amide | 443-980-9 | 221215-20-9 | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
616-189-00-3 | N-[5-(bis-(2-methoxy-ethyl)-amino]-2-(6-bromo-2-methyl-1,3-dioxo-2,3-dihydro-1H-isoindol-5-ylazo)-phenyl]acetamide | 444-780-4 | 452962-97-9 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-190-00-9 | N-decyl-4-nitrobenzamide | 445-880-0 | 64026-19-3 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-191-00-4 | 2-ethyl-N-methyl-N-(3-methylphenyl)butanamide | 446-190-2 | 406488-30-0 | Acute Tox. 4 * Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H302 H315 H319 H317 H411 | GHS07 GHS09 Wng | H302 H319 H315 H317 H411 |  |  |  | ATP01
616-192-00-X | 2-[2-(3-butoxypropyl)-1,1-dioxo-1,2,4-benzothiadiazin-3-yl]-5'-tert-butyl-2-(5,5-dimethyl-2,4-dioxo-1,3-oxazolidin-3-yl)-2'-[(2-ethylhexyl)thio]acetanilide | 448-060-0 | 727678-39-9 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-193-00-5 | N-[2-(2-butyl-4,6-dicyano-1,3-dioxo-2,3-dihydro-1H-isoindol-5-ylazo)-5-diethylamino-phenyl]acetamide | 449-940-7 | 368450-39-9 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-194-00-0 | 2,2-diethoxy-N,N-dimethylacetamide | 449-950-1 | 34640-92-1 | Eye Irrit. 2 | H319 | GHS07 Wng | H319 |  |  |  | ATP01
616-196-00-1 | disodium salt of 1-hydroxy-4-(β-(4-(1-hydroxy-3,6-disulfo-8-acetylamino-2-naphthylazo)phenoxy)ethoxy)-N-dodecyl-2-naphthamide | 419-990-4 | - | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | ATP01
616-197-00-7 | reaction mass of: potassium N-[3-(dimethyloxidoamino)propyl]-1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8,8-heptadecafluorooctane sulfonamidate; N-[3-(dimethyloxidoamino)propyl]-1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8,8-heptadecafluorooctane sulfonamide | 422-500-1 | - | STOT RE 2 * | H373 ** | GHS08 Wng | H373 ** |  |  |  | ATP01
616-198-00-2 | 1,3-bis[12-hydroxy-octadecamide-N-methylene]-benzene | 423-300-7 | - | Skin Sens. 1 Aquatic Chronic 4 | H317 H413 | GHS07 Wng | H317 H413 |  |  |  | ATP01
616-200-00-1 | reaction mass of N,N'-ethane-1,2-diylbis(hexanamide) and 12-hydroxy-N-[2-[(1-oxyhexyl)amino]ethyl]octadecanamide and N,N'-ethane-1,2-diylbis(12-hydroxyoctadecan amide) | 432-430-3 |  | Aquatic Chronic 4 | H413 |  | H413 |  |  |  | ATP01/ATP05
616-201-00-7 | 12-hydroxyoctadecanoic acid, reaction products with 1,3-benzenedimethanamine and hexamethylenediamine | 432-840-2 | 220926-97-6 | Acute Tox. 4 * Aquatic Chronic 4 | H332 H413 | GHS07 Wng | H332 H413 |  |  |  | ATP01
616-202-00-2 | reaction mass of: 2,2'-[(3,3'-dichloro[1,1'-biphenyl]-4,4'-diyl)bis(azo)]bis[N-(2,4-dimethylphenyl)]-3-oxo-butanamide; 2-[[3,3'-dichloro-4'-[[1[[(2,4-dimethylphenyl)amino]carbonyl]-2-oxopropyl]azo][1,1'-biphenyl]-4-yl]azo]-N-(2-methylphenyl)-3-oxo-butanamide; 2-[[3,3'-dichloro-4'-[[1[[(2,4-dimethylphenyl)amino]carbonyl]-2-oxopropyl]azo][1,1'-biphenyl]-4-yl]azo]-N-(2-carboxylphenyl)-3-oxo-butanamide | 434-330-5 | - | Carc. 2 Skin Sens. 1 Aquatic Chronic 4 | H351 H317 H413 | GHS08 GHS07 Wng | H351 H317 H413 |  |  |  | ATP01
616-203-00-8 | reaction mass of: N-[5-[bis-(2-methoxyethyl)amino]-2-(2-butyl-4,6-dicyano-1,3-dioxo-2,3-dihydro-1H-isoindol-5-yl-azo)phenyl]acetamide; N-[2-(2-butyl-4,6-dicyano-1,3-dioxo-2,3-dihydro-1H-isoindol-5-ylazo)5-diethylaminophenyl]acetamide | 442-280-0 | - | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-204-00-3 | N,N''-(methylenedi-4,1-phenylene)bis[N'-octylurea] | 451-060-3 | 122886-55-9 | Aquatic Chronic 4 | H413 | - | H413 |  |  |  | ATP01
616-205-00-9 | Metazachlor (ISO); 2-chloro-N-(2,6-dimethylphenyl)-N-(1H-pyrazol-1-ylmethyl)acetamide | 266-583-0 | 67129-08-2 | Carc. 2 Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 1 | H351 H317 H400 H410 | GHS07 GHS08 GHS09 Wng | H317 H351 H410 |  | M=100 M=100 |  | ATP03
616-206-00-4 | flufenoxuron (ISO); 1-(4-(2-cloro-α,α,α-p-trifluorotolyloxy)-2-fluorophenyl)-3-(2,6-difluorobenzolyl)urea | 417-680-3 | 101463-69-8 | Lact. Aquatic Acute 1 Aquatic Chronic 1 | H362 H400 H410 | GHS09 Wng | H362 H410 |  | M=10000 M=10000 |  | ATP05
616-207-00-X | polyhexamethylene biguanide hydrochloride; PHMB [1] PHMB [2] |  | 32289-58-0 [1] 27083-27-8 [2] | Carc. 2 Acute Tox. 2 Acute Tox. 4 STOT RE 1 Eye Dam. 1 Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 1 | H351 H330 H302 H372 (respiratory tract) (inhalation) H318 H317 H400 H410 | GHS08 GHS06 GHS05 GHS09 Dgr | H330 H302 H318 H317 H351 H372 (respiratory tract) (inhalation) H410 |  | M=10 M=10 |  | ATP05/ATP09
616-208-00-5 | N-ethyl-2-pyrrolidone; 1-ethylpyrrolidin-2-one | 220-250-6 | 2687-91-4 | Repr. 1B | H360D | GHS08 Dgr | H360D |  |  |  | ATP05
616-209-00-0 | amidosulfuron (ISO); 3-(4,6-dimethoxypyrimidin-2-yl)-1-((N-methyl-N-methylsulfonylamino)sulfonyl)urea | 407-380-0 | 120923-37-7 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=100 M=100 |  | ATP05
616-210-00-6 | tebufenpyrad (ISO); N-(4-tertbutylbenzyl)-4-chloro-3-ethyl-1-methyl-1Hpyrazole-5- carboxamide |  | 119168-77-3 | Acute Tox. 3 Acute Tox. 4 STOT RE 2 Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 1 | H301 H332 H373 (gastro-intestinal tract) (oral) H317 H400 H410 | GHS06 GHS08 GHS09 Dgr | H301 H332 H317 H373 (gastro-intestinal tract) (oral) H410 |  | M=10 M=10 |  | ATP05
616-211-00-1 | proquinazid (ISO); 6-iodo-2-propoxy-3-propylquinazolin-4(3H)-one |  | 189278-12-4 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  | M=1 M=10 |  | ATP05
616-212-00-7 | 3-iodo-2-propynyl butylcarbamate; 3-iodoprop-2-yn-1-yl butylcarbamate | 259-627-5 | 55406-53-6 | Acute Tox. 3 Acute Tox. 4 STOT RE 1 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H331 H302 H372 (larynx) H318 H317 H400 H410 | GHS06 GHS08 GHS05 GHS09 Dgr | H302 H331 H318 H317 H372 (larynx) H410 |  | M=10 M=1 |  | ATP06
616-213-00-2 | mandipropamid (ISO); 2-(4-chlorophenyl)-N-{2-[3-methoxy-4-(prop-2-yn-1-yloxy)phenyl]ethyl}-2-(prop-2-yn-1-yloxy)acetamide |  | 374726-62-2 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1 M=1 |  | ATP07
616-214-00-8 | metosulam (ISO); N-(2,6-dichloro-3-methylphenyl)-5,7-dimethoxy[1,2,4]triazolo[1,5-a]pyrimidine-2-sulfonamide |  | 139528-85-1 | Carc. 2 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H373 (eyes, kidneys) H400 H410 | GHS08 GHS09 Wng | H351 H373 (eyes, kidneys) H410 |  | M=1000 M=100 |  | ATP07
616-215-00-3 | dimethenamid-P (ISO); 2-chloro-N-(2,4-dimethyl-3-thienyl)-N-[(2S)-1-methoxypropan-2-yl]acetamide |  | 163515-14-8 | Acute Tox. 4 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H317 H400 H410 | GHS07 GHS09 Wng | H302 H317 H410 |  | M=10 M=10 |  | ATP07
616-216-00-9 | flonicamid (ISO); N-(cyanomethyl)-4-(trifluoromethyl)pyridine-3-carboxamide |  | 158062-67-0 | Acute Tox. 4 | H302 | GHS07 Wng | H302 |  |  |  | ATP07
616-217-00-4 | sulfoxaflor (ISO); [methyl(oxo){1-[6-(trifluoromethyl)-3-pyridyl]ethyl}-λ6-sulfanylidene]cyanamide |  | 946578-00-3 | Acute Tox. 4 Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  | M=1 M=1 |  | ATP07
616-218-00-X | benzovindiflupyr (ISO); N-[9-(dichloromethylene)-1,2,3,4-tetrahydro-1,4-methanonaphthalen-5-yl]-3-(difluoromethyl)-1-methyl-1H-pyrazole-4-carboxamide |  | 1072957-71-1 | Acute Tox. 3 Acute Tox. 3 Aquatic Acute 1 Aquatic Chronic 1 | H331 H301 H400 H410 | GHS06 GHS09 Dgr | H331 H301 H410 |  | M=100 M=100 |  | ATP09
616-219-00-5 | fluopyram (ISO); N-{2-[3-chloro-5-(trifluoromethyl)pyridin-2-yl]ethyl}-2-(trifluoromethyl)benzamide |  | 658066-35-4 | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | ATP09
616-220-00-0 | pencycuron (ISO); 1-[(4-chlorophenyl)methyl]-1-cyclopentyl-3-phenylurea | 266-096-3 | 66063-05-6 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1 M=1 |  | ATP09
616-221-00-6 | hexaflumuron (ISO); 1-(3,5-dichloro-4-(1,1,2,2-tetrafluoroethoxy)phenyl)-3-(2,6-difluorobenzoyl)urea | 401-400-1 | 86479-06-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1000 M=10000 |  | ATP10
616-222-00-1 | penthiopyrad (ISO); (RS)-N-[2-(1,3-dimethylbutyl)-3-thienyl]-1-methyl-3-(trifluoromethyl)pyrazole-4-carboxamide |  | 183675-82-3 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1 M=1 |  | ATP10
616-223-00-7 | carbetamide (ISO); (R)-1-(ethylcarbamoyl)ethyl carbanilate; (2R)-1-(ethylamino)-1-oxopropan-2-yl phenylcarbamate | 240-286-6 | 16118-49-3 | Carc. 2 Repr. 1B Acute Tox. 4 Aquatic Chronic 2 | H351 H360D H302 H411 | GHS08 GHS07 GHS09 Dgr | H302 H351 H360D H411 |  |  |  | ATP10
616-224-00-2 | amisulbrom (ISO); 3-(3-bromo-6-fluoro-2-methylindol-1-ylsulfonyl)-N,N-dimethyl-1H-1,2,4-triazole-1-sulfonamide |  | 348635-87-0 | Carc. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H319 H400 H410 | GHS08 GHS07 GHS09 Wng | H319 H351 H410 |  | M=10 M=10 |  | ATP13
616-225-00-8 | (RS)-2-methoxy-N-methyl-2-[α-(2,5-xylyloxy)-o-tolyl]acetamide; mandestrobin |  | 173662-97-0 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  | M=1 M=10 |  | ATP14
616-226-00-3 | carboxin (ISO); 2-methyl-N-phenyl-5,6-dihydro-1,4-oxathiine-3-carboxamide; 5,6-dihydro-2-methyl-1,4-oxathiine-3-carboxanilide | 226-031-1 | 5234-68-4 | STOT RE 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H373 (Kidneys) H317 H400 H410 | GHS07 GHS08 GHS09 Wng | H317 H373 (kidneys) H410 |  | M=1 M=1 |  | ATP14
616-227-00-9 | metaflumizone (ISO); (EZ)-2'-[2-(4-cyanophenyl)-1-(α,α,α -trifluoro-m-tolyl)ethylidene]-[4-(trifluoromethoxy)phenyl]carbanilohydrazide [E-isomer > 90%, Z-isomer <10% relative content] [1] (E)-2'-[2-(4-cyanophenyl)-1-(α,α,α -trifluoro-m-tolyl)ethylidene]-[4-(trifluoromethoxy)phenyl]carbanilohydrazide [2] |  | 139968-49-3 [1] 852403-68-0 [2] | Repr. 2 Lact. STOT RE 2 | H361fd H362 H373 | GHS08 Wng | H361fd H362 H373 |  |  |  | ATP14
616-228-00-4 | 3-(difluoromethyl)-1-methyl-N-(3',4',5'-tri­fluorobiphenyl-2-yl) pyrazole-4-carboxamide; fluxapyroxad |  | 907204-31-3 | Lact. Aquatic Acute 1 Aquatic Chronic 1 | H362 H400 H410 | GHS09 Wng | H362 H410 |  | M=1 M=1 |  | ATP15
616-230-00-5 | N-(hydroxymethyl)acrylamide; methylolacrylamide; [NMA] | 213-103-2 | 924-42-5 | Carc. 1B Muta. 1B STOT RE 1 | H350 H340 H372 (peripheral nervous system) | GHS08 Dgr | H340 H350 H372 (peripheral nervous system) |  |  |  | ATP15
616-231-00-0 | 5-fluoro-1,3-di­methyl-N-[2-(4-methylpentan-2-yl) phenyl]-1H-pyrazole-4-carboxamide; 2'-[(RS)-1,3-dimethylbutyl]-5-fluoro-1,3-di­methylpyrazole-4-carboxanilide; penflufen |  | 494793-67-8 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  | M=1 M=1 |  | ATP15
616-232-00-6 | iprovalicarb (ISO); isopropyl [(2S)-3-methyl-1-{[1-(4- methylphenyl)ethyl]amino}-1-oxobutan-2-yl]carbamate |  | 140923-17-7 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | ATP15
616-233-00-1 | silthiofam (ISO); N-allyl-4,5-dimethyl- 2-(trimethylsilyl)thiophene-3-carboxamide |  | 175217-20-6 | STOT RE 2 Aquatic Chronic 2 | H373 H411 | GHS08 GHS09 Wng | H373 H411 |  |  |  | ATP15
616-234-00-7 | N-methoxy-N-[1-methyl-2-(2,4,6-trichlorophenyl)-ethyl]-3-(difluoromethyl)-1-methylpyrazole-4-carboxamide; pydiflumetofen |  | 1228284-64-7 | Carc. 2 Repr. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H361f H400 H410 | GHS08 GHS09 Wng | H351 H361f H410 |  | M=1 M=1 |  | ATP17
616-235-00-2 | N-{2-[[1,1'-bi(cyclopropyl)]-2-yl]phenyl}-3-(difluoromethyl)-1-methyl-1H-pyrazole-4-carboxamide; sedaxane |  | 874967-67-6 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 2 | H351 H400 H411 | GHS08 GHS09 Wng | H351 H410 |  | M=1 |  | ATP17
616-237-00-3 | fluopicolide (ISO); 2,6-dichloro-N-[3-chloro-5-(trifluoromethyl)-2-pyridylmethyl]benzamide |  | 239110-15-7 | Repr. 2 | H361d | GHS08 Wng | H361d |  |  |  | ATP18
616-238-00-9 | N-(2-nitrophenyl)phosphoric triamide | 477-690-9 | 874819-71-3 | Repr. 1B STOT RE 2 | H360Fd H373 (kidney) | GHS08 Dgr | H360Fd H373 (kidney) |  |  |  | ATP18
616-239-00-4 | N-(5-chloro-2-isopropylbenzyl)-N-cyclopropyl-3-(difluoromethyl)-5-fluoro-1-methyl-1H-pyrazole-4-carboxamide; isoflucypram |  | 1255734-28-1 | Repr. 2 Acute Tox. 4 Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 1 | H361f H332 H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H361f H332 H317 H410 |  | inhalation: ATE = 2,2 mg/L (dusts or mists) M=10 M=1 |  | ATP18
616-240-00-X | Reaction mass of 3-(difluoromethyl)-1-methyl-N-[(1RS,4SR,9RS)-1,2,3,4-tetrahydro-9-isopropyl-1,4-methanonaphthalen-5-yl]pyrazole-4-carboxamide and 3-(difluoromethyl)-1-methyl-N-[(1RS,4SR,9SR)-1,2,3,4-tetrahydro-9-isopropyl-1,4-methanonaphthalen-5-yl]pyrazole-4-carboxamide [≥ 78% syn isomers ≤15% anti isomers relative content]; isopyrazam |  | 881685-58-1 | Carc. 2 Repr. 1B Skin Sens. 1B Aquatic Acute 1 Aquatic Chronic 1 | H351 H360D H317 H400 H410 | GHS08 GHS07 GHS09 Dgr | H351 H360D H317 H410 |  | Repr. 1B; H360D: C ≥ 3% M = 10 M = 10 |  | ATP18
616-241-00-5 | foramsulfuron (ISO); 2-{[(4,6-dimethoxypyrimidin-2-yl)carbamoyl]sulfamoyl}-4-formamido-N,N-dimethylbenzamide; 1-(4,6-dimethoxypyrimidin-2-yl)-3-(2-dimethylcarbamoyl-5-formamidophenylsulfonyl)urea |  | 173159-57-4 | Carc. 2 Aquatic Acute 1 Aquatic Chronic 1 | H351 H400 H410 | GHS08 GHS09 Wng | H351 H410 |  | M = 1000 M = 100 |  | ATP21
616-242-00-0 | picolinafen (ISO); N-(4-fluorophenyl)-6-[3-(trifluoromethyl)phenoxy]pyridine-2-carboxamide; 4′-fluoro-6-[(α,α,α-trifluoro-m-tolyl)oxy]picolinanilide |  | 137641-05-5 | STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H373 (blood system, thyroid) H400 H410 | GHS08 GHS09 Wng | H373 (blood system, thyroid) H410 |  | M = 1000 M = 1000 |  | ATP21
616-243-00-6 | N,N'-methylenediacrylamide | 203-750-9 | 110-26-9 | Muta. 1B | H340 | GHS08 Dgr | H340 |  |  |  | ATP22
617-001-00-2 | di-tert-butyl peroxide | 203-733-6 | 110-05-4 | Flam. Liq. 2 Org. Perox. E Muta. 2 | H225 H242 H341 | GHS02 GHS08 Dgr | H242 H225 H341 |  |  |  | CLP00/ATP03
617-002-00-8 | α,α-dimethylbenzyl hydroperoxide; cumene hydroperoxide | 201-254-7 | 80-15-9 | Org. Perox. E Acute Tox. 3 * Acute Tox. 4 * Acute Tox. 4 * STOT RE 2 * Skin Corr. 1B Aquatic Chronic 2 | H242 H331 H312 H302 H373 ** H314 H411 | GHS02 GHS06 GHS08 GHS05 GHS09 Dgr | H242 H331 H312 H302 H373 ** H314 H411 |  | Skin Corr. 1B; H314: C ≥ 10 % Skin Irrit. 2; H315: 3 % ≤ C < 10 % Eye Dam. 1; H318: 3 % ≤ C < 10 % Eye Irrit. 2; H319: 1 % ≤ C < 3 % STOT SE 3; H335: C < 10 % |  | CLP00
617-003-00-3 | dilauroyl peroxide | 203-326-3 | 105-74-8 | Org. Perox. D | H242 | GHS02 Dgr | H242 |  |  |  | CLP00
617-004-00-9 | 1,2,3,4-tetrahydro-1-naphthyl hydroperoxide | 212-230-0 | 771-29-9 | Org. Perox. D Acute Tox. 4 * Skin Corr. 1B Aquatic Acute 1 Aquatic Chronic 1 | H242 H302 H314 H400 H410 | GHS02 GHS05 GHS07 GHS09 Dgr | H242 H302 H314 H410 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
617-006-00-X | bis(α,α-dimethylbenzyl) peroxide | 201-279-3 | 80-43-3 | Org. Perox. F Repr. 1B Skin Irrit. 2 Eye Irrit. 2 Aquatic Chronic 2 | H242 H360D H315 H319 H411 | GHS02 GHS08 GHS07 GHS09 Dgr | H242 H315 H319 H360D H411 |  |  |  | CLP00/ATP15
617-007-00-5 | tert-butyl α,α-dimethylbenzyl peroxide | 222-389-8 | 3457-61-2 | Org. Perox. E Skin Irrit. 2 Aquatic Chronic 2 | H242 H315 H411 | GHS02 GHS07 GHS09 Wng | H242 H315 H411 |  |  |  | CLP00
617-008-00-0 | dibenzoyl peroxide; benzoyl peroxide | 202-327-6 | 94-36-0 | Org. Perox. B Eye Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H241 H319 H317 H400 H410 | GHS01 GHS02 GHS07 GHS09 Dgr | H241 H319 H317 H410 |  | M = 10 M = 10 |  | CLP00/ATP22
617-010-00-1 | 1-hydroperoxycyclohexyl 1-hydroxycyclohexyl peroxide; [> 91 % solution] [1] 1,1’-dioxybiscyclohexan-1-ol; [> 91 % solution] [2] cyclohexylidene hydroperoxide; [> 91 % solution] [3] cyclohexanone, peroxide; [> 91 % solution] [4] | 201-091-1 [1] 219-306-2 [2] 220-279-4 [3] 235-527-7 [4] | 78-18-2 [1] 2407-94-5 [2] 2699-11-8 [3] 12262-58-7 [4] | Org. Perox. A Acute Tox. 4 * Skin Corr. 1B | H240 H302 H314 | GHS01 GHS05 GHS07 Dgr | H240 H302 H314 |  | STOT SE 3; H335: C ≥ 5 % | C | CLP00/ATP01corr
617-010-01-9 | 1-hydroperoxycyclohexyl 1-hydroxycyclohexyl peroxide; [≤ 91 % solution] [1] 1,1'-dioxybiscyclohexan-1-ol; [≤ 91 % solution] [2] cyclohexylidene hydroperoxide; [≤ 91 % solution] [3] cyclohexanone, peroxide; [≤ 91 % solution] [4] | 201-091-1 [1] 219-306-2 [2] 220-279-4 [3] 235-527-7 [4] | 78-18-2 [1] 2407-94-5 [2] 2699-11-8 [3] 12262-58-7 [4] | Org. Perox. C Acute Tox. 4 * Skin Corr. 1B | H242 H302 H314 | GHS02 GHS05 GHS07 Dgr | H242 H302 H314 |  | STOT SE 3; H335: C ≥ 5 % | C T | CLP00
617-012-00-2 | 8-p-menthyl hydroperoxide; p-menthane hydroperoxide | 201-281-4 | 80-47-7 | Org. Perox. D Acute Tox. 4 * Skin Corr. 1B | H242 H332 H314 | GHS02 GHS05 GHS07 Dgr | H242 H314 H332 |  | STOT SE 3; H335: C ≥ 5 % |  | CLP00
617-013-00-8 | O,O-tert-butyl O-docosyl monoperoxyoxalate | 404-300-6 | 116753-76-5 | Org. Perox. C **** Aquatic Acute 1 Aquatic Chronic 1 | H242 H400 H410 | GHS02 GHS09 Dgr | H242 H410 |  |  |  | CLP00
617-014-00-3 | 6-(nonylamino)-6-oxo-peroxyhexanoic acid | 406-680-9 | 104788-63-8 | Org. Perox. C **** Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 | H242 H318 H317 H400 | GHS02 GHS05 GHS07 GHS09 Dgr | H242 H318 H317 H400 |  |  |  | CLP00
617-015-00-9 | bis(4-methylbenzoyl)peroxide | 407-950-9 | 895-85-2 | Org. Perox. B **** Aquatic Acute 1 Aquatic Chronic 1 | H241 H400 H410 | GHS01 GHS02 GHS09 Dgr | H241 H410 |  |  |  | CLP00
617-016-00-4 | 3-hydroxy-1,1-dimethylbutyl 2-ethyl-2-methylheptaneperoxoate | 413-910-1 |  | Flam. Liq. 3 Org. Perox. C **** Skin Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H226 H242 H315 H400 H410 | GHS02 GHS07 GHS09 Dgr | H242 H226 H315 H410 |  |  |  | CLP00
617-017-00-X | reaction mass of: 2,2'-bis(tert-pentylperoxy)-p-diisopropylbenzene; 2,2'-bis(tert-pentylperoxy)-m-diisopropylbenzene | 412-140-3 | 32144-25-5 | Org. Perox. D Aquatic Chronic 4 | H242 H413 | GHS02 Dgr | H242 H413 |  |  | T | CLP00/ATP01
617-018-00-5 | reaction mass of: 1-methyl-1-(3-(1-methylethyl)phenyl)ethyl-1-methyl-1-phenylethylperoxide, 63 % by weight; 1-methyl-1-(4-(1-methylethyl)phenyl)ethyl-1-methyl-1-phenylethylperoxide, 31 % by weight | 410-840-3 | 71566-50-2 | Org. Perox. C **** Aquatic Chronic 2 | H242 H411 | GHS02 GHS09 Dgr | H242 H411 |  |  | T | CLP00
617-019-00-0 | 6-(phthalimido)peroxyhexanoic acid | 410-850-8 | 128275-31-0 | Org. Perox. D Eye Dam. 1 Aquatic Acute 1 | H242 H318 H400 | GHS02 GHS05 GHS09 Dgr | H242 H318 H400 |  |  | T | CLP00
617-020-00-6 | 1,3-di(prop-2,2-diyl)benzene bis(neodecanoylperoxide) | 420-060-5 | 117663-11-3 | Flam. Liq. 3 Org. Perox. D **** Aquatic Chronic 2 | H226 H242 H411 | GHS02 GHS09 Dgr | H226 H242 H411 |  |  |  | CLP00
617-021-00-1 | methylethylketone peroxide trimer | 429-320-2 | - | Org. Perox. B **** Asp. Tox. 1 Skin Irrit. 2 Skin Sens. 1 | H241 H304 H315 H317 | GHS01 GHS02 GHS08 GHS07 Dgr | H241 H304 H315 H317 |  |  |  | ATP01
617-022-00-7 | reaction mass of: 1,2-dimethylpropylidene dihydroperoxide; dimethyl 1,2-benzenedicarboxylate | 442-480-8 | - | Org. Perox. C Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Chronic 2 | H242 H302 H314 H317 H411 | GHS02 GHS05 GHS07 GHS09 Dgr | H242 H302 H314 H317 H411 |  |  |  | ATP01
617-023-00-2 | tert-butyl hydroperoxide | 200-915-7 | 75-91-2 | Muta. 2 | H341 | GHS08 Wng | H341 |  |  |  | ATP09
617-024-00-8 | tert-butyl 2-ethylperoxyhexanoate | 221-110-7 | 3006-82-4 | Repr. 1B Skin Sens. 1 | H360FD H317 | GHS08 GHS07 Dgr | H360FD H317 |  |  |  | ATP22
647-001-00-8 | glucosidase, β- | 232-589-7 | 9001-22-3 | Resp. Sens. 1 | H334 | GHS08 Dgr | H334 |  |  |  | CLP00
647-002-00-3 | cellulase | 232-734-4 | 9012-54-8 | Resp. Sens. 1 | H334 | GHS08 Dgr | H334 |  |  |  | CLP00
647-003-00-9 | cellobiohydrolase, exo- | 253-465-9 | 37329-65-0 | Resp. Sens. 1 | H334 | GHS08 Dgr | H334 |  |  |  | CLP00
647-004-00-4 | cellulases with the exception of those specified elsewhere in this Annex |  |  | Resp. Sens. 1 | H334 | GHS08 Dgr | H334 |  |  | A | CLP00
647-005-00-X | bromelain, juice | 232-572-4 | 9001-00-7 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H335 H315 H319 H334 | GHS08 GHS07 Dgr | H319 H335 H315 H334 |  |  |  | CLP00
647-006-00-5 | ficin | 232-599-1 | 9001-33-6 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H335 H315 H319 H334 | GHS08 GHS07 Dgr | H319 H335 H315 H334 |  |  |  | CLP00
647-007-00-0 | papain | 232-627-2 | 9001-73-4 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H335 H315 H319 H334 | GHS08 GHS07 Dgr | H319 H335 H315 H334 |  |  |  | CLP00
647-008-00-6 | pepsin A | 232-629-3 | 9001-75-6 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H335 H315 H319 H334 | GHS08 GHS07 Dgr | H319 H335 H315 H334 |  |  |  | CLP00
647-009-00-1 | rennin | 232-645-0 | 9001-98-3 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H335 H315 H319 H334 | GHS08 GHS07 Dgr | H319 H335 H315 H334 |  |  |  | CLP00
647-010-00-7 | trypsin | 232-650-8 | 9002-07-7 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H335 H315 H319 H334 | GHS08 GHS07 Dgr | H319 H335 H315 H334 |  |  |  | CLP00
647-011-00-2 | chymotrypsin | 232-671-2 | 9004-07-3 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H335 H315 H319 H334 | GHS08 GHS07 Dgr | H319 H335 H315 H334 |  |  |  | CLP00
647-012-00-8 | subtilisin | 232-752-2 | 9014-01-1 | STOT SE 3 Skin Irrit. 2 Eye Dam. 1 Resp. Sens. 1 | H335 H315 H318 H334 | GHS08 GHS05 GHS07 Dgr | H335 H315 H318 H334 |  |  |  | CLP00
647-013-00-3 | proteinase, microbial neutral | 232-966-6 | 9068-59-1 | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H335 H315 H319 H334 | GHS08 GHS07 Dgr | H319 H335 H315 H334 |  |  |  | CLP00
647-014-00-9 | proteases with the exception of those specified elsewhere in this Annex |  |  | STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Resp. Sens. 1 | H335 H315 H319 H334 | GHS08 GHS07 Dgr | H319 H335 H315 H334 |  |  |  | CLP00
647-015-00-4 | amylase, α- | 232-565-6 | 9000-90-2 | Resp. Sens. 1 | H334 | GHS08 Dgr | H334 |  |  |  | CLP00
647-016-00-X | amylases with the exception of those specified elsewhere in this Annex |  |  | Resp. Sens. 1 | H334 | GHS08 Dgr | H334 |  |  |  | CLP00
647-017-00-5 | laccase | 420-150-4 | 80498-15-3 | Resp. Sens. 1 | H334 | GHS08 Dgr | H334 |  |  |  | ATP01
648-001-00-0 | Distillates (coal tar), benzole fraction; Light Oil; [A complex combination of hydrocarbons obtained by the distillation of coal tar. It consists of hydrocarbons having carbon numbers primarily in the range of C4 to C10 and distilling in the approximate range of 80 °C to 160 °C (175 °F to 320 °F).] | 283-482-7 | 84650-02-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
648-002-00-6 | Tar oils, brown-coal; Light Oil; [The distillate from lignite tar boiling in the range of approximately 80°C to 250°C (176°F to 482°F). Composed primarily of aliphatic and aromatic hydrocarbons and monobasic phenols.] | 302-674-4 | 94114-40-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-003-00-1 | Benzol forerunnings (coal); Light Oil Redistillate, low boiling; [The distillate from coke oven light oil having an approximate distillation range below 100°C (212°F). Composed primarily of C4 to C6 aliphatic hydrocarbons.] | 266-023-5 | 65996-88-5 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-004-00-7 | Distillates (coal tar), benzole fraction, BTX-rich; Light Oil Redistillate, low boiling; [A residue from the distillation of crude benzole to remove benzole fronts. Composed primarily of benzene, toluene and xylenes boiling in the range of approximately 75°C to 200°C (167°F to 392°F).] | 309-984-9 | 101896-26-8 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-005-00-2 | Aromatic hydrocarbons, C6-10, C8-rich; Light Oil Redistillate, low boiling | 292-697-5 | 90989-41-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-006-00-8 | Solvent naphtha (coal), light; Light Oil Redistillate, low boiling | 287-498-5 | 85536-17-0 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-007-00-3 | Solvent naphtha (coal), xylene-styrene cut; Light Oil Redistillate, intermediate boiling | 287-502-5 | 85536-20-5 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-008-00-9 | Solvent naphtha (coal), coumarone-styrene contg.; Light Oil Redistillate, intermediate boiling | 287-500-4 | 85536-19-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-009-00-4 | Naphtha (coal), distn. residues; Light Oil Redistillate, high boiling; [The residue remaining from the distillation of recovered naphtha. Composed primarily of naphthalene and condensation products of indene and styrene.] | 292-636-2 | 90641-12-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-010-00-X | Aromatic hydrocarbons, C8; Light Oil Redistillate, high boiling | 292-694-9 | 90989-38-1 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-012-00-0 | Aromatic hydrocarbons, C8-9, hydrocarbon resin polymn. by-product; Light Oil Redistillate, high boiling; [A complex combination of hydrocarbons obtained from the evaporation of solvent under vacuum from polymerized hydrocarbon resin. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C8 through C9 and boiling in the range of approximately 120°C to 215°C (248°F to 419°F).] | 295-281-1 | 91995-20-9 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-013-00-6 | Aromatic hydrocarbons, C9-12, benzene distn.; Light Oil Redistillate, high boiling | 295-551-9 | 92062-36-7 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-014-00-1 | Extract residues (coal), benzole fraction alk., acid ext.; Light Oil Extract Residues, low boiling; [The redistillate from the distillate, freed of tar acids and tar bases, from bituminous coal high temperature tar boiling in the approximate range of 90°C to 160°C (194°F to 320°F). It consists predominantly of benzene, toluene and xylenes.] | 295-323-9 | 91995-61-8 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-015-00-7 | Extract residues (coal tar), benzole fraction alk., acid ext.; Light Oil Extract Residues, low boiling; [A complex combination of hydrocarbons obtained by the redistillation of the distillate of high temperature coal tar (tar acid and tar base free). It consists predominantly of unsubstituted and substituted mononuclear aromatic hydrocarbons boiling in the range of 85°C to 195°C (185°F to 383°F).] | 309-868-8 | 101316-63-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-016-00-2 | Extract residues (coal), benzole fraction acid; Light Oil Extract Residues, low boiling; [An acid sludge by-product of the sulfuric acid refining of crude high temperature coal. Composed primarily of sulfuric acid and organic compounds.] | 298-725-2 | 93821-38-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-017-00-8 | Extract residues (coal), light oil alk., distn. overheads; Light Oil Extract Residues, low boiling; [The first fraction from the distillation of aromatic hydrocarbons, coumarone, naphthalene and indene rich prefractionator bottoms or washed carbolic oil boiling substantially below 145°C (293°F). Composed primarily of C7 and C8 aliphatic and aromatic hydrocarbons.] | 292-625-2 | 90641-02-4 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-018-00-3 | Extract residues (coal), light oil alk., acid ext., indene fraction; Light Oil Extract Residues, intermediate boiling | 309-867-2 | 101316-62-5 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-019-00-9 | Extract residues (coal), light oil alk., indene naphtha fraction; Light Oil Extract Residues, high boiling; [The distillate from aromatic hydrocarbons, coumarone, naphthalene and indene rich prefractionator bottoms or washed carbolic oils, having an approximate boiling range of 155°C to 180°C (311°F to 356°F). Composed primarily of indene, indan and trimethylbenzenes.] | 292-626-8 | 90641-03-5 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-020-00-4 | Solvent naphtha (coal); Light Oil Extract Residues, high boiling; [The distillate from either high temperature coal tar, coke oven light oil, or coal tar oil alkaline extract residue having an approximate distillation range of 130°C to 210°C (266°F to 410°F). Composed primarily of indene and other polycyclic ring systems containing a single aromatic ring. May contain phenolic compounds and aromatic nitrogen bases.] | 266-013-0 | 65996-79-4 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-021-00-X | Distillates (coal tar), light oils, neutral fraction; Light Oil Extract Residues, high boiling; [A distillate from the fractional distillation of high temperature coal tar. Composed primarily of alkyl-substituted one ring aromatic hydrocarbons boiling in the range of approximately 135°C to 210°C (275°F to 410°F). May also include unsaturated hydrocarbons such as indene and coumarone.] | 309-971-8 | 101794-90-5 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-022-00-5 | Distillates (coal tar), light oils, acid exts.; Light Oil Extract Residues, high boiling; [This oil is a complex reaction mass of aromatic hydrocarbons, primarily indene, naphthalene, coumarone, phenol, and o-, m- and p-cresol and boiling in the range of 140°C to 215°C (284°F to 419°F).] | 292-609-5 | 90640-87-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-023-00-0 | Distillates (coal tar), light oils; Carbolic Oil; [A complex combination of hydrocarbons obtained by distillation of coal tar. It consists of aromatic and other hydrocarbons, phenolic compounds and aromatic nitrogen compounds and distills at the approximate range of 150°C to 210°C (302°F to 410°F).] | 283-483-2 | 84650-03-3 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-024-00-6 | Tar oils, coal; Carbolic Oil; [The distillate from high temperature coal tar having an approximate distillation range of 130°C to 250°C (266°F to 410°F). Composed primarily of naphthalene, alkylnaphthalenes, phenolic compounds, and aromatic nitrogen bases.] | 266-016-7 | 65996-82-9 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-026-00-7 | Extract residues (coal), light oil alk., acid ext.; Carbolic Oil Extract Residue; [The oil resulting from the acid washing of alkali-washed carbolic oil to remove the minor amounts of basic compounds (tar bases). Composed primarily of indene, indan and alkylbenzenes.] | 292-624-7 | 90641-01-3 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-027-00-2 | Extract residues (coal), tar oil alk.; Carbolic Oil Extract Residue; [The residue obtained from coal tar oil by an alkaline wash such as aqueous sodium hydroxide after the removal of crude coal tar acids. Composed primarily of naphthalenes and aromatic nitrogen bases.] | 266-021-4 | 65996-87-4 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-028-00-8 | Extract oils (coal), light oil; Acid Extract; [The aqueous extract produced by an acidic wash of alkali-washed carbolic oil. Composed primarily of acid salts of various aromatic nitrogen bases including pyridine, quinoline and their alkyl derivatives.] | 292-622-6 | 90640-99-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-029-00-3 | Pyridine, alkyl derivs.; Crude Tar Bases; [The complex combination of polyalkylated pyridines derived from coal tar distillation or as high-boiling distillates approximately above 150°C (302°F) from the reaction of ammonia with acetaldehyde, formaldehyde or paraformaldehyde.] | 269-929-9 | 68391-11-7 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-030-00-9 | Tar bases, coal, picoline fraction; Distillate Bases; [Pyridine bases boiling in the range of approximately 125°C to 160°C (257°F 320°F) obtained by distillation of neutralized acid extract of the base-containing tar fraction obtained by the distillation of bituminous coal tars. Composed chiefly of lutidines and picolines.] | 295-548-2 | 92062-33-4 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-031-00-4 | Tar bases, coal, lutidine fraction; Distillate Bases | 293-766-2 | 91082-52-9 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-032-00-X | Extract oils (coal), tar base, collidine fraction; Distillate Bases; [The extract produced by the acidic extraction of bases from crude coal tar aromatic oils, neutralization, and distillation of the bases. Composed primarily of collidines, aniline, toluidines, lutidines, xylidines.] | 273-077-3 | 68937-63-3 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-033-00-5 | Tar bases, coal, collidine fraction; Distillate Bases; [The distillation fraction boiling in the range of approximately 181 °C to 186 °C (356 °F to 367 °F) from the crude bases obtained from the neutralized, acid-extracted base-containing tar fractions obtained by the distillation of bituminous coal tar. It contains chiefly aniline and collidines.] | 295-543-5 | 92062-28-7 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-034-00-0 | Tar bases, coal, aniline fraction; Distillate Bases; [The distillation fraction boiling in the range of approximately 180 °C to 200 °C (356 °F to 392 °F) from the crude bases obtained by dephenolating and debasing the carbolated oil from the distillation of coal tar. It contains chiefly aniline, collidines, lutidines and toluidines.] | 295-541-4 | 92062-27-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-035-00-6 | Tar bases, coal, toluidine fraction; Distillate Bases | 293-767-8 | 91082-53-0 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-036-00-1 | Distillates (petroleum), alkene-alkyne manuf. pyrolysis oil, mixed with high-temp. coal tar, indene fraction; Redistillates; [A complex combination of hydrocarbons obtained as a redistillate from the fractional distillation of bituminous coal high temperature tar and residual oils that are obtained by the pyrolytic production of alkenes and alkynes from petroleum products or natural gas. It consists predominantly of indene and boils in a range of approximately 160°C to 190°C (320°F to 374°F).] | 295-292-1 | 91995-31-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-037-00-7 | Distillates (coal), coal tar-residual pyrolysis oils, naphthalene oils; Redistillates; [The redistillate obtained from the fractional distillation of bituminous coal high temperature tar and pyrolysis residual oils and boiling in the range of approximately 190°C to 270°C (374°F to 518°F). Composed primarily of substituted dinuclear aromatics.] | 295-295-8 | 91995-35-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-038-00-2 | Extract oils (coal), coal tar-residual pyrolysis oils, naphthalene oil, redistillate; Redistillates; [The redistillate from the fractional distillation of dephenolated and debased methylnaphthalene oil obtained from bituminous coal high temperature tar and pyrolysis residual oils boiling in the approximate range of 220°C to 230°C (428°F to 446°F). It consists predominantly of unsubstituted and substituted dinuclear aromatic hydrocarbons.] | 295-329-1 | 91995-66-3 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-039-00-8 | Extract oils (coal), coal tar-residual pyrolysis oils, naphthalene oils; Redistillates; [A neutral oil obtained by debasing and dephenolating the oil obtained from the distillation of high temperature tar and pyrolysis residual oils which has a boiliing range of 225°C to 255°C (437°F to 491°F). Composed primarily of substituted dinuclear aromatic hydrocarbons.] | 310-170-0 | 122070-79-5 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-040-00-3 | Extract oils (coal), coal tar residual pyrolysis oils, naphthalene oil, distn. residues; Redistillates; [Residue from the distillation of dephenolated and debased methylnaphthalene oil (from bituminous coal tar and pyrolysis residual oils) with a boiling range of 240°C to 260°C (464°F to 500°F). Composed primarily of substituted dinuclear aromatic and heterocyclic hydrocarbons.] | 310-171-6 | 122070-80-8 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-041-00-9 | Absorption oils, bicyclo arom. and heterocyclic hydrocarbon fraction; Wash Oil Redistillate; [A complex combination of hydrocarbons obtained as a redistillate from the distillation of wash oil. It consists predominantly of 2-ringed aromatic and heterocyclic hydrocarbons boiling in the range of approximately 260 °C to 290 °C (500 °F to 554 °F).] | 309-851-5 | 101316-45-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-042-00-4 | Distillates (coal tar), upper, fluorene-rich; Wash Oil Redistillate; [A complex combination of hydrocarbons obtained by the crystallization of tar oil. It consists af aromatic and polycyclic hydrocarbons primarily fluorene and some acenaphthene.] | 284-900-0 | 84989-11-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-043-00-X | Creosote oil, acenaphthene fraction, acenaphthene-free; Wash Oil Redistillate; [The oil remaining after removal by a crystallization process of acenaphthene from acenaphthene oil from coal tar. Composed primarily of naphthalene and alkylnaphthalenes.] | 292-606-9 | 90640-85-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00/ATP01
648-044-00-5 | Distillates (coal tar), heavy oils; Heavy Anthracene Oil; [Distillate from the fractional distillation of coal tar of bituminous coal, with boiling range of 240 °C to 400 °C (464 °F to 752 °F). Composed primarily of tri- and polynuclear hydrocarbons and heterocyclic compounds.] | 292-607-4 | 90640-86-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
648-045-00-0 | Distillates (coal tar), upper; Heavy Anthracene Oil; [The distillate from coal tar having an approximate distillation range of 220 °C to 450 °C (428 °F to 842 °F). Composed primarily of three to four membered condensed ring aromatic hydrocarbons and other hydrocarbons.] | 266-026-1 | 65996-91-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-046-00-6 | Anthracene oil, acid ext.; Anthracene Oil Extract Residue; [A complex combination of hydrocarbons from the base-freed fraction obtained from the distillation of coal tar and boiling in the range of approximately 325 °C to 365 °C (617 °F to 689 °F). It contains predominantly anthracene and phenanthrene and their alkyl derivatives.] | 295-274-3 | 91995-14-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-047-00-1 | Distillates (coal tar); Heavy Anthracene Oil; [The distillate from coal tar having an approximate distillation range of 100 °C to 450 °C (212 °F to 842 °F). Composed primarily of two to four membered condensed ring aromatic hydrocarbons, phenolic compounds, and aromatic nitrogen bases.] | 266-027-7 | 65996-92-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-048-00-7 | Distillates (coal tar), pitch, heavy oils; Heavy Anthracene Oil; [The distillate from the distillation of the pitch obtained from bituminous high temperature tar. Composed primarily of tri- and polynuclear aromatic hydrocarbons and boiling in the range of approximately 300 °C to 470 °C (572 °F to 878 °F). The product may also contain heteroatoms.] | 295-312-9 | 91995-51-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-049-00-2 | Distillates (coal tar), pitch; Heavy Anthracene Oil; [The oil obtained from condensation of the vapors from the heat treatment of pitch. Composed primarily of two- to four-ring aromatic compounds boiling in the range of 200 °C to greater than 400 °C (392 °F to greater than 752 °F).] | 309-855-7 | 101316-49-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-050-00-8 | Distillates (coal tar), heavy oils, pyrene fraction; Heavy Anthracene Oil Redistillate; [The redistillate obtained from the fractional distillation of pitch distillate boiling in the range of approximately 350 °C to 400 °C (662 °F to 752 °F). Consists predominantly of tri- and polynuclear aromatics and heterocyclic hydrocarbons.] | 295-304-5 | 91995-42-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-051-00-3 | Distillates (coal tar), pitch, pyrene fraction; Heavy Anthracene Oil Redistillate; [The redistillate obtained from the fractional distillation of pitch distillate and boiling in the range of approximately 380 °C to 410 °C (716 to 770 °F). Composed primarily of tri- and polynuclear aromatic hydrocarbons and heterocyclic compounds.] | 295-313-4 | 91995-52-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-052-00-9 | Paraffin waxes (coal), brown-coal high-temp. tar, carbon-treated; Coal Tar Extract; [A complet combination of hydrocarbons obtained by the treatment of lignite carbonization tar with activated carbon for removal of trace constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C12.] | 308-296-6 | 97926-76-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-053-00-4 | Paraffin waxes (coal), brown-coal high-temp tar, clay-treated; Coal Tar Extract; [A complex combination of hydrocarbons obtained by the treatment of lignite carbonization tar with bentonite for removal of trace constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C12.] | 308-297-1 | 97926-77-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-054-00-X | Pitch; Pitch | 263-072-4 | 61789-60-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-055-00-5 | pitch, coal tar, high-temp.,; [The residue from the distillation of high temperature coal tar. A black solid with an approximate softening point from 30 °C to 180 °C (86 °F to 356 °F). Composed primarily of a complex mixture of three or more membered condensed ring aromatic hydrocarbons.] | 266-028-2 | 65996-93-2 | Carc. 1A Muta. 1B Repr. 1B | H350 H340 H360FD | GHS08 Dgr | H340 H350 H360FD |  |  |  | CLP00/ATP14
648-056-00-0 | Pitch, coal tar, high-temp., heat-treated; Pitch; [The heat treated residue from the distillation of high temperature coal tar. A black solid with an approximate softening point from 80 °C to 180 °C (176 °F to 356 °F). Composed primarily of a complex mixture of three or more membered condensed ring aromatic hydrocarbons.] | 310-162-7 | 121575-60-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-057-00-6 | Pitch, coal tar, high-temp., secondary; Pitch Redistillate; [The residue obtained during the distillation of high boiling fractions from bituminous coal high temperature tar and/or pitch coke oil, with a softening point of 140 °C to 170 °C (284 °F to 392 °F) according to DIN 52025. Composed primarily of tri- and polynuclear aromatic compounds which also contain heteroatoms.] | 302-650-3 | 94114-13-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-058-00-1 | Residues (coal tar), pitch distn.; Pitch Redistillate; [Residue from the fractional distillation of pitch distillate boiling in the range of approximately 400 °C to 470 °C (752 °F to 846 °F). Composed primarily of polynuclear aromatic hydrocarbons, and heterocyclic compounds.] | 295-507-9 | 92061-94-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-059-00-7 | Tar, coal, high-temp., distn. and storage residues; Coal Tar Solids Residue; [Coke- and ash-containing solid residues that separate on distillation and thermal treatment of bituminous coal high temperature tar in distillation installations and storage vessels. Consists predominantly of carbon and contains a small quantity of hetero compounds as well as ash components.] | 295-535-1 | 92062-20-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-060-00-2 | Tar, coal, storage residues; Coal Tar Solids Residue; [The deposit removed from crude coal tar storages. Composed primarily of coal tar and carbonaceous particulate matter.] | 293-764-1 | 91082-50-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-061-00-8 | Tar, coal, high-temp., residues; Coal Tar Solids Residue; [Solids formed during the coking of bituminous coal to produce crude bituminous coal high temperature tar. Composed primarily of coke and coal particles, highly aromatized compounds and mineral substances.] | 309-726-5 | 100684-51-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-062-00-3 | Tar, coal, high-temp., high-solids; Coal Tar Solids Residue; [The condensation product obtained by cooling, to approximately ambient temperature, the gas evolved in the high temperature (greater than 700 °C (1292 °F)) destructive distillation of coal. Composed primarily of a complex mixture of condensed ring aromatic hydrocarbons with a high solid content of coal-type materials.] | 273-615-7 | 68990-61-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-063-00-9 | Waste solids, coal-tar pitch coking; Coal Tar Solids Residue; [The combination of wastes formed by the coking of bituminous coal tar pitch. It consists predominantly of carbon.] | 295-549-8 | 92062-34-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-064-00-4 | Extract residues (coal), brown; Coal Tar Extract; [The residue from extraction of dried coal.] | 294-285-0 | 91697-23-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-065-00-X | Paraffin waxes (coal), brown-coal-high-temp. tar; Coal Tar Extract; [A complex combination of hydrocarbons obtained from lignite carbonization tar by solvent crystallisation (solvent deoiling), by sweating or an adducting process. It consists predominantly of straight and branched chain saturated hydrocarbons having carbon numbers predominantly greater than C12.] | 295-454-1 | 92045-71-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-066-00-5 | Paraffin waxes (coal), brown-coal-high-temp. tar, hydrotreated; Coal Tar Extract; [A complex combination of hydrocarbons obtained from lignite carbonization tar by solvent crystallisation (solvent deoiling), by sweating or an adducting process treated with hydrogen in the presence of a catalyst. It consists predominantly of straight and branched chain saturated hydrocarbons having carbon numbers predominantly greater than C12.] | 295-455-7 | 92045-72-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-067-00-0 | Paraffin waxes (coal), brown-coal high-temp tar, silicic acid-treated; Coal Tar Extract; [A complex combination of hydrocarbons obtained by the treatment of lignite carbonization tar with silicic acid for removal of trace constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C12.] | 308-298-7 | 97926-78-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-068-00-6 | Tar, coal, low-temp., distn. residues; Tar Oil, intermediate boiling; [Residues from fractional distillation of low temperature coal tar to remove oils that boil in a range up to approximately 300 °C (572 °F). Composed primarily of aromatic compounds.] | 309-887-1 | 101316-85-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-069-00-1 | Pitch, coal tar, low-temp; Pitch Residue; [A complex black solid or semi-solid obtained from the distillation of a low temperature coal tar. It has a softening point within the approximate range of 40 °C to 180 °C (104 °F to 356 °F). Composed primarily of a complex mixture of hydrocarbons.] | 292-651-4 | 90669-57-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-070-00-7 | Pitch, coal tar, low-temp., oxidized; Pitch Residue, oxidised; [The product obtained by air-blowing, at elevated temperature, low-temperature coal tar pitch. It has a softening-point within the approximate range of 70 °C to 180 °C (158 °F to 356 °F). Composed primarily of a complex mixture of hydrocarbons.] | 292-654-0 | 90669-59-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-071-00-2 | Pitch, coal tar, low-temp., heat-treated; Pitch Residue, oxidised; Pitch Residue, heat-treated; [A complex black solid obtained by the heat treatment of low temperature coal tar pitch. It has a softening point within the approximate range of 50 °C to 140 °C (122 °F to 284 °F). Composed primarily of a complex mixture of aromatic compounds.] | 292-653-5 | 90669-58-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-072-00-8 | Distillates (coal-petroleum), condensed-ring arom; Distillates; [The distillate from a mixture of coal and tar and aromatic petroleum streams having an approximate distillation range of 220 °C to 450 °C (428 °F to 842 °F). Composed primarily of 3- to 4-membered condensed ring aromatic hydrocarbons.] | 269-159-3 | 68188-48-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-073-00-3 | Aromatic hydrocarbons, C20-28, polycyclic, mixed coal-tar pitch-polyethylene-polypropylene pyrolysis-derived; Pyrolysis Products; [A complex combination hydrocarbons obtained from mixed coal tar pitch-polyethylene-polypropylene pyrolysis. Composed primarily of polycyclic aromatic hydrocarbons having carbon numbers predominantly in the range of C20 through C28 and having a softening point of 100 °C to 220 °C (212 °F to 428 °F) according to DIN 52025.] | 309-956-6 | 101794-74-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-074-00-9 | Aromatic hydrocarbons, C20-28, polycyclic, mixed coal-tar pitch-polyethylene pyrolysis-derived; Pyrolysis Products; [A complex combination of hydrocarbons obtained from mixed coal tar pitch-polyethylene pyrolysis. Composed primarily of polycyclic aromatic hydrocarbons having carbon numbers predominantly in the range of C20 through C28 and having a softening point of 100 °C to 220 °C (212 °F to 428 °F) according to DIN 52025.] | 309-957-1 | 101794-75-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-075-00-4 | Aromatic hydrocarbons, C20-28, polycyclic, mixed coal-tar pitch-polystyrene pyrolysis-derived; Pyrolysis Products; [A complex combination of hydrocarbons obtained from mixed coal tar pitch-polystyrene pyrolysis. Composed primarily of polycyclic aromatic hydrocarbons having carbon numbers predominantly in the range of C20 through C28 and having a softening point of 100 °C to 220 °C (212 °F to 428 °F) according to DIN 52025.] | 309-958-7 | 101794-76-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-076-00-X | Pitch, coal tar-petroleum; Pitch Residues; [The residue from the distillation of a mixture of coal tar and aromatic petroleum streams. A solid with a softening point from 40 °C to 180 °C (140 °F to 356 °F). Composed primarily of a complex combination of three or more membered condensed ring aromatic hydrocarbons.] | 269-109-0 | 68187-57-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-077-00-5 | Phenanthrene, distn. residues; Heavy Anthracene Oil Redistillate; [Residue from the distillation of crude phenanthrene boiling in the approximate range of 340 °C to 420 °C (644 °F to 788 °F). It consists predominantly of phenanthrene, anthracene and carbazole.] | 310-169-5 | 122070-78-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-078-00-0 | Distillates (coal tar), upper, fluorene-free; Wash Oil Redistillate; [A complex combination of hydrocarbons obtained by the crystallization of tar oil. It consists of aromatic polycyclic hydrocarbons, primarily diphenyl, dibenzofuran and acenaphthene.] | 284-899-7 | 84989-10-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-079-00-6 | Anthracene oil; [A complex combination of polycyclic aromatic hydrocarbons obtained from coal tar having an approximate distillation range of 300 °C to 400 °C (572 °F to 752 °F). Composed primarily of phenanthrene, anthracene and carbazole.] | 292-602-7 | 90640-80-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-080-00-1 | Residues (coal tar), creosote oil distn.; Wash Oil Redistillate; [The residue from the fractional distillation of wash oil boiling in the approximate range of 270°C to 330°C (518°F to 626°F). It consists predominantly of dinuclear aromatic and heterocyclic hydrocarbons.] | 295-506-3 | 92061-93-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00/ATP01
648-081-00-7 | Tar, coal; Coal tar; [The by-product from the destructive distillation of coal. Almost black semisolid. A complex combination of aromatic hydro-carbons, phenolic compounds, nitrogen bases and thiophene.] | 232-361-7 | 8007-45-2 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
648-082-00-2 | Tar, coal, high-temp.; Coal tar; [The condensation product obtained by cooling, to approximately ambient temperature, the gas evolved in the high temperature (greater than 700 °C (1292 °F)) destructive distillation of coal. A black viscous liquid denser than water. Composed primarily of a complex mixture of condensed ring aromatic hydrocarbons. May contain minor amounts of phenolic compounds and aromatic nitrogen bases.] | 266-024-0 | 65996-89-6 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
648-083-00-8 | Tar, coal, low-temp.; Coal oil; [The condensation product obtained by cooling, to approximately ambient temperature, the gas evolved in low temperature (less than 700 °C (1292 °F)) destructive distillation of coal. A black viscous liquid denser than water. Composed primarily of condensed ring aromatic hydrocarbons, phenolic compounds, aromatic nitrogen bases, and their alkyl derivatives.] | 266-025-6 | 65996-90-9 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
648-084-00-3 | Distillates (coal), coke-oven light oil, naphthalene cut; Naphthalene Oil; [The complex combination of hydrocarbons obtained from prefractionation (continuous distillation) of coke oven light oil. It consists predominantly of naphthalene, coumarone and indene and boils above 148°C (298°F).] | 285-076-5 | 85029-51-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-085-00-9 | Distillates (coal tar), naphthalene oils; Naphthalene Oil; [A complex combination of hydrocarbons obtained by the distillation of coal tar. It consists primarily of aromatic and other hydrocarbons, phenolic compounds and aromatic nitrogen compounds and distills in the approximate range of 200°C to 250°C (392°F to 482°F).] | 283-484-8 | 84650-04-4 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-086-00-4 | Distillates (coal tar), naphthalene oils, naphthalene-low; Naphthalene Oil Redistillate; [A complex combination of hydrocarbons obtained by crystallization of naphthalene oil. Composed primarily of naphthalene, alkyl naphthalenes and phenolic compounds.] | 284-898-1 | 84989-09-3 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-087-00-X | Distillates (coal tar), naphthalene oil crystn. mother liquor; Naphthalene Oil Redistillate; [A complex combination of organic compounds obtained as a filtrate from the crystallization of the naphthalene fraction from coal tar and boiling in the range of approximately 200°C to 230°C (392°F to 446°F). Contains chiefly naphthalene, thionaphthene and alkylnaphthalenes.] | 295-310-8 | 91995-49-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-088-00-5 | Extract residues (coal), naphthalene oil, alk.; Naphthalene Oil Extract Residue; [A complex combination of hydrocarbons obtained from the alkali washing of naphthalene oil to remove phenolic compounds (tar acids). It is composed of naphthalene and alkyl naphthalenes.] | 310-166-9 | 121620-47-1 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-089-00-0 | Extract residues (coal), naphthalene oil, alk., naphthalene-low; Naphthalene Oil Extract Residue; [A complex combination of hydrocarbons remaining after the removal of naphthalene from alkali-washed naphthalene oil by a crystallization process. It is composed primarily of naphthalene and alkyl naphthalenes.] | 310-167-4 | 121620-48-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-090-00-6 | Distillates (coal tar), naphthalene oils, naphthalene-free, alk. exts.; Naphthalene Oil Extract Residue; [The oil remaining after the removal of phenolic compounds (tar acids) from drained naphthalene oil by an alkali wash. Composed primarily of naphthalene and alkyl naphthalenes.] | 292-612-1 | 90640-90-7 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-091-00-1 | Extract residues (coal), naphthalene oil alk., distn. overheads; Naphthalene Oil Extract Residue; [The distillate from alkali-washed naphthalene oil having an approximate distillation range of 180°C to 220°C (356°F to 428°F). Composed primarily of naphthalene, alkylbenzenes, indene and indan.] | 292-627-3 | 90641-04-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-092-00-7 | Distillates (coal tar), naphthalene oils, methylnaphthalene fraction; Methylnaphthalene Oil; [A distillate from the fractional distillation of high temperature coal tar. Composed primarily of substituted two ring aromatic hydrocarbons and aromatic nitrogen bases boiling in the range of approximately 225°C to 255°C (437°F to 491°F).] | 309-985-4 | 101896-27-9 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-093-00-2 | Distillates (coal tar), naphthalene oils, indole-methylnaphthalene fraction; Methylnaphthalene Oil; [A distillate from the fractional distillation of high temperature coal tar. Composed primarily of indole and methylnaphthalene boiling in the range of approximately 235°C to 255°C (455°F to 491°F).] | 309-972-3 | 101794-91-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-094-00-8 | Distillates (coal tar), naphthalene oils, acid exts.; Methylnaphthalene Oil Extract Residue; [A complex combination of hydrocarbons obtained by debasing the methylnaphthalene fraction obtained by the distillation of coal tar and boiling in the range of approximately 230°C to 255°C (446°F to 491°F). Contains chiefly 1(2)-methylnaphthalene, naphthalene, dimethylnaphthalene and biphenyl.] | 295-309-2 | 91995-48-1 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-095-00-3 | Extract residues (coal), naphthalene oil alk., distn. residues; Methylnaphthalene Oil Extract Residue; [The residue from the distillation of alkali-washed naphthalene oil having an approximate distillation range of 220°C to 300°C (428°F to 572°F). Composed primarily of naphthalene, alkylnaphthalenes and aromatic nitrogen bases.] | 292-628-9 | 90641-05-7 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-096-00-9 | Extract oils (coal), acidic, tar-base free; Methylnaphthalene Oil Extract Residue; [The extract oil boiling in the range of approximately 220°C to 265°C (428°F to 509°F) from coal tar alkaline extract residue produced by an acidic wash such as aqueous sulfuric acid after distillation to remove tar bases. Composed primarily of alkylnaphthalenes.] | 284-901-6 | 84989-12-8 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-097-00-4 | Distillates (coal tar), benzole fraction, distn. residues; Wash Oil; [A complex combination of hydrocarbons obtained from the distillation of crude benzole (high temperature coal tar). It may be a liquid with the approximate distillation range of 150°C to 300°C (302°F to 572°F) or a semi-solid or solid with a melting point up to 70°C (158°F). It is composed primarily of naphthalene and alkyl naphthalenes.] | 310-165-3 | 121620-46-0 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-098-00-X | Creosote oil, acenaphthene fraction; Wash Oil; [A complex combination of hydrocarbons produced by the distillation of coal tar and boiling in the range of approximately 240°C to 280°C (464°F to 536°F). Composed primarily of acenaphthene, naphthalene and alkyl naphthalene.] | 292-605-3 | 90640-84-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00/ATP01
648-099-00-5 | Creosote oil; [A complex combination of hydrocarbons obtained by the distillation of coal tar. It consists primarily of aromatic hydrocarbons and may contain appreciable quantities of tar acids and tar bases. It distills at the approximate range of 200°C to 325°C (392°F to 617°F).] | 263-047-8 | 61789-28-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00/ATP01
648-100-00-9 | Creosote oil, high-boiling distillate; Wash Oil; [The high-boiling distillation fraction obtained from the high temperature carbonization of bituminous coal which is further refined to remove excess crystalline salts. It consists primarily of creosote oil with some of the normal polynuclear aromatic salts, which are components of coal tar distillates, removed. It is crystal free at approximately 5°C (41°F).] | 274-565-9 | 70321-79-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00/ATP01
648-101-00-4 | Creosote; [The distillate of coal tar produced by the high temperature carbonization of bituminous coal. It consists primarily of aromatic hydrocarbons, tar acids and tar bases.] | 232-287-5 | 8001-58-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
648-102-00-X | Extract residues (coal), creosote oil acid; Wash Oil Extract Residue; [A complex combination of hydrocarbons from the base-freed fraction from the distillation of coal tar, boiling in the range of approximately 250°C to 280°C (482°F to 536°F). It consists predominantly of biphenyl and isomeric diphenylnaphthalenes.] | 310-189-4 | 122384-77-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00/ATP01
648-103-00-5 | Anthracene oil, anthracene paste; Anthracene Oil Fraction; [The anthracene-rich solid obtained by the crystallization and centrifuging of anthracene oil. It is composed primarily of anthracene, carbazole and phenanthrene.] | 292-603-2 | 90640-81-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-104-00-0 | Anthracene oil, anthracene-low; Anthracene Oil Fraction; [The oil remaining after the removal, by a crystallization process, of an anthracene-rich solid (anthracene paste) from anthracene oil. It is composed primarily of two, three and four membered aromatic compounds.] | 292-604-8 | 90640-82-7 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-105-00-6 | Residues (coal tar), anthracene oil distn.; Anthracene Oil Fraction; [The residue from the fraction distillation of crude anthracene boiling in the approximate range of 340°C to 400°C (644°F to 752°F). It consists predominantly of tri- and polynuclear aromatic and heterocyclic hydrocarbons.] | 295-505-8 | 92061-92-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-106-00-1 | Anthracene oil, anthracene paste, anthracene fraction; Anthracene Oil Fraction; [A complex combination of hydrocarbons from the distillation of anthracene obtained by the crystallization of anthracene oil from bituminous high temperature tar and boiling in the range of 330°C to 350°C (626°F to 662°F). It contains chiefly anthracene, carbazole and phenanthrene.] | 295-275-9 | 91995-15-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-107-00-7 | Anthracene oil, anthracene paste, carbazole fraction; Anthracene Oil Fraction; [A complex combination of hydrocarbons from the distillation of anthracene obtained by crystallization of anthracene oil from bituminous coal high temperature tar and boiling in the approximate range of 350°C to 360°C (662°F to 680°F). It contains chiefly anthracene, carbazole and phenanthrene.] | 295-276-4 | 91995-16-3 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-108-00-2 | Anthracene oil, anthracene paste, distn. lights; Anthracene Oil Fraction; [A complex combination of hydrocarbons from the distillation of anthracene obtained by crystallization of anthracene oil from bituminous high temperature tar and boiling in the range of approximately 290°C to 340°C (554°F to 644°F). It contains chiefly trinuclear aromatics and their dihydro derivatives.] | 295-278-5 | 91995-17-4 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-109-00-8 | Tar oils, coal, low-temp.; Tar Oil, high boiling; [A distillate from low-temperature coal tar. Composed primarily of hydrocarbons, phenolic compounds and aromatic nitrogen bases boiling in the range of approximately 160°C to 340°C (320°F to 644°F).] | 309-889-2 | 101316-87-4 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-110-00-3 | Extract residues (coal), low temp. coal atar alk.; [The residue from low temperature coal tar oils after an alkaline wash, such as aqueous sodium hydroxide, to remove crude coal tar acids. Composed primarily of hydrocarbons and aromatic nitrogen bases.] | 310-191-5 | 122384-78-5 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-111-00-9 | Phenols, ammonia liquor ext.; Alkaline Extract; [The combination of phenols extracted, using isobutyl acetate, from the ammonia liquor condensed from the gas evolved in low-temperature (less than 700°C (1292°F)) destructive distillation of coal. It consists predominantly of a reaction mass of monohydric and dihydric phenols.] | 284-881-9 | 84988-93-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-112-00-4 | Distillates (coal tar), light oils, alk. exts.; Alkaline Extract; [The aqueous extract from carbolic oil produced by an alkaline wash such as aqueous sodium hydroxide. Composed primarily of the alkali salts of various phenolic compounds.] | 292-610-0 | 90640-88-3 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-113-00-X | Extracts, coal tar oil alk.; Alkaline Extract; [The extract from coal tar oil produced by an alkaline wash such as aqueous sodium hydroxide. Composed primarily of the alkali salts of various phenolic compounds.] | 266-017-2 | 65996-83-0 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-114-00-5 | Distillates (coal tar), naphthalene oils, alk. exts.; Alkaline Extract; [The aqueous extract from naphthalene oil produced by an alkaline wash such as aqueous sodium hydroxide. Composed primarily of the alkali salts of various phenolic compounds.] | 292-611-6 | 90640-89-4 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-115-00-0 | Extract residues (coal), tar oil alk., carbonated, limed; Crude Phenols; [The product obtained by treatment of coal tar oil alkaline extract with CO2 and CaO. Composed primarily of CaCO3, Ca(OH)2, Na2CO3 and other organic and inorganic impurities.] | 292-629-4 | 90641-06-8 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-116-00-6 | Tar acids, coal, crude; Crude Phenols; [The reaction product obtained by neutralizing coal tar oil alkaline extract with an acidic solution, such as aqueous sulfuric acid, or gaseous carbon dioxide, to obtain the free acids. Composed primarily of tar acids such as phenol, cresols, and xylenols.] | 266-019-3 | 65996-85-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-117-00-1 | Tar acids, brown-coal, crude; Crude Phenols; [An acidified alkaline extract of brown coal tar distillate. Composed primarily of phenol and phenol homologs.] | 309-888-7 | 101316-86-3 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-118-00-7 | Tar acids, brown-coal gasification; Crude Phenols; [A complex combination of organic compounds obtained from brown coal gasification. Composed primarily of C6-10 hydroxy aromatic phenols and their homologs.] | 295-536-7 | 92062-22-1 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-119-00-2 | Tar acids, distn. residues; Distillate Phenols; [A residue from the distillation of crude phenol from coal. It consists predominantly of phenols having carbon numbers in the range of C8 through C10 with a softening point of 60°C to 80°C (140°F to 176°F).] | 306-251-5 | 96690-55-0 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-120-00-8 | Tar acids, methylphenol fraction; Distillate Phenols; [The fraction of tar acid rich in 3- and 4-methylphenol, recovered by distillation of low-temperature coal tar crude tar acids.] | 284-892-9 | 84989-04-8 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-121-00-3 | Tar acids, polyalkylphenol fraction; Distillate Phenols; [The fraction of tar acids, recovered by distillation of low-temperature coal tar crude tar acids, having an approximate boiling range of 225°C to 320°C (437°F to 608°F). Composed primarily of polyalkylphenols.] | 284-893-4 | 84989-05-9 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-122-00-9 | Tar acids, xylenol fraction; Distillate Phenols; [The fraction of tar acids, rich in 2,4- and 2,5-dimethylphenol, recovered by distillation of low-temperature coal tar crude tar acids.] | 284-895-5 | 84989-06-0 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-123-00-4 | Tar acids, ethylphenol fraction; Distillate Phenols; [The fraction of tar acids, rich in 3- and 4-ethylphenol, recovered by distillation of low-temperature coal tar crude tar acids.] | 284-891-3 | 84989-03-7 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-124-00-X | Tar acids, 3,5-xylenol fraction; Distillate Phenols; [The fraction of tar acids, rich in 3,5-dimethylphenol, recovered by distillation of low-temperature coal tar acids.] | 284-896-0 | 84989-07-1 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-125-00-5 | Tar acids, residues, distillates, first-cut; Distillate Phenols; [The residue from the distillation in the range of 235°C to 355°C (481°F to 697°F) of light carbolic oil.] | 270-713-1 | 68477-23-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-126-00-0 | Tar acids, cresylic, residues; Distillate Phenols; [The residue from crude coal tar acids after removal of phenol, cresols, xylenols and any higher boiling phenols. A black solid with a melting point approximately 80°C (176°F). Composed primarily of polyalkylphenols, resin gums, and inorganic salts.] | 271-418-0 | 68555-24-8 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-127-00-6 | Phenols, C9-11; Distillate Phenols | 293-435-2 | 91079-47-9 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-128-00-1 | Tar acids, cresylic; Distillate Phenols; [A complex combination of organic compounds obtained from brown coal and boiling in the range of approximately 200°C to 230°C (392°F to 446°F). It contains chiefly phenols and pyridine bases.] | 295-540-9 | 92062-26-5 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-129-00-7 | Tar acids, brown-coal, C2-alkylphenol fraction; Distillate Phenols; [The distillate from the acidification of alkaline washed lignite tar distillate boiling in the range of approximately 200°C to 230°C (392°F to 446°F). Composed primarily of m- and p-ethylphenol as well as cresols and xylenols.] | 302-662-9 | 94114-29-1 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-130-00-2 | Extract oils (coal), naphthalene oils; Acid Extract; [The aqueous extract produced by an acidic wash of alkali-washed naphthalene oil. Composed primarily of acid salts of various aromatic nitrogen bases including pyridine, quinoline and their alkyl derivatives.] | 292-623-1 | 90641-00-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-131-00-8 | Tar bases, quinoline derivs.; Distillate Bases | 271-020-7 | 68513-87-1 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-132-00-3 | Tar bases, coal, quinoline derivs. fraction; Distillate Bases | 274-560-1 | 70321-67-4 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-133-00-9 | Tar bases, coal, distn. residues; Distillate Bases; [The distillation residue remaining after the distillation of the neutralized, acid-extracted base-containing tar fractions obtained by the distillation of coal tars. It contains chiefly aniline, collidines, quinoline and quinoline derivatives and toluidines.] | 295-544-0 | 92062-29-8 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-134-00-4 | Hydrocarbon oils, arom., mixed with polyethylene and polypropylene, pyrolyzed, light oil fraction; Heat Treatment Products; [The oil obtained from the heat treatment of a polyethylene/polypropylene reaction mass with coal tar pitch or aromatic oils. It consists predominantly of benzene and its homologs boiling in a range of approximately 70°C to 120°C (158°F to 248°F).] | 309-745-9 | 100801-63-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-135-00-X | Hydrocarbon oils, arom., mixed with polyethylene, pyrolyzed, light oil fraction; Heat Treatment Products; [The oil obtained from the heat treatment of polyethylene with coal tar pitch or aromatic oils. It consists predominantly of benzene and its homologs boiling in a range of 70°C to 120°C (158°F to 248°F).] | 309-748-5 | 100801-65-8 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-136-00-5 | Hydrocarbon oils, arom., mixed with polystyrene, pyrolyzed, light oil fraction; Heat Treatment Products; [The oil obtained from the heat treatment of polystyrene with coal tar pitch or aromatic oils. It consists predominantly of benzene and its homologs boiling in a range of approximately 70°C to 210°C (158°F to 410°F).] | 309-749-0 | 100801-66-9 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-137-00-0 | Extract residues (coal), tar oil alk., naphthalene distn. residues; Naphthalene Oil Extract Residue; [The residue obtained from chemical oil extracted after the removal of naphthalene by distillation composed primarily of two to four membered condensed ring aromatic hydrocarbons and aromatic nitrogen bases.] | 277-567-8 | 73665-18-6 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-138-00-6 | Creosote oil, low-boiling distillate; Wash Oil; [The low-boiling distillation fraction obtained from the high temperature carbonization of bituminous coal, which is further refined to remove excess crystalline salts. It consists primarily of creosote oil with some of the normal polynuclear aromatic salts, which are components of coal tar distillate, removed. It is crystal free at approximately 38°C (100°F).] | 274-566-4 | 70321-80-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00/ATP01
648-139-00-1 | Tar acids, cresylic, sodium salts, caustic solns.; Alkaline Extract | 272-361-4 | 68815-21-4 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-140-00-7 | Extract oils (coal), tar base; Acid Extract; [The extract from coal tar oil alkaline extract residue produced by an acidic wash such as aqueous sulfuric acid after distillation to remove naphthalene. Composed primarily of the acid salts of various aromatic nitrogen bases including pyridine, quinoline, and their alkyl derivatives.] | 266-020-9 | 65996-86-3 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-141-00-2 | Tar bases, coal, crude; Crude Tar Bases; [The reaction product obtained by neutralizing coal tar base extract oil with an alkaline solution, such as aqueous sodium hydroxide, to obtain the free bases. Composed primarily of such organic bases as acridine, phenanthridine, pyridine, quinoline and their alkyl derivatives.] | 266-018-8 | 65996-84-1 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J M | CLP00/ATP01
648-142-00-8 | Residues (coal), liq. solvent extn.; [A cohesive powder composed of coal mineral matter and undissolved coal remaining after extraction of coal by a liquid solvent.] | 302-681-2 | 94114-46-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-143-00-3 | Coal liquids, liq. solvent extn. soln.; [The product obtained by filtration of coal mineral matter and undissolved coal from coal extract solution produced by digesting coal in a liquid solvent. A black, viscous, highly complex liquid combination composed primarily of aromatic and partly hydro-genated aromatic hydrocarbons, aromatic nitrogen compounds, aromatic sulfur compounds, phenolic and other aromatic oxygen compounds and their alkyl derivatives.] | 302-682-8 | 94114-47-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-144-00-9 | Coal liquids, liq. solvent extn.; [The substantially solvent-free product obtained by the distillation of the solvent from filtered coal extract solution produced by digesting coal in a liquid solvent. A black semi-solid, composed primarily of a complex combination of condensed-ring aromatic hydrocarbons, aromatic nitrogen compounds, aromatic sulfur compounds, phenolic compounds and other aromatic oxygen compounds, and their alkyl derivatives.] | 302-683-3 | 94114-48-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | M | CLP00
648-145-00-4 | Tar brown-coal; [An oil distilled from brown-coal tar. Composed primarily of aliphatic, naphthenic and one- to three-ring aromatic hydrocarbons, their alkyl derivates, heteroaromatics and one- and two-ring phenols boiling in the range of approximately 150 °C to 360 °C (302 °F to 680 °F).] | 309-885-0 | 101316-83-0 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
648-146-00-X | Tar, brown-coal, low-temp.; [A tar obtained from low temperature carbonization and low temperature gasification of brown coal. Composed primarily of aliphatic, naphthenic and cyclic aromatic hydrocarbons, heteroaromatic hydrocarbons and cyclic phenols.] | 309-886-6 | 101316-84-1 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
648-147-00-5 | Light oil (coal), coke-oven; Crude benzole; [The volatile organic liquid extracted from the gas evolved in the high temperature (greater than 700°C (1292°F)) destructive distillation of coal. Composed primarily of benzene, toluene, and xylenes. May contain other minor hydrocarbon constituents.] | 266-012-5 | 65996-78-3 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-148-00-0 | Distillates (coal), liq. solvent extn., primary; [The liquid product of condensation of vapors emitted during the digestion of coal in a liquid solvent and boiling in the range of approximately 30°C to 300°C (86°F to 572°F). Composed primarily of partly hydrogenated condensed-ring aromatic hydrocarbons, aromatic compounds containing nitrogen, oxygen and sulfur, and their alkyl derivatives having carbon numbers predominantly in the range of C4 through C14.] | 302-688-0 | 94114-52-0 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-149-00-6 | Distillates (coal), solvent extn., hydrocracked; [Distillate obtained by hydrocracking of coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 30°C to 300°C (86°F to 572°F). Composed primarily of aromatic, hydrogenated aromatic and naphthenic compounds, their alkyl derivatives and alkanes with carbon numbers predominantly in the range of C4 through C14. Nitrogen, sulfur and oxygen-containing aromatic and hydrogenated aromatic compounds are also present.] | 302-689-6 | 94114-53-1 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-150-00-1 | Naphtha (coal), solvent extn., hydrocracked; [Fraction of the distillate obtained by hydrocracking of coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 30°C to 180°C (86°F to 356°F). Composed primarily of aromatic, hydrogenated aromatic and naphthenic compounds, their alkyl derivatives and alkanes with carbon numbers predominantly in the range of C4 to C9. Nitrogen, sulfur and oxygen-containing aromatic and hydrogenated aromatic compounds are also present.] | 302-690-1 | 94114-54-2 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-151-00-7 | Gasoline, coal solvent extn., hydrocracked naphtha; [Motor fuel produced by the reforming of the refined naphtha fraction of the products of hydrocracking of coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 30 °C to 180 °C (86 °F to 356 °F). Composed primarily of aromatic and naphthenic hydrocarbons, their alkyl derivatives and alkyl hydrocarbons having carbon numbers in the range of C4 through C9.] | 302-691-7 | 94114-55-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
648-152-00-2 | Distillates (coal), solvent extn., hydrocracked middle; [Distillate obtained from the hydrocracking of coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 180°C to 300°C (356°F to 572°F. Composed primarily of two-ring aromatic, hydrogenated aromatic and naphthenic compounds, their alkyl derivatives and alkanes having carbon numbers predominantly in the range of C9 through C14. Nitrogen, sulfur and oxygen-containing compounds are also present.] | 302-692-2 | 94114-56-4 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-153-00-8 | Distillates (coal), solvent extn., hydrocracked hydrogenated middle; [Distillate from the hydrogenation of hydrocracked middle distillate from coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 180°C to 280°C (356°F to 536°F). Composed primarily of hydrogenated two- ring carbon compounds and their alkyl derivatives having carbon numbers predominantly in the range of C9 through C14.] | 302-693-8 | 94114-57-5 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
648-154-00-3 | Fuels, jet aircraft, coal solvent extn., hydrocracked hydrogenated; [Jet engine fuel produced by hydrogenation of the middle distillate fraction of the products of hydrocracking of coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 180 °C to 225 °C (356 °F to 473 °F). Composed primarily of hydrogenated two-ring hydrocarbons and their alkyl derivatives having carbon numbers predominantly in the range of C10 through C12.] | 302-694-3 | 94114-58-6 | Carc. 2 | H351 | GHS08 Wng | H350 |  |  |  | CLP00
648-155-00-9 | Fuels, diesel, coal solvent extn., hydrocracked hydrogenated; [Diesel engine fuel produced by the hydrogenation of the middle distillate fraction of the products of hydrocracking of coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 200 °C to 280 °C (392 °F to 536 °F). Composed primarily of hydrogenated two-ring hydrocarbons and their alkyl derivatives having carbon numbers predominantly in the range of C11 through C14.] | 302-695-9 | 94114-59-7 | Carc. 2 | H351 | GHS08 Wng | H350 |  |  |  | CLP00
648-156-00-4 | Light oil (coal), semi-coking process; Fresh oil; [The volatile organic liquid condensed from the gas evolved in the low-temperature (less than 700°C (1292°F)) destructive distillation of coal. Composed primarily of C6-10 hydrocarbons.] | 292-635-7 | 90641-11-5 | Carc. 1B Muta. 1B | H350 H340 | GHS08 Dgr | H350 H340 |  |  | J | CLP00/ATP01
649-001-00-3 | Extracts (petroleum), light naphthenic distillate solvent | 265-102-1 | 64742-03-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-002-00-9 | Extracts (petroleum), heavy paraffinic distillate solvent | 265-103-7 | 64742-04-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-003-00-4 | Extracts (petroleum), light paraffinic distillate solvent | 265-104-2 | 64742-05-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-004-00-X | Extracts (petroleum), heavy naphthenic distillate solvent | 265-111-0 | 64742-11-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-005-00-5 | Extracts (petroleum), light vacuum gas oil solvent | 295-341-7 | 91995-78-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-006-00-0 | hydrocarbons C26-55, arom-rich | 307-753-7 | 97722-04-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-007-00-6 | fatty acids, tall-oil, reaction products with iminodiethanol and boric acid | 400-160-5 |  | Skin Irrit. 2 Aquatic Chronic 2 | H315 H411 | GHS07 GHS09 Wng | H315 H411 |  |  |  | CLP00
649-008-00-1 | Residues (petroleum), atm. tower; Heavy Fuel oil; [A complex residuum from the atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly greater than C20 and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 265-045-2 | 64741-45-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-009-00-7 | Gas oils (petroleum), heavy vacuum; Heavy Fuel oil; [A complex combination of hydrocarbons produced by the vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and boiling in the range of approximately 350 °C to 600 °C (662 °F to 1112 °F). This stream is likely to contain 5 wt. % or more of 4-to 6-membered condensed ring aromatic hydrocarbons.] | 265-058-3 | 64741-57-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-010-00-2 | Distillates (petroleum), heavy catalytic cracked; Heavy Fuel oil; [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C35 and boiling in the range of approximately 260 °C to 500 °C (500 °F to 932 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 265-063-0 | 64741-61-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-011-00-8 | Clarified oils (petroleum), catalytic cracked; Heavy Fuel oil; [A complex combination of hydrocarbons produced as the residual fraction from distillation of the products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly greater than C20 and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 265-064-6 | 64741-62-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-012-00-3 | Residues (petroleum), hydrocracked; Heavy Fuel oil; [A complex combination of hydrocarbons produced as the residual fraction from distillation of the products of a hydrocracking process. It consists of hydrocarbons having carbon numbers predominantly greater than C20 and boiling above approximately 350 °C (662 °F).] | 265-076-1 | 64741-75-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-013-00-9 | Residues (petroleum), thermal cracked; Heavy Fuel oil; [A complex combination of hydrocarbons produced as the residual fraction from distillation of the product from a thermal cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly greater than C20 and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 265-081-9 | 64741-80-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-014-00-4 | Distillates (petroleum), heavy thermal cracked; Heavy Fuel oil; [A complex combination of hydrocarbons from the distillation of the products from a thermal cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C15 through C36 and boiling in the range of approximately 260 °C to 480 °C (500 °F to 896 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 265-082-4 | 64741-81-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-015-00-X | Gas oils (petroleum), hydrotreated vacuum; Heavy Fuel oil; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C13 through C50 and boiling in the range of approximately 230 °C to 600 °C (446 °F to 1112 °F). This stream is likely to contain 5 wt.% or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 265-162-9 | 64742-59-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-016-00-5 | Residues (petroleum), hydrodesulfurized atmospheric tower; Heavy Fuel oil; [A complex combination of hydrocarbons obtained by treating an atmospheric tower residuum with hydrogen in the presence of a catalyst under conditions primarily to remove organic sulfur compounds. It consists of hydrocarbons having carbon numbers predominantly greater than C20 and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 265-181-2 | 64742-78-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-017-00-0 | Gas oils (petroleum), hydrodesulfurized heavy vacuum; Heavy Fuel oil; [A complex combination of hydrocarbons obtained from a catalytic hydrodesulfurization process. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and boiling in the range of approximately 350 °C to 600 °C (662 °F to 1112 °C). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 265-189-6 | 64742-86-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-018-00-6 | Residues (petroleum), steam-cracked; Heavy Fuel oil; [A complex combination of hydrocarbons obtained as the residual fraction from the distillation of the products of a steam cracking process (including steam cracking to produce ethylene). It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly greater than C14 and boiling above approximately 260 °C (500 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 265-193-8 | 64742-90-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-019-00-1 | Residues (petroleum), atmospheric; Heavy Fuel oil; [A complex residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly greater than C11 and boiling above approximately 200 °C (392 °F). This stream is likely to contain 5 wt. % or more of 4-to 6-membered condensed ring aromatic hydrocarbons.] | 269-777-3 | 68333-22-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-020-00-7 | Clarified oils (petroleum), hydrodesulfurized catalytic cracked; Heavy Fuel oil; [A complex combination of hydrocarbons obtained by treating catalytic cracked clarified oil with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists of hydrocarbons having carbon numbers predominantly greater than C20 and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt. % or more of 4-to 6-membered condensed ring aromatic hydrocarbons.] | 269-782-0 | 68333-26-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-021-00-2 | Distillates (petroleum), hydrodesulfurized intermediate catalytic cracked; Heavy Fuel oil; [A complex combination of hydrocarbons obtained by treating intermediate catalytic cracked distillates with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists of hydrocarbons having carbon numbers predominantly in the range of C11 through C30 and boiling in the range of approximately 205 °C to 450 °C (401 °F to 842 °F). It contains a relatively large proportion of tricyclic aromatic hydrocarbons.] | 269-783-6 | 68333-27-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-022-00-8 | Distillates (petroleum), hydrodesulfurized heavy catalytic cracked; Heavy Fuel oil; [A complex combination of hydrocarbons obtained by treatment of heavy catalytic cracked distillates with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C35 and boiling in the range of approximately 260 °C to 500 °C (500 °F to 932 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 269-784-1 | 68333-28-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-023-00-3 | Fuel oil, residues-straight-run gas oils, high-sulfur; Heavy Fuel oil | 270-674-0 | 68476-32-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-024-00-9 | Fuel oil, residual; Heavy Fuel oil; [The liquid product from various refinery streams, usually residues. The composition is complex and varies with the source of the crude oil.] | 270-675-6 | 68476-33-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-025-00-4 | Residues (petroleum), catalytic reformer fractionator residue distn.; Heavy Fuel oil; [A complex residuum from the distillation of catalytic reformer fractionator residue. It boils approximately above 399 °C (750 °F).] | 270-792-2 | 68478-13-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-026-00-X | Residues (petroleum), heavy coker gas oil and vacuum gas oil; Heavy Fuel oil; [A complex combination of hydrocarbons produced as the residual fraction from the distillation of heavy coker gas oil and vacuum gas oil. It predominantly consists of hydrocarbons having carbon numbers predominantly greater than C13 and boiling above approximately 230 °C (446 °F).] | 270-796-4 | 68478-17-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-027-00-5 | Residues (petroleum), heavy coker and light vacuum; Heavy Fuel oil; [A complex combination of hydrocarbons produced as the residual fraction from the distillation of heavy coker gas oil and light vacuum gas oil. It consists predominantly of hydrocarbons having carbon numbers predominantly greater than C13 and boiling above approximately 230 °C (446 °F).] | 270-983-0 | 68512-61-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-028-00-0 | Residues (petroleum), light vacuum; Heavy Fuel oil; [A complex residuum from the vacuum distillation of the residuum from the atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly greater than C13 and boiling above approximately 230 °C (446 °F).] | 270-984-6 | 68512-62-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-029-00-6 | Residues (petroleum), steam-cracked light; Heavy Fuel oil; [A complex residuum from the distillation of the products from a steam-cracking process. It consists predominantly of aromatic and unsaturated hydrocarbons having carbon numbers greater than C7 and boiling in the range of approximately 101 °C to 555 °C (214 °F to 1030 °F).] | 271-013-9 | 68513-69-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-030-00-1 | Fuel oil, No 6; Heavy Fuel oil; [A distillate oil having a minimum viscosity of 900 SUS at 37.7 °C (100 °F) to a maximum of 9000 SUS at 37.7 °C (100 °F).] | 271-384-7 | 68553-00-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-031-00-7 | Residues (petroleum), topping plant, low-sulfur; Heavy Fuel oil; [A low-sulfur complex combination of hydrocarbons produced as the residual fraction from the topping plant distillation of crude oil. It is the residuum after the straight-run gasoline cut, kerosene cut and gas oil cut have been removed.] | 271-763-7 | 68607-30-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-032-00-2 | Gas oils (petroleum), heavy atmospheric; Heavy Fuel oil; [A complex combination of hydrocarbons obtained by the distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C7 through C35 and boiling in the range of approximately 121 °C to 510 °C (250 °F to 950 °F).] | 272-184-2 | 68783-08-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-033-00-8 | Residues (petroleum), coker scrubber, Condensed-ring-arom.-contg.; Heavy Fuel oil; [A very complex combination of hydrocarbons produced as the residual fraction from the distillation of vaccum residuum and the products from a thermal cracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly greater than C20 and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt.% or more of 4- to 6-membered condensed rind aromatic hydrocarbons.] | 272-187-9 | 68783-13-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-034-00-3 | Distillates (petroleum), petroleum residues vacuum; Heavy Fuel oil; [A complex combination of hydrocarbons produced by the vacuum distillation of the residuum from the atmospheric distillation of crude oil.] | 273-263-4 | 68955-27-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-035-00-9 | Residues (petroleum), steam-cracked, resinous; Heavy Fuel oil; [A complex residuum from the distillation of steam-cracked petroleum residues.] | 273-272-3 | 68955-36-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-036-00-4 | Distillates (petroleum), intermediate vacuum; Heavy Fuel oil; [A complex combination of hydrocarbons produced by the vacuum, distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C14 through C42 and boiling in the range of approximately 250 °C to 545 °C (482 °F to 1013 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 274-683-0 | 70592-76-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-037-00-X | Distillates (petroleum), light vacuum; Heavy Fuel oil; [A complex combination of hydrocarbons produced by the vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C11 through C35 and boiling in the range of approximately 250 °C to 545 °C (482 °F to 1013 °F).] | 274-684-6 | 70592-77-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-038-00-5 | Distillates (petroleum), vacuum; Heavy Fuel oil; [A complex combination of hydrocarbons produced by the vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having numbers predominantly in the range of C15 through C50 and boiling in the range of approximately 270 °C to 600 °C (518 °F to 1112 °F). This stream is likely to contain 5 wt.% or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 274-685-1 | 70592-78-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-039-00-0 | Gas oils (petroleum), hydrodesulfurized coker heavy vacuum; Heavy Fuel oil; [A complex combination of hydrocarbons obtained by hydrodesulfurization of heavy coker distillate stocks, It consists predominantly of hydrocarbons having carbon numbers predominantly in the range C18 to C44 and boiling in the range of approximately 304 °C to 548 °C (579 °F to 1018 °F). Likely to contain 5 % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 285-555-9 | 85117-03-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-040-00-6 | Residues (petroleum), steam-cracked, distillates; Heavy Fuel oil; [A complex combination of hydrocarbons obtained during the production of refined petroleum tar by the distillation of steam cracked tar. It consists predominantly of aromatic and other hydrocarbons and organic sulfur compounds.] | 292-657-7 | 90669-75-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-041-00-1 | Residues (petroleum), vacuum, light; Heavy Fuel oil; [A complex residuum from the vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists predominantly of hydrocarbons having carbon numbers predominantly greater than C24 and boiling above approximately 390 °C (734 °F).] | 292-658-2 | 90669-76-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-042-00-7 | Fuel oil, heavy, high-sulfur; Heavy Fuel oil; [A complex combination of hydrocarbons obtained by the distillation of crude petroleum. It consists predominantly of aliphatic, aromatic and cycloaliphatic hydrocarbons having carbon numbers predominantly higher than C25 and boiling above approximately 400 °C (752 °F).] | 295-396-7 | 92045-14-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-043-00-2 | Residues (petroleum), catalytic cracking; Heavy Fuel oil; [A complex combination of hydrocarbons produced as the residual fraction from the distillation of the products from a catalytic cracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly greater than C11 and boiling above approximately 200 °C (392 °F).] | 295-511-0 | 92061-97-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-044-00-8 | Distillates (petroleum), intermediate catalytic cracked, thermally degraded; Heavy Fuel oil; [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process which has been used as a heat transfer fluid. It consists predominantly of hydrocarbons boiling in the range of approximately 220 °C to 450 °C (428 °F to 842 °F). This stream is likely to contain organic sulfur compounds.] | 295-990-6 | 92201-59-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-045-00-3 | Residual oils (petroleum); Heavy Fuel oil; [A complex combination of hydrocarbons, sulfur compounds and metal-containing organic compounds obtained as the residue from refinery fractionation cracking processes. It produces a finished oil with a viscosity above 2cSt. at 100 °C.] | 298-754-0 | 93821-66-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-046-00-9 | Residues, steam cracked, thermally treated; Heavy Fuel oil; [A complex combination of hydrocarbons obtained by the treatment and distillation of raw steam-cracked naphtha. It consists predominantly of unsaturated hydrocarbons boiling in the range above approximately 180 °C (356 °F).] | 308-733-0 | 98219-64-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-047-00-4 | Distillates (petroleum), hydrodesulfurized full-range middle; Heavy Fuel oil; [A complex combination of hydrocarbons obtained by treating a petroleum stock with hydrogen. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C9 through C25 and boiling in the range of approximately 150 °C to 400 °C (302 °F to 752 °F).] | 309-863-0 | 101316-57-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-048-00-X | Residues (petroleum), catalytic reformer fractionator; Heavy Fuel oil; [A complex combination of hydrocarbons produced as the residual fraction from distillation of the product from a catalytic reforming process. It consists of predominantly aromatic hydrocarbons having carbon numbers predominantly in the range of C10 through C25 and boiling in the range of approximately 160 °C to 400 °C (320 °F to 725 °F). This stream is likely to contain 5 wt. % or more of 4- or 6-membered condensed ring aromatic hydrocarbons.] | 265-069-3 | 64741-67-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-049-00-5 | Petroleum; Crude oil; [A complex combination of hydrocarbons, It consists predominantly of aliphatic, alicyclic and aromatic hydrocarbons. It may also contain small amounts of nitrogen, oxygen and sulfur compounds. This category encompasses light, medium, and heavy petroleums, as well as the oils extended from tar sands. Hydrocarbonaceous materials requiring major chemical changes for their recovery or conversion to petroleum refinery feedstocks such as crude shale oils; upgraded shale oils and liquid coal fuels are not included in this definition.] | 232-298-5 | 8002-05-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-050-00-0 | Distillates (petroleum), light paraffinic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons produced by vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated aliphatic hydrocarbons normally present in this distillation range of crude oil.] | 265-051-5 | 64741-50-0 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-051-00-6 | Distillates (petroleum), heavy paraffinic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons produced by vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated aliphatic hydrocarbons.] | 265-052-0 | 64741-51-1 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-052-00-1 | Distillates (petroleum), light naphthenic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons produced by vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-053-6 | 64741-52-2 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-053-00-7 | Distillates (petroleum), heavy naphthenic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons produced by vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-054-1 | 64741-53-3 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-054-00-2 | Distillates (petroleum), acid-treated heavy naphthenic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-117-3 | 64742-18-3 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-055-00-8 | Distillates (petroleum), acid-treated light naphthenic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-118-9 | 64742-19-4 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-056-00-3 | Distillates (petroleum), acid-treated heavy paraffinic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil having a viscosity of a least 100 SUS at 100 °F (19cSt at 40 °C).] | 265-119-4 | 64742-20-7 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-057-00-9 | Distillates (petroleum), acid-treated light paraffinic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil having a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C).] | 265-121-5 | 64742-21-8 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-058-00-4 | Distillates (petroleum), chemically neutralized heavy paraffinic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons obtained from a treating process to remove acidic materials. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of aliphatic hydrocarbons.] | 265-127-8 | 64742-27-4 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-059-00-X | Distillates (petroleum), chemically neutralized light paraffinic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity less than 100 SUS at 100 °F (19cSt at 40 °C).] | 265-128-3 | 64742-28-5 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-060-00-5 | Distillates (petroleum), chemically neutralized heavy naphthenic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-135-1 | 64742-34-3 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-061-00-0 | Distillates (petroleum), chemically neutralized light naphthenic; Unrefined or mildly refined baseoil; [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS a 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-136-7 | 64742-35-4 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-062-00-6 | Gases (petroleum), catalytic cracked naphtha depropanizer overhead, C3-rich acid-free; Petroleum gas; [A complex combination of hydrocarbons obtained from fractionation of catalytic cracked hydrocarbons and treated to remove acidic impurities. It consists of hydrocarbons having carbon numbers in the range of C2 through C4, predominantly C3.] | 270-755-0 | 68477-73-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-063-00-1 | Gases (petroleum), catalytic cracker; Petroleum gas; [A complex combination of hydrocarbons produced by the distillation of the products from a catalytic cracking process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 270-756-6 | 68477-74-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-064-00-7 | Gases (petroleum), catalytic cracker, C1-5-rich; Petroleum gas; [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of aliphatic hydrocarbons having carbon numbers in the range of C1 through C6, predominantly C1 through C5.] | 270-757-1 | 68477-75-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-065-00-2 | Gases (petroleum), catalytic polymd. naphtha stabilizer overhead, C2-4-rich; Petroleum gas; [A complex combination of hydrocarbons obtained from the fractionation stabilization of catalytic polymerized naphtha. It consists of aliphatic hydrocarbons having carbon numbers in the range of C2 through C6, predominantly C2 through C4.] | 270-758-7 | 68477-76-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-066-00-8 | Gases (petroleum), catalytic reformer, C1-4-rich; Petroleum gas; [A complex combination of hydrocarbons produced by distillation of products from a catalytic reforming process. It consists of hydrocarbons having carbon numbers in the range of C1 through C6, predominantly C1 through C4.] | 270-760-8 | 68477-79-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-067-00-3 | Gases (petroleum), C3-5 olefinic-paraffinic alkylation feed; Petroleum gas; [A complex combination of olefinic and paraffinic hydrocarbons having carbon numbers in the range of C3 through C5 which are used as alkylation feed. Ambient temperatures normally exceed the critical temperature of these combinations.] | 270-765-5 | 68477-83-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-068-00-9 | Gases (petroleum), C4-rich; Petroleum gas; [A complex combination of hydrocarbons produced by distillation of products from a catalytic fractionation process. It consists of aliphatic hydrocarbons having carbon numbers in the range of C3 through C5, predominantly C4.] | 270-767-6 | 68477-85-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-069-00-4 | Gases (petroleum), deethanizer overheads; Petroleum gas; [A complex combination of hydrocarbons produced from distillation of the gas and gasoline fractions from the catalytic cracking process. It contains predominantly ethane and ethylene.] | 270-768-1 | 68477-86-1 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-070-00-X | Gases (petroleum), deisobutanizer tower overheads; Petroleum gas; [A complex combination of hydrocarbons produced by the atmospheric distillation of a butane-butylene stream. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C3 through C4.] | 270-769-7 | 68477-87-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-071-00-5 | Gases (petroleum), depropanizer dry, propene-rich; Petroleum gas; [A complex combination of hydrocarbons produced by the distillation of products from the gas and gasoline fractions of a catalytic cracking process. It consists predominantly of propylene with some ethane and propane.] | 270-772-3 | 68477-90-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-072-00-0 | Gases (petroleum), depropanizer overheads; Petroleum gas; [A complex combination of hydrocarbons produced by distillation of products from the gas and gasoline fractions of a catalytic cracking process. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C4.] | 270-773-9 | 68477-91-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-073-00-6 | Gases (petroleum), gas recovery plant depropanizer overheads; Petroleum gas; [A complex combination of hydrocarbons obtained by fractionation of miscellaneous hydrocarbon streams. It consists predominantly of hydrocarbons having carbon numbers in the range of C1 through C4, predominantly propane.] | 270-777-0 | 68477-94-1 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-074-00-1 | Gases (petroleum), Girbotol unit feed; Petroleum gas; [A complex combination of hydrocarbons that is used as the feed into the Girbatol unit to remove hydrogen sulfide. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C4.] | 270-778-6 | 68477-95-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-075-00-7 | Gases (petroleum), isomerized naphtha fractionator, C4-rich, hydrogen sulfide-free; Petroleum gas | 270-782-8 | 68477-99-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-076-00-2 | Tail gas (petroleum), catalytic cracked clarified oil and thermal cracked vacuum residue fractionation reflux drum; Petroleum gas; [A complex combination of hydrocarbons obtained from fractionation of catalytic cracked clarified oil and thermal cracked vacuum residue. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 270-802-5 | 68478-21-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-077-00-8 | Tail gas (petroleum), catalytic cracked naphtha stabilization absorber; Petroleum gas; [A complex combination of hydrocarbons obtained from the stabilization of catalytic cracked naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 270-803-0 | 68478-22-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-078-00-3 | Tail gas (petroleum), catalytic cracker, catalytic reformer and hydrodesulfurizer combined fractionater; Petroleum gas; [A complex combination of hydrocarbons obtained from the fractionation of products from catalytic cracking, catalytic reforming and hydrodesulfurizing processes treated to remove acidic impurities. It consists predominantly of hydrocarbons having cabon numbers predominantly in the range of C1 through C5.] | 270-804-6 | 68478-24-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-079-00-9 | Tail gas (petroleum), catalytic reformed naphtha fractionation stabilizer; Petroleum gas; [A complex combination of hydrocarbons obtained from the fractionation stabilization of catalytic reformed naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 270-806-7 | 68478-26-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-080-00-4 | Tail gas (petroleum), saturate gas plant mixed stream, C4-rich; Petroleum gas; [A complex combination of hydrocarbons obtained from the fractionation stabilization of straight-run naphtha, distillation tail gas and catalytic reformed naphtha stabilizer tail gas. It consists of hydrocarbons having carbon numbers in the range of C3 through C6, predominantly butane and isobutane.] | 270-813-5 | 68478-32-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-081-00-X | Tail gas (petroleum), saturate gas recovery plant, C1-2-rich; Petroleum gas; [A complex combination of hydrocarbons obtained from fractionation of distillate tail gas, straight-run naphtha, catalytic reformed naphtha stabilizer tail gas. It consists predominantly of hydrocarbons having carbon numbers in the range of C1through C5, predominantly methane and ethane.] | 270-814-0 | 68478-33-1 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-082-00-5 | Tail gas (petroleum), vacuum residues thermal cracker; Petroleum gas; [A complex combination of hydrocarbons obtained from the thermal cracking of vacuum residues. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 270-815-6 | 68478-34-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-083-00-0 | Hydrocarbons, C3-4-rich, petroleum distillate; Petroleum gas; [A complex combination of hydrocarbons produced by distillation and condensation of crude oil. It consists of hydrocarbons having carbon numbers in the range of C3 through C5, predominantly C3 through C4.] | 270-990-9 | 68512-91-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-084-00-6 | Gases (petroleum), full-range straight-run naphtha dehexanizer off; petroleum gas; [A complex combination of hydrocarbons obtained by the fractionation of the full-range straight-run naphtha. It consists of hydrocarbons having carbon numbers predominantly in the range of C2 through C6.] | 271-000-8 | 68513-15-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-085-00-1 | Gases (petroleum), hydrocracking depropanizer off, hydrocarbon-rich; Petroleum gas; [A complex combination of hydrocarbon produced by the distillation of products from a hydrocracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4. It may also contain small amounts of hydrogen and hydrogen sulfide.] | 271-001-3 | 68513-16-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-086-00-7 | Gases (petroleum), light straight-run naphtha stabilizer off; Petroleum gas; [A complex combination of hydrocarbons obtained by the stabilization of light straight-run naphtha. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C6.] | 271-002-9 | 68513-17-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-087-00-2 | Residues (petroleum), alkylation splitter, C4-rich; Petroleum gas; [A complex residuum from the distillation of streams various refinery operations. It consists of hydrocarbons having carbon numbers in the range of C4 through C5, predominantly butane and boiling in the range of approximately – 11.7 °C to 27.8 °C (11 °F to 82 °F).] | 271-010-2 | 68513-66-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-088-00-8 | Hydrocarbons, C1-4; Petroleum gas; [A complex combination of hydrocarbons provided by thermal cracking and absorber operations and by distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C4 and boiling in the range of approximately minus 164 °C to minus 0.5 °C (– 263 °F to 31 °F).] | 271-032-2 | 68514-31-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-089-00-3 | Hydrocarbons, C1-4, sweetened; Petroleum gas; [A complex combination of hydrocarbons obtained by subjecting hydrocarbon gases to a sweetening process to convert mercaptans or to remove acidic impurities. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C4 and boiling in the range of approximately – 164 °C to – 0.5 °C (– 263 °F to 31 °F).] | 271-038-5 | 68514-36-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-090-00-9 | Hydrocarbons, C1-3; Petroleum gas; [A complex combination of hydrocarbons having carbon numbers predominantly in the range of C1 through C3 and boiling in the range of approximately minus 164 °C to minus 42 °C (– 263 °F to – 44 °F).] | 271-259-7 | 68527-16-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-091-00-4 | Hydrocarbons, C1-4, debutanizer fraction; Petroleum gas | 271-261-8 | 68527-19-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-092-00-X | Gases (petroleum), C1-5, wet; Petroleum gas; [A complex combination of hydrocarbons produced by the distillation of crude oil and/or the cracking of tower gas oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 271-624-0 | 68602-83-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-093-00-5 | Hydrocarbons, C2-4; Petroleum gas | 271-734-9 | 68606-25-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-094-00-0 | Hydrocarbons, C3; Petroleum gas | 271-735-4 | 68606-26-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-095-00-6 | Gases (petroleum), alkylation feed; Petroleum gas; [A complex combination of hydrocarbons produced by the catalytic cracking of gas oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C3 through C4.] | 271-737-5 | 68606-27-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-096-00-1 | Gases (petroleum), depropanizer bottoms fractionation off; Petroleum gas; [A complex combination of hydrocarbons obtained from the fractionation of depropanizer bottoms. It consists predominantly of butane, isobutane and butadiene.] | 271-742-2 | 68606-34-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-097-00-7 | Gases (petroleum), refinery blend; Petroleum gas; [A complex combination obtained from various processes. It consists of hydrogen, hydrogen sulfide and hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 272-183-7 | 68783-07-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-098-00-2 | Gases (petroleum), catalytic cracking; Petroleum gas; [A complex combination of hydrocarbons produced by the distillation of the products from a catalytic cracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C3 through C5.] | 272-203-4 | 68783-64-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-099-00-8 | Gases (petroleum), C2-4, sweetened; Petroleum gas; [A complex combination of hydrocarbons obtained by subjecting a petroleum distillate to a sweetening process to convert mercaptans or to remove acidic impurities. It consists predominantly of saturated and unsaturated hydrocarbons having carbon numbers predominantly in the range of C2 through C4 and boiling in the range of approximately – 51 °C to – 34 °C (– 60 °F to – 30 °F).] | 272-205-5 | 68783-65-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-100-00-1 | Gases (petroleum), crude oil fractionation off; Petroleum gas; [A complex combination of hydrocarbons produced by the fractionation of crude oil. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 272-871-7 | 68918-99-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-101-00-7 | Gases (petroleum), dehexanizer off; Petroleum gas; [A complex combination of hydrocarbons obtained by the fractionation of combined naphtha streams. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 272-872-2 | 68919-00-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-102-00-2 | Gases (petroleum), light straight run gasoline fractionation stabilizer off; Petroleum gas; [A complex combination of hydrocarbons obtained by the fractionation of light straight-run gasoline. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 272-878-5 | 68919-05-1 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-103-00-8 | Gases (petroleum), naphtha unifiner desulfurization stripper off; Petroleum gas; [A complex combination of hydrocarbons produced by a naphtha unifiner desulfurization process and stripped from the naphtha product. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 272-879-0 | 68919-06-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-104-00-3 | Gases (petroleum), straight-run naphtha catalytic reforming off; Petroleum gas; [A complex combination of hydrocarbons obtained by the catalytic reforming of straight-run naphtha and fractionation of the total effluent. It consists of methane, ethane, and propane.] | 272-882-7 | 68919-09-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-105-00-9 | Gases (petroleum), fluidized catalytic cracker splitter overheads; Petroleum gas; [A complex combination of hydrocarbons produced by the fractionation of the charge to the C3-C4 splitter. It consists predominantly of C3 hydrocarbons.] | 272-893-7 | 68919-20-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H350 H340 |  |  | K U | CLP00/ATP01corr
649-106-00-4 | Gases (petroleum), straight-run stabilizer off; Petroleum gas; [A complex combination of hydrocarbons obtained from the fractionation of the liquid from the first tower used in the distillation of crude oil. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 272-883-2 | 68919-10-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-107-00-X | Gases (petroleum), catalytic cracked naphtha debutanizer; Petroleum gas; [A complex combination of hydrocarbons obtained from fractionation of catalytic cracked naphtha. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 273-169-3 | 68952-76-1 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-108-00-5 | Tail gas (petroleum), catalytic cracked distillate and naphtha stabilizer; Petroleum gas; [A complex combination of hydrocarbons obtained by the fractionation of catalytic cracked naphtha and distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 273-170-9 | 68952-77-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H350 H340 |  |  | K U | CLP00/ATP01corr
649-109-00-0 | Tail gas (petroleum), thermal-cracked distillate, gas oil and naphtha absorber; petroleum gas; [A complex combination of hydrocarbons obtained from the separation of thermal-cracked distillates, naphtha and gas oil. It consists pedrominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 273-175-6 | 68952-81-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-110-00-6 | Tail gas (petroleum), thermal cracked hydrocarbon fractionation stabilizer, petroleum coking; Petroleum gas; [A complex combination of hydrocarbons obtained from the fractionation stabilization of thermal cracked hydrocarbons from petroleum coking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 273-176-1 | 68952-82-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-111-00-1 | Gases (petroleum, light steam-cracked, butadiene conc.; Petroleum gas; [A complex combination of hydrocarbons produced by the distillation of products from a thermal cracking process. It consists of hydrocarbons having a carbon number predominantly of C4.] | 273-265-5 | 68955-28-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-112-00-7 | Gases (petroleum), straight-run naphtha catalytic reformer stabilizer overhead; Petroleum gas; [A complex combination of hydrocarbons obtained by the catalytic reforming of straight-run naphtha and the fractionation of the total effluent. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C4.] | 273-270-2 | 68955-34-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-113-00-2 | Hydrocarbons, C4; Petroleum gas | 289-339-5 | 87741-01-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-114-00-8 | Alkanes, C1-4, C3-rich; Petroleum gas | 292-456-4 | 90622-55-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-115-00-3 | Gases (petroleum), steam-cracker C3-rich; Petroleum gas; [A complex combination of hydrocarbons produced by the distillation of products from a steam cracking process. It consists predominantly of propylene with some propane and boils in the range of approximately – 70 °C to 0 °C (– 94 °F to 32 °F).] | 295-404-9 | 92045-22-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-116-00-9 | Hydrocarbons, C4, steam-cracker distillate; Petroleum gas; [A complex combination of hydrocarbons produced by the distillation of the products of a steam cracking process. It consists predominantly of hydrocarbons having a carbon number of C4, predominantly 1-butene and 2-butene, containing also butane and isobutene and boiling in the range of approximately minus 12 °C to 5 °C (10.4 °F to 41 °F).] | 295-405-4 | 92045-23-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-117-00-4 | Petroleum gases, liquefied, sweetened, C4 fraction; Petroleum gas; [A complex combination of hydrocarbons obtained by subjecting a liquified petroleum gas mix to a sweetening process to oxidize mercaptans or to remove acidic impurities. It consists predominantly of C4 saturated and unsaturated hydrocarbons.] | 295-463-0 | 92045-80-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K S U | CLP00/ATP01corr
649-118-00-X | Hydrocarbons, C4, 1,3-butadiene- and isobutene-free; Petroleum gas | 306-004-1 | 95465-89-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-119-00-5 | Raffinates (petroleum), steam-cracked C4 fraction cuprous ammonium acetate extn., C3-5 and C3-5 unsatd., butadiene-free; Petroleum gas | 307-769-4 | 97722-19-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-120-00-0 | Gases (petroleum), amine system feed; Refinery gas; [The feed gas to the amine system for removal of hydrogen sulfide. It consists of hydrogen. Carbon monoxide, carbon dioxide, hydrogen sulfide and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5 may also be present.] | 270-746-1 | 68477-65-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-121-00-6 | Gases (petroleum), benzene unit hydrodesulfurizer off; Refinery gas; [Off gases produced by the benzene unit. It consists primarily of hydrogen. Carbon monoxide and hydrocarbons having carbon numbers predominantly in the range of C1 through C6, including benzene, may also be present.] | 270-747-7 | 68477-66-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-122-00-1 | Gases (petroleum), benzene unit recycle, hydrogen-rich; Refinery gas; [A complex combination of hydrocarbons obtained by recycling the gases of the benzene unit. It consists primarily of hydrogen with various small amounts of carbon monoxide and hydrocarbons having carbon numbers in the range of C1 through C6.] | 270-748-2 | 68477-67-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-123-00-7 | Gases (petroleum), blend oil, hydrogen-nitrogen-rich; Refinery gas; [A complex combination of hydrocarbons obtained by distillation of a blend oil. It consists primarily of hydrogen and nitrogen with various small amounts of carbon monoxide, carbon dioxide, and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 270-749-8 | 68477-68-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-124-00-2 | Gases (petroleum), catalytic reformed naphtha stripper overheads; Refinery gas; [A complex combination of hydrocarbons obtained from stabilization of catalytic reformed naphtha. Its consists of hydrogen and saturated hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 270-759-2 | 68477-77-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-125-00-8 | Gases (petroleum), C6-8 catalytic reformer recycle; Refinery gas; [A complex combination of hydrocarbons produced by distillation of products from catalytic reforming of C6-C8 feed and recycled to conserve hydrogen. It consists primarily of hydrogen. It may also contain various small amounts of carbon monoxide, carbon dioxide, nitrogen, and hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 270-761-3 | 68477-80-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-126-00-3 | Gases (petroleum), C6-8 catalytic reformer; Refinery gas; [A complex combination of hydrocarbons produced by distillation of products from catalytic reforming of C6-C8feed. It consists of hydrocarbons having carbon numbers in the range of C1 through C5 and hydrogen.] | 270-762-9 | 68477-81-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-127-00-9 | Gases (petroleum), C6-8 catalytic reformer recycle, hydrogen-rich; Refinery gas | 270-763-4 | 68477-82-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-128-00-4 | Gases (petroleum), C2-return stream; Refinery gas; [A complex combination of hydrocarbons obtained by the extraction of hydrogen from a gas stream which consists primarily of hydrogen with small amounts of nitrogen, carbon monoxide, methane, ethane, and ethylene. It contains predominantly hydrocarbons such as methane, ethane, and ethylene with small amounts of hydrogen, nitrogen and carbon monoxide.] | 270-766-0 | 68477-84-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H350 H340 |  |  | K U | CLP00/ATP01corr
649-129-00-X | Gases (petroleum), dry sour, gas-concn.-unit-off; Refinery gas; [The complex combination of dry gases from a gas concentration unit. It consists of hydrogen, hydrogen sulfide and hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] | 270-774-4 | 68477-92-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-130-00-5 | Gases (petroleum), gas concn. reabsorber distn.; Refinery gas; [A complex combination of hydrocarbons produced by distillation of products from combined gas streams in a gas concentration reabsorber. It consists predominantly of hydrogen, carbon monoxide, carbon dioxide, nitrogen, hydrogen sulfide and hydrocarbons having carbon numbers in the range of C1 through C3.] | 270-776-5 | 68477-93-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-131-00-0 | Gases (petroleum), hydrogen absorber off; Refinery gas; [A complex combination obtained by absorbing hydrogen from a hydrogen rich stream. It consists of hydrogen, carbon monoxide, nitrogen, and methane with small amounts of C2 hydrocarbons.] | 270-779-1 | 68477-96-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-132-00-6 | Gases (petroleum), hydrogen-rich; Refinery gas; [A complex combination separated as a gas from hydrocarbon gases by chilling. It consists primarily of hydrogen with various small amounts of carbon monoxide, nitrogen, methane, and C2 hydrocarbons.] | 270-780-7 | 68477-97-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-133-00-1 | Gases (petroleum), hydrotreater blend oil recycle, hydrogen-nitrogen-rich; Refinery gas; [A complex combination obtained from recycled hydrotreated blend oil. It consists primarily of hydrogen and nitrogen with various small amounts of carbon monoxide, carbon dioxide and hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 270-781-2 | 68477-98-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-134-00-7 | Gases (petroleum), recycle, hydrogen-rich; Refinery gas; [A complex combination obtained from recycled reactor gases. It consists primarily of hydrogen with various small amounts of carbon monoxide, carbon dioxide, nitrogen, hydrogen sulfide, and saturated aliphatic hydrocarbons having carbon numbers in the range of C1 through C5.] | 270-783-3 | 68478-00-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-135-00-2 | Gases (petroleum), reformer make-up, hydrogen-rich; Refinery gas; [A complex combination obtained from the reformers. It consists primarily of hydrogen with various small amounts of carbon monoxide and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 270-784-9 | 68478-01-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-136-00-8 | Gases (petroleum), reforming hydrotreater; Refinery gas; [A complex combination obtained from the reforming hydrotreating process. It consists primarily of hydrogen, methane, and ethane with various small amounts of hydrogen sulfide and aliphatic hydrocarbons having carbon numbers predominantly in the range of C3 through C5.] | 270-785-4 | 68478-02-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-137-00-3 | Gases (petroleum), reforming hydrotreater, hydrogen-methane-rich; Refinery gas; [A complex combination obtained from the reforming hydrotreating process. It consists primarily of hydrogen and methane with various small amounts of carbon monoxide, carbon dioxide, nitrogen and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C5.] | 270-787-5 | 68478-03-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-138-00-9 | Gases (petroleum), reforming hydrotreater make-up, hydrogen-rich; Refinery gas; [A complex combination obtained from the reforming hydrotreating process. It consists primarily of hydrogen with various small amounts of carbon monoxide and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 270-788-0 | 68478-04-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-139-00-4 | Gases (petroleum), thermal cracking distn.; Refinery gas; [A complex combination produced by distillation of products from a thermal cracking process. It consists of hydrogen, hydrogen sulfide, carbon monoxide, carbon dioxide and hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 270-789-6 | 68478-05-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-140-00-X | Tail gas (petroleum), catalytic cracker refractionation absorber; Refinery gas; [A complex combination of hydrocarbons obtained from refractionation of products from a catalytic cracking process. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] | 270-805-1 | 68478-25-1 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-141-00-5 | Tail gas (petroleum), catalytic reformed naphtha separator; Refinery gas; [A complex combination of hydrocarbons obtained from the catalytic reforming of straight run naphtha. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 270-807-2 | 68478-27-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-142-00-0 | Tail gas (petroleum), catalytic reformed naphtha stabilizer; Refinery gas; [A complex combination of hydrocarbons obtained from the stabilization of catalytic reformed naphtha. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 270-808-8 | 68478-28-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-143-00-6 | Tail gas (petroleum), cracked distillate hydrotreater separator; Refinery gas; [A complex combination of hydrocarbons obtained by treating cracked distillates with hydrogen in the presence of a catalyst. It consists of hydrogen and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 270-809-3 | 68478-29-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-144-00-1 | Tail gas (petroleum), hydrodesulfurized straight-run naphtha separator; Refinery gas; [A complex combination of hydrocarbons obtained from hydrodesulfurization of straight-run naphtha. It consists of hydrogen and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 270-810-9 | 68478-30-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-145-00-7 | Gases (petroleum), catalytic reformed straight-run naphtha stabilizer overheads; Refinery gas; [A complex combination of hydrocarbons obtained from the catalytic reforming of straight-run naphtha followed by fractionation of the total effluent. It consists of hydrogen, methane, ethane and propane.] | 270-999-8 | 68513-14-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-146-00-2 | Gases (petroleum), reformer effluent high-pressure flash drum off; Refinery gas; [A complex combination produced by the high-pressure flashing of the effluent from the reforming reactor. It consists primarily of hydrogen with various small amounts of methane, ethane, and propane.] | 271-003-4 | 68513-18-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-147-00-8 | Gases (petroleum), reformer effluent low-pressure flash drum off; Refinery gas; [A complex combination produced by low-pressure flashing of the effluent from the reforming reactor. It consists primarily of hydrogen with various small amounts of methane, ethane, and propane.] | 271-005-5 | 68513-19-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-148-00-3 | Gases (petroleum), oil refinery gas distn. off; Refinery gas; [A complex combination separated by distillation of a gas stream containing hydrogen, carbon monoxide, carbon dioxide and hydrocarbons having carbon numbers in the range of C1 through C6 or obtained by cracking ethane and propane. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C2, hydrogen, nitrogen, and carbon monoxide.] | 271-258-1 | 68527-15-1 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-149-00-9 | Gases (petroleum), benzene unit hydrotreater depentanizer overheads; Refinery gas; [A complex combination produced by treating the feed from the benzene unit with hydrogen in the presence of a catalyst followed by depentanizing. It consists primarily of hydrogen, ethane and propane with various small amounts of nitrogen, carbon monoxide, carbon dioxide and hydrocarbons having carbon numbers predominantly in the range of C1 through C6. It may contain trace amounts of benzene.] | 271-623-5 | 68602-82-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-150-00-4 | Gases (petroleum), secondary absorber off, fluidized catalytic cracker overheads fractionator; Refinery gas; [A complex combination produced by the fractionation of the overhead products from the catalytic cracking process in the fluidized catalytic cracker. It consists of hydrogen, nitrogen, and hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] | 271-625-6 | 68602-84-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-151-00-X | Petroleum products, refinery gases; Refinery gas; [A complex combination which consists primarily of hydrogen with various small amounts of methane, ethane, and propane.] | 271-750-6 | 68607-11-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-152-00-5 | Gases (petroleum), hydrocracking low-pressure separator; Refinery gas; [A complex combination obtained by the liquid-vapor separation of the hydrocracking process reactor effluent. It consists predominantly of hydrogen and saturated hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] | 272-182-1 | 68783-06-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-153-00-0 | Gases (petroleum), refinery; Refinery gas; [A complex combination obtained from various petroleum refining operations. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] | 272-338-9 | 68814-67-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-154-00-6 | Gases (petroleum), platformer products separator off; Refinery gas; [A complex combination obtained from the chemical reforming of naphthenes to aromatics. It consists of hydrogen and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C4.] | 272-343-6 | 68814-90-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-155-00-1 | Gases (petroleum), hydrotreated sour kerosine depentanizer stabilizer off; Refinery gas; [The complex combination obtained from the depentanizer stabilization of hydrotreated kerosine. It consists primarily of hydrogen, methane, ethane, and propane with various small amounts of nitrogen, hydrogen sulfide, carbon monoxide and hydrocarbons having carbon numbers predominantly in the range of C4 through C5.] | 272-775-5 | 68911-58-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-156-00-7 | Gases (petroleum), hydrotreated sour kerosine flash drum; Refinery gas; [A complex combination obtained from the flash drum of the unit treating sour kerosine with hydrogen in the presence of a catalyst. It consists primarily of hydrogen and methane with various small amounts of nitrogen, carbon monoxide, and hydrocarbons having carbon numbers predominantly in the range of C2 through C5.] | 272-776-0 | 68911-59-1 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-157-00-2 | Gases (petroleum), distillate unifiner desulfurization stripper off; Refinery gas; [A complex combination stripped from the liquid product of the unifiner desulfurization process. It consists of hydrogen sulfide, methane, ethane, and propane.] | 272-873-8 | 68919-01-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-158-00-8 | Gases (petroleum), fluidized catalytic cracker fractionation off; Refinery gas; [A complex combination produced by the fractionation of the overhead product of the fluidized catalytic cracking process. It consists of hydrogen, hydrogen sulfide, nitrogen, and hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 272-874-3 | 68919-02-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-159-00-3 | Gases (petroleum), fluidized catalytic cracker scrubbing secondary absorber off; Refinery gas; [A complex combination produced by scrubbing the overhead gas from the fluidized catalytic cracker. It consists of hydrogen, nitrogen, methane, ethane and propane.] | 272-875-9 | 68919-03-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-160-00-9 | Gases (petroleum), heavy distillate hydrotreater desulfurization stripper off; Refinery gas; [A complex combination stripped from the liquid product of the heavy distillate hydrotreater desulfurization process. It consists of hydrogen, hydrogen sulfide, and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 272-876-4 | 68919-04-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-161-00-4 | Gases (petroleum), platformer stabilizer off, light ends fractionation; Refinery gas; [A complex combination obtained by the fractionation of the light ends of the platinum reactors of the platformer unit. It consists of hydrogen, methane, ethane and propane.] | 272-880-6 | 68919-07-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-162-00-X | Gases (petroleum), preflash tower off, crude distn.; Refinery gas; [A complex combination produced from the first tower used in the distillation of crude oil. It consists of nitrogen and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 272-881-1 | 68919-08-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-163-00-5 | Gases (petroleum), tar stripper off; Refinery gas; [A complex combination obtained by the fractionation of reduced crude oil. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 272-884-8 | 68919-11-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-164-00-0 | Gases (petroleum), unifiner stripper off; Refinery gas; [A combination of hydrogen and methane obtained by fractionation of the products from the unifiner unit.] | 272-885-3 | 68919-12-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-165-00-6 | Tail gas (petroleum), catalytic hydrodesulfurized naphtha separator; Refinery gas; [A complex combination of hydrocarbons obtained from the hydrodesulfurization of naphtha. It consists of hydrogen, methane, ethane, and propane.] | 273-173-5 | 68952-79-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-166-00-1 | Tail gas (petroleum), straight-run naphtha hydrodesulfurizer; Refinery gas; [A complex combination obtained from the hydrodesulfurization of straight-run naphtha. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 273-174-0 | 68952-80-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-167-00-7 | Gases (petroleum), sponge absorber off, fluidized catalytic cracker and gas oil desulfurizer overhead fractionation; Refinery gas; [A complex combination obtained by the fractionation of products from the fluidized catalytic cracker and gas oil desulfurizer. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 273-269-7 | 68955-33-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-168-00-2 | Gases (petroleum), crude distn. and catalytic cracking; Refinery gas; [A complex combination produced by crude distillation and catalytic cracking processes. It consists of hydrogen, hydrogen sulfide, nitrogen, carbon monoxide and paraffinic and olefinic hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 273-563-5 | 68989-88-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-169-00-8 | Gases (petroleum), gas oil diethanolamine scrubber off; Refinery gas; [A complex combination produced by desulfurization of gas oils with diethanolamine. It consists predominantly of hydrogen sulfide, hydrogen and aliphatic hydrocarbons having carbon numbers in the range of C1 through C5.] | 295-397-2 | 92045-15-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-170-00-3 | Gases (petroleum), gas oil hydrodesulfurization effluent; Refinery gas; [A complex combination obtained by separation of the liquid phase from the effluent from the hydrogenation reaction. It consists predominantly of hydrogen, hydrogen sulfide and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] | 295-398-8 | 92045-16-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-171-00-9 | Gases (petroleum), gas oil hydrodesulfurization purge; Refinery gas; [A complex combination of gases obtained from the reformer and from the purges from the hydrogenation reactor. It consists predominantly of hydrogen and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 295-399-3 | 92045-17-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-172-00-4 | Gases (petroleum), hydrogenator effluent flash drum off; Refinery gas; [A complex combination of gases obtained from flash of the effluents after the hydrogenation reaction. It consists predominantly of hydrogen and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 295-400-7 | 92045-18-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-173-00-X | Gases (petroleum), naphtha steam cracking high-pressure residual; Refinery gas; [A complex combination obtained as a reaction mass of the non-condensable portions from the product of a naphtha steam cracking process as well as residual gases obtained during the preparation of subsequent products. It consists predominantly of hydrogen and paraffinic and olefinic hydrocarbons having carbon numbers predominantly in the range of C1 through C5 with which natural gas may also be mixed.] | 295-401-2 | 92045-19-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-174-00-5 | Gases (petroleum), residue visbaking off; Refinery gas; [A complex combination obtained from viscosity reduction of residues in a furnace. It consists predominantly of hydrogen sulfide and paraffinic and olefinic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 295-402-8 | 92045-20-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-175-00-0 | Foots oil (petroleum), acid-treated; Foots oil; [A complex combination of hydrocarbons obtained by treatment of Foot's oil with sulfuric acid. It consists predominantly of branched-chain hydrocarbons with carbon numbers predominantly in the range of C20 through C50.] | 300-225-7 | 93924-31-3 | Flam. Gas 1 Press. Gas Carc. 1B | H220 H350 | GHS02 GHS04 GHS08 Dgr | H220 H340 H350 |  |  | H K U | CLP00
649-176-00-6 | Foots oil (petroleum), clay-treated; Foots oil; [A complex combination of hydrocarbons obtained by treatment of Foot's oil with natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists predominantly of branched chain hydrocarbons with carbon numbers predominantly in the range of C20 through C50.] | 300-226-2 | 93924-32-4 | Flam. Gas 1 Press. Gas Carc. 1B | H220 H350 | GHS02 GHS04 GHS08 Dgr | H220 H340 H350 |  |  | H K U | CLP00
649-177-00-1 | Gases (petroleum), C3-4; Petroleum gas; [A complex combination of hydrocarbons produced by distillation of products from the cracking of crude oil. It consists of hydrocarbons having carbon numbers in the range of C3 through C4, predominantly of propane and propylene, and boiling in the range of approximately – 51 °C to – 1 °C (– 60 °F to 30 °F.)] | 268-629-5 | 68131-75-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-178-00-7 | Tail gas (petroleum), catalytic cracked distillate and catalytic cracked naphtha fractionation absorber; Petroleum gas; [The complex combination of hydrocarbons from the distillation of the products from catalytic cracked distillates and catalytic cracked naphtha. It consists predominantly of hydrocarbons having carbon numbers in the range of C1 through C4.] | 269-617-2 | 68307-98-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-179-00-2 | Tail gas (petroleum), catalytic polymn. naphtha fractionation stabilizer; Petroleum gas; [A complex combination of hydrocarbons from the fractionation stabilization products from polymerization of naphtha. It consists predominantly of hydrocarbons having carbon numbers in the range of C1 through C4.] | 269-618-8 | 68307-99-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-180-00-8 | Tail gas (petroleum), catalytic reformed naphtha fractionation stabilizer, hydrogen sulfide-free; Petroleum gas; [A complex combination of hydrocarbons obtained from fractionation stabilization of catalytic reformed naphtha and from which hydrogen sulfide has been removed by amine treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 269-619-3 | 68308-00-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-181-00-3 | Tail gas (petroleum), cracked distillate hydrotreater stripper; Petroleum gas; [A complex combination of hydrocarbons obtained by treating thermal cracked distillates with hydrogen in the presence of a catalyst. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 269-620-9 | 68308-01-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-182-00-9 | Tail gas (petroleum), straight-run distillate hydrodesulfurizer, hydrogen sulfide-free; Petroleum gas; [A complex combination of hydrocarbons obtained from catalytic hydrodesulfurization of straight run distillates and from which hydrogen sulfide has been removed by amine treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 269-630-3 | 68308-10-1 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-183-00-4 | Tail gas (petroleum), gas oil catalytic cracking absorber; Petroleum gas; [A complex combination of hydrocarbons obtained from the distillation of products from the catalytic cracking of gas oil. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 269-623-5 | 68308-03-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-184-00-X | Tail gas (petroleum), gas recovery plant; Petroleum gas; [A complex combination of hydrocarbons from the distillation of products from miscellaneous hydrocarbon streams. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 269-624-0 | 68308-04-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-185-00-5 | Tail gas (petroleum), gas recovery plant deethanizer; Petroleum gas; [A complex combination of hydrocarbons from the distillation of products from miscellaneous hydrocarbon streams. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 269-625-6 | 68308-05-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-186-00-0 | Tail gas (petroleum), hydrodesulfurized distillate and hydrodesulfurized naphtha fractionator, acid-free; Petroleum gas; [A complex combination of hydrocarbons obtained from fractionation of hydrodesulfurized naphtha and distillate hydrocarbon streams and treated to remove acidic impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 269-626-1 | 68308-06-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-187-00-6 | Tail gas (petroleum), hydrodesulfurized vacuum gas oil stripper, hydrogen sulfide-free; Petroleum gas; [A complex combination of hydrocarbons obtained from stripping stabilization of catalytic hydrodesulfurized vacuum gas oil and from which hydrogen sulfide has been removed by amine treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 269-627-7 | 68308-07-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-188-00-1 | Tail gas (petroleum), light straight-run naphtha stabilizer, hydrogen sulfide-free; Petroleum gas; [A complex combination of hydrocarbons obtained from fractionation stabilization of light straight run naphtha and from which hydrogen sulfide has been removed by amine treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 269-629-8 | 68308-09-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-189-00-7 | Tail gas (petroleum), propane-propylene alkylation feed prep deethanizer; Petroleum gas; [A complex combination of hydrocarbons obtained from the distillation of the reaction products of propane with propylene. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 269-631-9 | 68308-11-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-190-00-2 | Tail gas (petroleum), vacuum gas oil hydrodesulfurizer, hydrogen sulfide-free; Petroleum gas; [A complex combination of hydrocarbons obtained from catalytic hydrodesulfurization of vacuum gas oil and from which hydrogen sulfide has been removed by amine treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] | 269-632-4 | 68308-12-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-191-00-8 | Gases (petroleum), catalytic cracked overheads; Petroleum gas; [A complex combination of hydrocarbons produced by the distillation of products from the catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C3 through C5 and boiling in the range of approximately – 48 °C to 32 °C (– 54 °F to 90 °F).] | 270-071-2 | 68409-99-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-193-00-9 | Alkanes, C1-2; Petroleum gas | 270-651-5 | 68475-57-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-194-00-4 | Alkanes, C2-3; Petroleum gas | 270-652-0 | 68475-58-1 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-195-00-X | Alkanes, C3-4; petroleum gas | 270-653-6 | 68475-59-2 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H350 H340 |  |  | K U | CLP00/ATP01corr
649-196-00-5 | Alkanes, C4-5; Petroleum gas | 270-654-1 | 68475-60-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-197-00-0 | Fuel gases; Petroleum gas; [A combination of light gases. It consists predominantly of hydrogen and/or low molecular weight hydrocarbons.] | 270-667-2 | 68476-26-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-198-00-6 | Fuel gases, crude oil of distillates; Petroleum gas; [A complex combination of light gases produced by distillation of crude oil and by catalytic reforming of naphtha. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C4 and boiling in the range of approximately – 217 °C to – 12 °C (– 423 °F to 10 °F).] | 270-670-9 | 68476-29-9 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-199-00-1 | Hydrocarbons, C3-4; Petroleum gas | 270-681-9 | 68476-40-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-200-00-5 | Hydrocarbons, C4-5; Petroleum gas | 270-682-4 | 68476-42-6 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-201-00-0 | Hydrocarbons, C2-4, C3-rich; Petroleum gas | 270-689-2 | 68476-49-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-202-00-6 | Petroleum gases, liquefied; Petroleum gas; [A complex combination of hydrocarbons produced by the distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C3 through C7 and boiling in the range of approximately – 40 °C to 80 °C (– 40 °F to 176 °F).] | 270-704-2 | 68476-85-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K S U | CLP00/ATP01corr
649-203-00-1 | Petroleum gases, liquefied, sweetened; Petroleum gas; [A complex combination of hydrocarbons obtained by subjecting liquefied petroleum gas mix to a sweetening process to convert mercaptans or to remove acidic impurities. It consists of hydrocarbons having carbon numbers predominantly in the range of C3 through C7 and boiling in the range of approximately – 40 °C to 80 °C (– 40 °F to 176 °F).] | 270-705-8 | 68476-86-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K S U | CLP00/ATP01corr
649-204-00-7 | gases (petroleum), C3-4, isobutane-rich; Petroleum gas; [A complex combination of hydrocarbons from the distillation of saturated and unsaturated hydrocarbons usually ranging in carbon numbers from C3 through C6, predominantly butane and isobutane. It consists of saturated and unsaturated hydrocarbons having carbon numbers in the range of C3 through C4, predominantly isobutane.] | 270-724-1 | 68477-33-8 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-205-00-2 | Distillates (petroleum), C3-6, piperylene-rich; Petroleum gas; [A complex combination of hydrocarbons from the distillation of saturated and unsaturated aliphatic hydrocarbons usually ranging in the carbon numbers C3 through C6. It consists of saturated and unsaturated hydrocarbons having carbon numbers in the range of C3 through C6, predominantly piperylenes.] | 270-726-2 | 68477-35-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-206-00-8 | Gases (petroleum), butane splitter overheads; Petroleum gas; [A complex combination of hydrocarbons obtained from the distillation of the butane stream. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C3 through C4.] | 270-750-3 | 68477-69-0 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-207-00-3 | Gases (petroleum), C2-3-; Petroleum gas; [A complex combination of hydrocarbons produced by the distillation of products from a catalytic fractionation process. It contains predominantly ethane, ethylene, propane, and propylene.] | 270-751-9 | 68477-70-3 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-208-00-9 | Gases (petroleum), catalytic-cracked gas oil depropanizer bottoms, C4-rich acid-free; Petroleum gas; [A complex combination of hydrocarbons obtained from fractionation of catalytic cracked gas oil hydrocarbon stream and treated to remove hydrogen sulfide and other acidic components. It consists of hydrocarbons having carbon numbers in the range of C3 through C5, predominantly C4.] | 270-752-4 | 68477-71-4 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-209-00-4 | Gases (petroleum), catalytic-cracked naphtha debutanizer bottoms, C3-5-rich; Petroleum gas; [A complex combination of hydrocarbons obtained from the stabilization of catalytic cracked naphtha. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C3 through C5.] | 270-754-5 | 68477-72-5 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-210-00-X | Tail gas (petroleum), isomerized naphtha fractionation stabilizer; Petroleum gas; [A complex combination of hydrocarbons obtained from the fractionation stabilization products from isomerized naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] | 269-628-2 | 68308-08-7 | Flam. Gas 1 Press. Gas Carc. 1A Muta. 1B | H220 H350 H340 | GHS04 GHS02 GHS08 Dgr | H220 H340 H350 |  |  | K U | CLP00/ATP01corr
649-211-00-5 | Foots oil (petroleum), carbon-treated; Foots oil; [A complex combination of hydrocarbons obtained by the treatment of Foots oil with activated carbon for the removal of trace constituents and impurities. It consists predominantly of saturated straight chain hydrocarbons having carbon numbers predominantly greater than C12.] | 308-126-0 | 97862-76-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-212-00-0 | Distillates (petroleum), sweetened middle; Gasoil - unspecified; [A complex combination of hydrocarbons obtained by subjecting a petroleum distillate to a sweetening process to convert mercaptans or to remove acidic impurities. It consists of hydrocarbons having carbon numbers predominantly in the range of C9 through C20 and boiling in the range of approximately 150 °C to 345 °C (302 °F to 653 °F).] | 265-088-7 | 64741-86-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-213-00-6 | Gas oils (petroleum), solvent-refined; Gasoil - unspecified; [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C11 through C25 and boiling in the range of approximately 205 °C to 400 °C (401 °F to 752 °F).] | 265-092-9 | 64741-90-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-214-00-1 | Distillates (petroleum), solvent-refined middle; Gasoil - unspecified; [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C9 through C20 and boiling in the range of approximately 150 °C to 345 °C (302 °F to 653 °F).] | 265-093-4 | 64741-91-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-215-00-7 | Gas oils (petroleum), acid-treated; Gasoil - unspecified; [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C13 through C25 and boiling in the range of approximately 230 °C to 400 °C (446 °F to 752 °F).] | 265-112-6 | 64742-12-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-216-00-2 | Distillates (petroleum), acid-treated middle; Gasoil - unspecified; [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C11 through C20 and boiling in the range of approximately 205 °C to 345 °C (401 °F to 653 °F).] | 265-113-1 | 64742-13-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-217-00-8 | Distillates (petroleum), acid-treated light; Gasoil - unspecified; [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C9 through C16 and boiling in the range of approximately 150 °C to 290 °C (302 °F to 554 °F).] | 265-114-7 | 64742-14-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-218-00-3 | Gas oils (petroleum), chemically neutralized; Gasoil - unspecified; [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C13 through C25 and boiling in the range of approximately 230 °C to 400 °C (446 °F to 752 °F).] | 265-129-9 | 64742-29-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-219-00-9 | Distillates (petroleum), chemically neutralized middle; Gasoil - unspecified; [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C11 through C20 and boiling in the range of approximately 205 °C to 345 °C (401 °F to 653 °F).] | 265-130-4 | 64742-30-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-220-00-4 | Distillates (petroleum), clay-treated middle; Gasoil - unspecified; [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay, usually in a percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C9 through C20 and boiling in the range of approximately 150 °C to 345 °C (302 °F to 653 °F).] | 265-139-3 | 64742-38-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-221-00-X | Distillates (petroleum), hydrotreated middle; Gasoil - unspecified; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C11 through C25 and boiling in the range of approximately 205 °C to 400 °C (401 °F to 752 °F).] | 265-148-2 | 64742-46-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-222-00-5 | Gas oils (petroleum), hydrodesulfurized; Gasoil - unspecified; [A complex combination of hydrocarbons obtained from a petroleum stock by treating with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C13 through C25 and boiling in the range of approximately 230 °C to 400 °C (446 °F to 752 °F).] | 265-182-8 | 64742-79-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-223-00-0 | Distillates (petroleum), hydrodesulfurized middle; Gasoil - unspecified; [A complex combination of hydrocarbons obtained from a petroleum stock by treating with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists of hydrocarbons having carbon numbers predominantly in the range of C11 through C25 and boiling in the range of approximately 205 °C to 400 °C (401 °F to 752 °F).] | 265-183-3 | 64742-80-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-224-00-6 | Fuels, diesel; Gasoil - unspecified; [A complex combination of hydrocarbons produced by the distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C9 through C20 and boiling in the range of approximately 163 °C to 357 °C (325 °F to 675 °F).] | 269-822-7 | 68334-30-5 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  | N | CLP00
649-225-00-1 | Fuel oil, No 2; Gasoil - unspecified; [A distillate oil having a minimum viscosity of 32,6 SUS at 37,7 °C (100 °F) to a maximum of 37,9 SUS at 37,7 °C (100 °F).] | 270-671-4 | 68476-30-2 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | CLP00
649-226-00-7 | Fuel oil, No 4; Gasoil - unspecified; [A distillate oil having a minimum viscosity of 45 SUS at 37,7 °C (100 °F) to a maximum of 125 SUS at 37,7 °C (100 °F).] | 270-673-5 | 68476-31-3 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | CLP00
649-227-00-2 | Fuels, diesel, No 2; Gasoil - unspecified; [A distillate oil having a minimum viscosity of 32,6 SUS at 37,7 °C (100 °F).] | 270-676-1 | 68476-34-6 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | CLP00
649-228-00-8 | Distillates (petroleum), catalytic reformer fractionator residue, high-boiling; Gasoil - unspecified; [A complex combination of hydrocarbons from the distillation of catalytic reformer fracftionator residue. It boils in the range of approximately 343 °C to 399 °C (650 °F to 750 °F).] | 270-719-4 | 68477-29-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-229-00-3 | Distillates (petroleum), catalytic reformer fractionator residue, intermediate-boiling; Gasoil - unspecified; [A complex combination of hydrocarbons from the distillation of catalytic reformer fractionator residue. It boils in the range of approximately 288 °C to 371 °C (550 °F to 700 °F).] | 270-721-5 | 68477-30-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-230-00-9 | Distillates (petroleum), catalytic reformer fractionator residue, low-boiling; Gasoil - unspecified; [The complex combination of hydrocarbons from the distillation of catalytic reformer fractionator residue. It boils approximately below 288 °C (550 °F).] | 270-722-0 | 68477-31-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-231-00-4 | Distillates (petroleum), highly refined middle; Gasoil - unspecified; [A complex combination of hydrocarbons obtained by the subjection of a petroleum fraction to several of the following steps: filtration, centrifugation, atmospheric distillation, vacuum distillation, acidification, neutralization and clay treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C10 through C20.] | 292-615-8 | 90640-93-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-232-00-X | Distillates (petroleum) catalytic reformer, heavy arom. conc.; Gasoil - unspecified; [A complex combination of hydrocarbons obtained from the distillation of a catalytically reformed petroleum cut. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C10 through C16 and boiling in the range of approximately 200 °C to 300 °C (392 °F to 572 °F).] | 295-294-2 | 91995-34-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-233-00-5 | Gas oils, paraffinic; Gasoil - unspecified; [A distillate obtained from the redistillation of a complex combination of hydrocarbons obtained by the distillation of the effluents from a severe catalytic hydrotreatment of paraffins. It boils in the range of approximately 190 °C to 330 °C (374 °F to 594 °F).] | 300-227-8 | 93924-33-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-234-00-0 | Naphtha (petroleum), solvent-refined hydrodesulfurized heavy; Gasoil - unspecified | 307-035-3 | 97488-96-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-235-00-6 | Hydrocarbons, C16-20, hydrotreated middle distillate, distn. lights; Gasoil - unspecified; [A complex combination of hydrocarbons obtained as first runnings from the vacuum distillation of effluents from the treatment of a middle distillate with hydrogen. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C16 through C20 and boiling in the range of approximately 290 °C to 350 °C (554 °F to 662 °F). It produces a finished oil having a viscosity of 2cSt at 100 °C (212 °F).] | 307-659-6 | 97675-85-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-236-00-1 | Hydrocarbons, C12-20, hydrotreated paraffinic, distn. lights; Gasoil - unspecified; [A complex combination of hydrocarbons obtained as first runnings from the vacuum distillation of effluents from the treatment of heavy paraffins with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C12 through C20 and boiling in the range of approximately 230 °C to 350 °C (446 °F to 662 °F). It produces a finished oil having a viscosity of 2cSt at 100 °C (212 °F).] | 307-660-1 | 97675-86-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-237-00-7 | Hydrocarbons, C11-17, solvent-extd. light naphthenic; Gasoil - unspecified; [A complex combination of hydrocarbons obtained by extraction of the aromatics from a light naphthenic distillate having a visciosity of 2.2 cSt at 40 °C (104 °F). It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C11 through C17 and boiling in the range of approximately 200 °C to 300 °C (392 °F to 572 °F).] | 307-757-9 | 97722-08-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-238-00-2 | Gas oils, hydrotreated; Gasoil - unspecified; [A complex combination of hydrocarbons obtained from the redistillation of the effluents from the treatment of paraffins with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C17 through C27 and boiling in the range of approximately 330 °C to 340 °C (626 °F to 644 °F).] | 308-128-1 | 97862-78-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-239-00-8 | Distillates (petroleum), carbon-treated light paraffinic; Gasoil - unspecified; [A complex combination of hydrocarbons obtained by the treatment of a petroleum oil fraction with activated charcoal for the removal of traces of polar constituents and impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C12 through C28.] | 309-667-5 | 100683-97-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-240-00-3 | Distillates (petroleum), intermediate paraffinic, carbon-treated; Gasoil - unspecified; [A complex combination of hydrocarbons obtained by the treatment of petroleum with activated charcoal for the removal of trace polar constituents and impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C16 through C36.] | 309-668-0 | 100683-98-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-241-00-9 | Distillates (petroleum), intermediate paraffinic, clay-treated; Gasoil - unspecified; [A complex combination of hydrocarbons obtained by the treatment of petroleum with bleaching earth for the removal of trace polar constituents and impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C16 through C36.] | 309-669-6 | 100683-99-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-242-00-4 | Alkanes, C12-26-branched and linear | 292-454-3 | 90622-53-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-243-00-X | Lubricating greases; Grease; [A complex combination of hydrocarbons having carbon numbers predominantly in the range of C12 through C50. May contain organic salts of alkali metals, alkaline earth metals, and/or aluminium compounds.] | 278-011-7 | 74869-21-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-244-00-5 | Slack wax (petroleum); Slack wax; [A complex combination of hydrocarbons obtained from a petroleum fraction by solvent crystallization (solvent dewaxing) or as a distillation fraction from a very waxy crude. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C20.] | 265-165-5 | 64742-61-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-245-00-0 | Slack wax (petroleum), acid-treated; Slack wax; [A complex combination of hydrocarbons obtained as a raffinate by treatment of a petroleum slack wax fraction with sulfuric acid treating process. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C20.] | 292-659-8 | 90669-77-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-246-00-6 | Slack wax (petroleum), clay-treated; Slack wax; [A complex combination of hydrocarbons obtained by treatment of a petroleum slack wax fraction with natural or modified clay in either a contacting or percolation process. It consists predominantly of saturated straight and branched hydrocarbons having carbon numbers predominantly greater than C20.] | 292-660-3 | 90669-78-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-247-00-1 | Slack wax (petroleum), hydrotreated; Slack wax; [A complex combination of hydrocarbons obtained by treating slack wax with hydrogen in the presence of a catalyst. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C20.] | 295-523-6 | 92062-09-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-248-00-7 | Slack wax (petroleum), low-melting; Slack wax; [A complex combination of hydrocarbons obtained from a petroleum fraction by solvent deparaffination. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C12.] | 295-524-1 | 92062-10-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-249-00-2 | Slack wax (petroleum), low-melting, hydrotreated; Slack wax; [A complex combination of hydrocarbons obtained by treatment of low-melting petroleum slack wax with hydrogen in the presence of a catalyst. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C12.] | 295-525-7 | 92062-11-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-250-00-8 | Slack wax (petroleum), low-melting, carbon-treated; Slack wax; [A complex combination of hydrocarbons obtained by the treatment of low-melting slack wax with activated carbon for the removal of trace polar constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C12.] | 308-155-9 | 97863-04-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-251-00-3 | Slack wax (petroleum), low-melting, clay-treated; Slack wax; [A complex combination of hydrocarbons obtained by the treatment of low-melting petroleum slack wax with bentonite for removal of trace polar constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C12.] | 308-156-4 | 97863-05-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-252-00-9 | Slack wax (petroleum), low-melting, silicic acid-treated; Slack wax; [A complex combination of hydrocarbons obtained by the treatment of low-melting petroleum slack wax with silicic acid for the removal of trace polar constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C12.] | 308-158-5 | 97863-06-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-253-00-4 | Slack wax (petroleum), carbon-treated; Slack wax; [A complex combination of hydrocarbons obtained by treatment of petroleum slack wax with activated charcoal for the removal of trace polar constituents and impurities.] | 309-723-9 | 100684-49-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-254-00-X | Petrolatum; Petrolatum; [A complex combination of hydrocarbons obtained as a semi-solid from dewaxing paraffinic residual oil. It consists predominantly of saturated crystalline and liquid hydrocarbons having carbon numbers predominantly greater than C25.] | 232-373-2 | 8009-03-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-255-00-5 | Petrolatum (petroleum), oxidized; Petrolatum; [A complex combination of organic compounds, predominantly high molecular weight carboxylic acids, obtained by the air oxidation of petrolatum.] | 265-206-7 | 64743-01-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-256-00-0 | Petrolatum (petroleum), alumina-treated; Petrolatum; [A complex combination of hydrocarbons obtained when petrolatum is treated with Al2O3 to remove polar components and impurities. It consists predominantly of saturated, crystalline, and liquid hydrocarbons having carbon numbers predominantly greater than C25.] | 285-098-5 | 85029-74-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-257-00-6 | Petrolatum (petroleum), hydrotreated; Petrolatum; [A complex combination of hydrocarbons obtained as a semi-solid from dewaxed paraffinic residual oil treated with hydrogen in the presence of a catalyst. It consists predominantly of saturated microcrystalline and liquid hydrocarbons having carbon numbers predominantly greater than C20.] | 295-459-9 | 92045-77-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-258-00-1 | Petrolatum (petroleum), carbon-treated; Petrolatum; [A complex combination of hydrocarbons obtained by the treatment of petroleum petrolatum with activated carbon for the removal of trace polar constituents and impurities. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly greater than C20.] | 308-149-6 | 97862-97-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-259-00-7 | Petrolatum (petroleum), silicic acid-treated; Petrolatum; [A complex combination of hydrocarbons obtained by the treatment of petroleum petrolatum with silicic acid for the removal of trace polar constituents and impurities. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly greater than C20.] | 308-150-1 | 97862-98-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-260-00-2 | Petrolatum (petroleum), clay-treated; Petrolatum; [A complex combination of hydrocarbons obtained by treatment of petrolatum with bleaching earth for the removal of traces of polar constituents and impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of greater than C25.] | 309-706-6 | 100684-33-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | N | CLP00
649-261-00-8 | Gasoline, natural; Low boiling point naphtha; [A complex combination of hydrocarbons separated from natural gas by processes such as refrigeration or absorption. It consists predominantly of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C4 through C8 and boiling in the range of approximately minus 20°C to 120°C (-4°F to 248°F).] | 232-349-1 | 8006-61-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-262-00-3 | Naphtha; Low boiling point naphtha; [Refined, partly refined, or unrefined petroleum products produced by the distillation of natural gas. It consists of hydrocarbons having carbon numbers predominantly in the range of C5 through C6 and boiling in the range of approximately 100°C to 200°C (212°F to 392°F).] | 232-443-2 | 8030-30-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-263-00-9 | Ligroine; Low boiling point naphtha; [A complex combination of hydrocarbons obtained by the fractional distillation of petroleum. This fraction boils in a range of approximately 20°C to 135°C (58°F to 275°F).] | 232-453-7 | 8032-32-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-264-00-4 | Naphtha (petroleum), heavy straight-run; Low boiling point naphtha; [A complex combination of hydrocarbons produced by distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C6 through C12 and boiling in the range of approximately 65°C to 230°C (149°F to 446°F).] | 265-041-0 | 64741-41-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-265-00-X | Naphtha (petroleum), full-range straight-run; Low boiling point naphtha; [A complex combination of hydrocarbons produced by distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C4 through C11 and boiling in the range of approximately -20°C to 220°C (-4°F to 428°F).] | 265-042-6 | 64741-42-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-266-00-5 | Naphtha (petroleum), light straight-run; Low boiling point naphtha; [A complex combination of hydrocarbons produced by distillation of crude oil. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C4 through C10 and boiling in the range of approximately -20°C to 180°C (-4°F to 356°F).] | 265-046-8 | 64741-46-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-267-00-0 | Solvent naphtha (petroleum), light aliph.; Low boiling point naphtha; [A complex combination of hydrocarbons obtained from the distillation of crude oil or natural gasoline. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C5 through C10 and boiling in the range of approximately 35°C to 160°C (95°F to 320°F).] | 265-192-2 | 64742-89-8 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-268-00-6 | Distillates (petroleum), straight-run light; Low boiling point naphtha; [A complex combination of hydrocarbons produced by the distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C2 through C7 and boiling in the range of approximately -88°C to 99°C (-127°F to 210°F).] | 270-077-5 | 68410-05-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-269-00-1 | Gasoline, vapor-recovery; Low boiling point naphtha; [A complex combination of hydrocarbons separated from the gases from vapor recovery systems by cooling. It consists of hydrocarbons having carbon numbers predominantly in the range of C4 through C11 and boiling in the range of approximately -20°C to 196°C(-4°F to 384°F).] | 271-025-4 | 68514-15-8 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-270-00-7 | Gasoline, straight-run, topping-plant; Low boiling point naphtha; [A complex combination of hydrocarbons produced from the topping plant by the distillation of crude oil. It boils in the range of approximately 36.1°C to 193.3°C (97°F to 380°F).] | 271-727-0 | 68606-11-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-271-00-2 | Naphtha (petroleum), unsweetened; Low boiling point naphtha; [A complex combination of hydrocarbons produced from the distillation of naphtha streams from various refinery processes. It consists of hydrocarbons having carbon numbers predominantly in the range of C5 through C12 and boiling in the range of approximately 0°C to 230°C (25°F to 446°F).] | 272-186-3 | 68783-12-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-272-00-8 | Distillates (petroleum), light straight-run gasoline fractionation stabilizer overheads; Low boiling point naphtha; [A complex combination of hydrocarbons obtained by the fractionation of light straight-run gasoline. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C3 through C6.] | 272-931-2 | 68921-08-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-273-00-3 | Naphtha (petroleum), heavy straight run, arom.-contg.; Low boiling point naphtha; [A complex combination of hydrocarbons obtained from a distillation process of crude petroleum. It consists predominantly of hydrocarbons having carbon numbers in the range of C8 through C12 and boiling in the range of approximately 130°C to 210°C (266°F to 410°F).] | 309-945-6 | 101631-20-3 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-274-00-9 | Naphtha (petroleum), full-range alkylate; Low boiling point modified naphtha; [A complex combination of hydrocarbons produced by distillation of the reaction products of isobutane with monoolefinic hydrocarbons usually ranging in carbon numbers from C3 through C5 It consists of predominantly branched chain saturated hydrocarbons having carbon numbers predominantly in the range of C7 through C12 and boiling in the range of approximately 90°C to 220°C (194°F to 428°F).] | 265-066-7 | 64741-64-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-275-00-4 | Naphtha (petroleum), heavy alkylate; Low boiling point modified naphtha; [A complex combination of hydrocarbons produced by distillation of the reaction products of isobutane with monoolefinic hydrocarbons usually ranging in carbon numbers from C3 to C5. It consists of predominantly branched chain saturated hydrocarbons having carbon numbers predominantly in the range of C9 through C12 and boiling in the range of approximately 150°C to 220°C (302°F to 428°F).] | 265-067-2 | 64741-65-7 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-276-00-X | Naphtha (petroleum), light alkylate; Low boiling point modified naphtha; [A complex combination of hydrocarbons produced by distillation of the reaction products of isobutane with monoolefinic hydrocarbons usually ranging in carbon numbers from C3 through C5. It consists of predominantly branched chain saturated hydrocarbons having carbon numbers predominantly in the range of C7 through C10 and boiling in the range of approximately 90°C to 160°C (194°F to 320°F).] | 265-068-8 | 64741-66-8 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-277-00-5 | Naphtha (petroleum), isomerization; Low boiling point modified naphtha; [A complex combination of hydrocarbons obtained from catalytic isomerization of straight chain paraffinic C4 through C6 hydrocarbons. It consists predominantly of saturated hydrocarbons such as isobutane, isopentane, 2,2-dimethylbutane, 2-methylpentane, and 3-methylpentane.] | 265-073-5 | 64741-70-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-278-00-0 | Naphtha (petroleum), solvent-refined light; Low boiling point modified naphtha; [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C5 through C11 and boiling in the range of approximately 35°C to 190°C (95°F to 374°F).] | 265-086-6 | 64741-84-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-279-00-6 | Naphtha (petroleum), solvent-refined heavy; Low boiling point modified naphtha; [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C7 through C12 and boiling in the range of approximately 90°C to 230°C (194°F to 446°F).] | 265-095-5 | 64741-92-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-280-00-1 | Raffinates (petroleum), catalytic reformer ethylene glycol-water countercurrent exts.; Low boiling point modified naphtha; [A complex combination of hydrocarbons obtained as the raffinate from the UDEX extraction process on the catalytic reformer stream. It consists of saturated hydrocarbons having carbon numbers predominantly in the range of C6 through C9.] | 270-088-5 | 68410-71-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-281-00-7 | Raffinates (petroleum), reformer, Lurgi unit-sepd.; Low boiling point modified naphtha; [The complex combination of hydrocarbons obtained as a raffinate from a Lurgi separation unit. It consists predominantly of non-aromatic hydrocarbons with various small amounts of aromatic hydrocarbons having carbon numbers predominantly in the range of C6 through C8.] | 270-349-3 | 68425-35-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-282-00-2 | Naphtha (petroleum), full-range alkylate, butane-contg.; Low boiling point modified naphta; [A complex combination of hydrocarbons produced by the distillation of the reaction products of isobutane with monoolefinic hydrocarbons usually ranging in carbon numbers from C3 through C5. It consists of predominantly branched chain saturated hydrocarbons having carbon numbers predominantly in the range of C7 through C12 with some butanes and boiling in the range of approximately 35°C to 200°C (95°F to 428°F).] | 271-267-0 | 68527-27-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-283-00-8 | Distillates (petroleum), naphtha steam cracking-derived, solvent-refined light hydrotreated; Low boiling point modified naphtha; [A complex combination of hydrocarbons obtained as the raffinates from a solvent extraction process of hydrotreated light distillate from steam-cracked naphtha.] | 295-315-5 | 91995-53-8 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-284-00-3 | Naphtha (petroleum), C4-12, butane-alkylate, isooctane-rich; Low boiling point modified naphtha; [A complex combination of hydrocarbons obtained by alkylation of butanes. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C4 through C12, rich in isooctane, and boiling in the range of approximately 35°C to 210°C (95°F to 410°F).] | 295-430-0 | 92045-49-3 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-285-00-9 | Hydrocarbons, hydrotreated light naphtha distillates, solvent-refined; Low boiling point modified naphtha; [A combination of hydrocarbons obtained from the distillation of hydrotreated naphtha followed by a solvent extraction and distillation process. It consists predominantly of saturated hydrocarbons boiling in the range of approximately 94°C to 99°C (201°F to 210°F).] | 295-436-3 | 92045-55-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-286-00-4 | Naphtha (petroleum), isomerization, C6-fraction; Low boiling point modified naphtha; [A complex combination of hydrocarbons obtained by distillation of a gasoline which has been catalytically isomerized. It consists predominantly of hexane isomers boiling in the range of approximately 60°C to 66°C (140°F to 151°F).] | 295-440-5 | 92045-58-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-287-00-X | Hydrocarbons, C6-7, naphtha-cracking, solvent-refined; Low boiling point modified naphtha; [A complex combination of hydrocarbons obtained by the sorption of benzene from a catalytically fully hydrogenated benzene-rich hydrocarbon cut that was distillatively obtained from prehydrogenated cracked naphtha. It consists predominantly of paraffinic and naphthenic hydrocarbons having carbon numbers predominantly in the range of C6 through C7 and boiling in the range of approximately 70°C to 100°C (158°F to 212°F).] | 295-446-8 | 92045-64-2 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-288-00-5 | Hydrocarbons, C6-rich, hydrotreated light naphtha distillates, solvent-refined; Low boiling point modified naphtha; [A complex combination of hydrocarbons obtained by distillation of hydrotreated naphtha followed by solvent extraction. It consists predominantly of saturated hydrocarbons and boiling in the range of approximately 65°C to 70°C (149°F to 158°F).] | 309-871-4 | 101316-67-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-289-00-0 | Naphtha (petroleum), heavy catalytic cracked; Low boiling point cat-cracked naphtha; [A complex combination of hydrocarbons produced by a distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C6 through C12 and boiling in the range of approximately 65°C to 230°C (148°F to 446°F). It contains a relatively large proportion of unsaturated hydrocarbons.] | 265-055-7 | 64741-54-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-290-00-6 | Naphtha (petroleum), light catalytic cracked; Low boiling point cat-cracked naphtha; [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C4 through C11 and boiling in the range of approximately -20°C to 190°C (-4°F to 374°F). It contains a relatively large proportion of unsaturated hydrocarbons.] | 265-056-2 | 64741-55-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-291-00-1 | Hydrocarbons, C3-11, catalytic cracker distillates; Low boiling point cat-cracked naphtha; [A complex combination of hydrocarbons produced by the distillations of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C3 through C11 and boiling in a range approximately up to 204°C (400°F).] | 270-686-6 | 68476-46-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-292-00-7 | Naphtha (petroleum), catalytic cracked light distd.; Low boiling point cat-cracked naphtha; [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] | 272-185-8 | 68783-09-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-293-00-2 | Distillates (petroleum), naphtha steam cracking-derived, hydrotreated light arom.; Low boiling point cat-cracked naphtha.; [A complex combination of hydrocarbons obtained by treating a light distillate from steam-cracked naphtha. It consists predominantly of aromatic hydrocarbons.] | 295-311-3 | 91995-50-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-294-00-8 | Naphtha (petroleum), heavy catalytic cracked, sweetened; Low boiling point cat-cracked naphtha; [A complex combination of hydrocarbons obtained by subjecting a catalytic cracked petroleum distillate to a sweetening process to convert mercaptans or to remove acidic impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C6 through C12 and boiling in the range of approximately 60°C to 200°C (140°F to 392°F).] | 295-431-6 | 92045-50-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-295-00-3 | Naphtha (petroleum), light catalytic cracked sweetened; Low boiling point cat-cracked naphtha; [A complex combination of hydrocarbons obtained by subjecting naphtha from a catalytic cracking process to a sweetening process to convert mercaptans or to remove acidic impurities. It consists predominantly of hydrocarbons boiling in a range of approximately 35°C to 210°C (95°F to 410°F).] | 295-441-0 | 92045-59-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-296-00-9 | Hydrocarbons, C8-12, catalytic-cracking, chem. neutralized; Low boiling point cat-cracked naphtha; [A complex combination of hydrocarbons produced by the distillation of a cut from the catalytic cracking process, having undergone an alkaline washing. It consists predominantly of hydrocarbons having carbon numbers in the range of C8 through C12 and boiling in the range of approximately 130°C to 210°C (266°F to 410°F).] | 295-794-0 | 92128-94-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-297-00-4 | Hydrocarbons, C8-12, catalytic cracker distillates; Low boiling point cat-cracked naphtha; [A complex combination of hydrocarbons obtained by distillation of products from a catalytic cracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C8 through C12 and boiling in the range of approximately 140°C to 210°C (284°F to 410°F).] | 309-974-4 | 101794-97-2 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-298-00-X | Hydrocarbons, C8-12, catalytic cracking, chem. neutralized, sweetened; Low boiling point cat-cracked naphtha | 309-987-5 | 101896-28-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-299-00-5 | Naphtha (petroleum), light catalytic reformed; Low boiling point cat-reformed naphtha; [A complex combination of hydrocarbons produced from the distillation of products from a catalytic reforming process. It consists of hydrocarbons having carbon numbers predominantly in the range of C5 through C11 and boiling in the range of approximately 35°C to 190°C (95°F to 374°F). It contains a relatively large proportion of aromatic and branched chain hydrocarbons. This stream may contain 10 vol. % or more benzene.] | 265-065-1 | 64741-63-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-300-00-9 | Naphtha (petroleum), heavy catalytic reformed; Low boiling point cat-reformed naphtha; [A complex combination of hydrocarbons produced from the distillation of products from a catalytic reforming process. It consists of predominantly aromatic hydrocarbons having carbon numbers predominantly in the range of C7 through C12 and boiling in the range of approximately 90°C to 230°C (194°F to 446°F).] | 265-070-9 | 64741-68-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-301-00-4 | Distillates (petroleum), catalytic reformed depentanizer; Low boiling point cat-reformed naphtha; [A complex combination of hydrocarbons from the distillation of products from a catalytic reforming process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C3 through C6 and boiling in the range of approximately -49°C to 63°C (-57°F to 145°F).] | 270-660-4 | 68475-79-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-302-00-X | Hydrocarbons, C2-6, C6-8 catalytic reformer; Low boiling point cat-reformed naphtha | 270-687-1 | 68476-47-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-303-00-5 | Residues (petroleum), C6-8 catalytic reformer; Low boiling point cat-reformed naphtha; [A complex residuum from the catalytic reforming of C6-8 feed. It consists of hydrocarbons having carbon numbers predominantly in the range of C2 through C6.] | 270-794-3 | 68478-15-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-304-00-0 | Naphtha (petroleum), light catalytic reformed, arom.-free; Low boiling point cat-reformed naphtha; [A complex combination of hydrocarbons obtained from distillation of products from a catalytic reforming process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C5 through C8 and boiling in the range of approximately 35°C to 120°C (95°F to 248°F). It contains a relatively large proportion of branched chain hydrocarbons with the aromatic components removed.] | 270-993-5 | 68513-03-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-305-00-6 | Distillates (petroleum), catalytic reformed straight-run naphtha overheads; Low boiling point cat-reformed naphtha; [A complex combination of hydrocarbons obtained by the catalytic reforming of straight-run naphtha followed by the fractionation of the total effluent. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C6.] | 271-008-1 | 68513-63-3 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-306-00-1 | Petroleum products, hydrofiner-powerformer reformates; Low boiling point cat-reformed naphtha; [The complex combination of hydrocarbons obtained in a hydrofiner-powerformer process and boiling in a range of approximately 27°C to 210°C (80°F to 410°F).] | 271-058-4 | 68514-79-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-307-00-7 | Naphtha (petroleum), full-range reformed; Low boiling point cat-reformed naphtha; [A complex combination of hydrocarbons produced by the distillation of the products from a catalytic reforming process. It consists of hydrocarbons having carbon numbers predominantly in the range of C5 through C12 and boiling in the range of approximately 35°C to 230°C (95°F to 446°F).] | 272-895-8 | 68919-37-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-308-00-2 | Naphtha (petroleum), catalytic reformed; Low boiling point cat-reformed naphtha; [A complex combination of hydrocarbons produced by the distillation of products from a catalytic reforming process. It consists of hydrocarbons having carbon numbers predominantly in the range of C4 through C12 and boiling in the range of approximately 30°C to 220°C (90°F to 430°F). It contains a relatively large proportion of aromatic and branched chain hydrocarbons. This stream may contain 10 vol. % or more benzene.] | 273-271-8 | 68955-35-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-309-00-8 | Distillates (petroleum), catalytic reformed hydrotreated light, C8-12 arom. fraction; Low boiling point cat-reformed naphtha; [A complex combination of alkylbenzenes obtained by the catalytic reforming of petroleum naphtha. It consists predominantly of alkylbenzenes having carbon numbers predominantly in the range of C8 through C10 and boiling in the range of approximately 160°C to 180°C (320°F to 356°F).] | 285-509-8 | 85116-58-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-310-00-3 | Aromatic hydrocarbons, C8, catalytic reforming-derived; Low boiling point cat-reformed naphtha | 295-279-0 | 91995-18-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-311-00-9 | Aromatic hydrocarbons, C7-12, C8-rich; Low boiling point cat-reformed naphtha; [A complex combination of hydrocarbons obtained by separation from the platformate-containing fraction. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C7 through C12 (primarily C8) and can contain nonaromatic hydrocarbons, both boiling in the range of approximately 130°C to 200°C (266°F to 392°F).] | 297-401-8 | 93571-75-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-312-00-4 | Gasoline, C5-11, high-octane stabilised reformed; Low boiling point cat-reformed naphtha; [A complex high octane combination of hydrocarbons obtained by the catalytic dehydrogenation of a predominantly naphthenic naphtha. It consists predominantly of aromatics and non-aromatics having carbon numbers predominantly in the range of C5 through C11 and boiling in the range of approximately 45°C to 185°C (113°F to 365°F).] | 297-458-9 | 93572-29-3 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-313-00-X | Hydrocarbons, C7-12, C≥9-arom.-rich, reforming heavy fraction; Low boiling point cat-reformed naphtha; [A complex combination of hydrocarbons obtained by separation from the platformate-containing fraction. It consists predominantly of nonaromatic hydrocarbons having carbon numbers predominantly in the range of C7 through C12 and boiling in the range of approximately 120°C to 210°C (248°F to 380°F) and C9 and higher aromatic hydrocarbons.] | 297-465-7 | 93572-35-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-314-00-5 | Hydrocarbons, C5-11, nonaroms.-rich, reforming light fraction; Low boiling point cat-reformed naphtha; [A complex combination of hydrocarbons obtained by separation from the platformate-containing fraction. It consists predominantly of nonaromatic hydrocarbons having carbon numbers predominantly in the range of C5 through C11 and boiling in the range of approximately 35°C to 125°C (94°F to 257°F), benzene and toluene.] | 297-466-2 | 93572-36-2 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-315-00-0 | Foots oil (petroleum), silicic acid-treated; Foots oil; [A complex combination of hydrocarbons obtained by the treatment of Foots oil with silicic acid for removal of trace constituents and impurities. It consists predominantly of straight chain hydrocarbons having carbon numbers predominantly greater than C12.] | 308-127-6 | 97862-77-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-316-00-6 | Naphtha (petroleum), light thermal cracked; Low boiling point thermally cracked naphtha; [A complex combination of hydrocarbons from distillation of products from a thermal cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C4 through C8 and boiling in the range of approximately -10 °C to 130 °C (14 °F to 266 °F).] | 265-075-6 | 64741-74-8 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-317-00-1 | Naphtha (petroleum), heavy thermal cracked; Low boiling point thermally cracked naphtha; [A complex combination of hydrocarbons from distillation of the products from a thermal cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C6 through C12 and boiling in the range of approximately 65°C to 220°C (148°F to 428°F).] | 265-085-0 | 64741-83-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-318-00-7 | Distillates (petroleum), heavy arom.; Low boiling point thermally cracked naphtha; [The complex combination of hydrocarbons from the distillation of the products from the thermal cracking of ethane and propane. This higher boiling fraction consists predominantly of C5-7 aromatic hydrocarbons with some unsaturated aliphatic hydrocarbons having carbon number predominantly of C5. This stream may contain benzene.] | 267-563-4 | 67891-79-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-319-00-2 | Distillates (petroleum), light arom.; Low boiling point thermally cracked naphtha; [The complex combination of hydrocarbons from the distillation of the products from the thermal cracking of ethane and propane. This lower boiling fraction consists predominantly of C5-7 aromatic hydrocarbons with some unsaturated aliphatic hydrocarbons having a carbon number predominantly of C5. This stream may contain benzene.] | 267-565-5 | 67891-80-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-320-00-8 | Distillates (petroleum), naphtha-raffinate pyrolyzate-derived, gasoline-blending; Low boiling point thermally cracked naphtha; [The complex combination of hydrocarbons obtained by the pyrolysis fractionation at 816°C (1500°F) of naphtha and raffinate. It consists predominantly of hydrocarbons having a carbon number of C9 and boiling at approximately 204°C (400°F).] | 270-344-6 | 68425-29-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-321-00-3 | Aromatic hydrocarbons, C6-8, naphtha-raffinate pyrolyzate-derived; Low boiling point thermally cracked naphtha; [A complex combination of hydrocarbons obtained by the fractionation pyrolysis at 816°C (1500°F) of naphtha and raffinate. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C6 through C8, including benzene.] | 270-658-3 | 68475-70-7 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-322-00-9 | Distillates (petroleum), thermal cracked naphtha and gas oil; Low boiling point thermally cracked naphtha; [A complex combination of hydrocarbons produced by distillation of thermally cracked naphtha and/or gas oil. It consists predominantly of olefinic hydrocarbons having a carbon number of C5 and boiling in the range of approximately 33°C to 60°C (91°F to 140°F).] | 271-631-9 | 68603-00-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-323-00-4 | Distillates (petroleum), thermal cracked naphtha and gas oil, C5-dimer-contg.; Low boiling point thermally cracked naphtha; [A complex combination of hydrocarbons produced by the extractive distillation of thermal cracked naphtha and/or gas oil. It consists predominantly of hydrocarbons having a carbon number of C5 with some dimerized C5 olefins and boiling in the range of approximately 33°C to 184°C (91°F to 363°F).] | 271-632-4 | 68603-01-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-324-00-X | Distillates (petroleum), thermal cracked naphtha and gas oil, extractive; Low boiling point thermally cracked naphtha; [A complex combination of hydrocarbons produced by the extractive distillation of thermal cracked naphtha and/or gas oil. It consists of paraffinic and olefinic hydrocarbons, predominantly isoamylenes such as 2-methyl-1-butene and 2-methyl-2-butene and boiling in the range of approximately 31°C to 40°C (88°F to 104°F).] | 271-634-5 | 68603-03-2 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-325-00-5 | Distillates (petroleum), light thermal cracked, debutanized arom.; Low boiling point thermally cracked naphtha; [A complex combination of hydrocarbons produced by the distillation of products from a thermal cracking process. It consists predominantly of aromatic hydrocarbons, primarily benzene.] | 273-266-0 | 68955-29-3 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-326-00-0 | Naphtha (petroleum), light thermal cracked, sweetened; Low boiling point thermally cracked naphtha; [A complex combination of hydrocarbons obtained by subjecting a petroleum distillate from the high temperature thermal cracking of heavy oil fractions to a sweetening process to convert mercaptans. It consists predominantly of aromatics, olefins and saturated hydrocarbons boiling in the range of approximately 20°C to 100°C (68°F to 212°F).] | 295-447-3 | 92045-65-3 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-327-00-6 | Naphtha (petroleum), hydrotreated heavy; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C6 through C13 and boiling in the range of approximately 65°C to 230°C (149°F to 446°F).] | 265-150-3 | 64742-48-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-328-00-1 | Naphtha (petroleum), hydrotreated light; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C4 through C11 and boiling in the range of approximately minus 20°C to 190°C (-4°F to 374°F).] | 265-151-9 | 64742-49-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-329-00-7 | Naphtha (petroleum), hydrodesulfurized light; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained from a catalytic hydrodesulfurization process. It consists of hydrocarbons having carbon numbers predominantly in the range of C4 through C11 and boiling in the range of approximately -20°C to 190°C (-4°F to 374°F).] | 265-178-6 | 64742-73-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-330-00-2 | naphtha (petroleum), hydrodesulphurized heavy; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained from a catalytic hydrodesulfurization process. It consists of hydrocarbons having carbon numbers predominantly in the range of C7 through C12 and boiling in the range of approximately 90 °C to 230 °C (194 °F to 446 °F).] | 265-185-4 | 64742-82-1 | Carc. 1B Muta. 1B Asp. Tox. 1 STOT RE 1 | H350 H340 H304 H372 (central nervous system) | GHS08 Dgr | H340 H350 H372 (central nervous system) H304 |  |  | P | CLP00/ATP05
649-331-00-8 | Distillates (petroleum), hydrotreated middle, intermediate boiling; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained by the distillation of products from a middle distillate hydrotreating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C5 through C10 and boiling in the range of approximately 127°C to 188°C (262°F to 370°F).] | 270-092-7 | 68410-96-8 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-332-00-3 | Distillates (petroleum), light distillate hydrotreating process, low-boiling; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained by the distillation of products from the light distillate hydrotreating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C6 through C9 and boiling in the range of approximately 3°C to 194°C (37°F to 382°F).] | 270-093-2 | 68410-97-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-333-00-9 | Distillates (petroleum), hydrotreated heavy naphtha, deisohexanizer overheads; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained by distillation of the products from a heavy naphtha hydrotreating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C3 through C6 and boiling in the range of approximately -49°C to 68°C (-57°F to 155°F).] | 270-094-8 | 68410-98-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-334-00-4 | Solvent naphtha (petroleum), light arom., hydrotreated; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C8 through C10 and boiling in the range of approximately 135°C to 210°C (275°F to 410°F).] | 270-988-8 | 68512-78-7 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-335-00-X | Naphtha (petroleum), hydrodesulfurized thermal cracked light; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained by fractionation of hydrodesulfurized thermal cracker distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C5 to C11 and boiling in the range of approximately 23°C to 195°C (73°F to 383°F).] | 285-511-9 | 85116-60-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-336-00-5 | Naphtha (petroleum), hydrotreated light, cycloalkane-contg.; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained from the distillation of a petroleum fraction. It consists predominantly of alkanes and cycloalkanes boiling in the range of approximately -20°C to 190°C (-4°F to 374°F).] | 285-512-4 | 85116-61-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-337-00-0 | Naphtha (petroleum), heavy steam-cracked, hydrogenated; Low boiling point hydrogen treated naphtha | 295-432-1 | 92045-51-7 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-338-00-6 | Naphtha (petroleum), hydrodesulfurized full-range; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained from a catalytic hydrodesulfurization process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C4 through C11 and boiling in the range of approximately 30°C to 250°C (86°F to 482°F).] | 295-433-7 | 92045-52-8 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-339-00-1 | Naphtha (petroleum), hydrotreated light steam-cracked; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained by treating a petroleum fraction, derived from a pyrolysis process, with hydrogen in the presence of a catalyst. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C5 through C11 and boiling in the range of approximately 35°C to 190°C (95°F to 374°F).] | 295-438-4 | 92045-57-3 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-340-00-7 | Hydrocarbons, C4-12, naphtha-cracking, hydrotreated; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained by distillation from the product of a naphtha steam cracking process and subsequent catalytic selective hydrogenation of gum formers. It consists of hydrocarbons having carbon numbers predominantly in the range of C4 through C12 and boiling in the range of approximately 30°C to 230°C (86°F to 446°F).] | 295-443-1 | 92045-61-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-341-00-2 | Solvent naphtha (petroleum), hydrotreated light naphthenic; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists predominantly of cycloparaffinic hydrocarbons having carbon numbers predominantly in the range of C6 through C7 and boiling in the range of approximately 73°C to 85°C (163°F to 185°F).] | 295-529-9 | 92062-15-2 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-342-00-8 | Naphtha (petroleum), light steam-cracked, hydrogenated; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons produced from the separation and subsequent hydrogenation of the products of a steam-cracking process to produce ethylene. It consists predominantly of saturated and unsaturated paraffins, cyclic paraffins and cyclic aromatic hydrocarbons having carbon numbers predominantly in the range of C4 through C10 and boiling in the range of approximately 50°C to 200°C (122°F to 392°F). The proportion of benzene hydrocarbons may vary up to 30 wt. % and the stream may also contain small amounts of sulfur and oxygenated compounds.] | 296-942-7 | 93165-55-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-343-00-3 | Hydrocarbons, C6-11, hydrotreated, dearomatized; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained as solvents which have been subjected to hydrotreatment in order to convert aromatics to naphthenes by catalytic hydrogenation.] | 297-852-0 | 93763-33-8 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-344-00-9 | Hydrocarbons, C9-12, hydrotreated, dearomatized; Low boiling point hydrogen treated naphtha; [A complex combination of hydrocarbons obtained as solvents which have been subjected to hydrotreatment in order to convert aromatics to naphthenes by catalytic hydrogenation.] | 297-853-6 | 93763-34-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-345-00-4 | stoddard solvent; Low boiling point naphtha - unspecified; [A colourless, refined petroleum distillate that is free from rancid or objectionable odours and that boils in a range of approximately 148,8 °C to 204,4 °C (300 °F to 400 °F).] | 232-489-3 | 8052-41-3 | Carc. 1B Muta. 1B Asp. Tox. 1 STOT RE 1 | H350 H340 H304 H372 (central nervous system) | GHS08 Dgr | H340 H350 H372 (central nervous system) H304 |  |  | P | CLP00/ATP05
649-346-00-X | Natural gas condensates (petroleum); Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons separated as a liquid from natural gas in a surface separator by retrograde condensation. It consists mainly of hydrocarbons having carbon numbers predominantly in the range of C2 to C20. It is a liquid at atmospheric temperature and pressure.] | 265-047-3 | 64741-47-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-347-00-5 | Natural gas (petroleum), raw liq. mix; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons separated as a liquid from natural gas in a gas recycling plant by processes such as refrigeration or absorption. It consists mainly of saturated aliphatic hydrocarbons having carbon numbers in the range of C2 through C8.] | 265-048-9 | 64741-48-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-348-00-0 | Naphtha (petroleum), light hydrocracked; Low boiling naphtha - unspecified; [A complex combination of hydrocarbons from distillation of the products from a hydrocracking process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C4 through C10, and boiling in the range of approximately -20°C to 180°C (-4°F to 356°F).] | 265-071-4 | 64741-69-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-349-00-6 | Naphtha (petroleum), heavy hydrocracked; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons from distillation of the products from a hydrocracking process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C6 through C12, and boiling in the range of approximately 65°C to 230°C (148°F to 446°F).] | 265-079-8 | 64741-78-2 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-350-00-1 | Naphtha (petroleum), sweetened; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by subjecting a petroleum naphtha to a sweetening process to convert mercaptans or to remove acidic impurities. It consists of hydrocarbons having carbon numbers predominantly in the range of C4 through C12 and boiling in the range of approximately -10°C to 230°C (14°F to 446°F).] | 265-089-2 | 64741-87-3 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-351-00-7 | Naphtha (petroleum), acid-treated; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C7 through C12 and boiling in the range of approximately 90°C to 230°C (194°F to 446°F).] | 265-115-2 | 64742-15-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-352-00-2 | Naphtha (petroleum), chemically neutralized heavy; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C6 through C12 and boiling in the range of approximately 65°C to 230°C (149°F to 446°F).] | 265-122-0 | 64742-22-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-353-00-8 | Naphtha (petroleum), chemically neutralized light; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C4 through C11 and boiling in the range of approximately -20°C to 190°C (-4°F to 374°F).] | 265-123-6 | 64742-23-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-354-00-3 | Naphtha (petroleum), catalytic dewaxed; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained from the catalytic dewaxing of a petroleum fraction. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C5 through C12 and boiling in the range of approximately 35°C to 230°C (95°F to 446°F).] | 265-170-2 | 64742-66-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-355-00-9 | Naphtha (petroleum), light steam-cracked; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by the distillation of the products from a steam cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C4 through C11 and boiling in the range of approximately minus 20°C to 190°C (-4°F to 374°F). This stream is likely to contain 10 vol. % or more benzene.] | 265-187-5 | 64742-83-2 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-356-00-4 | Solvent naphtha (petroleum), light arom.; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained from distillation of aromatic streams. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C8 through C10 and boiling in the range of approximately 135°C to 210°C (275°F to 410°F).] | 265-199-0 | 64742-95-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-357-00-X | Aromatic hydrocarbons, C6-10, acid-treated, neutralized; Low boiling point naphtha - unspecified | 268-618-5 | 68131-49-7 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-358-00-5 | Distillates (petroleum), C3-5, 2-methyl-2-butene-rich; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons from the distillation of hydrocarbons usually ranging in carbon numbers from C3 through C5, predominantly isopentane and 3-methyl-1-butene. It consists of saturated and unsaturated hydrocarbons having carbon numbers in the range of C3 through C5, predominantly 2-methyl-2-butene.] | 270-725-7 | 68477-34-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-359-00-0 | Distillates (petroleum), polymd. steam-cracked petroleum distillates, C5-12 fraction; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained from the distillation of polymerized steam-cracked petroleum distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C5 through C12.] | 270-735-1 | 68477-50-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-360-00-6 | Distillates (petroleum), steam-cracked, C5-12 fraction; Low boiling point naphtha - unspecified; [A complex combination of organic compounds obtained by the distillation of products from a steam cracking process. It consists of unsaturated hydrocarbons having carbon numbers predominantly in the range of C5 through C12.] | 270-736-7 | 68477-53-2 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-361-00-1 | Distillates (petroleum), steam-cracked, C5-10 fraction, mixed with light steam-cracked petroleum naphtha C5 fraction; Low boiling point naphtha - unspecified | 270-738-8 | 68477-55-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-362-00-7 | Extracts (petroleum), cold-acid, C4-6; Low boiling point naphtha - unspecified; [A complex combination of organic compounds produced by cold acid unit extraction of saturated and unsaturated aliphatic hydrocarbons usually ranging in carbon numbers from C3 through C6, predominantly pentanes and amylenes. It consists predominantly of saturated and unsaturated hydrocarbons having carbon numbers in the range of C4 through C6, predominantly C5.] | 270-741-4 | 68477-61-2 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-363-00-2 | Distillates (petroleum), depentanizer overheads; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained from a catalytic cracked gas stream. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C4 through C6.] | 270-771-8 | 68477-89-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-364-00-8 | Residues (petroleum), butane splitter bottoms; Low boiling point naphtha - unspecified; [A complex residuum from the distillation of butane stream. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C4 through C6.] | 270-791-7 | 68478-12-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H340 H350 H304 |  |  | P | CLP00/ATP01corr
649-365-00-3 | Residual oils (petroleum), deisobutanizer tower; Low boiling point naphtha - unspecified; [A complex residuum from the atmospheric distillation of the butane-butylene stream. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C4 through C6.] | 270-795-9 | 68478-16-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-366-00-9 | Naphtha (petroleum), full-range coker; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons produced by the distillation of products from a fluid coker. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C4 through C15 and boiling in the range of approximately 43°C to 250°C (110°F-500°F).] | 270-991-4 | 68513-02-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-367-00-4 | Naphtha (petroleum), steam-cracked middle arom.; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons produced by the distillation of products from a steam-cracking process. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C7 through C12 and boiling in the range of approximately 130°C to 220°C (266°F to 428°F).] | 271-138-9 | 68516-20-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-368-00-X | Naphtha (petroleum), clay-treated full-range straight-run; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons resulting from treatment of full-range straight-run naphtha with natural or modified clay, usually in a percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C4 through C11 and boiling in the range of approximately -20°C to 220°C (-4°F to 429°F).] | 271-262-3 | 68527-21-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-369-00-5 | Naphtha (petroleum), clay-treated light straight-run; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons resulting from treatment of light straight-run naphtha with a natural or modified clay, usually in a percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C7 through C10 and boiling in the range of approximately 93°C to 180°C (200°F to 356°F).] | 271-263-9 | 68527-22-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-370-00-0 | Naphtha (petroleum), light steam-cracked arom.; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons produced by distillation of products from a steam-cracking process. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C7 through C9 and boiling in the range of approximately 110°C to 165°C (230°F to 329°F).] | 271-264-4 | 68527-23-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-371-00-6 | Naphtha (petroleum), light steam-cracked, debenzenized; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons produced by distillation of products from a steam-cracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C4 through C12 and boiling in the range of approximately 80°C to 218°C (176°F to 424°F).] | 271-266-5 | 68527-26-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-372-00-1 | Naphtha (petroleum), arom.-contg.; Low boiling point naphtha - unspecified | 271-635-0 | 68603-08-7 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-373-00-7 | Gasoline, pyrolysis, debutanizer bottoms; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained from the fractionation of depropanizer bottoms. It consists of hydrocarbons having carbon numbers predominantly greater than C5.] | 271-726-5 | 68606-10-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-374-00-2 | Naphtha (petroleum), light, sweetened; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by subjecting a petroleum distillate to a sweetening process to convert mercaptans or to remove acidic impurities. It consists predominantly of saturated and unsaturated hydrocarbons having carbon numbers predominantly in the range of C3 through C6 and boiling in the range of approximately -20°C to 100°C (-4°F to 212°F).] | 272-206-0 | 68783-66-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-375-00-8 | Natural gas condensates; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons separated and/or condensed from natural gas during transportation and collected at the wellhead and/or from the production, gathering, transmission, and distribution pipelines in deeps, scrubbers, etc. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C2 through C8.] | 272-896-3 | 68919-39-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-376-00-3 | Distillates (petroleum), naphtha unifiner stripper; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons produced by stripping the products from the naphtha unifiner. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C6.] | 272-932-8 | 68921-09-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-377-00-9 | Naphtha (petroleum), catalytic reformed light, arom.-free fraction; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons remaining after removal of aromatic compounds from catalytic reformed light naphtha in a selective absorption process. It consists predominantly of paraffinic and cyclic compounds having carbon numbers predominantly in the range of C5 to C8 and boiling in the range of approximately 66°C to 121°C (151°F to 250°F).] | 285-510-3 | 85116-59-2 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-378-00-4 | Gasoline; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons consisting primarily of paraffins, cycloparaffins, aromatic and olefinic hydrocarbons having carbon numbers predominantly greater than C3 and boiling in the range of 30°C to 260°C (86°F to 500°F).] | 289-220-8 | 86290-81-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-379-00-X | Aromatic hydrocarbons, C7-8, dealkylation products, distn. residues; Low boiling point naphtha - unspecified | 292-698-0 | 90989-42-7 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-380-00-5 | Hydrocarbons, C4-6, depentanizer lights, arom. hydrotreater; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained as first runnings from the depentanizer column before hydrotreatment of the aromatic charges. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C4 through C6, predominantly pentanes and pentenes, and boiling in the range of approximately 25°C to 40°C (77°F to 104°F).] | 295-298-4 | 91995-38-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-381-00-0 | Distillates (petroleum), heat-soaked steam-cracked naphtha, C5-rich; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by distillation of heat-soaked steam-cracked naphtha. It consists predominantly of hydrocarbons having carbon numbers in the range of C4 through C6, predominantly C5.] | 295-302-4 | 91995-41-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-382-00-6 | Extracts (petroleum), catalytic reformed light naphtha solvent; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained as the extract from the solvent extraction of a catalytically reformed petroleum cut. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C7 through C8 and boiling in the range of approximately 100°C to 200°C (212°F to 392°F).] | 295-331-2 | 91995-68-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-383-00-1 | Naphtha (petroleum), hydrodesulfurized light, dearomatized; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by distillation of hydrodesulfurized and dearomatized light petroleum fractions. It consists predominantly of C7 paraffins and cycloparaffins boiling in a range of approximately 90°C to 100°C (194°F to 212°F).] | 295-434-2 | 92045-53-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-384-00-7 | Naphtha (petroleum), light, C5-rich, sweetened; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by subjecting a petroleum naphtha to a sweetening process to convert mercaptans or to remove acidic impurities. It consists of hydrocarbons having carbon numbers predominantly in the range of C4 through C5, predominantly C5, and boiling in the range of approximately minus 10°C to 35°C (14°F to 95°F).] | 295-442-6 | 92045-60-8 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-385-00-2 | Hydrocarbons, C8-11, naphtha-cracking, toluene cut; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by distillation from prehydrogenated cracked naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C8 through C11 and boiling in the range of approximately 130°C to 205°C (266°F to 401°F).] | 295-444-7 | 92045-62-0 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-386-00-8 | Hydrocarbons, C4-11, naphtha-cracking, arom.-free; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained from prehydrogenated cracked naphtha after distillative separation of benzene- and toluene-containing hydrocarbon cuts and a higher boiling fraction. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C4 through C11 and boiling in the range of approximately 30°C to 205°C (86°F to 401°F).] | 295-445-2 | 92045-63-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-387-00-3 | Naphtha (petroleum), light heat-soaked, steam-cracked; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by the fractionation of steam cracked naphtha after recovery from a heat soaking process. It consists predominantly of hydrocarbons having a carbon number predominantly in the range of C4 through C6 and boiling in the range of approximately 0°C to 80°C (32°F to 176°F).] | 296-028-8 | 92201-97-3 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-388-00-9 | Distillates (petroleum), C6-rich; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained from the distillation of a petroleum feedstock. It consists predominantly of hydrocarbons having carbon numbers of C5 through C7, rich in C6, and boiling in the range of approximately 60°C to 70°C (140°F to 158°F).] | 296-903-4 | 93165-19-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-389-00-4 | Gasoline, pyrolysis, hydrogenated; Low boiling point naphtha-unspecified; [A distillation fraction from the hydrogenation of pyrolysis gasoline boiling in the range of approximately 20°C to 200°C (68°F to 392°F).] | 302-639-3 | 94114-03-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-390-00-X | Distillates (petroleum), steam-cracked, C8-12 fraction, polymd., distn. lights; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by distillation of the polymerized C8 through C12 fraction from steam-cracked petroleum distillates. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C8 through C12.] | 305-750-5 | 95009-23-7 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-391-00-5 | Extracts (petroleum) heavy naphtha solvent, clay-treated; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by the treatment of heavy naphthic solvent petroleum extract with bleaching earth. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C6 through C10 and boiling in the range of approximately 80°C to 180°C (175°F to 356°F).] | 308-261-5 | 97926-43-7 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-392-00-0 | Naphtha (petroleum), light steam-cracked, debenzenized, thermally treated; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by the treatment and distillation of debenzenized light steam-cracked petroleum naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C7 through C12 and boiling in the range of approximately 95°C to 200°C (203°F to 392°F).] | 308-713-1 | 98219-46-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-393-00-6 | Naphtha (petroleum), light steam-cracked, thermally treated; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by the treatment and distillation of light steam-cracked petroleum naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C5 through C6 and boiling in the range of approximately 35°C to 80°C (95°F to 176°F).] | 308-714-7 | 98219-47-7 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-394-00-1 | Distillates (petroleum), C7-9, C8-rich, hydrodesulfurized dearomatized; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by the distillation of petroleum light fraction, hydrodesulfurized and dearomatized. It consists predominantly of hydrocarbons having carbon numbers in the range of C7 through C9, predominantly C8 paraffins and cycloparaffins, boiling in the range of approximately 120°C to 130°C (248°F to 266°F).] | 309-862-5 | 101316-56-7 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-395-00-7 | Hydrocarbons, C6-8, hydrogenated sorption-dearomatized, toluene raffination; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained during the sorptions of toluene from a hydrocarbon fraction from cracked gasoline treated with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C6 through C8 and boiling in the range of approximately 80°C to 135°C (176°F to 275°F).] | 309-870-9 | 101316-66-9 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-396-00-2 | Naphtha (petroleum), hydrodesulfurised full-range coker; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by fractionation from hydrodesulfurised coker distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C5 to C11 and boiling in the range of approximately 23°C to 196°C (73°F to 385°F).] | 309-879-8 | 101316-76-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-397-00-8 | Naphtha (petroleum), sweetened light; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by subjecting a petroleum naphtha to a sweetening process to convert mercaptans or to remove acidic impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C5 through C8 and boiling in the range of approximately 20°C to 130°C (68°F to 266°F).] | 309-976-5 | 101795-01-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-398-00-3 | Hydrocarbons, C3-6, C5-rich, steam-cracked naphtha; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by distillation of steam-cracked naphtha. It consists predominantly of hydrocarbons having carbon numbers in the range of C3 through C6, predominantly C5.] | 310-012-0 | 102110-14-5 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-399-00-9 | Hydrocarbons, C5-rich, dicyclopentadiene-contg.; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by distillation of the products from a steam-cracking process. It consists predominantly of hydrocarbons having carbon numbers of C5 and dicyclopentadiene and boiling in the range of approximately 30°C to 170°C (86°F to 338°F).] | 310-013-6 | 102110-15-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-400-00-2 | Residues (petroleum), steam-cracked light, arom.; Low boiling point naphtha - unspecified; [A complex combination of hydrocarbons obtained by the distillation of the products of steam cracking or similar processes after taking off the very light products resulting in a residue starting with hydrocarbons having carbon numbers greater than C5. It consists predominantly of aromatic hydrocarbons having carbon numbers greater than C5 and boiling above approximately 40°C (104°F).] | 310-057-6 | 102110-55-4 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-401-00-8 | Hydrocarbons, C≥5, C5-6-rich; Low boiling point naphtha - unspecified | 270-690-8 | 68476-50-6 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-402-00-3 | Hydrocarbons, C5-rich; Low boiling point naphtha - unspecified | 270-695-5 | 68476-55-1 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-403-00-9 | Aromatic hydrocarbons, C8-10; Low boiling point naphtha - unspecified | 292-695-4 | 90989-39-2 | Carc. 1B Muta. 1B Asp. Tox. 1 | H350 H340 H304 | GHS08 Dgr | H350 H340 H304 |  |  | P | CLP00/ATP01
649-404-00-4 | Kerosine (petroleum); Straight run kerosine; [A complex combination of hydrocarbons produced by the distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C9 through C16 and boiling in the range of approximately 150 °C to 290 °C (320 °F to 554 °F).] | 232-366-4 | 8008-20-6 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-405-00-X | solvent naphtha (petroleum), medium aliph.; Straight run kerosine; [A complex combination of hydrocarbons obtained from the distillation of crude oil or natural gasoline. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C9 through C12 and boiling in the range of approximately 140 °C to 220 °C (284 °F to 428 °F).] | 265-191-7 | 64742-88-7 | Asp. Tox. 1 STOT RE 1 | H304 H372 (central nervous system) | GHS08 Dgr | H372 (central nervous system) H304 |  |  |  | CLP00/ATP05
649-406-00-5 | Solvent naphtha (petroleum) heavy aliph.; Straight run kerosine; [A complex combination of hydrocarbons obtained from the distillation of crude oil or natural gasoline. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C11 through C16 and boiling in the range of approximately 190 °C to 290 °C (374 °F to 554 °F).] | 265-200-4 | 64742-96-7 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-407-00-0 | Kerosine (petroleum), straight-run wide-cut; Straight run kerosine; [A complex combination of hydrocarbons obtained as a wide cut hydrocarbon fuel cut from atmospheric distillation and boiling in the range of approximately 70 °C to 220 °C (158 °F to 428 °F).] | 295-418-5 | 92045-37-9 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-408-00-6 | Distillates (petroleum), steam-cracked; Cracked kerosine; [A complex combination of hydrocarbons obtained by the distillation of the products from a steam cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C7 through C16 and boiling in the range of approximately 90 °C to 290 °C (190 °F to 554 °F).] | 265-194-3 | 64742-91-2 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-409-00-1 | Distillates (petroleum), cracked stripped steam-cracked petroleum distillates, C8-10 fraction; Cracked kerosine; [A complex combination of hydrocarbons obtained by distilling cracked stripped steam-cracked distillates. It consists of hydro-carbons having carbon numbers in the range of C8 through C10 and boiling in the range of approximately 129 °C to 194 °C (264 °F to 382 °F).] | 270-728-3 | 68477-39-4 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-410-00-7 | Distillates (petroleum), cracked stripped steam-cracked petroleum distillates, C10-12 fraction; Cracked kerosine; [A complex combination of hydrocarbons obtained by distilling cracked stripped steam-cracked distillates. It consists predominantly of aromatic hydrocarbons having carbon numbers in the range of C10 through C12.] | 270-729-9 | 68477-40-7 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-411-00-2 | Distillates (petroleum), steam-cracked, C8-12 fraction; Cracked kerosine; [A complex combination of organic compounds obtained by the distillation of products from a steam cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C8 through C12.] | 270-737-2 | 68477-54-3 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-412-00-8 | Kerosine (petroleum), hydrodesulfurized thermal cracked; Cracked kerosine; [A complex combination of hydrocarbons obtained by fractionation from hydrodesulfurized thermal cracker distillate. It consists predominantly of hydrocarbons predominantly in the range of C8 to C16 and boiling in the range of approximately 120 °C to 283 °C (284 °F to 541 °F).] | 285-507-7 | 85116-55-8 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-413-00-3 | Aromtic hydrocarbons, C≥10, steam-cracking, hydrotreated; Cracked kerosine; [A complex combination of hydrocarbons produced by the distillation of the products from a steam cracking process treated with hydrogen in the presence of a catalyst. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly greater than C10 and boiling in the range of approximately 150 °C to 320 °C (302 °F to 608 °F).] | 292-621-0 | 90640-98-5 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-414-00-9 | Naphtha (petroleum), steam-cracked, hydrotreated, C9-10-arom.-rich; Cracked kerosine; [A complex combination of hydrocarbons produced by the distillation of the products from a steam cracking process thereafter treated with hydrogen in the presence of a catalyst. It consists predominantly of aromatic hydrocarbons having carbon numbers in the range of C9 through C10 and boiling in the range of approximately 140 °C to 200 °C (284 °F to 392 °F).] | 292-637-8 | 90641-13-7 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-415-00-4 | Distillates (petroleum), thermal-cracked, alkylarom. hydrocarbon-rich; Cracked kerosine; [A complex combination of hydrocarbons obtained by distillation of thermal-cracking heavy tars. It consists predominantly of highly alkylated aromatic hydrocarbons boiling in the range of approximately 100 °C to 250 °C (212 °F to 482 °F.] | 309-866-7 | 101316-61-4 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-416-00-X | Distillates (petroleum), catalytic cracked heavy tar light; Cracked kerosine; [A complex combination of hydrocarbons obtained by distillation of catalytic cracking heavy tars. It consists predominantly of highly alkylated aromatic hydrocarbons boiling in the range of approximately 100 °C to 250 °C (212 °F to 482 °F).] | 309-938-8 | 101631-13-4 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-417-00-5 | Solvent naphtha (petroleum), hydrocracked heavy arom.; Cracked kerosine; [A complex combination of hydrocarbons obtained by the distillation of hydrocracked petroleum distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C9 through C16 and boiling in the range of approximately 235 °C to 290 °C (455 °F to 554 °F).] | 309-881-9 | 101316-80-7 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-418-00-0 | Distillates (petroleum), steam-cracked heavy tar light; Cracked kerosine; [A complex combination of hydrocarbons obtained by distillation of steam cracking heavy tars. It consists predominantly of highly alkylated aromatic hydrocarbons boiling in the range of approximately 100 °C to 250 °C (212 °F to 482 °F).] | 309-940-9 | 101631-15-6 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-419-00-6 | Distillates (petroleum), alkylate; Kerosine - unspecified; [A complex combination of hydrocarbons produced by distillation of the reaction products of isobutane with monoolefinic hydrocarbons usually ranging in carbon numbers from C3 through C5. It consists of predominantly branched chain saturated hydro-carbons having carbon numbers predominantly in the range of C11 through C17 and boiling in the range of approximately 205 °C to 320 °C (401 °F to 608 °F).] | 265-074-0 | 64741-73-7 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-420-00-1 | Extracts (petroleum), heavy naphtha solvent; Kerosine - unspecified; [A complex combination of hydrocarbons obtained as the extract from a solvent extraction process. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C7 through C12 and boiling in the range of approximately 90 °C to 220 °C (194 °F to 428 °F).] | 265-099-7 | 64741-98-6 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-421-00-7 | Distillates (petroleum), chemically neutralized light; Kerosine - unspecified; [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C9 through C16 and boiling in the range of approximately 150 °C to 290 °C (302 °F to 554 °F).] | 265-132-5 | 64742-31-0 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-422-00-2 | Distillates (petroleum), hydrotreated light; Kerosine - unspecified; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C9 through C16 and boiling in the range of approximately 150 °C to 290 °C (302 °F to 554 °F).] | 265-149-8 | 64742-47-8 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-423-00-8 | Kerosine (petroleum), hydrodesulfurized; Kerosine - unspecified; [A complex combination of hydrocarbons obtained from a petroleum stock by treating with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists of hydrocarbons having carbon numbers predominantly in the range of C9 through C16 and boiling in the range of approximately 150 °C to 290 °C (302 °F to 554 °F).] | 265-184-9 | 64742-81-0 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-424-00-3 | Solvent naphtha (petroleum), heavy arom.; Kerosine - unspecified; [A complex combination of hydrocarbons obtained from distillation of aromatic streams. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C9 through C16 and boiling in the range of approximately 165 °C to 290 °C (330 °F to 554 °F).] | 265-198-5 | 64742-94-5 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-425-00-9 | Naphtha (petroleum), heavy coker; Kerosine - unspecified; [A complex combination of hydrocarbons from the distillation of products from a fluid coker. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C6 through C15 and boiling in the range of approximately 157 °C to 288 °C (315 °F to 550 °F).] | 269-778-9 | 68333-23-3 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-426-00-4 | Naphtha (petroleum), catalytic reformed hydrodesulfurized heavy, arom. fraction; Kerosine - unspecified; [A complex combination of hydrocarbons produced by fractionation from catalytically reformed hydrodesulfurized naphtha. It consists predominantly of aromatic hydrocarbons having carbon numbers predominently in the range of C7 to C 13 and boiling in the range of approximately 98 °C to 218 °C (208 °F to 424 °F).] | 285-508-2 | 85116-57-0 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-427-00-X | Kerosine (petroleum), sweetened; Kerosine - unspecified; [A complex combination of hydrocarbons obtained by subjecting a petroleum distillate to a sweetening process to convert mercaptans or to remove acidic impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C9 through C16 and boiling in the range of 130 °C to 290 °C (266 °F to 554 °F).] | 294-799-5 | 91770-15-9 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-428-00-5 | Kerosine (petroleum), solvent-refined sweetened; Kerosine - unspecified; [A complex combination of hydrocarbons obtained from a petroleum stock by solvent refining and sweetening and boiling in the range of approximately 150 °C to 260 °C (302 °F to 500 °F).] | 295-416-4 | 92045-36-8 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-429-00-0 | Hydrocarbons, C9-16, hydrotreated, dearomatized; Kerosine - unspecified; [A complex combination of hydrocarbons obtained as solvents which have been subjected to hydrotreatment in order to convert aromatics to naphthenes by catalytic hydrogenation.] | 297-854-1 | 93763-35-0 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-430-00-6 | Kerosine (petroleum), solvent-refined hydrodesulfurized; Kerosine - unspecified | 307-033-2 | 97488-94-3 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-431-00-1 | Distillates (petroleum), hydrodesulfurized full-range middle coker; Kerosine - unspecified; [A complex combination of hydrocarbons obtained by fractionation from hydrodesulfurized coker distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C8 through C16 and boiling in the range of approximately 120 °C to 283 °C (248 °F to 541 °F).] | 309-864-6 | 101316-58-9 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-432-00-7 | Solvent naphtha (petroleum), hydrodesulfurized heavy arom.; Kerosine - unspecified; [A complex combination of hydrocarbons obtained by the catalytic hydrodesulfurization of a petroleum fraction. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C10 through C13 and boiling in the range of approximately 180 °C to 240 °C (356 °F to 464 °F).] | 309-882-4 | 101316-81-8 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-433-00-2 | Solvent naphtha (petroleum), hydrodesulfurized medium; Kerosine - unspecified; [A complex combination of hydrocarbons obtained by the catalytic hydrodesulfurization of a petroleum fraction. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C10 through C13 and boiling in the range of approximately 175 °C to 220 °C (347 °F to 428 °F).] | 309-884-5 | 101316-82-9 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-434-00-8 | Kerosine (petroleum), hydrotreated; Kerosine - unspecified; [A complex combination of hydrocarbons obtained from the distillation of petroleum and subsequent hydrotreatment. It consists predominantly of alkanes, cycloalkanes and alkylbenzenes having carbon numbers predominantly in the range of C12 through C16 and boiling in the range of approximately 230 °C to 270 °C (446 °F to 518 °F).] | 309-944-0 | 101631-19-0 | Asp. Tox. 1 | H304 | GHS08 Dgr | H304 |  |  |  | CLP00
649-435-00-3 | Distillates (petroleum), light catalytic cracked; Cracked gasoil; [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C9 through C25 and boiling in the range of approximately 150 °C to 400 °C (302 °F to 752 °F). It contains a relatively large proportion of bicyclic aromatic hydrocarbons.] | 265-060-4 | 64741-59-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-436-00-9 | Distillates (petroleum), intermediate catalytic cracked; Cracked gasoil; [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C11 through C30 and boiling in the range of approximately 205 °C to 450 °C (401 °F to 842 °F). It contains a relatively large proportion of tricyclic aromatic hydrocarbons.] | 265-062-5 | 64741-60-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-437-00-4 | Distillates (petroleum), light hydrocracked; Cracked gasoil; [A complex combination of hydrocarbons from distillation of the products from a hydrocracking process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C10 through C18 and boiling in the range of approximately 160 °C to 320 °C (320 °F to 608 °F).] | 265-078-2 | 64741-77-1 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | CLP00
649-438-00-X | Distillates (petroleum), light thermal cracked; Cracked gasoil; [A complex combination of hydrocarbons from the distillation of the products from a thermal cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C10 through C22 and boiling in the range of approximately 160 °C to 370 °C (320 °F to 698 °F).] | 265-084-5 | 64741-82-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-439-00-5 | Distillates (petroleum), hydrodesulfurized light catalytic cracked; Cracked gasoil; [A complex combination of hydrocarbons obtained by treating light catalytic cracked distillates with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists of hydrocarbons having carbon numbers predominantly in the range of C9 through C25 and boiling in the range of approximately 150 °C to 400 °C (302 °F to 752 °F). It contains a relatively large proportion of bicyclic aromatic hydrocarbons.] | 269-781-5 | 68333-25-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-440-00-0 | Distillates (petroleum), light steam-cracked naphtha; Cracked gasoil; [A complex combination of hydrocarbons from the multiple distillation of products from a steam cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C10 through C18.] | 270-662-5 | 68475-80-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-441-00-6 | Distillates (petroleum), cracked steam-cracked petroleum distillates; Cracked gasoil; [A complex combination of hydrocarbons produced by distilling cracked steam cracked distillate and/or its fractionation products. It consists of hydrocarbons having carbon numbers predominently in the range of C10 to low molecular weight polymers.] | 270-727-8 | 68477-38-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-442-00-1 | Gas oils (petroleum), steam-cracked; Cracked gasoil; [A complex combination of hydrocarbons produced by distillation of the products from a steam cracking process. It consists of hydrocarbons having carbon numbers predominantly greater than C9 and boiling in the range of from approximately 205 °C to 400 °C (400 °F to 752 °F).] | 271-260-2 | 68527-18-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-443-00-7 | Distillates (petroleum), hydrodesulfurized thermal cracked middle; Cracked gasoil; [A complex combination of hydrocarbons obtained by fractionation from hydrodesulfurized themal cracker distillate stocks. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C11 to C25 and boiling in the range of approximately 205 °C to 400 °C (401 °F to 752 °F).] | 285-505-6 | 85116-53-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-444-00-2 | Gas oils (petroleum), thermal-cracked, hydrodesulfurized; Cracked gasoil | 295-411-7 | 92045-29-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-445-00-8 | Residues (petroleum), hydrogenated steam-cracked naphtha; Cracked gasoil; [A complex combination of hydrocarbons obtained as a residual fraction from the distillation of hydrotreated steam-cracked naphtha. It consists predominantly of hydrocarbons boiling in the range of approximately 200 °C to 350 °C (32 °F to 662 °F).] | 295-514-7 | 92062-00-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-446-00-3 | Residues (petroleum), steam-cracked naphtha distn.; Cracked gasoil; [A complex combination of hydrocarbons obtained as a column bottom from the separation of effluents from steam cracking naphtha at a high temperature. It boils in the range of approximately 147 °C to 300 °C (297 °F to 572 °F) and produces a finished oil having a viscosity of 18cSt at 50 °C.] | 295-517-3 | 92062-04-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-447-00-9 | Distillates (petroleum), light catalytic cracked, thermally degraded; Cracked gasoil; [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process which has been used as a heat transfer fluid. It consists predominantly of hydrocarbons boiling in the range of approximately 190 °C to 340 °C (374 °F to 644 °F). This stream is likely to contain organic sulfur compounds.] | 295-991-1 | 92201-60-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-448-00-4 | Residues (petroleum), steam-cracked heat-soaked naphtha; Cracked gasoil; [A complex combination of hydrocarbons obtained as residue from the distillation of steam cracked heat soaked naphtha and boiling in the range of approximately 150 °C to 350 °C (302 °F to 662 °F).] | 297-905-8 | 93763-85-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-449-00-X | Hydrocarbons, C16-20, solvent-dewaxed hydrocracked paraffinic distn. residue; Cracked gasoil; [A complex combination of hydrocarbons obtained by solvent dewaxing of a distillation residue from a hydrocracked paraffinic distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C16 through C20 and boiling in the range of approximately 360 °C to 500 °C (680 °F to 932 °F). It produces a finished oil having a viscosity of 4,5 cSt at approximately 100 °C (212 °F).] | 307-662-2 | 97675-88-2 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | CLP00
649-450-00-5 | Gas oils (petroleum), light vacuum, thermal-cracked hydrodesulfurized; Cracked gasoil; [A complex combination of hydrocarbons obtained by catalytic dehydrosulfurization of thermal-cracked light vacuum petroleum. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C14 through C20 and boiling in the range of approximately 270 °C to 370 °C (518 °F to 698 °F).] | 308-278-8 | 97926-59-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-451-00-0 | Distillates (petroleum), hydrodesulfurized middle coker; Cracked gasoil; [A complex combination of hydrocarbons by fractionation from hydrodesulfurised coker distillate stocks. Is consists of hydro-carbons having carbon numbers predominantly in the range of C12 through C21 and boiling in the range of approximately 200 °C to 360 °C (392 °F to 680 °F).] | 309-865-1 | 101316-59-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-452-00-6 | Distillates (petroleum), heavy steam-cracked; Cracked gasoil; [A complex combination of hydrocarbons obtained by distillation of steam cracking heavy residues. It consists predominantly of highly alkylated heavy aromatic hydrocarbons boiling in the range of approximately 250 °C to 400 °C (482 °F to 752 °F).] | 309-939-3 | 101631-14-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
649-453-00-1 | Distillates (petroleum), heavy hydrocracked; Baseoil - unspecified; [A complex combination of hydrocarbons from the distillation of the products from a hydrocracking process. It consists predominantly of saturated hydrocarbons having carbon numbers in the range of C15-C39 and boiling in the range of approximately 260 °C to 600 °C (500 °F to 1112 °F).] | 265-077-7 | 64741-76-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-454-00-7 | Distillates (petroleum), solvent-refined heavy paraffinic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C).] | 265-090-8 | 64741-88-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-455-00-2 | Distillates (petroleum), solvent-refined light paraffinic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C).] | 265-091-3 | 64741-89-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-456-00-8 | Residual oils (petroleum), solvent deasphalted; Baseoil - unspecified; [A complex combination of hydrocarbons obtained as the solvent soluble fraction from C3-C4 solvent deasphalting of a residuum. It consists of hydrocarbons having carbon numbers predominantly higher than C25 and boiling above approximately 400 °C (752 °F).] | 265-096-0 | 64741-95-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-457-00-3 | Distillates (petroleum), solvent-refined heavy naphthenic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt a 40 °C). It contains relatively few normal paraffins.] | 265-097-6 | 64741-96-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-458-00-9 | Distillates (petroleum), solvent-refined light naphthenic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-098-1 | 64741-97-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-459-00-4 | Residual oils (petroleum,) solvent-refined; Baseoil - unspecified; [A complex combination by hydrocarbons obtained as the solvent insoluble fraction from solvent refining of a residuum using a polar organic solvent such as phenol or furfural. It consists of hydrocarbons having carbon numbers predominantly higher than C25 and boiling above approximately 400 °C (752 °F).] | 265-101-6 | 64742-01-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-460-00-X | Distillates (petroleum), clay-treated paraffinic; Baseoil - unspecified; [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated hydrocarbons.] | 265-137-2 | 64742-36-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-461-00-5 | Distillates (petroleum), clay-treated light paraffinic; Baseoil - unspecified; [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated hydrocarbons.] | 265-138-8 | 64742-37-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-462-00-0 | Residual oils (petroleum), clay-treated; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treatment of a residual oil with a natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydro-carbons having carbon numbers predominantly higher than C25 and boiling above approximately 400 °C (752 °F).] | 265-143-5 | 64742-41-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-463-00-6 | Distillates (petroleum), clay-treated heavy naphthenic; Baseoil - unspecified; [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-146-1 | 64742-44-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-464-00-1 | Distillates (petroleum), clay-treated light naphthenic; Baseoil - unspecified; [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-147-7 | 64742-45-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-465-00-7 | Distillates (petroleum), hydrotreated heavy naphthenic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-155-0 | 64742-52-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-466-00-2 | Distillates (petroleum), hydrotreated light naphthenic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-156-6 | 64742-53-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-467-00-8 | Distillates (petroleum), hydrotreated heavy paraffinic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated hydrocarbons.] | 265-157-1 | 64742-54-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-468-00-3 | Distillates (petroleum), hydrotreated light paraffinic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated hydrocarbons.] | 265-158-7 | 64742-55-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-469-00-9 | Distillates (petroleum), solvent-dewaxed light paraffinic; Baseoil - unspecified; [A complex comination of hydrocarbons obtained by removal of normal paraffins from a petroleum fraction by solvent crystallization. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C).] | 265-159-2 | 64742-56-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-470-00-4 | Residual oils (petroleum), hydrotreated; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly greater than C25 and boiling above approximately 400 °C (752 °F).] | 265-160-8 | 64742-57-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-471-00-X | Residual oils (petroleum), solvent-dewaxed; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by removal of long, branched chain hydrocarbons from a residual oil by solvent crystallization. It consists of hydrocarbons having carbon numbers predominantly greater than C25 and boiling above approximately 400 °C (752 °F).] | 265-166-0 | 64742-62-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-472-00-5 | Distillates (petroleum), solvent-dewaxed heavy naphthenic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by removal of normal paraffins from a petroleum fraction by solvent crystallization. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 . through C50 and produces a finished oil of not less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-167-6 | 64742-63-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-473-00-0 | Distillates (petroleum), solvent-dewaxed light naphthenic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by removal of normal paraffins from a petroleum fraction by solvent crystallization. It consists of hydrocarbons having carbon numbers predominantly in the range C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-168-1 | 64742-64-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-474-00-6 | Distillates (petroleum), solvent-dewaxed heavy paraffinic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by removal of normal paraffins from a petroleum fraction by solvent crystallization. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity not less than 100 SUS at 100 °F (19cSt at 40 °C).] | 265-169-7 | 64742-65-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-475-00-1 | Naphthenic oils (petroleum), catalytic dewaxed heavy; Baseoil - unspecified; [A complex combination of hydrocarbons obtained from a catalytic dewaxing process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-172-3 | 64742-68-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-476-00-7 | Naphthenic oils (petroleum), catalytic dewaxed light; Baseoil - unspecified; [A complex combination of hydrocarbons obtained from a catalytic dewaxing process. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-173-9 | 64742-69-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-477-00-2 | Paraffin oils (petroleum), catalytic dewaxed heavy; Baseoil - unspecified; [A complex combination of hydrocarbons obtained from a catalytic dewaxing process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C).] | 265-174-4 | 64742-70-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-478-00-8 | Paraffin oils (petroleum), catalytic dewaxed light; Baseoil - unspecified; [A complex combination of hydrocarbons obtained from a catalytic dewxing process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C).] | 265-176-5 | 64742-71-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-479-00-3 | Naphthenic oils (petroleum), complex dewaxed heavy; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by removing straight chain paraffin hydrocarbons as a solid by treatment with an agent such as urea. It consists of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil having a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-179-1 | 64742-75-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-480-00-9 | Naphthenic oils (petroleum), complex dewaxed light; Baseoil - unspecified; [A complex combination of hydrocarbons obtained from a catalytic dewaxing process. It consists of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil having a viscosity less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 265-180-7 | 64742-76-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-481-00-4 | Lubricating oils (petroleum), C20-50, hydrotreated neutral oil-based, high-viscosity; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treating light vacuum gas oil, heavy vacuum gas oil, and solvent deasphalted residual oil with hydrogen in the presence of a catalyst in a two stage process with dewaxing being carried out between the two stages. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil having a viscosity of approximately 112cSt at 40 °C. It contains a relatively large proportion of saturated hydrocarbons.] | 276-736-3 | 72623-85-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-482-00-X | Lubricating oils (petroleum), C15-30, hydrotreated neutral oil-based; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treating light vacuum gas oil and heavy vacuum gas oil with hydrogen in the presence of a catalyst in a two stage process with dewaxing being carried out between the two stages. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil having a viscosity of approximately 15cSt at 40 °C. It contains a relatively large proportion of saturated hydrocabons.] | 276-737-9 | 72623-86-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-483-00-5 | Lubricating oils (petroleum), C20-50, hydrotreated neutral oil-based; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treating light vacuum gas oil, heavy vacuum gas oil and solvent deasphalted residual oil with hydrogen in the presence of a catalyst in a two stage process with dewaxing being carried out between the two stages. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of approximately 32cSt at 40 °C. It contains a relatively large proportion of saturated hydrocarbons.] | 276-738-4 | 72623-87-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-484-00-0 | Lubricating oils; Baseoil - unspecified; [A complex combination of hydrocarbons obtained from solvent extraction and dewaxing processes. It consists predominantly of saturated hydrocarbons having carbon numbers in the range C15 through C50.] | 278-012-2 | 74869-22-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-485-00-6 | Distillates (petroleum), complex dewaxed heavy paraffinci; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by dewaxing heavy paraffinic distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil with a viscosity of equal to or greater than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 292-613-7 | 90640-91-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-486-00-1 | Distillates (petroleum), complex dewaxed light paraffinic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by dewaxing light paraffinic distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C12 through C30 and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] | 292-614-2 | 90640-92-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-487-00-7 | Distillates (petroleum), solvent dewaxed heavy paraffinic, clay-treated; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treating dewaxed heavy paraffinic distillate with neutral or modified clay in either a contacting or percolation process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C20 through C50.] | 292-616-3 | 90640-94-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-488-00-2 | Hydrocarbons, C20-50, solvent dewaxed heavy paraffinic, hydrotreated; Baseoil - unspecified; [A complex combination of hydrocarbons produced by treating dewaxed heavy paraffinic distillate with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C20 through C50.] | 292-617-9 | 90640-95-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-489-00-8 | Distillates (petroleum), solvent dewaxed light paraffinic, clay-treated; Baseoil - unspecified; [A complex combination of hydrocarbons resulting from treatment of dewaxed light paraffinic distillate with natural or modified clay in either a contacting or percolation process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C15 through C30.] | 292-618-4 | 90640-96-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-490-00-3 | Distillates (petroleum), solvent dewaxed light paraffinic, hydrotreated; Baseoil - unspecified; [A complex combination of hydrocarbons produced by treating a dewaxed light paraffinic distillate with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C15 through C30.] | 292-620-5 | 90640-97-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-491-00-9 | Residual oils (petroleum), hydrotreated solvent dewaxed; Baseoil - unspecified | 292-656-1 | 90669-74-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-492-00-4 | Residual oils (petroleum), catalytic dewaxed; Baseoil - unspecified | 294-843-3 | 91770-57-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-493-00-X | Distillates (petroleum), dewaxed heavy paraffinic, hydrotreated; Baseoil - unspecified; [A complex combination of hydrocarbons obtained from an intensive treatment of dewaxed distillate by hydrogenation in the presence of a catalyst. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C25 through C39 and produces a finished oil with a viscosity of approximately 44 cSt at 50 °C.] | 295-300-3 | 91995-39-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-494-00-5 | Distillates (petroleum), dewaxed light paraffinic, hydrotreated; Baseoil - unspecified; [A complex combination of hydrocarbons obtained from an intensive treatment of dewaxed distillate by hydrogenation in the presence of a catalyst. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C21 through C29 and produces a finished oil with a viscosity of approximately 13 cSt at 50 °C.] | 295-301-9 | 91995-40-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-495-00-0 | Distillates (petroleum), hydrocracked solvent-refined, dewaxed; Baseoil - unspecified; [A complex combination of liquid hydrocarbons obtained by recrystallization of dewaxed hydrocracked solvent-refined petroleum distillates.] | 295-306-6 | 91995-45-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-496-00-6 | Distillates (petroleum), solvent-refined light naphthenic, hydrotreated; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst and removing the aromatic hydrocarbons by solvent extraction. It consists predominantly of naphthenic hydrocarbons having carbon numbers predominantly in the range of C15 through C30 and produces a finished oil with a viscosity of between 13-15cSt at 40 °C.] | 295-316-0 | 91995-54-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-497-00-1 | Lubricating oils (petroleum), C17-35, solvent-extd., dewaxed, hydrotreated; Baseoil - unspecified | 295-423-2 | 92045-42-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-498-00-7 | Lubricating oils (petroleum), hydrocracked nonarom. solvent-deparaffined; Baseoil - unspecified | 295-424-8 | 92045-43-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-499-00-2 | Residual oils (petroleum), hydrocracked acid-treated solvent-dewaxed; Baseoil - unspecified; [A complex combination of hydrocarbons produced by solvent removal of paraffins from the residue of the distillation of acid-treated, hydrocracked heavy paraffins and boiling approximately above 380 °C (716 °F).] | 295-499-7 | 92061-86-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-500-00-6 | Paraffin oils (petroleum), solvent-refined dewaxed heavy; Baseoil - unspecified; [A complex combination of hydrocarbons obtained from sulfur-containing paraffinic crude oil. It consists predominantly of a solvent refined deparaffinated lubricating oil with a viscosity of 65cSt at 50 °C.] | 295-810-6 | 92129-09-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-501-00-1 | Lubricating oils (petroleum), base oils, paraffinic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by refining of crude oil. It consists predominantly of aromatics, naphthenics and paraffinics and produces a finished oil with a viscosity of 120 SUS at 100 °F (23cSt at 40 °C).] | 297-474-6 | 93572-43-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-502-00-7 | Hydrocarbons, hydrocracked paraffinic distn. residues, solvent-dewaxed; Baseoil - unspecified | 297-857-8 | 93763-38-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-503-00-2 | Hydrocarbons, C20-50, residual oil hydrogenation vacuum distillate; Baseoil - unspecified | 300-257-1 | 93924-61-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-504-00-8 | Distillates (petroleum), solvent-refined hydrotreated heavy, hydrogenated; Baseoil - unspecified | 305-588-5 | 94733-08-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-505-00-3 | Distillates (petroleum), solvent-refined hydrocracked light; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by solvent dearomatization of the residue of hydrocracked petroleum. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C18 through C27 and boiling in the range of approximately 370 °C to 450 °C (698 °F to 842 °F).] | 305-589-0 | 94733-09-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-506-00-9 | Lubricating oils (petroleum), C18-40, solvent-dewaxed hydrocracked distillate-based; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by solvent deparaffination of the distillation residue from hydrocracked petroleum. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C18 through C40 and boiling in the range of approximately 370 °C to 550 °C (698 °F to 1022 °F).] | 305-594-8 | 94733-15-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-507-00-4 | Lubricating oils (petroleum), C18-40, solvent-dewaxed hydrogenated raffinate-based; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by solvent deparaffination of the hydrogenated raffinate obtained by solvent extraction of a hydrotreated petroleum distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C18 through C40 and boiling in the range of approximately 370 °C to 550 °C (698 °F to 1022 °F).] | 305-595-3 | 94733-16-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-508-00-X | Hydrocarbons, C13-30, arom.-rich, solvent-extd. naphthenic distillate; Baseoil - unspecified | 305-971-7 | 95371-04-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-509-00-5 | Hydrocarbons, C16-32, arom. rich, solvent-extd. naphthenic distillate; Baseoil - unspecified | 305-972-2 | 95371-05-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-510-00-0 | Hydrocarbons, C37-68, dewaxed deasphalted hydrotreated vacuum distn. residues; Baseoil - unspecified | 305-974-3 | 95371-07-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-511-00-6 | Hydrocarbons, C37-65, hydrotreated deasphalted vacuum distn. residues; Baseoil - unspecified | 305-975-9 | 95371-08-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-512-00-1 | Distillates (petroleum), hydrocracked solvent-refined light; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by the solvent treatment of a distillate from hydrocracked petroleum distillates. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C18 through C27 and boiling in the range of approximately 370 °C to 450 °C (698 °F to 842 °F.] | 307-010-7 | 97488-73-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-513-00-7 | Distillates (petroleum), solvent-refined hydrogenated heavy; Baseoil - unspecified; [A complex combination of hydrocarbons, obtained by the treatment of a hydrogenated petroleum distillate with a solvent. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C19 through C40 and boiling in the range of approximately 390 °C to 550 °C (734 °F to 1022 °F).] | 307-011-2 | 97488-74-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-514-00-2 | Lubricating oils (petroleum), C18-27, hydrocracked solvent-dewaxed; Baseoil - unspecified | 307-034-8 | 97488-95-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-515-00-8 | Hydrocarbons, C17-30, hydrotreated solvent-deasphalted atm. distn. residue, distn. lights; Baseoil - unspecified; [A complex combination of hydrocarbons obtained as first runnings from the vacuum distillation of effluents from the treatment of a solvent deasphalted short residue with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C17 through C30 and boiling in the range of approximately 300 °C to 400 °C (572 °F to 752 °F). It produces a finished oil having a viscosity of 4cSt at approximately 100 °C (212 °F).] | 307-661-7 | 97675-87-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-516-00-3 | Hydrocarbons, C17-40, hydrotreated solvent-deasphalted distn. residue, vacuum distn. lights; Baseoil - unspecified; [A complex combination of hydrocarbons obtained as first runnings from the vacuum distillation of effluents from the catalytic hydrotreatment of a solvent deasphalted short residue having a viscosity of 8cSt at approximately 100 °C (212 °F). It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C17 through C40 and boiling in the range of approximately 300 °C to 500 °C (592 °F to 932 °F).] | 307-755-8 | 97722-06-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-517-00-9 | Hydrocarbons, C13-27, solvent-extd. light naphthenic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by extraction of the aromatics from a light naphthenic distillate having a viscosity of 9.5cSt at 40 °C (104 °F). It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C13 through C27 and boiling in the range of approximately 240 °C to 400 °C (464 °F to 752 °F.] | 307-758-4 | 97722-09-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-518-00-4 | Hydrocarbons, C14-29, solvent-extd. light naphthenic; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by extraction of the aromatics from a light naphthenic distillate having a viscosity of 16cSt at 40 °C (104 °F). It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C14 through C29 and boiling in the range of approximately 250 °C to 425 °C (482 °F to 797 °F).] | 307-760-5 | 97722-10-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-519-00-X | Hydrocarbons, C27-42, dearomatized; Baseoil - unspecified | 308-131-8 | 97862-81-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-520-00-5 | Hydrocarbons, C17-30, hydrotreated distillates, distn. lights; Baseoil - unspecified | 308-132-3 | 97862-82-3 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-521-00-0 | Hydrocarbons, C27-45, naphthenic vacuum distn.; Baseoil - unspecified | 308-133-9 | 97862-83-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-522-00-6 | Hydrocarbons, C27-45, dearomatized; Baseoil - unspecified | 308-287-7 | 97926-68-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-523-00-1 | Hydrocarbons, C20-58, hydrotreated; Baseoil - unspecified | 308-289-8 | 97926-70-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-524-00-7 | Hydrocarbons, C27-42, naphthenic; Baseoil - unspecified | 308-290-3 | 97926-71-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-525-00-2 | Residual oils (petroleum), carbon-treated solvent-dewaxed; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by the treatment of solvent-dewaxed petroleum residual oils with activated charcoal for the removal of trace polar constituents and impurities.] | 309-710-8 | 100684-37-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-526-00-8 | Residual oils (petroleum), clay-treated solvent-dewaxed; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by treatment of solvent-dewaxed petroleum residual oils with bleaching earth for the removal of trace polar constituents and impurities.] | 309-711-3 | 100684-38-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-527-00-3 | Lubricating oils (petroleum), C >25, solvent-extd., deasphalted, dewaxed, hydrogenated; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of vacuum distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly greater than C25 and produces a finished oil with a viscosity in the order of 32cSt to 37cSt at 100 °C (212 °F).] | 309-874-0 | 101316-69-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-528-00-9 | Lubricating oils (petroleum), C17-32, solvent-extd., dewaxed, hydrogenated; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of atmospheric distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C17 through C32 and produced a finished oil with a viscosity in the order of 17cSt to 23cSt at 40 °C (104 °F.] | 309-875-6 | 101316-70-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-529-00-4 | Lubricating oils (petroleum), C20-35, solvent-extd., dewaxed, hydrogenated; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of atmospheric distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C20 through C35 and produces a finished oil with a viscosity in the order of 37cSt to 44cSt at 40 °C (104 °F).] | 309-876-1 | 101316-71-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-530-00-X | Lubricating oils (petroleum), C24-50, solvent-extd., dewaxed, hydrogenated; Baseoil - unspecified; [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of atmospheric distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C24 through C50 and produces a finished oil with a viscosity in the order of 16cSt to 75cSt at 40 °C (104 °F).] | 309-877-7 | 101316-72-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-531-00-5 | Extracts (petroleum), heavy naphthenic distillate solvent, arom. conc.; Distillate aromatic extract (treated); [An aromatic concentrate produced by adding water to heavy naphthenic distillate solvent extract and extraction solvent.] | 272-175-3 | 68783-00-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-532-00-0 | Extracts (petroleum), solvent-refined heavy paraffinic distillate solvent; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained as the extract from the re-extraction of solvent-refined heavy paraffinic distillate. It consists of saturated and aromatic hydrocarbons having carbon numbers predominantly in the range of C20 through C50.] | 272-180-0 | 68783-04-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-533-00-6 | Extracts (petroleum), heavy paraffinic distillates, solvent-deasphalted; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained as the extract from a solvent extraction of heavy paraffinic distillate.] | 272-342-0 | 68814-89-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-534-00-1 | Extracts (petroleum), heavy naphthenic distillate solvent, hydrotreated; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained by treating a heavy naphthenic distillate solvent extract with hydrogen in the presence of a catalyst. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C20 through C50 and produces a finished oil of at least 19cSt at 40 °C (100 SUS at 100 °F).] | 292-631-5 | 90641-07-9 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-535-00-7 | Extracts (petroleum), heavy paraffinic distillate solvent, hydrotreated; Distillate aromatic extract (treated); [A complex combination of hydrocarbons produced by treating a heavy paraffinic distillate solvent extract with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C21 through C33 and boiling in the range of approximately 350 °C to 480 °C (662 °F to 896 °F). | 292-632-0 | 90641-08-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-536-00-2 | Extracts (petroleum), light paraffinic distillate solvent, hydrotreated; Distillate aromatic extract (treated); [A complex combination of hydrocarbons produced by treating a light paraffinic distillate solvent extract with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C17 through C26 and boiling in the range of approximately 280 °C to 400 °C (536 °F to 752 °F).] | 292-633-6 | 90641-09-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-537-00-8 | Extracts (petroleum), hydrotreated light paraffinic distillate solvent; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained as the extract from solvent extraction of intermediate paraffinic top solvent distillate that is treated with hydrogen in the presence of a catalyst. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C16 through C36.] | 295-335-4 | 91995-73-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-538-00-3 | Extracts (petroleum), light naphthenic distillate solvent, hydrodesulfurized; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained by treating the extract, obtained from a solvent extraction process, with hydrogen in the presence of a catalyst under conditions primarily to remove sulfur compounds. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C15 through C30. This stream is likely to contain 5 wt.% or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] | 295-338-0 | 91995-75-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-539-00-9 | Extracts (petroleum), light paraffinic distillate solvent, acid-treated; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained as a fraction of the distillation of an extract from the solvent extraction of light paraffinic top petroleum distillates that is subjected to a sulfuric acid refining. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C16 through C32.] | 295-339-6 | 91995-76-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-540-00-4 | Extracts (petroleum), light paraffinic distillate solvent, hydrodesulfurized; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained by solvent extraction of a light paraffin distillate and treated with hydrogen to convert the organic sulfur to hydrogen sulfide which is eliminated. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C15 through C40 and produces a finished oil with a viscosity of greater than 10cSt at 40 °C.] | 295-340-1 | 91995-77-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-541-00-X | Extracts (petroleum), light vacuum gas oil solvent, hydrotreated; Distillate aromatic extract (treated); [A complex combination of hydrocarbons, obtained by solvent extraction from light vacuum petroleum gas oils and treated with hydrogen in the presence of a catalyst. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C13 through C30.] | 295-342-2 | 91995-79-8 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-542-00-5 | Extracts (petroleum), heavy paraffinic distillate solvent, clay-treated; Distillate aromatic extract (treated); [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay in either a contact or percolation process to remove the trace amounts of polar compounds and impurities present. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C20 through C50. This stream is likely to contain 5 wt.% or more 4-6 membered ring aromatic hydrocarbons.] | 296-437-1 | 92704-08-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-543-00-0 | Extracts (petroleum), heavy naphthenic distillate solvent, hydrodesulfurized; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained from a petroleum stock by treating with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C15 through C50 and produces a finished oil with a viscosity of greater than 19cSt at 40 °C.] | 297-827-4 | 93763-10-1 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-544-00-6 | Extracts (petroleum), solvent-dewaxed heavy paraffinic distillate solvent, hydrodesulfurized; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained from a solvent dewaxed petroleum stock by treating with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C15 through C50 and produces a finished oil with a viscosity of greater than 19cSt at 40 °C.] | 297-829-5 | 93763-11-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-545-00-1 | Extracts (petroleum), light paraffinic distillate solvent, carbon-treated; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained as a fraction from distillation of an extract recovered by solvent extraction of light paraffinic top petroleum distillate treated with activated charcoal to remove traces of polar constituents and impurities. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C16 through C32.] | 309-672-2 | 100684-02-4 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-546-00-7 | Extracts (petroleum), light paraffinic distillate solvent, clay-treated; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained as a fraction from distillation of an extract recovered by solvent extraction of light paraffinic top petroleum distillates treated with bleaching earth to remove traces of polar constituents and impurities. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C16 through C32.] | 309-673-8 | 100684-03-5 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-547-00-2 | Extracts (petroleum), light vacuum, gas oil solvent, carbon-treated; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained by solvent extraction of light vacuum petroleum gas oil treated with activated charcoal for the removal of trace polar constituents and impurities. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C13 through C30.] | 309-674-3 | 100684-04-6 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-548-00-8 | Extracts (petroleum), light vacuum gas oil solvent, clay-treated; Distillate aromatic extract (treated); [A complex combination of hydrocarbons obtained by solvent extraction of light vacuum petroleum gas oils treated with bleaching earth for removal of trace polar constituents and impurities. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C13 through C30.] | 309-675-9 | 100684-05-7 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-549-00-3 | Foots oil (petroleum); Foots oil; [A complex combination of hydrocarbons obtained as the oil fraction from a solvent deoiling or a wax sweating process. It consists predominantly of branched chain hydrocarbons having carbon numbers predominantly in the range of C20 through C50.] | 265-171-8 | 64742-67-2 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
649-550-00-9 | Foots oil (petroleum), hydrotreated; Foots oil | 295-394-6 | 92045-12-0 | Carc. 1B | H350 | GHS08 Dgr | H350 |  |  | L | CLP00
650-002-00-6 | turpentine, oil | 232-350-7 | 8006-64-2 | Flam. Liq. 3 Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Asp. Tox. 1 Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 2 | H226 H332 H312 H302 H304 H315 H319 H317 H411 | GHS02 GHS08 GHS07 GHS09 Dgr | H226 H332 H312 H302 H304 H319 H315 H317 H411 |  |  |  | CLP00
650-003-00-1 | fenson (ISO); 4-chlorophenyl benzenesulphonate | 201-274-6 | 80-38-6 | Acute Tox. 4 * Eye Irrit. 2 Aquatic Chronic 2 | H302 H319 H411 | GHS07 GHS09 Wng | H302 H319 H411 |  |  |  | CLP00
650-004-00-7 | norbormide (ISO); 5-(α-hydroxy-α-2-pyridylbenzyl)-7-(α-2-pyridylbenzylidene)bicyclo [2.2.1] hept-5-ene-2,3-dicarboximide | 213-589-6 | 991-42-4 | Acute Tox. 4 * | H302 | GHS07 Wng | H302 |  |  |  | CLP00
650-005-00-2 | (2R,6aS,12aS)-1,2,6,6a,12,12a-hexahydro-2-isopropenyl-8,9-dimethoxychromeno[3,4-b]furo[2,3-h]chromen-6-one, rotenone | 201-501-9 | 83-79-4 | Acute Tox. 3 * STOT SE 3 Skin Irrit. 2 Eye Irrit. 2 Aquatic Acute 1 Aquatic Chronic 1 | H301 H335 H315 H319 H400 H410 | GHS06 GHS09 Dgr | H301 H319 H335 H315 H410 |  |  |  | CLP00
650-006-00-8 | benquinox (ISO); p-benzoquinone 1-benzoylhydrazone 4-oxime | 207-807-9 | 495-73-8 | Acute Tox. 3 * Acute Tox. 4 * | H301 H312 | GHS06 Dgr | H301 H312 |  |  |  | CLP00
650-007-00-3 | chlordimeform (ISO); N2-(4-chloro-o-tolyl)-N1,N1-dimethylformamidine | 228-200-5 | 6164-98-3 | Carc. 2 Acute Tox. 4 * Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H312 H302 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H312 H302 H410 |  |  |  | CLP00
650-008-00-9 | drazoxolon (ISO); 4-(2-chlorophenylhydrazone)-3-methyl-5-isoxazolone | 227-197-8 | 5707-69-7 | Acute Tox. 3 * Aquatic Acute 1 Aquatic Chronic 1 | H301 H400 H410 | GHS06 GHS09 Dgr | H301 H410 |  |  |  | CLP00
650-009-00-4 | chlordimeform hydrochloride; N'-(4-chloro-o-tolyl)-N,N-dimethylformamidine monohydrochloride; N2-(4-chloro-o-tolyl)-N1,N1-dimethylformamidine hydorchloride | 243-269-1 | 19750-95-9 | Carc. 2 Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H351 H302 H400 H410 | GHS08 GHS07 GHS09 Wng | H351 H302 H410 |  |  |  | CLP00
650-010-00-X | benzyl violet 4B; α-[4-(4-dimethylamino-α-{}{4-[ethyl(3-sodiosulphonatobenzyl)amino] phenyl}}benzylidene)cyclohexa-2,5-dienylidene(ethyl)ammonio]toluene-3-sulphonate | 216-901-9 | 1694-09-3 | Carc. 2 | H351 | GHS08 Wng | H351 |  |  |  | CLP00
650-012-00-0 | erionite |  | 12510-42-8 | Carc. 1A | H350 | GHS08 Dgr | H350 |  |  |  | CLP00
650-013-00-6 | asbestos [1] asbestos [2] asbestos [3] asbestos [4] asbestos [5] asbestos [6] asbestos [7] |  | 12001-28-4 [1] 132207-32-0 [2] 12172-73-5 [3] 77536-66-4 [4] 77536-68-6 [5] 77536-67-5 [6] 12001-29-5 [7] | Carc. 1A STOT RE 1 | H350 H372 ** | GHS08 Dgr | H350 H372 ** |  |  |  | CLP00
650-014-00-1 | diethyl 2,4-dihydroxycyclodisiloxane-2,4-diylbis(trimethylene)diphosphonate, tetrasodium salt, reaction products with disodium metasilicate | 401-770-4 |  | Acute Tox. 4 * Skin Corr. 1B | H302 H314 | GHS05 GHS07 Dgr | H314 H302 |  |  |  | CLP00
650-015-00-7 | rosin; colophony [1] rosin; colophony [2] rosin; colophony [3] | 232-475-7 [1] 232-484-6 [2] 277-299-1 [3] | 8050-09-7 [1] 8052-10-6 [2] 73138-82-6 [3] | Skin Sens. 1 | H317 | GHS07 Wng | H317 |  |  |  | CLP00
650-016-00-2 | Mineral wool, with the exception of those specified elsewhere in this Annex; [Man-made vitreous (silicate) fibres with random orientation with alkaline oxide and alkali earth oxide (Na2O+K2O+CaO+MgO+BaO) content greater than 18 % by weight] | - | - | Carc. 2 | H351 | GHS08 Wng | H351 |  |  | A Q R | CLP00/ATP01
650-017-00-8 | Refractory Ceramic Fibres, Special Purpose Fibres, with the exception of those specified elsewhere in this Annex; [Man-made vitreous (silicate) fibres with random orientation with alkaline oxide and alkali earth oxide (Na2O+K2O+CaO+ MgO+BaO) content less or equal to 18 % by weight] | - | - | Carc. 1B | H350i | GHS08 Dgr | H350i |  |  | A R | CLP00/ATP01
650-018-00-3 | reaction product of: acetophenone, formaldehyde, cyclohexylamine, methanol and acetic acid | 406-230-1 |  | Flam. Liq. 3 Carc. 2 Acute Tox. 4 * Skin Corr. 1B Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H226 H351 H332 H314 H317 H400 H410 | GHS02 GHS08 GHS05 GHS07 GHS09 Dgr | H226 H351 H314 H332 H317 H410 |  |  |  | CLP00
650-031-00-4 | bis(4-hydroxy-N-methylanilinium) sulphate | 200-237-1 | 55-55-0 | Acute Tox. 4 * STOT RE 2 * Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H302 H373 ** H317 H400 H410 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H317 H410 |  |  |  | CLP00
650-032-00-X | cyproconazole (ISO); (2RS,3RS;2RS,3SR)-2-(4-chlorophenyl)-3-cyclopropyl-1-(1H-1,2,4-triazol-1-yl)butan-2-ol |  | 94361-06-5 | Repr. 1B Acute Tox. 3 STOT RE 2 Aquatic Acute 1 Aquatic Chronic 1 | H360D H301 H373 (liver) H400 H410 | GHS08 GHS06 GHS09 Dgr | H301 H360D H373 (liver) H410 |  | M=10 M=1 |  | CLP00/ATP10
650-041-00-9 | triasulfuron (ISO); 1-[2-(2-chloroethoxy)phenylsulfonyl]-3-(4-methoxy-6-methyl-1,3,5-triazin-2-yl)urea |  | 82097-50-5 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
650-042-00-4 | reaction product of: polyethylene-polyamine-(C16-C18)-alkylamides with monothio-(C2)-alkyl phosphonates | 417-450-2 |  | Skin Irrit. 2 Eye Irrit. 2 Skin Sens. 1 Aquatic Chronic 3 | H315 H319 H317 H412 | GHS07 Wng | H319 H315 H317 H412 |  |  |  | CLP00
650-043-00-X | reaction product of: 3,5-bis-tert-butylsalicylic acid and aluminiumsulfate | 420-310-3 |  | Acute Tox. 4 * Aquatic Acute 1 Aquatic Chronic 1 | H302 H400 H410 | GHS07 GHS09 Wng | H302 H410 |  |  |  | CLP00
650-044-00-5 | mixed linear and branched C14-15 alcohols ethoxylated, reaction product with epichlorohydrin | 420-480-9 | 158570-99-1 | Skin Irrit. 2 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H317 H400 H410 | GHS07 GHS09 Wng | H315 H317 H410 |  |  |  | CLP00
650-045-00-0 | reaction product of: 1,2,3-propanetricarboxylic acid, 2-hydroxy, diethyl ester, 1-propanol and zirconium tetra-n-propanolate | 417-110-3 |  | Flam. Liq. 2 Skin Irrit. 2 Eye Dam. 1 Aquatic Chronic 2 | H225 H315 H318 H411 | GHS02 GHS05 GHS09 Dgr | H225 H315 H318 H411 |  |  |  | CLP00
650-046-00-6 | di(tetramethylammonium)(29H,31H-phthalocyanin-N29,N30,N31,N32)disulfonamide disulfonate, cuprate(2-)complex, derivates | 416-180-2 | 12222-04-7 | Acute Tox. 4 * STOT RE 2 * Aquatic Chronic 2 | H302 H373 ** H411 | GHS08 GHS07 GHS09 Wng | H302 H373 ** H411 |  |  |  | CLP00
650-047-00-1 | dibenzylphenylsulfonium hexafluoroantimonate | 417-760-8 | 134164-24-2 | Acute Tox. 4 * STOT RE 1 Eye Dam. 1 Skin Sens. 1 Aquatic Chronic 2 | H302 H372 ** H318 H317 H411 | GHS08 GHS05 GHS07 GHS09 Dgr | H372 ** H302 H318 H317 H411 |  |  |  | CLP00
650-048-00-7 | reaction product of: borax, hydrogen peroxide, acetic acid anhydride and acetic acid | 420-070-1 |  | Org. Perox. D **** Acute Tox. 4 * Acute Tox. 4 * Acute Tox. 4 * Skin Corr. 1A Aquatic Acute 1 | H242 H332 H312 H302 H314 H400 | GHS02 GHS05 GHS07 GHS09 Dgr | H242 H332 H312 H302 H314 H400 |  |  |  | CLP00
650-049-00-2 | 2-alkoyloxyethyl hydrogen maleate, where alkoyl represents (by weight) 70 to 85 % unsaturated octadecoyl, 0.5 to 10 % saturated octadecoyl, and 2 to 18 % saturated hexadecoyl | 417-960-5 |  | Skin Irrit. 2 Eye Dam. 1 Skin Sens. 1 Aquatic Acute 1 Aquatic Chronic 1 | H315 H318 H317 H400 H410 | GHS05 GHS07 GHS09 Dgr | H315 H318 H317 H410 |  |  |  | CLP00
650-050-00-8 | reaction mass of: 1-methyl-3-hydroxypropyl 3,5-[1,1-dimethylethyl]-4-hydroxydihydro-cinnamate and/or 3-hydroxybutyl 3,5-[1,1-dimethylethyl]-4-hydroxydihydrocinnamate; 1,3-butanediol bis[3-(3'-(1,1-dimethylethyl)4'-hydroxy-phenyl)propionate] isomers; 1,3-butanediol bis[3-(3',5'-(1,1-dimethylethyl)-4'-hydroxyphenyl)propionate] isomers | 423-600-8 |  | Aquatic Chronic 2 | H411 | GHS09 | H411 |  |  |  | CLP00
650-055-00-5 | silver sodium zirconium hydrogenphosphate | 422-570-3 | 155925-27-2 | Aquatic Acute 1 Aquatic Chronic 1 | H400 H410 | GHS09 Wng | H410 |  |  |  | CLP00
650-056-00-0 | dibutylbis(pentane-2,4-dionato-O,O')tin | 245-152-0 | 22673-19-4 | Repr. 1B STOT RE 1 | H360FD H372 (immune system) | GHS08 Dgr | H360FD H372 (immune system) |  |  |  | ATP14
650-057-00-6 | Margosa, ext. [cold-pressed oil of Azadirachta indica seeds without shells extracted with super-critical carbon dioxide] | 283-644-7 | 84696-25-3 | Aquatic Chronic 3 | H412 |  | H412 |  |  |  | ATP15
650-058-00-1 | Margosa, ext. [from the kernels of Azadirachta indica extracted with water and further processed with organic solvents] | 283-644-7 | 84696-25-3 | Repr. 2 Skin Sens. 1 Aquatic Chronic 1 | H361d H317 H410 | GHS08 GHS07 GHS09 Wng | H361d H317 H410 |  | M = 10 |  | ATP18

## SVHC Candidates
Candidate List of SVHC for authorisation
Note: Group entries are split in different rows.
https://echa.europa.eu/candidate-list-table
Last table update: 2026-05-04
Substance name | Description | EC No. | CAS No. | Reason for inclusion | Date of inclusion | Decision | IUCLID dataset | Support document | Response to comments | Remarks
n-hexane |  | 203-777-6 | 110-54-3 | Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 04-Feb-2026 | https://echa.europa.eu/documents/10162/2b029c1a-8669-aee2-adfa-6af3f5d8b86c |  | https://echa.europa.eu/documents/10162/fee14698-29ec-7b04-fbb9-a48ea68eb15b | https://echa.europa.eu/documents/10162/9c49de65-bd78-154a-05fa-b95839cc64a0 | N-hexane is identified as a substance of very high concern in accordance with Article 57(f) of Regulation (EC) 1907/2006 (REACH) because it is a substance with neurotoxic properties for which there is scientific evidence of probable serious effects to human health which gives rise to an equivalent level of concern to those substances listed in points (a) to (e) of Article 57 of REACH.
4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]diphenol and its salts |  | - | - | Toxic for reproduction (Article 57c) | 04-Feb-2026 | https://echa.europa.eu/documents/10162/71f21fe4-c687-cf7d-42d9-fcf0b7b384eb |  | https://echa.europa.eu/documents/10162/e04eb45f-b900-941d-6db0-b1b195eca48f | https://echa.europa.eu/documents/10162/64c4fd72-5c3f-caf1-a92c-33f0949e97dd
4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]diphenol (Bisphenol AF) |  | 216-036-7 | 1478-61-1 | Toxic for reproduction (Article 57c) | 04-Feb-2026 | https://echa.europa.eu/documents/10162/71f21fe4-c687-cf7d-42d9-fcf0b7b384eb |  | https://echa.europa.eu/documents/10162/e04eb45f-b900-941d-6db0-b1b195eca48f | https://echa.europa.eu/documents/10162/64c4fd72-5c3f-caf1-a92c-33f0949e97dd
Benzyltriphenylphosphonium, salt with 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]bis[phenol] (1:1) |  | 278-305-5 | 75768-65-9 | Toxic for reproduction (Article 57c) | 04-Feb-2026 | https://echa.europa.eu/documents/10162/71f21fe4-c687-cf7d-42d9-fcf0b7b384eb |  | https://echa.europa.eu/documents/10162/e04eb45f-b900-941d-6db0-b1b195eca48f | https://echa.europa.eu/documents/10162/64c4fd72-5c3f-caf1-a92c-33f0949e97dd
disodium 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]diphenolate |  | 425-060-9 | - | Toxic for reproduction (Article 57c) | 04-Feb-2026 | https://echa.europa.eu/documents/10162/71f21fe4-c687-cf7d-42d9-fcf0b7b384eb |  | https://echa.europa.eu/documents/10162/e04eb45f-b900-941d-6db0-b1b195eca48f | https://echa.europa.eu/documents/10162/64c4fd72-5c3f-caf1-a92c-33f0949e97dd
N,N-dibutyl-N-methylbutan-1-aminium 4-[1,1,1,3,3,3-hexafluoro-2-(4-hydroxyphenyl)propan-2-yl]phenolate |  | 443-330-4 | - | Toxic for reproduction (Article 57c) | 04-Feb-2026 | https://echa.europa.eu/documents/10162/71f21fe4-c687-cf7d-42d9-fcf0b7b384eb |  | https://echa.europa.eu/documents/10162/e04eb45f-b900-941d-6db0-b1b195eca48f | https://echa.europa.eu/documents/10162/64c4fd72-5c3f-caf1-a92c-33f0949e97dd
Reaction mass of 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]diphenol and benzyltriphenylphosphonium, salt with 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]diphenol (1:1) |  | - | - | Toxic for reproduction (Article 57c) | 04-Feb-2026 | https://echa.europa.eu/documents/10162/71f21fe4-c687-cf7d-42d9-fcf0b7b384eb |  | https://echa.europa.eu/documents/10162/e04eb45f-b900-941d-6db0-b1b195eca48f | https://echa.europa.eu/documents/10162/64c4fd72-5c3f-caf1-a92c-33f0949e97dd
Reaction products of benzene, chlorine and sulfur chloride (S2Cl2) with 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]diphenol |  | 469-080-6 | 921213-47-0 | Toxic for reproduction (Article 57c) | 04-Feb-2026 | https://echa.europa.eu/documents/10162/71f21fe4-c687-cf7d-42d9-fcf0b7b384eb |  | https://echa.europa.eu/documents/10162/e04eb45f-b900-941d-6db0-b1b195eca48f | https://echa.europa.eu/documents/10162/64c4fd72-5c3f-caf1-a92c-33f0949e97dd
benzyl(diethylamino)diphenylphosphonium 4-[1,1,1,3,3,3-hexafluoro-2-(4-hydroxyphenyl)propan-2-yl]phenolate |  | 479-100-5 | 577705-90-9 | Toxic for reproduction (Article 57c) | 04-Feb-2026 | https://echa.europa.eu/documents/10162/71f21fe4-c687-cf7d-42d9-fcf0b7b384eb |  | https://echa.europa.eu/documents/10162/e04eb45f-b900-941d-6db0-b1b195eca48f | https://echa.europa.eu/documents/10162/64c4fd72-5c3f-caf1-a92c-33f0949e97dd
Reaction mass of 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]diphenol and benzyl(diethylamino)diphenylphosphonium 4-[1,1,1,3,3,3-hexafluoro-2-(4-hydroxyphenyl)propan-2-yl]phenolate (1:1) |  | - | - | Toxic for reproduction (Article 57c) | 04-Feb-2026 | https://echa.europa.eu/documents/10162/71f21fe4-c687-cf7d-42d9-fcf0b7b384eb |  | https://echa.europa.eu/documents/10162/e04eb45f-b900-941d-6db0-b1b195eca48f | https://echa.europa.eu/documents/10162/64c4fd72-5c3f-caf1-a92c-33f0949e97dd
Phosphonium, tributyl(2-methoxypropyl)-, salt with 4,4'-[2,2,2-trifluoro-1-(trifluoromethyl)ethylidene]bis[phenol] (1:1) |  | 468-740-0 | 126049-00-1 | Toxic for reproduction (Article 57c) | 04-Feb-2026 | https://echa.europa.eu/documents/10162/71f21fe4-c687-cf7d-42d9-fcf0b7b384eb |  | https://echa.europa.eu/documents/10162/e04eb45f-b900-941d-6db0-b1b195eca48f | https://echa.europa.eu/documents/10162/64c4fd72-5c3f-caf1-a92c-33f0949e97dd
1,1'-(ethane-1,2-diyl)bis[pentabromobenzene] (DBDPE) |  | 284-366-9 | 84852-53-9 | vPvB (Article 57e) | 05-Nov-2025 | https://echa.europa.eu/documents/10162/9f352981-26b8-8bf6-72e2-932a47850a9d |  | https://echa.europa.eu/documents/10162/75729030-8bba-a59d-7ec9-906009190aa8#https://echa.europa.eu/documents/10162/fda5eadf-8499-26f2-b4b8-181a27f1511e#https://echa.europa.eu/documents/10162/118fcbb9-52d9-d504-40d9-781aa930f929 | https://echa.europa.eu/documents/10162/55f8e9d3-c992-9dd6-9424-973400f81a44
tetra(sodium/potassium) 7-[(E)-{2-acetamido-4-[(E)-(4-{[4-chloro-6-({2-[(4-fluoro-6-{[4-(vinylsulfonyl)phenyl]amino}-1,3,5-triazine-2-yl)amino]propyl}amino)-1,3,5-triazine-2-yl]amino}-5-sulfonato-1-naphthyl)diazenyl]-5-methoxyphenyl}diazenyl]-1,3,6-naphthalenetrisulfonate (Reactive Brown 51) | substance having a complex composition with <80% of the above constituents and other reaction side products | 466-490-7 | - | Toxic for reproduction (Article 57c) | 25-Jun-2025 | https://echa.europa.eu/documents/10162/b5b30988-fcd2-a29e-57c1-535381aa1e47 |  | https://echa.europa.eu/documents/10162/78400f11-0361-2ae0-1851-04a654158f5a | https://echa.europa.eu/documents/10162/e7385b11-47d5-6c73-2073-54d54a44e69f
decamethyltetrasiloxane |  | 205-491-7 | 141-62-8 | vPvB (Article 57e) | 25-Jun-2025 | https://echa.europa.eu/documents/10162/956b44c5-c892-b00a-3518-e82fe59b9309 |  | https://echa.europa.eu/documents/10162/e9f427a1-e41c-d61a-726a-89666f3de869 | https://echa.europa.eu/documents/10162/aa51d34c-bdfa-ad9c-4568-25a7d14e1e21
1,1,1,3,5,5,5-heptamethyl-3-[(trimethylsilyl)oxy]trisiloxane |  | 241-867-7 | 17928-28-8 | vPvB (Article 57e) | 25-Jun-2025 | https://echa.europa.eu/documents/10162/96a9ca97-6130-7853-b7f7-2f72da9c7428 |  | https://echa.europa.eu/documents/10162/7a912253-5076-6e2f-457b-707f72b74cb8 | https://echa.europa.eu/documents/10162/c5339392-0837-e090-4833-131259b7e4f2
reaction mass of: triphenylthiophosphate and tertiary butylated phenyl derivatives |  | 421-820-9 | 192268-65-8 | PBT (Article 57d) | 21-Jan-2025 | https://echa.europa.eu/documents/10162/7dad1173-ac2e-f860-e1f4-66b1c2155dce | https://echa.europa.eu/documents/10162/19d0c561-c440-7d58-1f1a-82f635a56284 | https://echa.europa.eu/documents/10162/76f7f5c4-e1b0-cfb7-870a-37f9be06da6b | https://echa.europa.eu/documents/10162/26ae499e-7ca7-ce7f-454b-7652380d3500
Perfluamine |  | 206-420-2 | 338-83-0 | vPvB (Article 57e) | 21-Jan-2025 | https://echa.europa.eu/documents/10162/0aff23c4-f42e-752f-808f-cb297c0895f1 | https://echa.europa.eu/documents/10162/fdaa58f9-036f-fa45-32a5-ff298257a105 | https://echa.europa.eu/documents/10162/648286a4-3084-fe50-8ea2-7a0ba0f0019c | https://echa.europa.eu/documents/10162/b9748360-658a-8add-40e9-3f27d5d9b3f6
Octamethyltrisiloxane |  | 203-497-4 | 107-51-7 | vPvB (Article 57e) | 21-Jan-2025 | https://echa.europa.eu/documents/10162/686168b6-edc1-8e37-4113-637e22730db9 | https://echa.europa.eu/documents/10162/381a5fab-2de2-6fa7-86dd-9fbfd173a162 | https://echa.europa.eu/documents/10162/528cc7c8-c20c-8445-761c-f75c1dcdbdc1 | https://echa.europa.eu/documents/10162/2105ba6b-e4da-84b5-a171-99973f093537
O,O,O-triphenyl phosphorothioate |  | 209-909-9 | 597-82-0 | PBT (Article 57d) | 21-Jan-2025 | https://echa.europa.eu/documents/10162/fe683d44-f2a2-cd2c-d2fd-a714cb339f69 | https://echa.europa.eu/documents/10162/52c84e84-f317-913a-4c58-01406b51abbb | https://echa.europa.eu/documents/10162/dcdb37a7-9c5f-435a-68f8-637c3ba941aa | https://echa.europa.eu/documents/10162/a6c1c31e-06ff-7611-679a-6cdb078b7bcc
6-[(C10-C13)-alkyl-(branched, unsaturated)-2,5-dioxopyrrolidin-1-yl]hexanoic acid |  | 701-118-1 | 2156592-54-8 | Toxic for reproduction (Article 57c) | 21-Jan-2025 | https://echa.europa.eu/documents/10162/1db77454-c1a6-4d42-b904-853060611229 | https://echa.europa.eu/documents/10162/13385f79-6113-2f54-c5d2-9e867039c6ea | https://echa.europa.eu/documents/10162/525810de-10b9-69cb-54d4-cb73dd6aa9fe | https://echa.europa.eu/documents/10162/fd5ba6c6-e3cb-d132-46bd-afa234c28fd8
Triphenyl phosphate |  | 204-112-2 | 115-86-6 | Endocrine disrupting properties (Article 57(f) - environment) | 07-Nov-2024 | https://echa.europa.eu/documents/10162/5af3b452-563f-ceae-608e-d9206a978b9c | https://echa.europa.eu/documents/10162/ef88904b-3e7b-2326-7937-2569a1970af6 | https://echa.europa.eu/documents/10162/19b0eb6a-bb2f-33ad-1611-0b32b43c49f0 | https://echa.europa.eu/documents/10162/1c702acc-fc2f-f5b1-ea9d-ba39d4a3ce5a
Bis(α,α-dimethylbenzyl) peroxide |  | 201-279-3 | 80-43-3 | Toxic for reproduction (Article 57c) | 27-Jun-2024 | https://echa.europa.eu/documents/10162/733088cd-8dd7-aa78-b278-ad3f5b79e942 | https://echa.europa.eu/documents/10162/d6f3d76b-597a-1d2e-89d6-d0e81a739b92 | https://echa.europa.eu/documents/10162/236dc0e0-5123-fc87-1117-ee76d8da8372 | https://echa.europa.eu/documents/10162/48e7c30f-cd5a-ef64-55ca-c2a6b0133048
Oligomerisation and alkylation reaction products of 2-phenylpropene and phenol |  | 700-960-7 | - | vPvB (Article 57e) | 23-Jan-2024 | https://echa.europa.eu/documents/10162/0a8c79d0-6ebd-6198-323e-aa97e22da3e1 | https://echa.europa.eu/documents/10162/634c637f-1400-f2b3-3e6d-ae8b75fe8a6f | https://echa.europa.eu/documents/10162/9a9d530a-3944-ebca-a9c3-9d25fe47a9e1#https://echa.europa.eu/documents/10162/6c4b3401-858d-0b65-e5b1-1484fce3c84c#https://echa.europa.eu/documents/10162/7e782af9-ee43-9c63-b3b9-4c979a746bb0 | https://echa.europa.eu/documents/10162/7cd990a0-7122-92c8-f360-7e9030aee114 | The name phenol, methylstyrenated and EC number 270-966-8 and CAS number 68512-30-1 have been historically used to identify the substance “Oligomerisation and alkylation reaction products of 2-phenylpropene and phenol”. These identifiers are reported in the candidate list as additional information.
Phenol, methylstyrenated |  | 270-966-8 | 68512-30-1 | vPvB (Article 57e) | 23-Jan-2024 | https://echa.europa.eu/documents/10162/0a8c79d0-6ebd-6198-323e-aa97e22da3e1 | https://echa.europa.eu/documents/10162/634c637f-1400-f2b3-3e6d-ae8b75fe8a6f | https://echa.europa.eu/documents/10162/9a9d530a-3944-ebca-a9c3-9d25fe47a9e1#https://echa.europa.eu/documents/10162/6c4b3401-858d-0b65-e5b1-1484fce3c84c#https://echa.europa.eu/documents/10162/7e782af9-ee43-9c63-b3b9-4c979a746bb0 | https://echa.europa.eu/documents/10162/7cd990a0-7122-92c8-f360-7e9030aee114 | The name phenol, methylstyrenated and EC number 270-966-8 and CAS number 68512-30-1 have been historically used to identify the substance “Oligomerisation and alkylation reaction products of 2-phenylpropene and phenol”. These identifiers are reported in the candidate list as additional information.
Bumetrizole (UV-326) |  | 223-445-4 | 3896-11-5 | vPvB (Article 57e) | 23-Jan-2024 | https://echa.europa.eu/documents/10162/05ea7d09-6615-9ac3-be76-5c9cd96b8eb7 | https://echa.europa.eu/documents/10162/c5b482b0-1ade-976d-5b4a-d8cf5227bcc7 | https://echa.europa.eu/documents/10162/94233c79-4de1-80b8-97a0-8ad81d969d84 | https://echa.europa.eu/documents/10162/abe497c8-9cd7-82e3-540d-beea800982dc
2-(dimethylamino)-2-[(4-methylphenyl)methyl]-1-[4-(morpholin-4-yl)phenyl]butan-1-one |  | 438-340-0 | 119344-86-4 | Toxic for reproduction (Article 57c) | 23-Jan-2024 | https://echa.europa.eu/documents/10162/fd512d85-da9a-f5f3-a976-86ab4eca63dc | https://echa.europa.eu/documents/10162/035bf1c5-2e51-8bca-e90f-aeed72d4bce8 | https://echa.europa.eu/documents/10162/05406482-711d-a786-c26d-47115cab58c1 | https://echa.europa.eu/documents/10162/1221ef83-3bd4-dbdd-08d2-1be0a9a49023
2-(2H-benzotriazol-2-yl)-4-(1,1,3,3-tetramethylbutyl)phenol (UV-329) |  | 221-573-5 | 3147-75-9 | vPvB (Article 57e) | 23-Jan-2024 | https://echa.europa.eu/documents/10162/ee0c2587-d57b-01d1-9f91-2309c8bf0d89 | https://echa.europa.eu/documents/10162/fca389fe-e982-c59c-8c1e-cbadb70c65e6 | https://echa.europa.eu/documents/10162/fd8ac81a-0677-7b9d-8cbe-fdd091a689b1 | https://echa.europa.eu/documents/10162/f459fc07-ef0e-e1f4-10b8-29bf4a6a2ca6
2,4,6-tri-tert-butylphenol |  | 211-989-5 | 732-26-3 | Toxic for reproduction (Article 57c)#PBT (Article 57d) | 23-Jan-2024 | https://echa.europa.eu/documents/10162/4e4d1a17-caca-03c5-8f48-42cc69a75cc3 | https://echa.europa.eu/documents/10162/b998098f-c61f-6c7a-1500-2e10d3f291c2 | https://echa.europa.eu/documents/10162/3c90b7d6-d44d-3158-09be-62500f784117 | https://echa.europa.eu/documents/10162/cfa927dc-22e0-671c-4ea8-38ff1479c011
diphenyl(2,4,6-trimethylbenzoyl)phosphine oxide |  | 278-355-8 | 75980-60-8 | Toxic for reproduction (Article 57c) | 14-Jun-2023 | https://echa.europa.eu/documents/10162/a9f6af4d-579d-12ab-1e99-dc94b2be02b5 | https://echa.europa.eu/documents/10162/bba7bd61-c532-61e7-f84a-46be8495747b | https://echa.europa.eu/documents/10162/2c882400-790a-8df7-5da1-5434e2ddaa2b | https://echa.europa.eu/documents/10162/95a153af-c61f-9216-eee2-76123989bc6e
bis(4-chlorophenyl) sulphone |  | 201-247-9 | 80-07-9 | vPvB (Article 57e) | 14-Jun-2023 | https://echa.europa.eu/documents/10162/9632ec09-e901-d718-e54c-78ea9a7ef1d2 | https://echa.europa.eu/documents/10162/63a9b09d-1319-9c0d-3390-6d41aad19e45 | https://echa.europa.eu/documents/10162/78530af8-b2e3-a962-c6bc-07405104b4b7 | https://echa.europa.eu/documents/10162/59aaa00c-81fa-6259-a53a-aae021b40f38
reaction mass of 2,2,3,3,5,5,6,6-octafluoro-4-(1,1,1,2,3,3,3-heptafluoropropan-2-yl)morpholine and 2,2,3,3,5,5,6,6-octafluoro-4-(heptafluoropropyl)morpholine |  | 473-390-7 | - | vPvB (Article 57e) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/63bacec8-0474-2eb7-f806-df033037e179 | https://echa.europa.eu/documents/10162/9e6b1274-70f2-bf1e-90d9-a12700e18bd5 | https://echa.europa.eu/documents/10162/3f294a7d-108a-5793-8f18-17bc85716e08 | https://echa.europa.eu/documents/10162/6b868b22-2b62-1e00-19cf-b3747c19d7b6
Perfluoroheptanoic acid and its salts |  | - | - | Toxic for reproduction (Article 57c)#PBT (Article 57d)#vPvB (Article 57e)#Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/a74c795f-dd93-830a-08a8-c6c585400742 | https://echa.europa.eu/documents/10162/1f041880-8c8b-ee7a-2bef-fdb84e6b39da | https://echa.europa.eu/documents/10162/0eb7e7f6-a1a1-831e-4d0d-e13c1b13c28a | https://echa.europa.eu/documents/10162/bbe905ea-43ce-50e7-694a-171e6327fa98
Perfluoroheptanoic acid |  | 206-798-9 | 375-85-9 | Toxic for reproduction (Article 57c)#PBT (Article 57d)#vPvB (Article 57e)#Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/a74c795f-dd93-830a-08a8-c6c585400742 | https://echa.europa.eu/documents/10162/1f041880-8c8b-ee7a-2bef-fdb84e6b39da | https://echa.europa.eu/documents/10162/0eb7e7f6-a1a1-831e-4d0d-e13c1b13c28a | https://echa.europa.eu/documents/10162/bbe905ea-43ce-50e7-694a-171e6327fa98
Sodium perfluoroheptanoate |  | 243-518-4 | 20109-59-5 | Toxic for reproduction (Article 57c)#PBT (Article 57d)#vPvB (Article 57e)#Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/a74c795f-dd93-830a-08a8-c6c585400742 | https://echa.europa.eu/documents/10162/1f041880-8c8b-ee7a-2bef-fdb84e6b39da | https://echa.europa.eu/documents/10162/0eb7e7f6-a1a1-831e-4d0d-e13c1b13c28a | https://echa.europa.eu/documents/10162/bbe905ea-43ce-50e7-694a-171e6327fa98
potassium perfluoroheptanoate |  | - | 21049-36-5 | Toxic for reproduction (Article 57c)#PBT (Article 57d)#vPvB (Article 57e)#Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/a74c795f-dd93-830a-08a8-c6c585400742 | https://echa.europa.eu/documents/10162/1f041880-8c8b-ee7a-2bef-fdb84e6b39da | https://echa.europa.eu/documents/10162/0eb7e7f6-a1a1-831e-4d0d-e13c1b13c28a | https://echa.europa.eu/documents/10162/bbe905ea-43ce-50e7-694a-171e6327fa98
Ammonium perfluoroheptanoate |  | 228-098-2 | 6130-43-4 | Toxic for reproduction (Article 57c)#PBT (Article 57d)#vPvB (Article 57e)#Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/a74c795f-dd93-830a-08a8-c6c585400742 | https://echa.europa.eu/documents/10162/1f041880-8c8b-ee7a-2bef-fdb84e6b39da | https://echa.europa.eu/documents/10162/0eb7e7f6-a1a1-831e-4d0d-e13c1b13c28a | https://echa.europa.eu/documents/10162/bbe905ea-43ce-50e7-694a-171e6327fa98
Melamine |  | 203-615-4 | 108-78-1 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/653a0781-6b4b-d085-5393-edc530459a4f | https://echa.europa.eu/documents/10162/400552f2-c5b1-b2d7-2572-6ffa0cafa310 | https://echa.europa.eu/documents/10162/a94feb1d-4719-61ce-f815-0ee82ba17762 | https://echa.europa.eu/documents/10162/9e7356e0-1da3-6710-d5eb-e1409a3a9ffa
Isobutyl 4-hydroxybenzoate |  | 224-208-8 | 4247-02-3 | Endocrine disrupting properties (Article 57(f) - human health) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/f7f66e6b-8952-54fc-eef6-ab16d5d294bb | https://echa.europa.eu/documents/10162/3ee697d2-1942-3ab9-45ae-2579795169ce | https://echa.europa.eu/documents/10162/7805bba2-8ee9-6579-bf7b-88f7e2030c9a#https://echa.europa.eu/documents/10162/d85e48f5-7ef7-c2d9-73d3-8d65ecd73f12#https://echa.europa.eu/documents/10162/e3035965-6913-bd85-1d3e-cac410bf2fcb | https://echa.europa.eu/documents/10162/7241207b-b1c8-f174-d346-5051ce1583ff
bis(2-ethylhexyl) tetrabromophthalate covering any of the individual isomers and/or combinations thereof |  | - | - | vPvB (Article 57e) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/03b75dc6-489d-44a4-3a91-9e8854a4584e | https://echa.europa.eu/documents/10162/042c08dc-3385-5cc7-b851-e60ecead7a0b | https://echa.europa.eu/documents/10162/170b9b2f-a1d0-9661-d36d-faf0846f6763 | https://echa.europa.eu/documents/10162/e7d3971a-6f01-b503-13d9-d82cecf4b0ad
Bis(2-ethylhexyl) tetrabromophthalate |  | 247-426-5 | 26040-51-7 | vPvB (Article 57e) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/03b75dc6-489d-44a4-3a91-9e8854a4584e | https://echa.europa.eu/documents/10162/042c08dc-3385-5cc7-b851-e60ecead7a0b | https://echa.europa.eu/documents/10162/170b9b2f-a1d0-9661-d36d-faf0846f6763 | https://echa.europa.eu/documents/10162/e7d3971a-6f01-b503-13d9-d82cecf4b0ad
Barium diboron tetraoxide |  | 237-222-4 | 13701-59-2 | Toxic for reproduction (Article 57c) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/301582c5-3f0d-8698-f00a-fcfd56737350 | https://echa.europa.eu/documents/10162/ece3647e-bf00-b6bc-b668-9c1688f9e0a1 | https://echa.europa.eu/documents/10162/dd52fbc0-7d62-2708-7285-e486c67d52c2 | https://echa.europa.eu/documents/10162/4ddfad88-87cd-15bc-0297-0a784b23a1bf
4,4'-sulphonyldiphenol |  | 201-250-5 | 80-09-1 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/4ccf3271-2d8a-a026-a689-52f26f2610ee | https://echa.europa.eu/documents/10162/ee7a9523-4533-90fe-3e9f-84f4e6cdce32 | https://echa.europa.eu/documents/10162/cffe2e0d-2c25-bfb1-f5ff-eca0b3b30f32 | https://echa.europa.eu/documents/10162/9a0c40de-ee96-49b4-9ac5-d6dee8fa5985
2,2',6,6'-tetrabromo-4,4'-isopropylidenediphenol |  | 201-236-9 | 79-94-7 | Carcinogenic (Article 57a) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/dfb4de3c-5868-4ef8-d675-ba8daf759cf3 | https://echa.europa.eu/documents/10162/d89e7bbc-bac9-a092-4c94-30a067b46656 | https://echa.europa.eu/documents/10162/9238ce7d-52aa-e2a1-5474-b5adf5ca9003 | https://echa.europa.eu/documents/10162/62473a52-ba74-48e6-90eb-108788cf592b
1,1'-[ethane-1,2-diylbisoxy]bis[2,4,6-tribromobenzene] |  | 253-692-3 | 37853-59-1 | vPvB (Article 57e) | 17-Jan-2023 | https://echa.europa.eu/documents/10162/382d3710-462b-09b3-4a65-a3df830db6e3 | https://echa.europa.eu/documents/10162/dfd5c516-be38-2da8-5a2c-5815a938e9ab | https://echa.europa.eu/documents/10162/f23e510a-2a13-8942-c9f9-4e9951015862 | https://echa.europa.eu/documents/10162/91e92169-98d3-e2a4-d0c1-d63f4fd2da0b
N-(hydroxymethyl)acrylamide |  | 213-103-2 | 924-42-5 | Carcinogenic (Article 57a)#Mutagenic (Article 57b) | 10-Jun-2022 | https://echa.europa.eu/documents/10162/02904d92-61a9-be52-f565-a6ec65c0fcbd | https://echa.europa.eu/documents/10162/7b2404fe-f91e-0b39-32b7-b853e82ab2af | https://echa.europa.eu/documents/10162/f5c6a946-c14d-b45d-b9ca-cc0bd85403c4 | https://echa.europa.eu/documents/10162/32f6be21-a6b3-9730-ddc5-2bb7050499a5
tris(2-methoxyethoxy)vinylsilane |  | 213-934-0 | 1067-53-4 | Toxic for reproduction (Article 57c) | 17-Jan-2022 | https://echa.europa.eu/documents/10162/1b4d5632-3958-a563-bbd4-d318e59e520b | https://echa.europa.eu/documents/10162/b175f54b-e658-de9b-0dec-dad099bb1f17 | https://echa.europa.eu/documents/10162/21bc25f0-b6f0-1174-18ba-a9081578097e | https://echa.europa.eu/documents/10162/3914417a-18b5-5334-8d25-fc1a40d0782e
S-(tricyclo(5.2.1.0'2,6)deca-3-en-8(or 9)-yl O-(isopropyl or isobutyl or 2-ethylhexyl) O-(isopropyl or isobutyl or 2-ethylhexyl) phosphorodithioate | X4261 | 401-850-9 | 255881-94-8 | PBT (Article 57d) | 17-Jan-2022 | https://echa.europa.eu/documents/10162/a70cd813-f087-c8df-03ce-1359f8f7e516 | https://echa.europa.eu/documents/10162/fab08a5e-d762-bc39-8999-fab001c36b7c | https://echa.europa.eu/documents/10162/4f98791d-869c-d7a4-5791-9380fa825a31 | https://echa.europa.eu/documents/10162/22c2d1af-3748-206d-3031-02de14d58047
6,6'-di-tert-butyl-2,2'-methylenedi-p-cresol |  | 204-327-1 | 119-47-1 | Toxic for reproduction (Article 57c) | 17-Jan-2022 | https://echa.europa.eu/documents/10162/fe17e78e-729f-770e-5f4a-5da23bb2c361 | https://echa.europa.eu/documents/10162/2409e362-247f-8ca4-c856-40231ac2dc88 | https://echa.europa.eu/documents/10162/7f58d058-f965-6fe2-864c-4df6a605c8ce | https://echa.europa.eu/documents/10162/92936d26-abe2-f0cb-9bdd-f5be90304c5b
(±)-1,7,7-trimethyl-3-[(4-methylphenyl)methylene]bicyclo[2.2.1]heptan-2-one covering any of the individual isomers and/or combinations thereof (4-MBC) |  | - | - | Endocrine disrupting properties (Article 57(f) - human health) | 17-Jan-2022 | https://echa.europa.eu/documents/10162/b05be4ea-eff8-5c93-f021-9961a9abc10d | https://echa.europa.eu/documents/10162/94537711-ea88-61c3-4a4a-d134d900b81c | https://echa.europa.eu/documents/10162/41008a30-53db-84bd-6d4e-7f31d9aa78dc | https://echa.europa.eu/documents/10162/36e3ce7f-fa01-9134-239a-9dae6e783fd1
(±)-1,7,7-trimethyl-3-[(4-methylphenyl)methylene]bicyclo[2.2.1]heptan-2-one |  | 253-242-6 | 36861-47-9 | Endocrine disrupting properties (Article 57(f) - human health) | 17-Jan-2022 | https://echa.europa.eu/documents/10162/b05be4ea-eff8-5c93-f021-9961a9abc10d | https://echa.europa.eu/documents/10162/94537711-ea88-61c3-4a4a-d134d900b81c | https://echa.europa.eu/documents/10162/41008a30-53db-84bd-6d4e-7f31d9aa78dc | https://echa.europa.eu/documents/10162/36e3ce7f-fa01-9134-239a-9dae6e783fd1
(3E)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | - | 1782069-81-1 | Endocrine disrupting properties (Article 57(f) - human health) | 17-Jan-2022 | https://echa.europa.eu/documents/10162/b05be4ea-eff8-5c93-f021-9961a9abc10d | https://echa.europa.eu/documents/10162/94537711-ea88-61c3-4a4a-d134d900b81c | https://echa.europa.eu/documents/10162/41008a30-53db-84bd-6d4e-7f31d9aa78dc | https://echa.europa.eu/documents/10162/36e3ce7f-fa01-9134-239a-9dae6e783fd1
(1R,3E,4S)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | - | 95342-41-9 | Endocrine disrupting properties (Article 57(f) - human health) | 17-Jan-2022 | https://echa.europa.eu/documents/10162/b05be4ea-eff8-5c93-f021-9961a9abc10d | https://echa.europa.eu/documents/10162/94537711-ea88-61c3-4a4a-d134d900b81c | https://echa.europa.eu/documents/10162/41008a30-53db-84bd-6d4e-7f31d9aa78dc | https://echa.europa.eu/documents/10162/36e3ce7f-fa01-9134-239a-9dae6e783fd1
(1S,3E,4R)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | - | 852541-30-1 | Endocrine disrupting properties (Article 57(f) - human health) | 17-Jan-2022 | https://echa.europa.eu/documents/10162/b05be4ea-eff8-5c93-f021-9961a9abc10d | https://echa.europa.eu/documents/10162/94537711-ea88-61c3-4a4a-d134d900b81c | https://echa.europa.eu/documents/10162/41008a30-53db-84bd-6d4e-7f31d9aa78dc | https://echa.europa.eu/documents/10162/36e3ce7f-fa01-9134-239a-9dae6e783fd1
(1R,3Z,4S)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | - | 852541-21-0 | Endocrine disrupting properties (Article 57(f) - human health) | 17-Jan-2022 | https://echa.europa.eu/documents/10162/b05be4ea-eff8-5c93-f021-9961a9abc10d | https://echa.europa.eu/documents/10162/94537711-ea88-61c3-4a4a-d134d900b81c | https://echa.europa.eu/documents/10162/41008a30-53db-84bd-6d4e-7f31d9aa78dc | https://echa.europa.eu/documents/10162/36e3ce7f-fa01-9134-239a-9dae6e783fd1
(1R,4S)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | - | 741687-98-9 | Endocrine disrupting properties (Article 57(f) - human health) | 17-Jan-2022 | https://echa.europa.eu/documents/10162/b05be4ea-eff8-5c93-f021-9961a9abc10d | https://echa.europa.eu/documents/10162/94537711-ea88-61c3-4a4a-d134d900b81c | https://echa.europa.eu/documents/10162/41008a30-53db-84bd-6d4e-7f31d9aa78dc | https://echa.europa.eu/documents/10162/36e3ce7f-fa01-9134-239a-9dae6e783fd1
(1S,3Z,4R)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | - | 852541-25-4 | Endocrine disrupting properties (Article 57(f) - human health) | 17-Jan-2022 | https://echa.europa.eu/documents/10162/b05be4ea-eff8-5c93-f021-9961a9abc10d | https://echa.europa.eu/documents/10162/94537711-ea88-61c3-4a4a-d134d900b81c | https://echa.europa.eu/documents/10162/41008a30-53db-84bd-6d4e-7f31d9aa78dc | https://echa.europa.eu/documents/10162/36e3ce7f-fa01-9134-239a-9dae6e783fd1
Phenol, alkylation products (mainly in para position) with C12-rich branched alkyl chains from oligomerisation, covering any individual isomers and/ or combinations thereof (PDDP) |  | - | - | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/6c954d43-13d4-b12b-4c0a-bc6c1f35fc72 | https://echa.europa.eu/documents/10162/6823656a-b163-5985-75c4-a768a67b9d3c | https://echa.europa.eu/documents/10162/6a83bb36-00c4-3b6a-74a4-76a1223aacd9 | https://echa.europa.eu/documents/10162/7c788805-871e-9568-78a8-402f94541ae0
4-isododecylphenol |  | - | 27459-10-5 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/6c954d43-13d4-b12b-4c0a-bc6c1f35fc72 | https://echa.europa.eu/documents/10162/6823656a-b163-5985-75c4-a768a67b9d3c | https://echa.europa.eu/documents/10162/6a83bb36-00c4-3b6a-74a4-76a1223aacd9 | https://echa.europa.eu/documents/10162/7c788805-871e-9568-78a8-402f94541ae0
Phenol, tetrapropylene- |  | - | 57427-55-1 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/6c954d43-13d4-b12b-4c0a-bc6c1f35fc72 | https://echa.europa.eu/documents/10162/6823656a-b163-5985-75c4-a768a67b9d3c | https://echa.europa.eu/documents/10162/6a83bb36-00c4-3b6a-74a4-76a1223aacd9 | https://echa.europa.eu/documents/10162/7c788805-871e-9568-78a8-402f94541ae0
Phenol, 4-dodecyl, branched |  | - | 210555-94-5 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/6c954d43-13d4-b12b-4c0a-bc6c1f35fc72 | https://echa.europa.eu/documents/10162/6823656a-b163-5985-75c4-a768a67b9d3c | https://echa.europa.eu/documents/10162/6a83bb36-00c4-3b6a-74a4-76a1223aacd9 | https://echa.europa.eu/documents/10162/7c788805-871e-9568-78a8-402f94541ae0
Phenol, (tetrapropenyl) derivatives |  | - | 74499-35-7 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/6c954d43-13d4-b12b-4c0a-bc6c1f35fc72 | https://echa.europa.eu/documents/10162/6823656a-b163-5985-75c4-a768a67b9d3c | https://echa.europa.eu/documents/10162/6a83bb36-00c4-3b6a-74a4-76a1223aacd9 | https://echa.europa.eu/documents/10162/7c788805-871e-9568-78a8-402f94541ae0
Phenol, 4-isododecyl- |  | - | 27147-75-7 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/6c954d43-13d4-b12b-4c0a-bc6c1f35fc72 | https://echa.europa.eu/documents/10162/6823656a-b163-5985-75c4-a768a67b9d3c | https://echa.europa.eu/documents/10162/6a83bb36-00c4-3b6a-74a4-76a1223aacd9 | https://echa.europa.eu/documents/10162/7c788805-871e-9568-78a8-402f94541ae0
Phenol, dodecyl-, branched |  | 310-154-3 | 121158-58-5 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/6c954d43-13d4-b12b-4c0a-bc6c1f35fc72 | https://echa.europa.eu/documents/10162/6823656a-b163-5985-75c4-a768a67b9d3c | https://echa.europa.eu/documents/10162/6a83bb36-00c4-3b6a-74a4-76a1223aacd9 | https://echa.europa.eu/documents/10162/7c788805-871e-9568-78a8-402f94541ae0
orthoboric acid, sodium salt |  | - | - | Toxic for reproduction (Article 57c) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/2ffc10a7-fe57-2c75-9af2-001cb5f5aa1e | https://echa.europa.eu/documents/10162/0b757d38-e10b-6692-2e56-b6d26b366dc8 | https://echa.europa.eu/documents/10162/c4c5aa8d-8f61-51ef-d120-6f54b5d9a0c0 | https://echa.europa.eu/documents/10162/4b353aba-79ac-3605-690c-e6f1ca320647
boric acid (H3BO3), sodium salt, hydrate |  | - | 25747-83-5 | Toxic for reproduction (Article 57c) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/2ffc10a7-fe57-2c75-9af2-001cb5f5aa1e | https://echa.europa.eu/documents/10162/0b757d38-e10b-6692-2e56-b6d26b366dc8 | https://echa.europa.eu/documents/10162/c4c5aa8d-8f61-51ef-d120-6f54b5d9a0c0 | https://echa.europa.eu/documents/10162/4b353aba-79ac-3605-690c-e6f1ca320647
Trisodium orthoborate |  | 238-253-6 | 14312-40-4 | Toxic for reproduction (Article 57c) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/2ffc10a7-fe57-2c75-9af2-001cb5f5aa1e | https://echa.europa.eu/documents/10162/0b757d38-e10b-6692-2e56-b6d26b366dc8 | https://echa.europa.eu/documents/10162/c4c5aa8d-8f61-51ef-d120-6f54b5d9a0c0 | https://echa.europa.eu/documents/10162/4b353aba-79ac-3605-690c-e6f1ca320647
Boric acid, sodium salt |  | 215-604-1 | 1333-73-9 | Toxic for reproduction (Article 57c) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/2ffc10a7-fe57-2c75-9af2-001cb5f5aa1e | https://echa.europa.eu/documents/10162/0b757d38-e10b-6692-2e56-b6d26b366dc8 | https://echa.europa.eu/documents/10162/c4c5aa8d-8f61-51ef-d120-6f54b5d9a0c0 | https://echa.europa.eu/documents/10162/4b353aba-79ac-3605-690c-e6f1ca320647
Boric acid (H3BO3), disodium salt |  | - | 22454-04-2 | Toxic for reproduction (Article 57c) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/2ffc10a7-fe57-2c75-9af2-001cb5f5aa1e | https://echa.europa.eu/documents/10162/0b757d38-e10b-6692-2e56-b6d26b366dc8 | https://echa.europa.eu/documents/10162/c4c5aa8d-8f61-51ef-d120-6f54b5d9a0c0 | https://echa.europa.eu/documents/10162/4b353aba-79ac-3605-690c-e6f1ca320647
Orthoboric acid, sodium salt |  | 237-560-2 | 13840-56-7 | Toxic for reproduction (Article 57c) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/2ffc10a7-fe57-2c75-9af2-001cb5f5aa1e | https://echa.europa.eu/documents/10162/0b757d38-e10b-6692-2e56-b6d26b366dc8 | https://echa.europa.eu/documents/10162/c4c5aa8d-8f61-51ef-d120-6f54b5d9a0c0 | https://echa.europa.eu/documents/10162/4b353aba-79ac-3605-690c-e6f1ca320647
boric acid (H3BO3), sodium salt (1:1) |  | - | 14890-53-0 | Toxic for reproduction (Article 57c) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/2ffc10a7-fe57-2c75-9af2-001cb5f5aa1e | https://echa.europa.eu/documents/10162/0b757d38-e10b-6692-2e56-b6d26b366dc8 | https://echa.europa.eu/documents/10162/c4c5aa8d-8f61-51ef-d120-6f54b5d9a0c0 | https://echa.europa.eu/documents/10162/4b353aba-79ac-3605-690c-e6f1ca320647
Medium-chain chlorinated paraffins (MCCP) | UVCB substances consisting of more than or equal to 80% linear chloroalkanes with carbon chain lengths within the range from C14 to C17 | - | - | PBT (Article 57d)#vPvB (Article 57e) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/af3efea2-1518-3bbe-0bf5-3867131c2f4c | https://echa.europa.eu/documents/10162/799da465-c1f4-00ea-e771-53d405fee8af | https://echa.europa.eu/documents/10162/98611952-49d5-b0be-d4b9-3df6579315c9 | https://echa.europa.eu/documents/10162/2b9b5c86-8980-3ce8-5bda-77928392642c#https://echa.europa.eu/documents/10162/a47cba77-d86d-006e-80e5-bbd3a103a193#https://echa.europa.eu/documents/10162/5a99cbfc-fd8b-df56-af12-c39a5c6cc8ce
di-, tri- and tetrachlorotetradecane |  | - | - | PBT (Article 57d)#vPvB (Article 57e) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/af3efea2-1518-3bbe-0bf5-3867131c2f4c | https://echa.europa.eu/documents/10162/799da465-c1f4-00ea-e771-53d405fee8af | https://echa.europa.eu/documents/10162/98611952-49d5-b0be-d4b9-3df6579315c9 | https://echa.europa.eu/documents/10162/2b9b5c86-8980-3ce8-5bda-77928392642c#https://echa.europa.eu/documents/10162/a47cba77-d86d-006e-80e5-bbd3a103a193#https://echa.europa.eu/documents/10162/5a99cbfc-fd8b-df56-af12-c39a5c6cc8ce
Alkanes, C14-17, chloro |  | 287-477-0 | 85535-85-9 | PBT (Article 57d)#vPvB (Article 57e) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/af3efea2-1518-3bbe-0bf5-3867131c2f4c | https://echa.europa.eu/documents/10162/799da465-c1f4-00ea-e771-53d405fee8af | https://echa.europa.eu/documents/10162/98611952-49d5-b0be-d4b9-3df6579315c9 | https://echa.europa.eu/documents/10162/2b9b5c86-8980-3ce8-5bda-77928392642c#https://echa.europa.eu/documents/10162/a47cba77-d86d-006e-80e5-bbd3a103a193#https://echa.europa.eu/documents/10162/5a99cbfc-fd8b-df56-af12-c39a5c6cc8ce
Tetradecane, chloro derivs. |  | - | 198840-65-2 | PBT (Article 57d)#vPvB (Article 57e) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/af3efea2-1518-3bbe-0bf5-3867131c2f4c | https://echa.europa.eu/documents/10162/799da465-c1f4-00ea-e771-53d405fee8af | https://echa.europa.eu/documents/10162/98611952-49d5-b0be-d4b9-3df6579315c9 | https://echa.europa.eu/documents/10162/2b9b5c86-8980-3ce8-5bda-77928392642c#https://echa.europa.eu/documents/10162/a47cba77-d86d-006e-80e5-bbd3a103a193#https://echa.europa.eu/documents/10162/5a99cbfc-fd8b-df56-af12-c39a5c6cc8ce
Alkanes, C14-16, chloro |  | - | 1372804-76-6 | PBT (Article 57d)#vPvB (Article 57e) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/af3efea2-1518-3bbe-0bf5-3867131c2f4c | https://echa.europa.eu/documents/10162/799da465-c1f4-00ea-e771-53d405fee8af | https://echa.europa.eu/documents/10162/98611952-49d5-b0be-d4b9-3df6579315c9 | https://echa.europa.eu/documents/10162/2b9b5c86-8980-3ce8-5bda-77928392642c#https://echa.europa.eu/documents/10162/a47cba77-d86d-006e-80e5-bbd3a103a193#https://echa.europa.eu/documents/10162/5a99cbfc-fd8b-df56-af12-c39a5c6cc8ce
glutaral |  | 203-856-5 | 111-30-8 | Respiratory sensitising properties (Article 57(f) - human health) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/f8bab22e-f605-bbed-66eb-1332835436f9 | https://echa.europa.eu/documents/10162/b5ed1907-5313-68e0-c7a7-751d9de47330 | https://echa.europa.eu/documents/10162/c02d3fc8-772e-14d0-1eda-2db650aa0b8f | https://echa.europa.eu/documents/10162/6e8721b5-cda7-662b-17e5-2f395fe0d521
4,4'-(1-methylpropylidene)bisphenol |  | 201-025-1 | 77-40-7 | Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/ab77aafb-7b98-5cbb-3416-fc28e393a48e | https://echa.europa.eu/documents/10162/6755b6f5-11e8-ba1f-8168-6325fad64bce | https://echa.europa.eu/documents/10162/dd9e01dd-8aac-8ecd-4c39-4da6c92c2ec8 | https://echa.europa.eu/documents/10162/141706c5-95f0-5197-6ab0-3287bc0e7f46
2-(4-tert-butylbenzyl)propionaldehyde and its individual stereoisomers |  | - | - | Toxic for reproduction (Article 57c) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/142cc638-9c2b-b2d2-09da-b38806ba7565 | https://echa.europa.eu/documents/10162/17f27f86-efe0-c723-32b5-ff774ae31a16 | https://echa.europa.eu/documents/10162/fc185a73-227f-328a-c6c2-e77f848f53c6 | https://echa.europa.eu/documents/10162/276ad049-1ceb-e417-b4a7-2c72322e0ce5
(2R)-3-(4-tert-butylphenyl)-2-methylpropanal |  | - | 75166-31-3 | Toxic for reproduction (Article 57c) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/142cc638-9c2b-b2d2-09da-b38806ba7565 | https://echa.europa.eu/documents/10162/17f27f86-efe0-c723-32b5-ff774ae31a16 | https://echa.europa.eu/documents/10162/fc185a73-227f-328a-c6c2-e77f848f53c6 | https://echa.europa.eu/documents/10162/276ad049-1ceb-e417-b4a7-2c72322e0ce5
2-(4-tert-butylbenzyl)propionaldehyde |  | 201-289-8 | 80-54-6 | Toxic for reproduction (Article 57c) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/142cc638-9c2b-b2d2-09da-b38806ba7565 | https://echa.europa.eu/documents/10162/17f27f86-efe0-c723-32b5-ff774ae31a16 | https://echa.europa.eu/documents/10162/fc185a73-227f-328a-c6c2-e77f848f53c6 | https://echa.europa.eu/documents/10162/276ad049-1ceb-e417-b4a7-2c72322e0ce5
(2S)-3-(4-tert-butylphenyl)-2-methylpropanal |  | - | 75166-30-2 | Toxic for reproduction (Article 57c) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/142cc638-9c2b-b2d2-09da-b38806ba7565 | https://echa.europa.eu/documents/10162/17f27f86-efe0-c723-32b5-ff774ae31a16 | https://echa.europa.eu/documents/10162/fc185a73-227f-328a-c6c2-e77f848f53c6 | https://echa.europa.eu/documents/10162/276ad049-1ceb-e417-b4a7-2c72322e0ce5
2,2-bis(bromomethyl)propane-1,3-diol (BMP); 2,2-dimethylpropan-1-ol, tribromo derivative/3-bromo-2,2-bis(bromomethyl)-1-propanol (TBNPA); 2,3-dibromo-1-propanol (2,3-DBPA) |  | - | - | Carcinogenic (Article 57a) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/27d5b22f-c74a-00c5-e988-d0d4db720822 | https://echa.europa.eu/documents/10162/12706f21-c551-9d0e-7b28-0c07bfcd89be | https://echa.europa.eu/documents/10162/22d07e1a-0a10-7050-3dd7-c28264224db0 | https://echa.europa.eu/documents/10162/97e7c76f-99ad-ae48-dda3-fe76c5098e3f
3-bromo-2,2-bis(bromomethyl)-1-propanol (TBNPA) |  | - | 1522-92-5 | Carcinogenic (Article 57a) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/27d5b22f-c74a-00c5-e988-d0d4db720822 | https://echa.europa.eu/documents/10162/12706f21-c551-9d0e-7b28-0c07bfcd89be | https://echa.europa.eu/documents/10162/22d07e1a-0a10-7050-3dd7-c28264224db0 | https://echa.europa.eu/documents/10162/97e7c76f-99ad-ae48-dda3-fe76c5098e3f
2,2-dimethylpropan-1-ol, tribromo derivative (TBNPA) |  | 253-057-0 | 36483-57-5 | Carcinogenic (Article 57a) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/27d5b22f-c74a-00c5-e988-d0d4db720822 | https://echa.europa.eu/documents/10162/12706f21-c551-9d0e-7b28-0c07bfcd89be | https://echa.europa.eu/documents/10162/22d07e1a-0a10-7050-3dd7-c28264224db0 | https://echa.europa.eu/documents/10162/97e7c76f-99ad-ae48-dda3-fe76c5098e3f
2,3-dibromo-1-propanol (2,3-DBPA) |  | 202-480-9 | 96-13-9 | Carcinogenic (Article 57a) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/27d5b22f-c74a-00c5-e988-d0d4db720822 | https://echa.europa.eu/documents/10162/12706f21-c551-9d0e-7b28-0c07bfcd89be | https://echa.europa.eu/documents/10162/22d07e1a-0a10-7050-3dd7-c28264224db0 | https://echa.europa.eu/documents/10162/97e7c76f-99ad-ae48-dda3-fe76c5098e3f
2,2-bis(bromomethyl)propane-1,3-diol (BMP) |  | 221-967-7 | 3296-90-0 | Carcinogenic (Article 57a) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/27d5b22f-c74a-00c5-e988-d0d4db720822 | https://echa.europa.eu/documents/10162/12706f21-c551-9d0e-7b28-0c07bfcd89be | https://echa.europa.eu/documents/10162/22d07e1a-0a10-7050-3dd7-c28264224db0 | https://echa.europa.eu/documents/10162/97e7c76f-99ad-ae48-dda3-fe76c5098e3f
1,4-dioxane |  | 204-661-8 | 123-91-1 | Carcinogenic (Article 57a)#Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 08-Jul-2021 | https://echa.europa.eu/documents/10162/17ab47c9-f60e-ecbc-da78-474468076133 | https://echa.europa.eu/documents/10162/031ccf23-e060-c12b-5978-757b21f51ef5 | https://echa.europa.eu/documents/10162/ce76bdd9-006a-19fc-05eb-4bb60daa0256 | https://echa.europa.eu/documents/10162/9b5464f1-ef30-7e57-5184-10291c44328e
Dioctyltin dilaurate, stannane, dioctyl-, bis(coco acyloxy) derivs., and any other stannane, dioctyl-, bis(fatty acyloxy) derivs. wherein C12 is the predominant carbon number of the fatty acyloxy moiety |  | - | - | Toxic for reproduction (Article 57c) | 19-Jan-2021 | https://echa.europa.eu/documents/10162/7fc5894b-3785-4c7f-bba8-5a172ad287a7 | https://echa.europa.eu/documents/10162/3daf4bfd-78b4-7e5a-717f-f8169e350d65 | https://echa.europa.eu/documents/10162/879129fe-2615-0c9f-e91b-00b900d26c17 | https://echa.europa.eu/documents/10162/77d1fdea-6e60-9311-7e9e-301686ef38c6
dioctyltin dilaurate; stannane, dioctyl-, bis(coco acyloxy) derivs. |  | - | - | Toxic for reproduction (Article 57c) | 19-Jan-2021 | https://echa.europa.eu/documents/10162/7fc5894b-3785-4c7f-bba8-5a172ad287a7 | https://echa.europa.eu/documents/10162/3daf4bfd-78b4-7e5a-717f-f8169e350d65 | https://echa.europa.eu/documents/10162/879129fe-2615-0c9f-e91b-00b900d26c17 | https://echa.europa.eu/documents/10162/77d1fdea-6e60-9311-7e9e-301686ef38c6
Dioctyltin dilaurate |  | 222-883-3 | 3648-18-8 | Toxic for reproduction (Article 57c) | 19-Jan-2021 | https://echa.europa.eu/documents/10162/7fc5894b-3785-4c7f-bba8-5a172ad287a7 | https://echa.europa.eu/documents/10162/3daf4bfd-78b4-7e5a-717f-f8169e350d65 | https://echa.europa.eu/documents/10162/879129fe-2615-0c9f-e91b-00b900d26c17 | https://echa.europa.eu/documents/10162/77d1fdea-6e60-9311-7e9e-301686ef38c6
Stannane, dioctyl-, bis(coco acyloxy) derivs. |  | 293-901-5 | 91648-39-4 | Toxic for reproduction (Article 57c) | 19-Jan-2021 | https://echa.europa.eu/documents/10162/7fc5894b-3785-4c7f-bba8-5a172ad287a7 | https://echa.europa.eu/documents/10162/3daf4bfd-78b4-7e5a-717f-f8169e350d65 | https://echa.europa.eu/documents/10162/879129fe-2615-0c9f-e91b-00b900d26c17 | https://echa.europa.eu/documents/10162/77d1fdea-6e60-9311-7e9e-301686ef38c6
Bis(2-(2-methoxyethoxy)ethyl)ether |  | 205-594-7 | 143-24-8 | Toxic for reproduction (Article 57c) | 19-Jan-2021 | https://echa.europa.eu/documents/10162/854c5708-b10f-7b75-52e7-968e7c448992 | https://echa.europa.eu/documents/10162/b3398d8d-f30d-1db5-1cbc-094de67753cf | https://echa.europa.eu/documents/10162/c99c2be3-8ba4-0a7c-4711-423639029884 | https://echa.europa.eu/documents/10162/3bdbc35b-91e1-3a8a-302a-2ba6313cc50e
Dibutylbis(pentane-2,4-dionato-O,O')tin |  | 245-152-0 | 22673-19-4 | Toxic for reproduction (Article 57c) | 25-Jun-2020 | https://echa.europa.eu/documents/10162/10861be8-f130-266d-ca8f-c6c7139fe6d9 | https://echa.europa.eu/documents/10162/6f88822c-faae-8f80-7121-9e64752668dc | https://echa.europa.eu/documents/10162/741c3cb4-34b1-3dc3-a12b-07d651f6eb53 | https://echa.europa.eu/documents/10162/cd37b495-7065-d09f-74b4-40b39eef4e82
Butyl 4-hydroxybenzoate |  | 202-318-7 | 94-26-8 | Endocrine disrupting properties (Article 57(f) - human health) | 25-Jun-2020 | https://echa.europa.eu/documents/10162/9227ca75-c14c-29ab-81ff-ee86058dd7a7 | https://echa.europa.eu/documents/10162/59ed15a3-66ee-b8c8-1384-54277d0223eb | https://echa.europa.eu/documents/10162/c61a924a-4b3a-d122-1343-d5557deb3029#https://echa.europa.eu/documents/10162/9c45b069-0fbf-c247-08c0-d0567c58a43e | https://echa.europa.eu/documents/10162/d4447c8f-c2f4-df26-c06e-c0a4e93b20e9
2-methylimidazole |  | 211-765-7 | 693-98-1 | Toxic for reproduction (Article 57c) | 25-Jun-2020 | https://echa.europa.eu/documents/10162/31fcbd73-6d35-06d9-35cd-42dc09362aa1 | https://echa.europa.eu/documents/10162/99d34015-c23f-882d-ba9e-8404434bb456 | https://echa.europa.eu/documents/10162/03e937e3-2f19-ad9d-6bd5-ab9ceae0d0c6 | https://echa.europa.eu/documents/10162/23a73a48-dbec-b764-c3e0-f3f45a7cd36d
1-vinylimidazole |  | 214-012-0 | 1072-63-5 | Toxic for reproduction (Article 57c) | 25-Jun-2020 | https://echa.europa.eu/documents/10162/4b662f2b-020e-e3a9-c398-3a6d79b65c35 | https://echa.europa.eu/documents/10162/24d19609-35d1-edaa-4aac-07db7025edde | https://echa.europa.eu/documents/10162/ce50a489-acc9-3908-e6af-d43fa4af292b | https://echa.europa.eu/documents/10162/36a36d48-2f35-cfe6-4bca-07870d3c65fd
Perfluorobutane sulfonic acid (PFBS) and its salts |  | - | - | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
bis(4-t-butylphenyl)iodonium perfluorobutanesulfonate |  | 432-660-4 | 194999-85-4 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
tetrabutyl-phosphonium nonafluoro-butane-1-sulfonate |  | 444-440-5 | 220689-12-3 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
dimethyl(phenyl)sulfanium perfluorobutanesulfonate |  | 452-310-4 | 220133-51-7 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
1-(4-butoxy-1-naphthalenyl)tetrahydrothiophenium 1,1,2,2,3,3,4,4,4-nonafluoro-1-butanesulfonate |  | 468-770-4 | - | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
Triphenylsulfanium perfluorobutane sulfonate |  | 478-340-8 | 144317-44-2 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
1,1,2,2,3,3,4,4,4-nonafluorobutane-1-sulphonic acid |  | 206-793-1 | 375-73-5 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
magnesium perfluorobutanesulfonate |  | - | 507453-86-3 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
lithium perfluorobutanesulfonate |  | - | 131651-65-5 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
morpholinium perfluorobutanesulfonate |  | - | 503155-89-3 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
Potassium 1,1,2,2,3,3,4,4,4-nonafluorobutane-1-sulphonate |  | 249-616-3 | 29420-49-3 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
Ammonium 1,1,2,2,3,3,4,4,4-nonafluorobutane-1-sulphonate |  | 269-513-7 | 68259-10-9 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
N,N,N-triethylethanaminium 1,1,2,2,3,3,4,4,4-nonafluorobutane-1-sulfonate |  | - | 25628-08-4 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/079c04a0-2464-4168-f132-a22ffb04d910 | https://echa.europa.eu/documents/10162/53791353-6577-a9e8-5600-c5ef51864652 | https://echa.europa.eu/documents/10162/891ab33d-d263-cc4b-0f2d-d84cfb7f424a | https://echa.europa.eu/documents/10162/84d51ffb-f3ac-f182-10f2-31c4d9386ab9 | The combined intrinsic properties justifying the inclusion as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: very high persistence, high mobility in water and soil, high potential for long-range transport, and difficulty of remediation and water purification as well as moderate bioaccumulation in humans. The observed probable serious effects for human health and the environment are thyroid hormonal disturbances and reproductive toxicity seen in rodents, and effects on liver, kidney and haematological system in rats, hormonal disturbances and effects on reproduction in marine medaka fish and effects on expression of hormone receptors in tadpoles. Together, these elements lead to a very high potential for irreversible effects.
Diisohexyl phthalate |  | 276-090-2 | 71850-09-4 | Toxic for reproduction (Article 57c) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/802ed481-8bf1-0096-61a4-0bbd2f9b573c | https://echa.europa.eu/documents/10162/c9adbb8d-3989-2380-cb54-f34940de0002 | https://echa.europa.eu/documents/10162/cc11a6e1-0b26-8a5d-189d-417eb06a60d4 | https://echa.europa.eu/documents/10162/9f133792-ce5e-da2c-6a2c-8616ec336f6e
2-methyl-1-(4-methylthiophenyl)-2-morpholinopropan-1-one | ACETOCURE 97; GENOCURE*PMP; IGM 4817; IRGACURE 907; SPEEDCURE 97 | 400-600-6 | 71868-10-5 | Toxic for reproduction (Article 57c) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/7d1e26b8-37af-4f0b-0d0f-5edb0b6e5a7f | https://echa.europa.eu/documents/10162/e31382a6-3e8a-a858-da3d-614c10f6c75d | https://echa.europa.eu/documents/10162/e4917be1-32fd-dbd6-e305-90f269d446dd | https://echa.europa.eu/documents/10162/4bd29548-f917-95c6-3c59-d6afd1e0901d
2-benzyl-2-dimethylamino-4'-morpholinobutyrophenone | CG 25-369; IRGACURE 369; TK 11-319 | 404-360-3 | 119313-12-1 | Toxic for reproduction (Article 57c) | 16-Jan-2020 | https://echa.europa.eu/documents/10162/ba42270a-93fe-7cde-6d41-7f7af4e00e66 | https://echa.europa.eu/documents/10162/dfd0d651-604f-f33f-0831-3672c1cbb580 | https://echa.europa.eu/documents/10162/6d5c6d89-bd66-3c04-4a31-f580476b88b4 | https://echa.europa.eu/documents/10162/54890824-f79f-c051-6c29-2efb018be01a
Tris(4-nonylphenyl, branched and linear) phosphite (TNPP) |  | - | - | Endocrine disrupting properties (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/a661cbe6-c805-b9c7-1287-53dee5cd1df5#https://echa.europa.eu/documents/10162/f162b35e-2aac-6b5a-a3de-3290bedd8805 | https://echa.europa.eu/documents/10162/8da7df6b-e62f-47f7-d9d1-154f8db267f8#https://echa.europa.eu/documents/10162/0a66176d-3f11-34ed-74e1-6062ae73f264 | https://echa.europa.eu/documents/10162/ceb2fc2c-78f2-627f-29ae-b37c0e91bc69#https://echa.europa.eu/documents/10162/f186ad0a-3831-025c-2628-3c0948e174d5 | https://echa.europa.eu/documents/10162/70879c79-2312-31bd-09e5-558df06f19e2#https://echa.europa.eu/documents/10162/a524f414-c041-2388-f47a-302a46efdc79 | In addition to its intrinsic properties (see ECHA decision D(2024)7663-DC of 21 January 2025), tris(4-nonylphenyl, branched and linear) phosphite (TNPP) also meets the criteria of Article 57(f) of Regulation (EC) 1907/2006 (REACH) as a substance which is an endocrine disruptor to the environment when it contains ≥ 0.1% w/w of 4-nonylphenol, branched and linear (4-NP) (see ED decision ED/71/2019 of 16 July 2019).
tris(nonylphenyl) phosphite |  | 247-759-6 | 26523-78-4 | Endocrine disrupting properties (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/a661cbe6-c805-b9c7-1287-53dee5cd1df5#https://echa.europa.eu/documents/10162/f162b35e-2aac-6b5a-a3de-3290bedd8805 | https://echa.europa.eu/documents/10162/8da7df6b-e62f-47f7-d9d1-154f8db267f8#https://echa.europa.eu/documents/10162/0a66176d-3f11-34ed-74e1-6062ae73f264 | https://echa.europa.eu/documents/10162/ceb2fc2c-78f2-627f-29ae-b37c0e91bc69#https://echa.europa.eu/documents/10162/f186ad0a-3831-025c-2628-3c0948e174d5 | https://echa.europa.eu/documents/10162/70879c79-2312-31bd-09e5-558df06f19e2#https://echa.europa.eu/documents/10162/a524f414-c041-2388-f47a-302a46efdc79 | In addition to its intrinsic properties (see ECHA decision D(2024)7663-DC of 21 January 2025), tris(4-nonylphenyl, branched and linear) phosphite (TNPP) also meets the criteria of Article 57(f) of Regulation (EC) 1907/2006 (REACH) as a substance which is an endocrine disruptor to the environment when it contains ≥ 0.1% w/w of 4-nonylphenol, branched and linear (4-NP) (see ED decision ED/71/2019 of 16 July 2019).
Phenol, 4-nonyl-, phosphite (3:1) |  | - | 3050-88-2 | Endocrine disrupting properties (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/a661cbe6-c805-b9c7-1287-53dee5cd1df5#https://echa.europa.eu/documents/10162/f162b35e-2aac-6b5a-a3de-3290bedd8805 | https://echa.europa.eu/documents/10162/8da7df6b-e62f-47f7-d9d1-154f8db267f8#https://echa.europa.eu/documents/10162/0a66176d-3f11-34ed-74e1-6062ae73f264 | https://echa.europa.eu/documents/10162/ceb2fc2c-78f2-627f-29ae-b37c0e91bc69#https://echa.europa.eu/documents/10162/f186ad0a-3831-025c-2628-3c0948e174d5 | https://echa.europa.eu/documents/10162/70879c79-2312-31bd-09e5-558df06f19e2#https://echa.europa.eu/documents/10162/a524f414-c041-2388-f47a-302a46efdc79 | In addition to its intrinsic properties (see ECHA decision D(2024)7663-DC of 21 January 2025), tris(4-nonylphenyl, branched and linear) phosphite (TNPP) also meets the criteria of Article 57(f) of Regulation (EC) 1907/2006 (REACH) as a substance which is an endocrine disruptor to the environment when it contains ≥ 0.1% w/w of 4-nonylphenol, branched and linear (4-NP) (see ED decision ED/71/2019 of 16 July 2019).
Phenol, p-isononyl-, phosphite (3:1) |  | - | 31631-13-7 | Endocrine disrupting properties (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/a661cbe6-c805-b9c7-1287-53dee5cd1df5#https://echa.europa.eu/documents/10162/f162b35e-2aac-6b5a-a3de-3290bedd8805 | https://echa.europa.eu/documents/10162/8da7df6b-e62f-47f7-d9d1-154f8db267f8#https://echa.europa.eu/documents/10162/0a66176d-3f11-34ed-74e1-6062ae73f264 | https://echa.europa.eu/documents/10162/ceb2fc2c-78f2-627f-29ae-b37c0e91bc69#https://echa.europa.eu/documents/10162/f186ad0a-3831-025c-2628-3c0948e174d5 | https://echa.europa.eu/documents/10162/70879c79-2312-31bd-09e5-558df06f19e2#https://echa.europa.eu/documents/10162/a524f414-c041-2388-f47a-302a46efdc79 | In addition to its intrinsic properties (see ECHA decision D(2024)7663-DC of 21 January 2025), tris(4-nonylphenyl, branched and linear) phosphite (TNPP) also meets the criteria of Article 57(f) of Regulation (EC) 1907/2006 (REACH) as a substance which is an endocrine disruptor to the environment when it contains ≥ 0.1% w/w of 4-nonylphenol, branched and linear (4-NP) (see ED decision ED/71/2019 of 16 July 2019).
Phenol, p-sec-nonyl-, phosphite |  | - | 106599-06-8 | Endocrine disrupting properties (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/a661cbe6-c805-b9c7-1287-53dee5cd1df5#https://echa.europa.eu/documents/10162/f162b35e-2aac-6b5a-a3de-3290bedd8805 | https://echa.europa.eu/documents/10162/8da7df6b-e62f-47f7-d9d1-154f8db267f8#https://echa.europa.eu/documents/10162/0a66176d-3f11-34ed-74e1-6062ae73f264 | https://echa.europa.eu/documents/10162/ceb2fc2c-78f2-627f-29ae-b37c0e91bc69#https://echa.europa.eu/documents/10162/f186ad0a-3831-025c-2628-3c0948e174d5 | https://echa.europa.eu/documents/10162/70879c79-2312-31bd-09e5-558df06f19e2#https://echa.europa.eu/documents/10162/a524f414-c041-2388-f47a-302a46efdc79 | In addition to its intrinsic properties (see ECHA decision D(2024)7663-DC of 21 January 2025), tris(4-nonylphenyl, branched and linear) phosphite (TNPP) also meets the criteria of Article 57(f) of Regulation (EC) 1907/2006 (REACH) as a substance which is an endocrine disruptor to the environment when it contains ≥ 0.1% w/w of 4-nonylphenol, branched and linear (4-NP) (see ED decision ED/71/2019 of 16 July 2019).
tris(4-nonylphenyl, branched) phosphite |  | - | - | Endocrine disrupting properties (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/a661cbe6-c805-b9c7-1287-53dee5cd1df5#https://echa.europa.eu/documents/10162/f162b35e-2aac-6b5a-a3de-3290bedd8805 | https://echa.europa.eu/documents/10162/8da7df6b-e62f-47f7-d9d1-154f8db267f8#https://echa.europa.eu/documents/10162/0a66176d-3f11-34ed-74e1-6062ae73f264 | https://echa.europa.eu/documents/10162/ceb2fc2c-78f2-627f-29ae-b37c0e91bc69#https://echa.europa.eu/documents/10162/f186ad0a-3831-025c-2628-3c0948e174d5 | https://echa.europa.eu/documents/10162/70879c79-2312-31bd-09e5-558df06f19e2#https://echa.europa.eu/documents/10162/a524f414-c041-2388-f47a-302a46efdc79 | In addition to its intrinsic properties (see ECHA decision D(2024)7663-DC of 21 January 2025), tris(4-nonylphenyl, branched and linear) phosphite (TNPP) also meets the criteria of Article 57(f) of Regulation (EC) 1907/2006 (REACH) as a substance which is an endocrine disruptor to the environment when it contains ≥ 0.1% w/w of 4-nonylphenol, branched and linear (4-NP) (see ED decision ED/71/2019 of 16 July 2019).
4-tert-butylphenol |  | 202-679-0 | 98-54-4 | Endocrine disrupting properties (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/ca2542ab-a35b-1c00-5490-708d59332c38#https://echa.europa.eu/documents/10162/99a7e1c6-2aa4-0c35-acc3-35ec26bb1ee8 | https://echa.europa.eu/documents/10162/fedde563-03f7-51f0-85fd-1fceb0eaa18f | https://echa.europa.eu/documents/10162/41de4b52-51f3-84e9-3f9c-fa6f868139ad | https://echa.europa.eu/documents/10162/2c66d98c-d2a2-e146-265e-b8a1615e5c17
2-methoxyethyl acetate |  | 203-772-9 | 110-49-6 | Toxic for reproduction (Article 57c) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/b7bc601c-ad77-4e90-0cdf-2701a59478fb | https://echa.europa.eu/documents/10162/eea50979-c40b-f00f-b767-891e93941ee1 | https://echa.europa.eu/documents/10162/874297f1-6d65-b808-a62e-5d23cc5f1e5c | https://echa.europa.eu/documents/10162/7357fc26-f169-a309-95b1-7693dada001b
2,3,3,3-tetrafluoro-2-(heptafluoropropoxy)propionic acid, its salts and its acyl halides | covering any of their individual isomers and combinations thereof | - | - | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/fc76aefc-fc86-a5fc-b5c4-e358467ca832 | https://echa.europa.eu/documents/10162/c07bc76c-4d62-7799-d0d4-17e43508aeb0 | https://echa.europa.eu/documents/10162/53fa6a5b-e95f-3128-ea9d-fa27f43b18bc | https://echa.europa.eu/documents/10162/5d5500fc-b935-f95c-02c7-d7722d1b795d | The combined intrinsic properties justifying the inclusion for the Member State Committee as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: Persistence, mobility, potential for long-range transport, observed adverse effects (at least the following probable effects for human health: effects on the liver, the kidney, and the haematological and immune systems and effects on development; at least the following probable effects for the environment: population relevant effects on birds and mammals); as well as low adsorption potential and high water solubility rendering the substance fully bioavailable for uptake via (drinking) water. Together, these elements lead to a very high potential for irreversible effects.
2,3,3,3-tetrafluoro-2-(heptafluoropropoxy)propionyl fluoride |  | 218-173-8 | 2062-98-8 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/fc76aefc-fc86-a5fc-b5c4-e358467ca832 | https://echa.europa.eu/documents/10162/c07bc76c-4d62-7799-d0d4-17e43508aeb0 | https://echa.europa.eu/documents/10162/53fa6a5b-e95f-3128-ea9d-fa27f43b18bc | https://echa.europa.eu/documents/10162/5d5500fc-b935-f95c-02c7-d7722d1b795d | The combined intrinsic properties justifying the inclusion for the Member State Committee as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: Persistence, mobility, potential for long-range transport, observed adverse effects (at least the following probable effects for human health: effects on the liver, the kidney, and the haematological and immune systems and effects on development; at least the following probable effects for the environment: population relevant effects on birds and mammals); as well as low adsorption potential and high water solubility rendering the substance fully bioavailable for uptake via (drinking) water. Together, these elements lead to a very high potential for irreversible effects.
2,3,3,3-tetrafluoro-2-(heptafluoropropoxy)propionic acid |  | 236-236-8 | 13252-13-6 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/fc76aefc-fc86-a5fc-b5c4-e358467ca832 | https://echa.europa.eu/documents/10162/c07bc76c-4d62-7799-d0d4-17e43508aeb0 | https://echa.europa.eu/documents/10162/53fa6a5b-e95f-3128-ea9d-fa27f43b18bc | https://echa.europa.eu/documents/10162/5d5500fc-b935-f95c-02c7-d7722d1b795d | The combined intrinsic properties justifying the inclusion for the Member State Committee as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: Persistence, mobility, potential for long-range transport, observed adverse effects (at least the following probable effects for human health: effects on the liver, the kidney, and the haematological and immune systems and effects on development; at least the following probable effects for the environment: population relevant effects on birds and mammals); as well as low adsorption potential and high water solubility rendering the substance fully bioavailable for uptake via (drinking) water. Together, these elements lead to a very high potential for irreversible effects.
ammonium 2,3,3,3-tetrafluoro-2-(heptafluoropropoxy)propanoate |  | - | 62037-80-3 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/fc76aefc-fc86-a5fc-b5c4-e358467ca832 | https://echa.europa.eu/documents/10162/c07bc76c-4d62-7799-d0d4-17e43508aeb0 | https://echa.europa.eu/documents/10162/53fa6a5b-e95f-3128-ea9d-fa27f43b18bc | https://echa.europa.eu/documents/10162/5d5500fc-b935-f95c-02c7-d7722d1b795d | The combined intrinsic properties justifying the inclusion for the Member State Committee as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: Persistence, mobility, potential for long-range transport, observed adverse effects (at least the following probable effects for human health: effects on the liver, the kidney, and the haematological and immune systems and effects on development; at least the following probable effects for the environment: population relevant effects on birds and mammals); as well as low adsorption potential and high water solubility rendering the substance fully bioavailable for uptake via (drinking) water. Together, these elements lead to a very high potential for irreversible effects.
potassium 2,3,3,3-tetrafluoro-2-(heptafluoropropoxy)propionate |  | 266-578-3 | 67118-55-2 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/fc76aefc-fc86-a5fc-b5c4-e358467ca832 | https://echa.europa.eu/documents/10162/c07bc76c-4d62-7799-d0d4-17e43508aeb0 | https://echa.europa.eu/documents/10162/53fa6a5b-e95f-3128-ea9d-fa27f43b18bc | https://echa.europa.eu/documents/10162/5d5500fc-b935-f95c-02c7-d7722d1b795d | The combined intrinsic properties justifying the inclusion for the Member State Committee as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: Persistence, mobility, potential for long-range transport, observed adverse effects (at least the following probable effects for human health: effects on the liver, the kidney, and the haematological and immune systems and effects on development; at least the following probable effects for the environment: population relevant effects on birds and mammals); as well as low adsorption potential and high water solubility rendering the substance fully bioavailable for uptake via (drinking) water. Together, these elements lead to a very high potential for irreversible effects.
Propanoic acid, 2,3,3,3-tetrafluoro-2-(heptafluoropropoxy)-, (-)- |  | - | 75579-40-7 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/fc76aefc-fc86-a5fc-b5c4-e358467ca832 | https://echa.europa.eu/documents/10162/c07bc76c-4d62-7799-d0d4-17e43508aeb0 | https://echa.europa.eu/documents/10162/53fa6a5b-e95f-3128-ea9d-fa27f43b18bc | https://echa.europa.eu/documents/10162/5d5500fc-b935-f95c-02c7-d7722d1b795d | The combined intrinsic properties justifying the inclusion for the Member State Committee as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: Persistence, mobility, potential for long-range transport, observed adverse effects (at least the following probable effects for human health: effects on the liver, the kidney, and the haematological and immune systems and effects on development; at least the following probable effects for the environment: population relevant effects on birds and mammals); as well as low adsorption potential and high water solubility rendering the substance fully bioavailable for uptake via (drinking) water. Together, these elements lead to a very high potential for irreversible effects.
Propanoic acid, 2,3,3,3-tetrafluoro-2-(heptafluoropropoxy)-, (+)- |  | - | 75579-39-4 | Equivalent level of concern having probable serious effects to human health (Article 57(f) - human health)#Equivalent level of concern having probable serious effects to the environment (Article 57(f) - environment) | 16-Jul-2019 | https://echa.europa.eu/documents/10162/fc76aefc-fc86-a5fc-b5c4-e358467ca832 | https://echa.europa.eu/documents/10162/c07bc76c-4d62-7799-d0d4-17e43508aeb0 | https://echa.europa.eu/documents/10162/53fa6a5b-e95f-3128-ea9d-fa27f43b18bc | https://echa.europa.eu/documents/10162/5d5500fc-b935-f95c-02c7-d7722d1b795d | The combined intrinsic properties justifying the inclusion for the Member State Committee as a substance for which there is scientific evidence of probable serious effects to human health and the environment which give rise to an equivalent level of concern are the following: Persistence, mobility, potential for long-range transport, observed adverse effects (at least the following probable effects for human health: effects on the liver, the kidney, and the haematological and immune systems and effects on development; at least the following probable effects for the environment: population relevant effects on birds and mammals); as well as low adsorption potential and high water solubility rendering the substance fully bioavailable for uptake via (drinking) water. Together, these elements lead to a very high potential for irreversible effects.
Pyrene |  | 204-927-3 | 129-00-0 | PBT (Article 57d)#vPvB (Article 57e) | 15-Jan-2019 | https://echa.europa.eu/documents/10162/d6b1b306-6f05-9765-204a-689334527ef9 | https://echa.europa.eu/documents/10162/bf9b02a4-6e4b-b1bf-2c58-6d61f5998817 | https://echa.europa.eu/documents/10162/47121daf-04a7-6d4a-b0b6-595794d3e66c | https://echa.europa.eu/documents/10162/e824171c-1263-b1a2-4438-c5add37de2fa
Phenanthrene |  | 201-581-5 | 85-01-8 | vPvB (Article 57e) | 15-Jan-2019 | https://echa.europa.eu/documents/10162/91535066-71da-d9b4-71cc-8d4003b7f983 | https://echa.europa.eu/documents/10162/efc23200-6f59-12c6-4b18-533a80aa7b4f | https://echa.europa.eu/documents/10162/c62a1c4a-ceec-3a00-52a4-0a298f7d8bb3 | https://echa.europa.eu/documents/10162/12db6a91-9ac9-c4f6-4046-d6896ad9b3b2
Fluoranthene |  | 205-912-4 | 206-44-0 | PBT (Article 57d)#vPvB (Article 57e) | 15-Jan-2019 | https://echa.europa.eu/documents/10162/21272691-d1a0-3334-fde9-a0782d9354a2 | https://echa.europa.eu/documents/10162/375e39c4-c59b-bfbd-d156-a1bf2a959108 | https://echa.europa.eu/documents/10162/0d1ee6d4-1a47-0737-35c7-3503f0fca417 | https://echa.europa.eu/documents/10162/b525e382-e27e-c609-2abf-ddd77d07ec44
Benzo[k]fluoranthene |  | 205-916-6 | 207-08-9 | Carcinogenic (Article 57a)#PBT (Article 57d)#vPvB (Article 57e) | 15-Jan-2019 | https://echa.europa.eu/documents/10162/afa7d75c-b83f-fef0-ae10-41e70a1e57f7 | https://echa.europa.eu/documents/10162/ae8cba56-9884-9e35-34c7-0369a4703600 | https://echa.europa.eu/documents/10162/06cc1281-efd9-9845-0215-e6b0c94c94db | https://echa.europa.eu/documents/10162/f72dc11d-5ff2-c4d8-c19c-40ec8dadddac
2,2-bis(4'-hydroxyphenyl)-4-methylpentane | AP-5 | 401-720-1 | 6807-17-6 | Toxic for reproduction (Article 57c) | 15-Jan-2019 | https://echa.europa.eu/documents/10162/562afdb1-00aa-32ab-1022-721917185b12 | https://echa.europa.eu/documents/10162/8ec5e70d-db99-d21c-c54b-9de01c41f435 | https://echa.europa.eu/documents/10162/76cf3fb1-07eb-ff40-f41c-251f8c657b33 | https://echa.europa.eu/documents/10162/26b55c4f-92c0-e4b6-fa1a-973bd625ae0f
1,7,7-trimethyl-3-(phenylmethylene)bicyclo[2.2.1]heptan-2-one | 3-benzylidene camphor; 3-BC | 239-139-9 | 15087-24-8 | Endocrine disrupting properties (Article 57(f) - environment) | 15-Jan-2019 | https://echa.europa.eu/documents/10162/db2f223c-20a1-5318-d6b1-7a147d7f08c3#https://echa.europa.eu/documents/10162/56d708da-69e5-a7d1-e970-780352b8ed6a | https://echa.europa.eu/documents/10162/22f73d01-dfc1-15d5-a8fe-a29b460fa108 | https://echa.europa.eu/documents/10162/84e7e600-0f73-6a4a-b246-19482576aa20#https://echa.europa.eu/documents/10162/2e1c96ea-2918-90d2-9c23-e07effac4ec9#https://echa.europa.eu/documents/10162/0c80cbbd-998f-b415-f9be-9e3ced41f333 | https://echa.europa.eu/documents/10162/ca1d3937-2301-7a43-4b9e-1008ca84d3f6
Terphenyl, hydrogenated |  | 262-967-7 | 61788-32-7 | vPvB (Article 57e) | 27-Jun-2018 | https://echa.europa.eu/documents/10162/fc78f7e5-8f8a-de84-1141-499c5e021d31 | https://echa.europa.eu/documents/10162/b30abea4-5ba0-e195-f595-a617687ad1c1 | https://echa.europa.eu/documents/10162/fb7a9167-7d65-dd4c-2aa2-b8182d4a8e37 | https://echa.europa.eu/documents/10162/421a040e-320b-5d39-8582-118bf23054f9
Octamethylcyclotetrasiloxane | D4 | 209-136-7 | 556-67-2 | PBT (Article 57d)#vPvB (Article 57e) | 27-Jun-2018 | https://echa.europa.eu/documents/10162/2be7bcbf-f797-c28c-2c67-939664155c7c | https://echa.europa.eu/documents/10162/1c5dc0f3-cd5c-7b71-044a-9167e3b03d99 | https://echa.europa.eu/documents/10162/115f70a9-a387-1525-d49f-b715e84996e4 | https://echa.europa.eu/documents/10162/e654e7d3-031c-0f5f-eae5-673c1dd03432
Lead |  | 231-100-4 | 7439-92-1 | Toxic for reproduction (Article 57c) | 27-Jun-2018 | https://echa.europa.eu/documents/10162/61ac8d81-6ea2-6ad0-ffef-95037c9182ce | https://echa.europa.eu/documents/10162/79eec48e-c0b6-0d58-5218-ddb32ce0d22d | https://echa.europa.eu/documents/10162/07a87920-1b8f-b0d9-b6a7-1c0b1c16c8c4 | https://echa.europa.eu/documents/10162/aec05f99-6be1-d615-732f-ec4f68ab66e2
Ethylenediamine | EDA | 203-468-6 | 107-15-3 | Respiratory sensitising properties (Article 57(f) - human health) | 27-Jun-2018 | https://echa.europa.eu/documents/10162/2ed8a7b4-8e9f-2e3e-407d-819dbdaa1623 | https://echa.europa.eu/documents/10162/0fa2b79d-e520-b857-6572-61c49095b5a3 | https://echa.europa.eu/documents/10162/f0a61fa4-64d1-2d19-35ed-d98134aec10d | https://echa.europa.eu/documents/10162/060c2b77-ce67-26fd-5117-43c52d6d7188
Dodecamethylcyclohexasiloxane | D6 | 208-762-8 | 540-97-6 | PBT (Article 57d)#vPvB (Article 57e) | 27-Jun-2018 | https://echa.europa.eu/documents/10162/b11de9bc-6e60-01f9-cdeb-20fa3b7e864e | https://echa.europa.eu/documents/10162/9061869f-7e55-aaac-9ea5-761b9050a952 | https://echa.europa.eu/documents/10162/a9682f4b-fc3e-cd99-3db9-b0f9f383c3c5 | https://echa.europa.eu/documents/10162/45532e5e-8bbe-a3ee-bd1c-6c1640984c27 | Dodecamethylcyclohexasiloxane (D6) meets the criteria of Article 57 (d) of Regulation (EC) 1907/2006 (REACH) as a substance which is persistent, bioaccumulative and toxic when it contains ≥ 0.1 % w/w octamethylcyclotetrasiloxane (D4) (EC No. 209-136-7). In addition to its intrinsic properties, it also meets the criteria of Article 57 (e) of Regulation (EC) 1907/2006 (REACH) as a substance which is very persistent and very bioaccumulative (vPvB) when it contains ≥ 0.1 % w/w decamethylcyclopentasiloxane (D5) (EC No. 208-764-9) or ≥ 0.1% w/w octamethylcyclotetrasiloxane (D4) (EC No. 209-136-7).
Disodium octaborate |  | 234-541-0 | 12008-41-2 | Toxic for reproduction (Article 57c) | 27-Jun-2018 | https://echa.europa.eu/documents/10162/74a5fdff-6a00-d3b4-4056-a5b24ee9ba6e | https://echa.europa.eu/documents/10162/62353089-57f4-c7c7-64f9-291364883323 | https://echa.europa.eu/documents/10162/24e48bb3-3a68-1dfa-ef39-6e9f4ef92720 | https://echa.europa.eu/documents/10162/da7779d3-8314-7357-b82b-e6044a5dd6be
Dicyclohexyl phthalate | DCHP | 201-545-9 | 84-61-7 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - human health) | 27-Jun-2018 | https://echa.europa.eu/documents/10162/8c434af5-cfbe-c87e-aa51-65893e385d1f#https://echa.europa.eu/documents/10162/2fe151f3-d425-cff3-fe31-d13c11afcc7a | https://echa.europa.eu/documents/10162/1be86650-c2d6-326e-1991-735cbf66ecbf | https://echa.europa.eu/documents/10162/919d3291-e835-609a-0800-03f64cda86b5#https://echa.europa.eu/documents/10162/b2995010-92b6-183f-2a3c-e5b823b537c9#https://echa.europa.eu/documents/10162/73c5f83c-3fe6-e1a0-b756-9dfcebf1eb85 | https://echa.europa.eu/documents/10162/a0f00120-5e32-7c74-0b7f-9c92f84cb38c
Decamethylcyclopentasiloxane | D5 | 208-764-9 | 541-02-6 | PBT (Article 57d)#vPvB (Article 57e) | 27-Jun-2018 | https://echa.europa.eu/documents/10162/1b8ab766-b3ff-840f-a0bd-f1ade391745d | https://echa.europa.eu/documents/10162/32855182-9c49-d41f-150a-b6a6aa012fff | https://echa.europa.eu/documents/10162/ddd97c9a-fe79-6f50-ff1a-c1d4bf305aab | https://echa.europa.eu/documents/10162/4c47a1f3-3513-d688-6f46-32a4023415a5 | Decamethylcyclopentasiloxane (D5) meets the criteria of Article 57 (d) of Regulation (EC) 1907/2006 (REACH) as a substance which is persistent, bioaccumulative and toxic when it contains ≥ 0.1 % w/w octamethylcyclotetrasiloxane (D4) (EC No: 209-136-7).
Benzo[ghi]perylene |  | 205-883-8 | 191-24-2 | PBT (Article 57d)#vPvB (Article 57e) | 27-Jun-2018 | https://echa.europa.eu/documents/10162/ef81b8a3-7ec8-1380-d2ff-db1ceae26073 | https://echa.europa.eu/documents/10162/16752063-ed16-d532-0418-8613beeceab0 | https://echa.europa.eu/documents/10162/b45e78cd-5d70-0dd9-e895-e9fa89ab43d0 | https://echa.europa.eu/documents/10162/2d44cd37-2ac3-7cfc-a643-670a075df973
Benzene-1,2,4-tricarboxylic acid 1,2 anhydride | trimellitic anhydride; TMA | 209-008-0 | 552-30-7 | Respiratory sensitising properties (Article 57(f) - human health) | 27-Jun-2018 | https://echa.europa.eu/documents/10162/996c1b49-d6f7-5899-c94a-43a04d98ef94#https://echa.europa.eu/documents/10162/a6547852-e697-84ec-512c-5ba264ecf09e | https://echa.europa.eu/documents/10162/358de08d-e71f-b578-e833-71fec1df75d2 | https://echa.europa.eu/documents/10162/c2b11ef6-5ca8-ba78-da4b-38baa59ac416#https://echa.europa.eu/documents/10162/e3f5cdaa-1633-1c3e-a50f-9fae7da39104#https://echa.europa.eu/documents/10162/4a628698-2072-0f8c-336e-4a20965aa7a5 | https://echa.europa.eu/documents/10162/63cd8fa3-4a72-4ddb-b4b0-100ef95048e0
Reaction products of 1,3,4-thiadiazolidine-2,5-dithione, formaldehyde and 4-heptylphenol, branched and linear (RP-HP) | with ≥0.1% w/w 4-heptylphenol, branched and linear (4-HPbl) | - | - | Endocrine disrupting properties (Article 57(f) - environment) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/339e0894-361d-ab6c-615f-7e201376d128 | https://echa.europa.eu/documents/10162/c8ea7b6a-06bc-eeaf-ae4d-3d980a6078bc | https://echa.europa.eu/documents/10162/2064f792-8443-873c-27de-eae18f593b59#https://echa.europa.eu/documents/10162/b33d78f4-946c-1d4d-e923-1ade7869c464 | https://echa.europa.eu/documents/10162/662821c8-16be-e02e-e52f-6e214b6041d5
Formaldehyde, reaction products with branched and linear heptylphenol, carbon disulfide and hydrazine |  | 300-298-5 | 93925-00-9 | Endocrine disrupting properties (Article 57(f) - environment) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/339e0894-361d-ab6c-615f-7e201376d128 | https://echa.europa.eu/documents/10162/c8ea7b6a-06bc-eeaf-ae4d-3d980a6078bc | https://echa.europa.eu/documents/10162/2064f792-8443-873c-27de-eae18f593b59#https://echa.europa.eu/documents/10162/b33d78f4-946c-1d4d-e923-1ade7869c464 | https://echa.europa.eu/documents/10162/662821c8-16be-e02e-e52f-6e214b6041d5
Reaction product of 1,3,4-thiadiazolidine-2,5-dithione, formaldehyde and phenol, heptyl derivs. |  | - | 1471311-26-8 | Endocrine disrupting properties (Article 57(f) - environment) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/339e0894-361d-ab6c-615f-7e201376d128 | https://echa.europa.eu/documents/10162/c8ea7b6a-06bc-eeaf-ae4d-3d980a6078bc | https://echa.europa.eu/documents/10162/2064f792-8443-873c-27de-eae18f593b59#https://echa.europa.eu/documents/10162/b33d78f4-946c-1d4d-e923-1ade7869c464 | https://echa.europa.eu/documents/10162/662821c8-16be-e02e-e52f-6e214b6041d5
Chrysene |  | 205-923-4 | 218-01-9 | Carcinogenic (Article 57a)#PBT (Article 57d)#vPvB (Article 57e) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/a7438bc1-a7fd-caef-d74a-ca33fc0f3610 | https://echa.europa.eu/documents/10162/2b898da9-812d-8a07-e059-e81d2c2883cb | https://echa.europa.eu/documents/10162/92791ee1-40a6-f4dd-a2fe-927f6f1cb478 | https://echa.europa.eu/documents/10162/e3f6792d-7e59-654f-b90e-2149525906cd
Cadmium nitrate |  | 233-710-6 | 10325-94-7 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/d1e3a2d5-13a6-c5da-62cf-2d581d94b87d | https://echa.europa.eu/documents/10162/8b38c499-dad0-ee30-11ab-198b55b2e69d | https://echa.europa.eu/documents/10162/ac791555-9e63-3208-e176-cd0fdb4fbdf0 | https://echa.europa.eu/documents/10162/264dda83-c2e5-7c7b-0c1f-57a7cc32afe4
Cadmium hydroxide |  | 244-168-5 | 21041-95-2 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/deb481e6-b2f3-9792-7d2b-260a1f383dc1 | https://echa.europa.eu/documents/10162/d6ad39ad-bbf7-562c-2042-a823ed29cc1e | https://echa.europa.eu/documents/10162/38688481-353e-bac2-e83e-68f4c772aac7 | https://echa.europa.eu/documents/10162/0252778f-258a-2b3d-7765-0291f9fe3c45
Cadmium carbonate |  | 208-168-9 | 513-78-0 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/53cefde3-40e9-8d3d-b52d-87a9d828871d | https://echa.europa.eu/documents/10162/31f9a47b-191c-526e-0446-f10162fd312c | https://echa.europa.eu/documents/10162/acc7ace6-0118-6d74-4446-38f5bbe0b6b6 | https://echa.europa.eu/documents/10162/c26799d3-6187-0494-1cc4-af9c9645b1b3
Benz[a]anthracene |  | 200-280-6 | 56-55-3 | Carcinogenic (Article 57a)#PBT (Article 57d)#vPvB (Article 57e) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/6adbea83-2790-92a4-06cd-ce39c4bf3211 | https://echa.europa.eu/documents/10162/7d8ac604-d716-4a97-8682-d3d5e5c4158f | https://echa.europa.eu/documents/10162/75ad552f-1e59-2599-e379-55ecec998d3f | https://echa.europa.eu/documents/10162/efef93e7-63f5-75b8-9f6e-d260b0c595ab
1,6,7,8,9,14,15,16,17,17,18,18-Dodecachloropentacyclo[12.2.1.16,9.02,13.05,10]octadeca-7,15-diene (“Dechlorane Plus”™) | covering any of its individual anti- and syn-isomers or any combination thereof | - | - | vPvB (Article 57e) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/48495b62-d4d3-e6fc-bf86-f42fd4f2e6dc | https://echa.europa.eu/documents/10162/78fb98fd-a8e7-31a3-f910-f2bad587d8a6 | https://echa.europa.eu/documents/10162/9aa77dde-f0fc-4422-2c00-55b620e57552#https://echa.europa.eu/documents/10162/97b3c3bf-f38a-f3e2-6b53-45654bcc02dc | https://echa.europa.eu/documents/10162/72b812a4-77eb-be74-c5bb-694b20017d58
(1S,2S,5R,6R,9S,10S,13R,14R)-1,6,7,8,9,14,15,16,17,17,18,18-Dodecachloropentacyclo[12.2.1.1⁶,⁹.0²,¹³.0⁵,¹⁰]octadeca-7,15-diene |  | - | 135821-03-3 | vPvB (Article 57e) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/48495b62-d4d3-e6fc-bf86-f42fd4f2e6dc | https://echa.europa.eu/documents/10162/78fb98fd-a8e7-31a3-f910-f2bad587d8a6 | https://echa.europa.eu/documents/10162/9aa77dde-f0fc-4422-2c00-55b620e57552#https://echa.europa.eu/documents/10162/97b3c3bf-f38a-f3e2-6b53-45654bcc02dc | https://echa.europa.eu/documents/10162/72b812a4-77eb-be74-c5bb-694b20017d58
1,6,7,8,9,14,15,16,17,17,18,18-dodecachloropentacyclo[12.2.1.16,9.02,13.05,10]octadeca-7,15-diene |  | 236-948-9 | 13560-89-9 | vPvB (Article 57e) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/48495b62-d4d3-e6fc-bf86-f42fd4f2e6dc | https://echa.europa.eu/documents/10162/78fb98fd-a8e7-31a3-f910-f2bad587d8a6 | https://echa.europa.eu/documents/10162/9aa77dde-f0fc-4422-2c00-55b620e57552#https://echa.europa.eu/documents/10162/97b3c3bf-f38a-f3e2-6b53-45654bcc02dc | https://echa.europa.eu/documents/10162/72b812a4-77eb-be74-c5bb-694b20017d58
(1S,2S,5S,6S,9R,10R,13R,14R)-1,6,7,8,9,14,15,16,17,17,18,18-Dodecachloropentacyclo[12.2.1.1⁶,⁹.0²,¹³.0⁵,¹⁰]octadeca-7,15-diene |  | - | 135821-74-8 | vPvB (Article 57e) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/48495b62-d4d3-e6fc-bf86-f42fd4f2e6dc | https://echa.europa.eu/documents/10162/78fb98fd-a8e7-31a3-f910-f2bad587d8a6 | https://echa.europa.eu/documents/10162/9aa77dde-f0fc-4422-2c00-55b620e57552#https://echa.europa.eu/documents/10162/97b3c3bf-f38a-f3e2-6b53-45654bcc02dc | https://echa.europa.eu/documents/10162/72b812a4-77eb-be74-c5bb-694b20017d58
rel-(1R,4S,4aS,6aR,7R,10S,10aS,12aR)-1,2,3,4,7,8,9,10,13,13,14,14-dodecachloro-1,4,4a,5,6,6a,7,10,10a,11,12,12a-dodecahydro-1,4:7,10-dimethanodibenzo[a,e]cyclooctene |  | - | - | vPvB (Article 57e) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/48495b62-d4d3-e6fc-bf86-f42fd4f2e6dc | https://echa.europa.eu/documents/10162/78fb98fd-a8e7-31a3-f910-f2bad587d8a6 | https://echa.europa.eu/documents/10162/9aa77dde-f0fc-4422-2c00-55b620e57552#https://echa.europa.eu/documents/10162/97b3c3bf-f38a-f3e2-6b53-45654bcc02dc | https://echa.europa.eu/documents/10162/72b812a4-77eb-be74-c5bb-694b20017d58
rel-(1R,4S,4aS,6aS,7S,10R,10aR,12aR)-1,2,3,4,7,8,9,10,13,13,14,14-dodecachloro-1,4,4a,5,6,6a,7,10,10a,11,12,12a-dodecahydro-1,4:7,10-dimethanodibenzo[a,e]cyclooctene |  | - | - | vPvB (Article 57e) | 15-Jan-2018 | https://echa.europa.eu/documents/10162/48495b62-d4d3-e6fc-bf86-f42fd4f2e6dc | https://echa.europa.eu/documents/10162/78fb98fd-a8e7-31a3-f910-f2bad587d8a6 | https://echa.europa.eu/documents/10162/9aa77dde-f0fc-4422-2c00-55b620e57552#https://echa.europa.eu/documents/10162/97b3c3bf-f38a-f3e2-6b53-45654bcc02dc | https://echa.europa.eu/documents/10162/72b812a4-77eb-be74-c5bb-694b20017d58
Perfluorohexane-1-sulphonic acid and its salts | PFHxS | - | - | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
potassium perfluorohexane-1-sulphonate |  | 223-393-2 | 3871-99-6 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
ammonium perfluorohexane-1-sulphonate |  | 269-511-6 | 68259-08-5 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
tridecafluorohexanesulphonic acid, compound with 2,2'-iminodiethanol (1:1) |  | 274-462-9 | 70225-16-0 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
perfluorohexane-1-sulphonic acid |  | 206-587-1 | 355-46-4 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Sulfonium, [4-[(2-methyl-1-oxo-2-propenyl)oxy]phenyl]diphenyl-, salt with 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonic acid (1:1), polymer with 2-ethyltricyclo[3.3.1.13,7]dec-2-yl 2-methyl-2-propenoate, 3-hydroxytricyclo[3.3.1.13,7]dec-1-yl 2-methyl-2-propenoate and tetrahydro-2-oxo-3-furanyl 2-methyl-2-propenoate |  | - | 911027-69-5 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
1-Hexanesulfonic acid, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-, cesium salt (1:1) |  | - | 92011-17-1 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Dibenzo[k,n][1,4,7,10,13]tetraoxathiacyclopentadecinium, 19-[4-(1,1-dimethylethyl)phenyl]-6,7,9,10,12,13-hexahydro-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 928049-42-7 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
1-Hexanesulfonic acid, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-, lithium salt (1:1) |  | - | 55120-77-9 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Quinolinium, 1-(carboxymethyl)-4-[2-[4-[4-(2,2-diphenylethenyl)phenyl]-1,2,3,3a,4,8b-hexahydrocyclopent[b]indol-7-yl]ethenyl]-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 1462414-59-0 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Iodonium, diphenyl-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 153443-35-7 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Methanaminium, N,N,N-trimethyl-, salt with 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonic acid (1:1) |  | - | 189274-31-5 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
1-Hexanesulfonic acid, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-, compd.with 2-methyl-2-propanamine (1:1) |  | - | 202189-84-2 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Iodonium, bis[4-(1,1-dimethylethyl)phenyl]-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 213740-81-9 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
1-Hexanesulfonic acid, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-, gallium salt (9CI) |  | - | 341035-71-0 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Sulfonium, bis(4-methylphenyl)phenyl-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 341548-85-4 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
1-Hexanesulfonic acid, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-, scandium(3+) salt (3:1) |  | - | 350836-93-0 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
1-Hexanesulfonic acid, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-, zinc salt |  | - | 70136-72-0 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
1-Hexanesulfonic acid, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-, neodymium(3+) salt (3:1) |  | - | 41184-65-0 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
1-Hexanesulfonic acid, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-, yttrium(3+) salt (3:1) |  | - | 41242-12-0 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Sulfonium, (thiodi-4,1-phenylene)bis[diphenyl-, salt with 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonic acid (1:2) |  | - | 421555-73-9 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Iodonium, bis[4-(1,1-dimethylpropyl)phenyl]-, salt with 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonic |  | - | 421555-74-0 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Sulfonium, tris[4-(1,1-dimethylethyl)phenyl]-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 425670-70-8 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Phosphonium, triphenyl(phenylmethyl)-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 1000597-52-3 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
N,N,N-tributylbutan-1-aminium tridecafluorohexane-1-sulfonate |  | - | 108427-54-9 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
N,N,N-triethylethanaminium tridecafluorohexane-1-sulfonate |  | - | 108427-55-0 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
1-Hexanesulfonic acid, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-, compd. With pyrrolidine (1:1) |  | - | 1187817-57-7 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Ethanaminium, N-[4-[[4-(diethylamino)phenyl][4-(ethylamino)-1-naphthalenyl]methylene]-2,5-cyclohexadien-1-ylidene]-N-ethyl-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 1310480-24-0 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Methanaminium, N-[4-[[4-(dimethylamino)phenyl][4-(ethylamino)-1-naphthalenyl]methylene]-2,5-cyclohexadien-1-ylidene]-N-methyl-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 1310480-27-3 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Methanaminium, N-[4-[[4-(dimethylamino)phenyl][4-(phenylamino)-1-naphthalenyl]methylene]-2,5-cyclohexadien-1-ylidene]-N-methyl-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 1310480-28-4 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Beta-Cyclodextrin, compd. with 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonic acid ion(1-)(1:1) |  | - | 1329995-45-0 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Gamma-Cyclodextrin, compd. with 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonic acid ion(1-)(1:1) |  | - | 1329995-69-8 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Sulfonium, triphenyl-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 144116-10-9 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Sulfonium, [4-[(2-methyl-1-oxo-2-propen-1-yl)oxy]phenyl]diphenyl-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 911027-68-4 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Sulfonium, (4-methylphenyl)diphenyl-, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonate (1:1) |  | - | 910606-39-2 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
Iodonium, bis[(1,1-dimethylethyl)phenyl]-, salt with 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-1-hexanesulfonic acid (1:1) (9CI) |  | - | 866621-50-3 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
1-Hexanesulfonic acid, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-, compd. with N,N-diethylethanamine (1:1) |  | - | 72033-41-1 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
1-Hexanesulfonic acid, 1,1,2,2,3,3,4,4,5,5,6,6,6-tridecafluoro-, sodium salt |  | - | 82382-12-5 | vPvB (Article 57e) | 07-Jul-2017 | https://echa.europa.eu/documents/10162/20a23653-34b1-bb48-4887-7ea77bedc637 | https://echa.europa.eu/documents/10162/a118329d-6a4f-f820-bbc0-94c640c0a9fa | https://echa.europa.eu/documents/10162/149c1af7-6375-9ad5-a141-ec53d76b4707#https://echa.europa.eu/documents/10162/1f48372e-97dd-db9f-4335-8cec7ae55eee | https://echa.europa.eu/documents/10162/9cbe10e5-9e87-4992-4097-ea127ca394e2
p-(1,1-dimethylpropyl)phenol |  | 201-280-9 | 80-46-6 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/c11b5b68-67f4-8044-53a6-26759a106c80 | https://echa.europa.eu/documents/10162/c6fa630c-af6f-0df1-8f5d-ec0b890e58a3 | https://echa.europa.eu/documents/10162/02d37927-a27f-fa9b-d8de-bc3c3a46799f | https://echa.europa.eu/documents/10162/d19dd624-e601-fbae-dd35-036f9335aef9
Nonadecafluorodecanoic acid (PFDA) and its sodium and ammonium salts |  | - | - | Toxic for reproduction (Article 57c)#PBT (Article 57d) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/df3daa02-0c97-2c3a-2c7b-90c267642086 | https://echa.europa.eu/documents/10162/a1000164-695a-311d-3eb0-71c09e92dc84 | https://echa.europa.eu/documents/10162/48c9acdc-7474-b256-e8cb-b0c6c1e5f2d0 | https://echa.europa.eu/documents/10162/71f60a29-6e2a-f20f-e79a-58d11056df2f
Nonadecafluorodecanoic acid |  | 206-400-3 | 335-76-2 | Toxic for reproduction (Article 57c)#PBT (Article 57d) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/df3daa02-0c97-2c3a-2c7b-90c267642086 | https://echa.europa.eu/documents/10162/a1000164-695a-311d-3eb0-71c09e92dc84 | https://echa.europa.eu/documents/10162/48c9acdc-7474-b256-e8cb-b0c6c1e5f2d0 | https://echa.europa.eu/documents/10162/71f60a29-6e2a-f20f-e79a-58d11056df2f
Ammonium nonadecafluorodecanoate |  | 221-470-5 | 3108-42-7 | Toxic for reproduction (Article 57c)#PBT (Article 57d) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/df3daa02-0c97-2c3a-2c7b-90c267642086 | https://echa.europa.eu/documents/10162/a1000164-695a-311d-3eb0-71c09e92dc84 | https://echa.europa.eu/documents/10162/48c9acdc-7474-b256-e8cb-b0c6c1e5f2d0 | https://echa.europa.eu/documents/10162/71f60a29-6e2a-f20f-e79a-58d11056df2f
sodium nonadecafluorodecanoate |  | - | 3830-45-3 | Toxic for reproduction (Article 57c)#PBT (Article 57d) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/df3daa02-0c97-2c3a-2c7b-90c267642086 | https://echa.europa.eu/documents/10162/a1000164-695a-311d-3eb0-71c09e92dc84 | https://echa.europa.eu/documents/10162/48c9acdc-7474-b256-e8cb-b0c6c1e5f2d0 | https://echa.europa.eu/documents/10162/71f60a29-6e2a-f20f-e79a-58d11056df2f
4-heptylphenol, branched and linear | substances with a linear and/or branched alkyl chain with a carbon number of 7 covalently bound predominantly in position 4 to phenol, covering also UVCB- and well-defined substances which include any of the individual isomers or a combination thereof | - | - | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-heptylphenol |  | 217-862-0 | 1987-50-4 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
Phenol, heptyl derivs. |  | 276-743-1 | 72624-02-3 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
Phenol, 4-tert-heptyl- |  | - | 288864-02-8 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
Phenol, 4-(1-ethyl-1,2-dimethylpropyl)- |  | - | 30784-27-1 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(2,3,3-trimethylbutan-2-yl)phenol |  | - | 72861-06-4 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(2,4-dimethylpentan-2-yl)phenol |  | - | 33104-11-9 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(3-ethylpentan-3-yl)phenol |  | - | 37872-24-5 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(2-methylhexan-2-yl)phenol |  | - | 30784-31-7 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(3,3-dimethylpentan-2-yl)phenol |  | - | 911371-06-7 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(3-methylhexan-2-yl)phenol |  | - | 854904-93-1 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(2,4-dimethylpentan-3-yl)phenol |  | - | 1824346-00-0 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(4-methylhexan-2-yl)phenol |  | - | 71945-81-8 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(5-methylhexan-2-yl)phenol |  | - | 857629-71-1 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(2,2-dimethylpentan-3-yl)phenol |  | - | 861010-65-3 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(5-methylhexan-3-yl)phenol |  | - | 854904-92-0 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(heptan-3-yl)phenol |  | - | 6465-74-3 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(heptan-2-yl)phenol |  | - | 6863-24-7 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(heptan-4-yl)phenol |  | - | 6465-71-0 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(3-ethylpentyl)phenol |  | - | 911370-98-4 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(3-methylhexyl)phenol |  | - | 102570-52-5 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(4-methylhexyl)phenol |  | - | 1139800-98-8 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(5-methylhexyl)phenol |  | - | 100532-36-3 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(2,3-dimethylpentan-2-yl)phenol |  | - | 861011-60-1 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(3-methylhexan-3-yl)phenol |  | - | 30784-32-8 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4-(4,4-dimethylpentan-2-yl)phenol |  | - | 911371-07-8 | Endocrine disrupting properties (Article 57(f) - environment) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/0f8c5cf3-ccb7-3df6-c351-1c2df00cbc91 | https://echa.europa.eu/documents/10162/32f515fd-4785-d96a-5394-ace770b72b97 | https://echa.europa.eu/documents/10162/9ec084eb-b304-e40b-acd4-23761f9bb6cd#https://echa.europa.eu/documents/10162/8501ed51-1cc7-1ab6-7b4f-4d6e6e0c65df | https://echa.europa.eu/documents/10162/d70e2c1c-c5b8-4570-4f7c-57d6a3f2c39f
4,4'-isopropylidenediphenol | Bisphenol A; BPA | 201-245-8 | 80-05-7 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 12-Jan-2017 | https://echa.europa.eu/documents/10162/36834f25-582c-0855-37fb-bd20b409382c#https://echa.europa.eu/documents/10162/eeed2c09-2263-25ad-49cd-a0926736c877#https://echa.europa.eu/documents/10162/ede153a4-db00-daf6-120f-6b6ccce0c539 | https://echa.europa.eu/documents/10162/a6dc6937-ccec-1024-8282-8304a76549fa | https://echa.europa.eu/documents/10162/b10d6a00-8e47-9b14-4f61-c779a8dc8450#https://echa.europa.eu/documents/10162/908badc9-e65d-3bae-933a-3512a9262e59#https://echa.europa.eu/documents/10162/769b2777-19cd-9fff-33c4-54fe6d8290d5 | https://echa.europa.eu/documents/10162/9abf8b9b-9f39-f9fb-ad06-6f90ed3a7a5e#https://echa.europa.eu/documents/10162/fef2066e-2d2b-677b-71c9-2ea81f517ada#https://echa.europa.eu/documents/10162/4b8caa8c-125d-c1d2-fb59-06b2f0545fbb
Benzo[def]chrysene (Benzo[a]pyrene) |  | 200-028-5 | 50-32-8 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Toxic for reproduction (Article 57c)#PBT (Article 57d)#vPvB (Article 57e) | 20-Jun-2016 | https://echa.europa.eu/documents/10162/4b054c5b-8511-4a30-8ef8-35ab143b4fd0 | https://echa.europa.eu/documents/10162/75ad6d26-78d9-4903-b06a-cf69953713dc | https://echa.europa.eu/documents/10162/985e117f-38e2-4d45-9d0c-94413dd0462e | https://echa.europa.eu/documents/10162/040bf73f-9e4f-57c7-d32b-43bd8880702d
Perfluorononan-1-oic-acid and its sodium and ammonium salts |  | - | - | Toxic for reproduction (Article 57c)#PBT (Article 57d) | 17-Dec-2015 | https://echa.europa.eu/documents/10162/725df6cb-070c-48c9-89a5-500ee2dabe16 | https://echa.europa.eu/documents/10162/2aa67f6e-621e-4337-84aa-076f0de11f4e | https://echa.europa.eu/documents/10162/48ae5fe3-9436-4a10-a533-ed642b92ce47 | https://echa.europa.eu/documents/10162/d6543a35-481d-c564-314b-e2080bb787b6
Perfluorononan-1-oic-acid |  | 206-801-3 | 375-95-1 | Toxic for reproduction (Article 57c)#PBT (Article 57d) | 17-Dec-2015 | https://echa.europa.eu/documents/10162/725df6cb-070c-48c9-89a5-500ee2dabe16 | https://echa.europa.eu/documents/10162/2aa67f6e-621e-4337-84aa-076f0de11f4e | https://echa.europa.eu/documents/10162/48ae5fe3-9436-4a10-a533-ed642b92ce47 | https://echa.europa.eu/documents/10162/d6543a35-481d-c564-314b-e2080bb787b6
sodium heptadecafluorononanoate |  | - | 21049-39-8 | Toxic for reproduction (Article 57c)#PBT (Article 57d) | 17-Dec-2015 | https://echa.europa.eu/documents/10162/725df6cb-070c-48c9-89a5-500ee2dabe16 | https://echa.europa.eu/documents/10162/2aa67f6e-621e-4337-84aa-076f0de11f4e | https://echa.europa.eu/documents/10162/48ae5fe3-9436-4a10-a533-ed642b92ce47 | https://echa.europa.eu/documents/10162/d6543a35-481d-c564-314b-e2080bb787b6
ammonium heptadecafluorononanoate |  | - | 4149-60-4 | Toxic for reproduction (Article 57c)#PBT (Article 57d) | 17-Dec-2015 | https://echa.europa.eu/documents/10162/725df6cb-070c-48c9-89a5-500ee2dabe16 | https://echa.europa.eu/documents/10162/2aa67f6e-621e-4337-84aa-076f0de11f4e | https://echa.europa.eu/documents/10162/48ae5fe3-9436-4a10-a533-ed642b92ce47 | https://echa.europa.eu/documents/10162/d6543a35-481d-c564-314b-e2080bb787b6
Nitrobenzene |  | 202-716-0 | 98-95-3 | Toxic for reproduction (Article 57c) | 17-Dec-2015 | https://echa.europa.eu/documents/10162/bcde2926-a00e-4389-8e48-67122bceba67 | https://echa.europa.eu/documents/10162/d51c9722-aed2-47a0-8356-831db2066b0a | https://echa.europa.eu/documents/10162/d71ba080-df7e-42e6-a76f-4eec3880ebe5 | https://echa.europa.eu/documents/10162/d0e3c2a3-d273-70ee-3191-d939681c22f4
2-(2H-benzotriazol-2-yl)-4-(tert-butyl)-6-(sec-butyl)phenol (UV-350) |  | 253-037-1 | 36437-37-3 | vPvB (Article 57e) | 17-Dec-2015 | https://echa.europa.eu/documents/10162/aca017f7-ce9f-4bce-a991-7c9e911a3ddd | https://echa.europa.eu/documents/10162/2721734c-c482-493b-82e8-3b4d131a6507 | https://echa.europa.eu/documents/10162/13121440-8260-4e04-bc10-5dcad34c7ac0 | https://echa.europa.eu/documents/10162/118b1880-d7bc-b1ed-3539-bf39a3fda3bb
2,4-di-tert-butyl-6-(5-chlorobenzotriazol-2-yl)phenol (UV-327) |  | 223-383-8 | 3864-99-1 | vPvB (Article 57e) | 17-Dec-2015 | https://echa.europa.eu/documents/10162/86a55090-a222-4389-83d6-4859f9efb875 | https://echa.europa.eu/documents/10162/b9c11f35-3120-4aa0-abec-d80c9aeb3786 | https://echa.europa.eu/documents/10162/5d71b975-bba0-482d-9a5e-a102c7a0fc0d | https://echa.europa.eu/documents/10162/2a1e557b-3091-610d-3175-3b08da084dbd
1,3-propanesultone |  | 214-317-9 | 1120-71-4 | Carcinogenic (Article 57a) | 17-Dec-2015 | https://echa.europa.eu/documents/10162/513d6dfe-b82b-a513-687a-d7a8ce6e8a59 | https://echa.europa.eu/documents/10162/ab2bd960-50e9-44b4-a02d-4a3c9a622bdc | https://echa.europa.eu/documents/10162/08f7169e-1697-4323-b13a-b037046eff96 | https://echa.europa.eu/documents/10162/22a7e514-1415-8815-b17d-a91d4e671812
5-sec-butyl-2-(2,4-dimethylcyclohex-3-en-1-yl)-5-methyl-1,3-dioxane [1], 5-sec-butyl-2-(4,6-dimethylcyclohex-3-en-1-yl)-5-methyl-1,3-dioxane [2] | covering any of the individual stereoisomers of [1] and [2] or any combination thereof | - | - | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,3-Dioxane, 2-[(1R,2R)-2,4-dimethyl-3-cyclohexen-1-yl]-5-methyl-5-(1-methylpropyl)-, cis- |  | - | 676367-05-8 | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,3-Dioxane, 2-(2,4-dimethyl-3-cyclohexen-1-yl)-5-methyl-5-(1-methylpropyl)- |  | - | 117933-89-8 | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,3-Dioxane, 2-[(1R,2R)-2,4-dimethyl-3-cyclohexen-1-yl]-5-methyl-5-(1-methylpropyl)-, trans- |  | - | 676367-09-2 | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,3-Dioxane, 2-[(1R,2R)-2,4-dimethyl-3-cyclohexen-1-yl]-5-methyl-5-(1-methylpropyl)-, trans-rel- |  | - | 343934-05-4 | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,3-Dioxane, 2-[(1R,2S)-2,4-dimethyl-3-cyclohexen-1-yl]-5-methyl-5-(1-methylpropyl)-, cis- |  | - | 676367-04-7 | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,3-Dioxane, 2-[(1R,2S)-2,4-dimethyl-3-cyclohexen-1-yl]-5-methyl-5-(1-methylpropyl)-, trans- |  | - | 676367-08-1 | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,3-Dioxane, 2-[(1S,2R)-2,4-dimethyl-3-cyclohexen-1-yl]-5-methyl-5-(1-methylpropyl)-, cis- |  | - | 676367-03-6 | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,3-Dioxane, 2-[(1S,2R)-2,4-dimethyl-3-cyclohexen-1-yl]-5-methyl-5-(1-methylpropyl)-, trans- |  | - | 676367-07-0 | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,3-Dioxane, 2-[(1S,2S)-2,4-dimethyl-3-cyclohexen-1-yl]-5-methyl-5-(1-methylpropyl)-, cis- |  | - | 676367-02-5 | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,3-Dioxane, 2-[(1S,2S)-2,4-dimethyl-3-cyclohexen-1-yl]-5-methyl-5-(1-methylpropyl)-, trans- |  | - | 676367-06-9 | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
Reaction mass of 5-sec-butyl-2-(2,4-dimethylcyclohex-3-en-1-yl)-5-methyl-1,3-dioxane and 5-sec-butyl-2-(4,6-dimethylcyclohex-3-en-1-yl)-5-methyl-1,3-dioxane |  | 413-720-9 | - | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
Reaction mass of 5-[(2R)-butan-2-yl]-2-[(1R,2R)-2,4-dimethylcyclohex-3-en-1-yl]-5-methyl-1,3-dioxane and 5-[(2R)-butan-2-yl]-2-[(1R,6R)-4,6-dimethylcyclohex-3-en-1-yl]-5-methyl-1,3-dioxane and 5-[(2S)-butan-2-yl]-2-[(1R,2R)-2,4-dimethylcyclohex-3-en-1-yl]-5-methyl-1,3-dioxane and 5-[(2S)-butan-2-yl]-2-[(1S,2R)-2,4-dimethylcyclohex-3-en-1-yl]-5-methyl-1,3-dioxane and 5-[(2S)-butan-2-yl]-2-[(1S,6R)-4,6-dimethylcyclohex-3-en-1-yl]-5-methyl-1,3-dioxane |  | - | - | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
5-sec-butyl-2-(2,4-dimethylcyclohex-3-en-1-yl)-5-methyl-1,3-dioxane |  | - | - | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
5-sec-butyl-2-(4,6-dimethylcyclohex-3-en-1-yl)-5-methyl-1,3-dioxane |  | - | - | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,3-Dioxane, 2-[(1R,2R)-2,4-dimethyl-3-cyclohexen-1-yl]-5-methyl-5-(1-methylpropyl)-, cis-rel- |  | - | 343934-04-3 | vPvB (Article 57e) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/f7c78fc1-bc3e-4913-9a8f-1f57e03326e4 | https://echa.europa.eu/documents/10162/c87c2930-58af-20dd-8394-afcd0d96ee3f | https://echa.europa.eu/documents/10162/a4fac134-09e6-43c1-a65f-dfaee5f85731 | https://echa.europa.eu/documents/10162/8831bc31-c2e0-f702-ef41-fb3a49ec0789
1,2-benzenedicarboxylic acid, di-C6-10-alkyl esters or mixed decyl and hexyl and octyl diesters | with ≥ 0.3% of dihexyl phthalate (EC No. 201-559-5) | - | - | Toxic for reproduction (Article 57c) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/af844510-0d70-6a9e-4fd5-714632a0f988 | https://echa.europa.eu/documents/10162/0d1991c8-23cf-4779-9e08-b9208bd40338#https://echa.europa.eu/documents/10162/f5495229-3f8c-4321-817c-dd498912310e | https://echa.europa.eu/documents/10162/2b4a5011-39c9-4eb0-9309-9aa6b10981ab | https://echa.europa.eu/documents/10162/c79982ce-adc9-d5c8-e42a-d8bc025c9f93
1,2-Benzenedicarboxylic acid, mixed decyl and hexyl and octyl diesters |  | 272-013-1 | 68648-93-1 | Toxic for reproduction (Article 57c) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/af844510-0d70-6a9e-4fd5-714632a0f988 | https://echa.europa.eu/documents/10162/0d1991c8-23cf-4779-9e08-b9208bd40338#https://echa.europa.eu/documents/10162/f5495229-3f8c-4321-817c-dd498912310e | https://echa.europa.eu/documents/10162/2b4a5011-39c9-4eb0-9309-9aa6b10981ab | https://echa.europa.eu/documents/10162/c79982ce-adc9-d5c8-e42a-d8bc025c9f93
1,2-Benzenedicarboxylic acid, di-C6-10-alkyl esters |  | 271-094-0 | 68515-51-5 | Toxic for reproduction (Article 57c) | 15-Jun-2015 | https://echa.europa.eu/documents/10162/af844510-0d70-6a9e-4fd5-714632a0f988 | https://echa.europa.eu/documents/10162/0d1991c8-23cf-4779-9e08-b9208bd40338#https://echa.europa.eu/documents/10162/f5495229-3f8c-4321-817c-dd498912310e | https://echa.europa.eu/documents/10162/2b4a5011-39c9-4eb0-9309-9aa6b10981ab | https://echa.europa.eu/documents/10162/c79982ce-adc9-d5c8-e42a-d8bc025c9f93
Reaction mass of 2-ethylhexyl 10-ethyl-4,4-dioctyl-7-oxo-8-oxa-3,5-dithia-4-stannatetradecanoate and 2-ethylhexyl 10-ethyl-4-[[2-[(2-ethylhexyl)oxy]-2-oxoethyl]thio]-4-octyl-7-oxo-8-oxa-3,5-dithia-4-stannatetradecanoate (reaction mass of DOTE and MOTE) |  | - | - | Toxic for reproduction (Article 57c) | 17-Dec-2014 | https://echa.europa.eu/documents/10162/917a0639-8b9d-43ce-802c-b4e82a7f20e3 | https://echa.europa.eu/documents/10162/371b4814-2cbd-429d-88b9-905c19d0d299 | https://echa.europa.eu/documents/10162/a410b50c-11f9-49ca-9e8f-54f2a674b032 | https://echa.europa.eu/documents/10162/ba3a1dda-f1bb-0323-6067-257f1d310652
Cadmium sulphate |  | 233-331-6 | 10124-36-4, 31119-53-6 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Toxic for reproduction (Article 57c)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 17-Dec-2014 | https://echa.europa.eu/documents/10162/9e3c41d5-088a-47e7-944a-2ccb112493b3 | https://echa.europa.eu/documents/10162/bbe26e5b-facf-4668-bc8c-38a175be4c62 | https://echa.europa.eu/documents/10162/037ede8d-4f3a-030c-8663-8b04336272ad | https://echa.europa.eu/documents/10162/8418e3af-5df2-d281-49bf-5ce691521e86
Sulfuric acid, cadmium salt (1:1), hydrate |  | 626-360-4 | 15244-35-6 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Toxic for reproduction (Article 57c)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 17-Dec-2014 | https://echa.europa.eu/documents/10162/9e3c41d5-088a-47e7-944a-2ccb112493b3 | https://echa.europa.eu/documents/10162/bbe26e5b-facf-4668-bc8c-38a175be4c62 | https://echa.europa.eu/documents/10162/037ede8d-4f3a-030c-8663-8b04336272ad | https://echa.europa.eu/documents/10162/8418e3af-5df2-d281-49bf-5ce691521e86
Sulfuric acid, cadmium salt, hydrate (3:3:8) |  | 616-572-5 | 7790-84-3 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Toxic for reproduction (Article 57c)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 17-Dec-2014 | https://echa.europa.eu/documents/10162/9e3c41d5-088a-47e7-944a-2ccb112493b3 | https://echa.europa.eu/documents/10162/bbe26e5b-facf-4668-bc8c-38a175be4c62 | https://echa.europa.eu/documents/10162/037ede8d-4f3a-030c-8663-8b04336272ad | https://echa.europa.eu/documents/10162/8418e3af-5df2-d281-49bf-5ce691521e86
Cadmium fluoride |  | 232-222-0 | 7790-79-6 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Toxic for reproduction (Article 57c)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 17-Dec-2014 | https://echa.europa.eu/documents/10162/e333469a-fca7-48c2-a88f-52eaa8e7ad7e | https://echa.europa.eu/documents/10162/0bba6429-659f-448b-9623-c6737291a3b1 | https://echa.europa.eu/documents/10162/6eb02b71-b691-4b53-b0f1-c8c68b2b3083 | https://echa.europa.eu/documents/10162/c185b4bd-150f-f36b-d3a4-48cf45d1e5fb
2-ethylhexyl 10-ethyl-4,4-dioctyl-7-oxo-8-oxa-3,5-dithia-4-stannatetradecanoate (DOTE) |  | 239-622-4 | 15571-58-1 | Toxic for reproduction (Article 57c) | 17-Dec-2014 | https://echa.europa.eu/documents/10162/8f7e275c-c02c-4357-90d9-59a37b14337a | https://echa.europa.eu/documents/10162/0a3e8564-b077-46a7-a90e-2f59f458932f | https://echa.europa.eu/documents/10162/04d23d27-3484-48c9-862c-0637e30642a1 | https://echa.europa.eu/documents/10162/80eac049-4076-f96d-cd11-0d8477cb12dc
2-benzotriazol-2-yl-4,6-di-tert-butylphenol (UV-320) |  | 223-346-6 | 3846-71-7 | PBT (Article 57d)#vPvB (Article 57e) | 17-Dec-2014 | https://echa.europa.eu/documents/10162/27d9317d-3aeb-41d5-9116-1e07696d90cd | https://echa.europa.eu/documents/10162/96387d49-cb01-423c-bd6e-927a7b6c434f | https://echa.europa.eu/documents/10162/33c375a1-c4de-4f14-a143-cff986175d76 | https://echa.europa.eu/documents/10162/6ac88f0e-1962-78d3-a7ed-cfb4f89b9aa7
2-(2H-benzotriazol-2-yl)-4,6-ditertpentylphenol (UV-328) |  | 247-384-8 | 25973-55-1 | PBT (Article 57d)#vPvB (Article 57e) | 17-Dec-2014 | https://echa.europa.eu/documents/10162/c6a0c43b-adba-4351-a093-66b36d470de3 | https://echa.europa.eu/documents/10162/0ca781ed-adf7-4d40-95b7-a7ea17543f5c | https://echa.europa.eu/documents/10162/31f2c070-ad6f-4de9-b17d-402eedb05eac | https://echa.europa.eu/documents/10162/e324f17f-e945-74ac-a7ea-9471cc2bfdec
Sodium peroxometaborate |  | 231-556-4 | 7632-04-4 | Toxic for reproduction (Article 57c) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/55e1c126-a528-440f-8c10-49a31b82eda3 | https://echa.europa.eu/documents/10162/59b315b3-efb1-4ffa-b4bb-3c2c1259da45 | https://echa.europa.eu/documents/10162/46e192c7-457f-43fc-b93f-14d2ce0347f9 | https://echa.europa.eu/documents/10162/993ec24b-d88d-8fc4-97af-6c71fca118fa
Sodium perborate, perboric acid, sodium salt |  | - | - | Toxic for reproduction (Article 57c) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/b5ae52f9-bb66-62b5-9950-242db987ec5a | https://echa.europa.eu/documents/10162/ec145926-d7d2-a584-9c82-e325a05367a3 | https://echa.europa.eu/documents/10162/42fee5a3-e00e-e554-30cb-66a735d4fe9e | https://echa.europa.eu/documents/10162/3af316eb-892a-abf5-48c2-b26c6f0351b5
Perboric acid, sodium salt, tetrahydrate () |  | - | 37244-98-7 | Toxic for reproduction (Article 57c) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/b5ae52f9-bb66-62b5-9950-242db987ec5a | https://echa.europa.eu/documents/10162/ec145926-d7d2-a584-9c82-e325a05367a3 | https://echa.europa.eu/documents/10162/42fee5a3-e00e-e554-30cb-66a735d4fe9e | https://echa.europa.eu/documents/10162/3af316eb-892a-abf5-48c2-b26c6f0351b5
Perboric acid (HBO(O2)), sodium salt, tetrahydrate |  | - | 10486-00-7 | Toxic for reproduction (Article 57c) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/b5ae52f9-bb66-62b5-9950-242db987ec5a | https://echa.europa.eu/documents/10162/ec145926-d7d2-a584-9c82-e325a05367a3 | https://echa.europa.eu/documents/10162/42fee5a3-e00e-e554-30cb-66a735d4fe9e | https://echa.europa.eu/documents/10162/3af316eb-892a-abf5-48c2-b26c6f0351b5
Sodium perborate monohydrate |  | - | 10332-33-9 | Toxic for reproduction (Article 57c) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/b5ae52f9-bb66-62b5-9950-242db987ec5a | https://echa.europa.eu/documents/10162/ec145926-d7d2-a584-9c82-e325a05367a3 | https://echa.europa.eu/documents/10162/42fee5a3-e00e-e554-30cb-66a735d4fe9e | https://echa.europa.eu/documents/10162/3af316eb-892a-abf5-48c2-b26c6f0351b5
Perboric acid, sodium salt |  | 234-390-0 | 11138-47-9 | Toxic for reproduction (Article 57c) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/b5ae52f9-bb66-62b5-9950-242db987ec5a | https://echa.europa.eu/documents/10162/ec145926-d7d2-a584-9c82-e325a05367a3 | https://echa.europa.eu/documents/10162/42fee5a3-e00e-e554-30cb-66a735d4fe9e | https://echa.europa.eu/documents/10162/3af316eb-892a-abf5-48c2-b26c6f0351b5
Sodium perborate |  | 239-172-9 | 15120-21-5 | Toxic for reproduction (Article 57c) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/b5ae52f9-bb66-62b5-9950-242db987ec5a | https://echa.europa.eu/documents/10162/ec145926-d7d2-a584-9c82-e325a05367a3 | https://echa.europa.eu/documents/10162/42fee5a3-e00e-e554-30cb-66a735d4fe9e | https://echa.europa.eu/documents/10162/3af316eb-892a-abf5-48c2-b26c6f0351b5
Borate(2-), tetrahydroxybis[μ-(peroxy-κO1:κO2)]di-, sodium, hydrate (1:2:6) |  | - | 125022-34-6 | Toxic for reproduction (Article 57c) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/b5ae52f9-bb66-62b5-9950-242db987ec5a | https://echa.europa.eu/documents/10162/ec145926-d7d2-a584-9c82-e325a05367a3 | https://echa.europa.eu/documents/10162/42fee5a3-e00e-e554-30cb-66a735d4fe9e | https://echa.europa.eu/documents/10162/3af316eb-892a-abf5-48c2-b26c6f0351b5
Borate(2-), tetrahydroxybis[μ-(peroxy-κO1:κO2)]di-, sodium (1:2) |  | - | 90568-23-3 | Toxic for reproduction (Article 57c) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/b5ae52f9-bb66-62b5-9950-242db987ec5a | https://echa.europa.eu/documents/10162/ec145926-d7d2-a584-9c82-e325a05367a3 | https://echa.europa.eu/documents/10162/42fee5a3-e00e-e554-30cb-66a735d4fe9e | https://echa.europa.eu/documents/10162/3af316eb-892a-abf5-48c2-b26c6f0351b5
Perboric acid (H3BO2(O2)), monosodium salt, trihydrate |  | - | 13517-20-9 | Toxic for reproduction (Article 57c) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/b5ae52f9-bb66-62b5-9950-242db987ec5a | https://echa.europa.eu/documents/10162/ec145926-d7d2-a584-9c82-e325a05367a3 | https://echa.europa.eu/documents/10162/42fee5a3-e00e-e554-30cb-66a735d4fe9e | https://echa.europa.eu/documents/10162/3af316eb-892a-abf5-48c2-b26c6f0351b5
Cadmium chloride |  | 233-296-7 | 10108-64-2 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Toxic for reproduction (Article 57c)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/e2bc867c-de85-4d6d-a9d6-d856ab51685c | https://echa.europa.eu/documents/10162/a7d0fcf7-deb2-402e-af3b-c3b1759615f9 | https://echa.europa.eu/documents/10162/a210c4dd-6ef4-4f76-9fd9-5b35a7482c1b | https://echa.europa.eu/documents/10162/dc9015af-635b-c359-b9d4-56d7725b9d58
cadmium chloride (CdCl2), hydrate (2:5) |  | 640-998-0 | 7790-78-5 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Toxic for reproduction (Article 57c)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/e2bc867c-de85-4d6d-a9d6-d856ab51685c | https://echa.europa.eu/documents/10162/a7d0fcf7-deb2-402e-af3b-c3b1759615f9 | https://echa.europa.eu/documents/10162/a210c4dd-6ef4-4f76-9fd9-5b35a7482c1b | https://echa.europa.eu/documents/10162/dc9015af-635b-c359-b9d4-56d7725b9d58
1,2-Benzenedicarboxylic acid, dihexyl ester, branched and linear |  | 271-093-5 | 68515-50-4 | Toxic for reproduction (Article 57c) | 16-Jun-2014 | https://echa.europa.eu/documents/10162/34533e6e-84d2-4a40-815d-7a02dac17041 | https://echa.europa.eu/documents/10162/4ffa71cf-34d5-4968-9b46-58d2398a211a | https://echa.europa.eu/documents/10162/f2fc61d4-7e23-4a9a-a52d-10ae36e9987e | https://echa.europa.eu/documents/10162/e4456f95-477d-9cf9-2fcb-3db6f05ca8ab
Trixylyl phosphate |  | 246-677-8 | 25155-23-1 | Toxic for reproduction (Article 57c) | 16-Dec-2013 | https://echa.europa.eu/documents/10162/464f639f-6e07-4966-b63a-081ac8040e63 | https://echa.europa.eu/documents/10162/e764d667-57dc-4956-8628-6245f36d5552 | https://echa.europa.eu/documents/10162/aa428e86-ed97-4450-8e78-cdbb06e46c8f | https://echa.europa.eu/documents/10162/3e37f35a-48a7-7ce9-a97d-7659def7e080
Lead di(acetate) |  | 206-104-4 | 301-04-2 | Toxic for reproduction (Article 57c) | 16-Dec-2013 | https://echa.europa.eu/documents/10162/ff4184dc-c595-4dc8-8241-941c073f2221 | https://echa.europa.eu/documents/10162/86ace337-8153-47fb-b0b0-8193c621c433 | https://echa.europa.eu/documents/10162/49fa197b-8a06-4d92-abf4-f40ad0a6f815 | https://echa.europa.eu/documents/10162/c0a30142-49b9-a857-0d20-e1a6f6da0525
Imidazolidine-2-thione (2-imidazoline-2-thiol) |  | 202-506-9 | 96-45-7 | Toxic for reproduction (Article 57c) | 16-Dec-2013 | https://echa.europa.eu/documents/10162/e01d8301-5596-4905-98ad-f17eb9455241 | https://echa.europa.eu/documents/10162/85a2311e-2252-4f8a-b987-451846cee286 | https://echa.europa.eu/documents/10162/24751f9c-5206-4190-bcc9-3ca9418e0e31 | https://echa.europa.eu/documents/10162/f6d7dd12-593f-57e5-8026-a2f9cccf561f
Disodium 4-amino-3-[[4'-[(2,4-diaminophenyl)azo][1,1'-biphenyl]-4-yl]azo] -5-hydroxy-6-(phenylazo)naphthalene-2,7-disulphonate (C.I. Direct Black 38) |  | 217-710-3 | 1937-37-7 | Carcinogenic (Article 57a) | 16-Dec-2013 | https://echa.europa.eu/documents/10162/6ab9d8cb-2d82-4dba-a68b-8752a00c4c76 | https://echa.europa.eu/documents/10162/d04de1ae-26e0-45ac-969c-259a0cdd5338 | https://echa.europa.eu/documents/10162/ef1f3daa-fc8b-4bf5-8bd1-71e80a2adf62 | https://echa.europa.eu/documents/10162/2261c554-cdf2-4ed1-9205-e54d68d6e20f
Disodium 3,3'-[[1,1'-biphenyl]-4,4'-diylbis(azo)]bis(4-aminonaphthalene-1-sulphonate) (C.I. Direct Red 28) |  | 209-358-4 | 573-58-0 | Carcinogenic (Article 57a) | 16-Dec-2013 | https://echa.europa.eu/documents/10162/f9d799e8-5a32-44d1-b383-4c9695085cbf | https://echa.europa.eu/documents/10162/fd83a2e5-21ca-4ec4-a4d2-2465637f90c0 | https://echa.europa.eu/documents/10162/cf3a4398-3305-424e-ab1e-a02fbc548431 | https://echa.europa.eu/documents/10162/e306453a-92c1-0aa9-469d-acde10267249
Dihexyl phthalate |  | 201-559-5 | 84-75-3 | Toxic for reproduction (Article 57c) | 16-Dec-2013 | https://echa.europa.eu/documents/10162/d90bb47a-7e33-43c3-82ad-18a55d39d502 | https://echa.europa.eu/documents/10162/eecf78d5-f23d-44b0-a3db-0357c698ae4b | https://echa.europa.eu/documents/10162/65336c46-c178-4b71-a927-15437e8b303a | https://echa.europa.eu/documents/10162/fa40d148-4f2d-9cbe-7295-c2ada7db7356
Cadmium sulphide |  | 215-147-8 | 1306-23-6 | Carcinogenic (Article 57a)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 16-Dec-2013 | https://echa.europa.eu/documents/10162/204bc9fa-0673-4753-bd0e-e3503b4a1956 | https://echa.europa.eu/documents/10162/f30a9a9a-9ce1-429e-96fb-7f24fe4de850 | https://echa.europa.eu/documents/10162/2aceee75-6b50-416c-aa30-d3131c5eeee3 | https://echa.europa.eu/documents/10162/8cce0a23-b9bf-a13d-afa1-a42e946b4390
Pentadecafluorooctanoic acid (PFOA) |  | 206-397-9 | 335-67-1 | Toxic for reproduction (Article 57c)#PBT (Article 57d) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/092663e6-b14a-4a06-aadf-fc0e56bc0a23 | https://echa.europa.eu/documents/10162/facda3f9-a050-4f29-91f1-358175b67a39 | https://echa.europa.eu/documents/10162/8059e342-1092-410f-bd85-80118a5526f5 | https://echa.europa.eu/documents/10162/730a8511-d866-e91a-d3ad-1f700ea36b8b
Dipentyl phthalate (DPP) |  | 205-017-9 | 131-18-0 | Toxic for reproduction (Article 57c) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/365998db-beb8-47b6-b080-ca5549397cd9 | https://echa.europa.eu/documents/10162/8a80cf7b-cce9-4152-89f4-e8ad58ce6fde | https://echa.europa.eu/documents/10162/11bbf18f-02ac-4b12-a715-056df904057c | https://echa.europa.eu/documents/10162/7f0bea3d-3432-0528-6f36-53f3a2bef27a
Cadmium oxide |  | 215-146-2 | 1306-19-0 | Carcinogenic (Article 57a)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/2fb75d7e-485c-4edb-8c72-4204f224bbca | https://echa.europa.eu/documents/10162/a03b1b4e-3090-455e-9cb0-1dc52005f7c0 | https://echa.europa.eu/documents/10162/f09f23c5-98fc-4e56-abc2-feae96de0f3d | https://echa.europa.eu/documents/10162/cf1074c6-b588-f914-90aa-85f885d3c8a4
Monteponite (CdO) |  | - | 12139-21-8 | Carcinogenic (Article 57a)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/2fb75d7e-485c-4edb-8c72-4204f224bbca | https://echa.europa.eu/documents/10162/a03b1b4e-3090-455e-9cb0-1dc52005f7c0 | https://echa.europa.eu/documents/10162/f09f23c5-98fc-4e56-abc2-feae96de0f3d | https://echa.europa.eu/documents/10162/cf1074c6-b588-f914-90aa-85f885d3c8a4
Cadmium |  | 231-152-8 | 7440-43-9 | Carcinogenic (Article 57a)#Specific target organ toxicity after repeated exposure (Article 57(f) - human health) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/49a335ae-1ec4-40e5-88fb-3b08702da95c | https://echa.europa.eu/documents/10162/b252eefb-8012-46bb-8a47-74ff66a0a8d4 | https://echa.europa.eu/documents/10162/a048359b-de39-4b7e-8602-51272a55aeae | https://echa.europa.eu/documents/10162/a852ca1e-50fb-c3a4-f14d-2fee19c69701
Ammonium pentadecafluorooctanoate (APFO) |  | 223-320-4 | 3825-26-1 | Toxic for reproduction (Article 57c)#PBT (Article 57d) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/b06436fd-6367-4a2f-b6a2-2b8ffa183ae9 | https://echa.europa.eu/documents/10162/df602a5b-2df3-490f-ad1c-e2e782ad7dc9 | https://echa.europa.eu/documents/10162/5e2c1e53-be98-4104-8b96-9cd88655a92a | https://echa.europa.eu/documents/10162/8cf84645-353d-7cb2-bdc3-c44d12e19c84
4-Nonylphenol, branched and linear, ethoxylated | substances with a linear and/or branched alkyl chain with a carbon number of 9 covalently bound in position 4 to phenol, ethoxylated covering UVCB- and well-defined substances, polymers and homologues, which include any of the individual isomers and/or combinations thereof | 799-990-1 | - | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Nonylphenol, ethoxylated |  | 500-024-6 | 9016-45-9 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
2-[2-(4-nonylphenoxy)ethoxy]ethanol |  | 243-816-4 | 20427-84-3 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
2-(nonylphenoxy)ethanol |  | 248-762-5 | 27986-36-3 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
14-(nonylphenoxy)-3,6,9,12-tetraoxatetradecan-1-ol |  | 247-555-7 | 26264-02-8 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
2-[2-(nonylphenoxy)ethoxy]ethanol |  | 248-291-5 | 27176-93-8 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
23-(nonylphenoxy)-3,6,9,12,15,18,21-heptaoxatricosan-1-ol |  | 248-293-6 | 27177-05-5 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
4-Nonylphenol, ethoxylated | 1 - 2.5 moles ethoxylated | 500-045-0 | 26027-38-3 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
26-(4-Nonylphenoxy)-3,6,9,12,15,18,21,24- octaoxahexacosan -1-ol |  | - | 14409-72-4 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
4-Nonylphenol, branched, ethoxylated | 1 - 2.5 moles ethoxylated | 500-315-8 | 127087-87-0 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Nonylphenol, branched, ethoxylated | 1 - 2.5 moles ethoxylated | 500-209-1 | 68412-54-4 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
2-{2-[4-(3,6-dimethylheptan-3-yl)phenoxy]ethoxy}ethanol |  | - | 1119449-38-5 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
2-[2-[2-[2-(4-nonylphenoxy)ethoxy]ethoxy]ethoxy]ethanol |  | 230-770-5 | 7311-27-5 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
20-(4-nonylphenoxy)-3,6,9,12,15,18-hexaoxaicosan-1-ol |  | 248-743-1 | 27942-27-4 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Poly(oxy-1,2-ethanediyl), α-(nonylphenyl)-ω-hydroxy- |  | - | 9016-45-9 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
4-Nonylphenol, branched, ethoxylated |  | - | 127087-87-0 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
26-(nonylphenoxy)-3,6,9,12,15,18,21,24-octaoxahexacosan-1-ol; |  | 255-695-5 | 42173-90-0 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
4-t-Nonylphenol-diethoxylate |  | - | 156609-10-8 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
2-(4-nonylphenoxy)ethanol |  | - | 104-35-8 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
3,6,9,12-Tetraoxatetradecan-1-ol, 14-(4-nonylphenoxy)-, branched |  | 293-926-1 | 91648-64-5 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
3,6,9,12,15,18,21,24,27-Nonaoxanonacosan-1-ol, 29-(isononylphenoxy)- |  | - | 65455-72-3 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
3,6,9,12,15-Pentaoxaheptadecan-1-ol, 17-(nonylphenoxy)- |  | - | 27177-01-1 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
3,6,9,12-Tetraoxatetradecan-1-ol, 14-(4-nonylphenoxy)- |  | - | 20636-48-0 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Nonylphenol, ethoxylated (EO = 10) |  | - | - | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Nonylphenol, ethoxylated (EO = 4) |  | - | - | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Nonylphenol, ethoxylated (polymer) |  | - | - | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
2-[4-(3,6-dimethylheptan-3-yl)phenoxy]ethanol |  | - | 1119449-37-4 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Isononylphenol, ethoxylated |  | - | 37205-87-1 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
17-(4-nonylphenoxy)-3,6,9,12,15-pentaoxaheptadecan-1-ol |  | - | 34166-38-6 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
2-(isononylphenoxy)ethanol |  | 284-987-5 | 85005-55-6 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Nonylphenolpolyglycolether |  | - | - | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
44-(nonylphenoxy)-3,6,9,12,15,18,21,24,27,30,33,36,39,42-tetradecaoxatetratetracontanol |  | 260-678-0 | 57321-10-5 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
20-(isononylphenoxy)-3,6,9,12,15,18-hexaoxaicosan-1-ol |  | 265-785-6 | 65455-69-8 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
29-(nonylphenoxy)-3,6,9,12,15,18,21,24,27-nonaoxanonacosanol |  | 248-294-1 | 27177-08-8 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
20-(nonylphenoxy)-3,6,9,12,15,18-hexaoxaicosan-1-ol |  | 248-292-0 | 27177-03-3 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
26-(nonylphenoxy)-3,6,9,12,15,18,21,24-octaoxahexacosan-1-ol |  | 247-816-5 | 26571-11-9 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Nonylphenol, ethoxylated (8-EO) |  | - | 9016-45-9 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Nonylphenol, ethoxylated (6,5-EO) |  | - | 9016-45-9 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Nonylphenol, ethoxylated (10-EO) |  | - | 9016-45-9 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Poly(oxy-1,2-ethanediyl), α-(nonylphenyl)-ω-hydroxy-, branched |  | - | 68412-54-4 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Nonylphenol, ethoxylated (15-EO) |  | - | 9016-45-9 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Poly (oxy-1,2-ethanediyl), alpha -(nonylphenyl)-omega-hydroxy-, branched |  | - | 68412-54-4 | Endocrine disrupting properties (Article 57(f) - environment) | 20-Jun-2013 | https://echa.europa.eu/documents/10162/fd80f890-fb73-4201-8e2f-06dc04eaf5a9 | https://echa.europa.eu/documents/10162/90d219f2-cd83-4e2b-bb9e-3ed6d5ff91bf | https://echa.europa.eu/documents/10162/9af34d5f-cd2f-4e63-859c-529bb39da7ae#https://echa.europa.eu/documents/10162/33836bb4-e6b9-b686-1dc9-e8618ee32db5 | https://echa.europa.eu/documents/10162/0ef7af92-b74c-b38b-aa95-3730bf89372d
Trilead dioxide phosphonate |  | 235-252-2 | 12141-20-7 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/af3610a5-f180-4fcd-95b7-636e43da9198 | https://echa.europa.eu/documents/10162/57804b85-bfe3-4fb3-ba6e-85681ae9a946 | https://echa.europa.eu/documents/10162/0b3bffdb-ee49-4d1d-9ef7-7e1aeed13818 | https://echa.europa.eu/documents/10162/e890ef44-027b-69a9-123e-020a2e1ed3fd
Trilead bis(carbonate) dihydroxide |  | 215-290-6 | 1319-46-6 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/bd13bb56-cf5a-4eee-b7a2-367ba3838e86 | https://echa.europa.eu/documents/10162/e8388e88-796e-48d6-b857-92e1de8cba7b | https://echa.europa.eu/documents/10162/4641a546-af85-9bd7-f550-8a6269bd02f8 | https://echa.europa.eu/documents/10162/2620a423-631b-ddac-0985-f657f444b4be
Tricosafluorododecanoic acid |  | 206-203-2 | 307-55-1 | vPvB (Article 57e) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dd827b99-e744-4d49-84c3-00b82732059b | https://echa.europa.eu/documents/10162/10982bb6-eb4f-4b49-8ae2-934e051c4227 | https://echa.europa.eu/documents/10162/44fd0731-2211-0216-b984-e3cd0f32d5ca | https://echa.europa.eu/documents/10162/ed8c31e6-8c70-9a9f-6861-06cd4d493c09
Tetralead trioxide sulphate |  | 235-380-9 | 12202-17-4 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/f56110e1-a0d6-4a1e-8c7d-9940d5ec8d91 | https://echa.europa.eu/documents/10162/08f78ab7-6708-418d-ab77-adf1100ab429 | https://echa.europa.eu/documents/10162/735b9afe-5a6a-41d2-9b7b-7758ec3c5702 | https://echa.europa.eu/documents/10162/1339c831-6e8e-da3a-a622-55892c08ebc0
Tetraethyllead |  | 201-075-4 | 78-00-2 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/0b417b76-b533-42a1-9bd2-519f1dc1990d | https://echa.europa.eu/documents/10162/118d71ba-5734-45ac-9dfe-ab74edca7803 | https://echa.europa.eu/documents/10162/136a8d6e-ebac-4000-8183-ad9176fd134b | https://echa.europa.eu/documents/10162/35c5243b-558d-85ff-94f8-e54babec4986
Sulfurous acid, lead salt, dibasic |  | 263-467-1 | 62229-08-7 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/486d5d7c-c872-48cf-9392-e86ecff34c04 | https://echa.europa.eu/documents/10162/44b75647-0300-4836-a312-f11ba9574ece | https://echa.europa.eu/documents/10162/cc6c0ca7-538e-4a5b-88fe-b1a635241a45 | https://echa.europa.eu/documents/10162/679974d0-99da-fb52-5344-70942eaf0b21
Silicic acid, lead salt |  | 234-363-3 | 11120-22-2 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/a7314719-5b08-47ef-9ee7-daca71e06f54 | https://echa.europa.eu/documents/10162/8a921407-3b0f-471a-a684-131dfba77523 | https://echa.europa.eu/documents/10162/13a5f189-b83e-e1b4-b201-084b82fd7a68 | https://echa.europa.eu/documents/10162/426137d8-a88b-d850-0610-3eab583446ad
Silicic acid (H2Si2O5), barium salt (1:1), lead-doped | with lead (Pb) content above the applicable generic concentration limit for ’toxicity for reproduction’ Repr. 1A (CLP) or category 1 (DSD),the substance is a member of the group entry of lead compounds, with index number 082-001-00-6 in Regulation (EC) No 1272/2008 | 272-271-5 | 68784-75-8 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/6bd3a623-dcbc-43ff-a5fa-8f710fc4053b | https://echa.europa.eu/documents/10162/d0de0c48-5768-4589-8dd4-dbcafce8e7dd | https://echa.europa.eu/documents/10162/ca9a6d0f-2618-bf67-6c3b-63bc706422bc | https://echa.europa.eu/documents/10162/602a0aad-a902-c5b7-421b-56a30dc89a4a
Pyrochlore, antimony lead yellow | - | 232-382-1 | 8012-00-8 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/9e101a56-02e0-481f-babb-e4f62d36912b | https://echa.europa.eu/documents/10162/3da568db-d1e1-4082-b0c7-78246ccb1489 | https://echa.europa.eu/documents/10162/1014fa88-d28e-4385-8530-fe2c38032c27 | https://echa.europa.eu/documents/10162/d19da13c-65ed-6bcc-6922-e499ef3caba7
Pentalead tetraoxide sulphate |  | 235-067-7 | 12065-90-6 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/664946a7-1fba-487d-8d28-f47de7fb3a5c | https://echa.europa.eu/documents/10162/adc19a74-9ef0-4920-93c5-0b25d600b8c4 | https://echa.europa.eu/documents/10162/5a790e77-efd4-4c56-84bb-b00a92b732b4 | https://echa.europa.eu/documents/10162/690d1df2-8058-e37f-235f-ca184e88ea88
Pentacosafluorotridecanoic acid |  | 276-745-2 | 72629-94-8 | vPvB (Article 57e) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/0df4a67b-03ac-4468-b6b5-6526237f92ba | https://echa.europa.eu/documents/10162/111dde32-185d-4e50-9112-5c1c019b920a | https://echa.europa.eu/documents/10162/4894dac8-608d-c1ab-2473-f67e2e314d39 | https://echa.europa.eu/documents/10162/b2524815-0623-1a5e-6bea-1f908ea9123f
Orange lead (lead tetroxide) |  | 215-235-6 | 1314-41-6 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/97af2122-f294-472c-9cc8-978f116c6929 | https://echa.europa.eu/documents/10162/aab36e2a-90e2-46b0-b940-07e4cf119f05 | https://echa.europa.eu/documents/10162/28065aee-33ca-5892-0d45-484e326f6700 | https://echa.europa.eu/documents/10162/9d2032aa-9839-6ab3-dc80-3b9227d496b0
o-toluidine |  | 202-429-0 | 95-53-4 | Carcinogenic (Article 57a) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/953517bb-4055-4a5b-9ed7-4bb69c3b96e7 | https://echa.europa.eu/documents/10162/c4e9db39-8285-48d9-87b3-29c15c861802 | https://echa.europa.eu/documents/10162/2963ade6-f093-4c8c-b9b2-604f538a3131 | https://echa.europa.eu/documents/10162/40cff30c-61e5-7020-5613-2c1960a359f0
o-aminoazotoluene |  | 202-591-2 | 97-56-3 | Carcinogenic (Article 57a) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/d7aaf254-b61a-48a2-9f03-9c49bd0164c1 | https://echa.europa.eu/documents/10162/d1a365c6-34d1-4fda-81c7-09c7f4b07a77 | https://echa.europa.eu/documents/10162/b94dedbc-8af4-4223-b68e-fcddccee3119 | https://echa.europa.eu/documents/10162/6b9044f3-d941-fb0d-096b-286a4dd77761
n-pentyl-isopentyl phthalate |  | - | 776297-69-9 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/d54fc5db-a2c6-4b1c-9000-e398645ad294 | https://echa.europa.eu/documents/10162/c33869fc-43f1-44c2-812f-0c6ad82104c6 | https://echa.europa.eu/documents/10162/73ef6184-8578-44dc-ac23-8034f1db5471 | https://echa.europa.eu/documents/10162/604ad38f-b652-c9af-2d99-88eba89b1809
N-methylacetamide |  | 201-182-6 | 79-16-3 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/2d6d4c98-5c55-41ff-89e0-6557aed1c003 | https://echa.europa.eu/documents/10162/cd5a3079-1069-4f2f-815c-a628714f76a7 | https://echa.europa.eu/documents/10162/df25f106-5e4c-4f4b-baed-fc2251e9d233 | https://echa.europa.eu/documents/10162/45f25c3e-6382-15e2-b6cc-d981bb3caa0f
N,N-dimethylformamide |  | 200-679-5 | 68-12-2 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/52c73750-c192-457c-9c48-2bd9776cec88 | https://echa.europa.eu/documents/10162/e3a1814f-6a8c-4c02-858e-74f98d525824 | https://echa.europa.eu/documents/10162/64875a92-e9c1-3d65-6722-0fab6c7283b3 | https://echa.europa.eu/documents/10162/f9d62a48-e855-fb93-e3dc-aa5187fabd7f
Methyloxirane (Propylene oxide) |  | 200-879-2 | 75-56-9 | Carcinogenic (Article 57a)#Mutagenic (Article 57b) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/005a9afb-2a53-41df-8711-5c28765c6e6b | https://echa.europa.eu/documents/10162/38636e85-49ef-4b37-890c-65ba488e7521 | https://echa.europa.eu/documents/10162/60dcbe57-a21b-4420-9e9a-33780223b72e | https://echa.europa.eu/documents/10162/0d327156-85d3-bcf9-7b7c-96e26eaa10c9
Methoxyacetic acid |  | 210-894-6 | 625-45-6 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/afb7a0e1-9222-4385-88cf-71c0b06191d0 | https://echa.europa.eu/documents/10162/2a9e3fcd-2243-4357-b837-9532235a307f | https://echa.europa.eu/documents/10162/0be0390a-a33a-d8f0-7aa5-5fee34deff20 | https://echa.europa.eu/documents/10162/280be7f8-c3f1-1589-ddaa-a203f59faad7
Lead titanium zirconium oxide |  | 235-727-4 | 12626-81-2 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/01818c6d-bcc9-4136-9eec-888823ae5fb0 | https://echa.europa.eu/documents/10162/b655b355-0ced-4795-a2d2-76c912c7acca | https://echa.europa.eu/documents/10162/319be01e-c34d-7d76-1b3a-7e463c763751 | https://echa.europa.eu/documents/10162/fedc8ac1-2c73-5364-2e71-4bde22939b55
Lead titanium trioxide |  | 235-038-9 | 12060-00-3 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/933f22c2-cfa6-4126-b0a5-3ba33ce75487 | https://echa.europa.eu/documents/10162/d69baa79-6ed9-4c3e-9aed-584577c19696 | https://echa.europa.eu/documents/10162/a1fbea0a-5718-38cc-c08a-66341ecb969e | https://echa.europa.eu/documents/10162/c5034009-0357-8d59-0128-23c99dddeb4b
Lead oxide sulfate |  | 234-853-7 | 12036-76-9 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/23bfbbe6-9985-42fe-b401-a7df81ec12cc | https://echa.europa.eu/documents/10162/50822d59-490c-4e01-88c3-c8ff8d2d389a | https://echa.europa.eu/documents/10162/0362e7d9-bcb5-41bc-945c-19365594480f | https://echa.europa.eu/documents/10162/58651835-2c5c-28bb-0653-9e13a2f98666
Lead monoxide (lead oxide) |  | 215-267-0 | 1317-36-8 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/ea71029e-f8de-4e5d-a393-6961aefd6e14 | https://echa.europa.eu/documents/10162/86c143a2-eb69-44bd-9710-73ab2aca57fc | https://echa.europa.eu/documents/10162/229a3cca-ceeb-101d-1247-2e824784f473 | https://echa.europa.eu/documents/10162/e456243a-65e4-5d07-6332-fd323cdd406e
Lead dinitrate |  | 233-245-9 | 10099-74-8 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/e1c55ba4-1f9a-4633-b4c3-fdaba005b095 | https://echa.europa.eu/documents/10162/97c17507-b6bc-42d7-8aac-59c0c0e9c21b | https://echa.europa.eu/documents/10162/be34d0ca-8d1e-488f-8b49-35063b0054b3 | https://echa.europa.eu/documents/10162/c3be9101-a646-b11c-7aaa-427d650ea242
Lead cyanamidate |  | 244-073-9 | 20837-86-9 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/a6f8212c-590f-469a-b6c1-f40f234e4518 | https://echa.europa.eu/documents/10162/0225706b-6dbf-4b6f-94f8-fa3e8fc24ff6 | https://echa.europa.eu/documents/10162/7f3131c3-b63f-49e8-9717-68dc7883b549 | https://echa.europa.eu/documents/10162/1f7fc65b-4cfc-ebc2-f9ac-73d043977553
Lead bis(tetrafluoroborate) |  | 237-486-0 | 13814-96-5 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/56f43ea4-a623-4d5e-b741-fee1fef7e771 | https://echa.europa.eu/documents/10162/0d9b7449-9ec1-4c1e-a117-294cccc4aa9d | https://echa.europa.eu/documents/10162/3d78e51f-bd96-fd06-4cd4-9d268d040478 | https://echa.europa.eu/documents/10162/5257f061-6e42-e233-84f3-1904b0529ce0
Hexahydromethylphthalic anhydride | including cis- and trans- stereo isomeric forms and all possible combinations of the isomers | - | - | Respiratory sensitising properties (Article 57(f) - human health) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/81052549-dd49-40b5-9388-b5a72a1a3fb3 | https://echa.europa.eu/documents/10162/bd282c4d-91b6-4537-9f27-b65a5fccbd48#https://echa.europa.eu/documents/10162/670b8e64-298e-4ed7-939f-67da04164a42#https://echa.europa.eu/documents/10162/7e96ce49-1e77-4333-8e0e-35892531aa5c#https://echa.europa.eu/documents/10162/081b759a-1aae-42d9-9f62-4051fb9187bc | https://echa.europa.eu/documents/10162/abbc92cc-984a-16e1-efce-37b665ecb79a | https://echa.europa.eu/documents/10162/723209fe-d033-bc5e-b307-9066ccf0426d
Hexahydro-1-methylphthalic anhydride |  | 256-356-4 | 48122-14-1 | Respiratory sensitising properties (Article 57(f) - human health) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/81052549-dd49-40b5-9388-b5a72a1a3fb3 | https://echa.europa.eu/documents/10162/bd282c4d-91b6-4537-9f27-b65a5fccbd48#https://echa.europa.eu/documents/10162/670b8e64-298e-4ed7-939f-67da04164a42#https://echa.europa.eu/documents/10162/7e96ce49-1e77-4333-8e0e-35892531aa5c#https://echa.europa.eu/documents/10162/081b759a-1aae-42d9-9f62-4051fb9187bc | https://echa.europa.eu/documents/10162/abbc92cc-984a-16e1-efce-37b665ecb79a | https://echa.europa.eu/documents/10162/723209fe-d033-bc5e-b307-9066ccf0426d
Hexahydromethylphthalic anhydride |  | 247-094-1 | 25550-51-0 | Respiratory sensitising properties (Article 57(f) - human health) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/81052549-dd49-40b5-9388-b5a72a1a3fb3 | https://echa.europa.eu/documents/10162/bd282c4d-91b6-4537-9f27-b65a5fccbd48#https://echa.europa.eu/documents/10162/670b8e64-298e-4ed7-939f-67da04164a42#https://echa.europa.eu/documents/10162/7e96ce49-1e77-4333-8e0e-35892531aa5c#https://echa.europa.eu/documents/10162/081b759a-1aae-42d9-9f62-4051fb9187bc | https://echa.europa.eu/documents/10162/abbc92cc-984a-16e1-efce-37b665ecb79a | https://echa.europa.eu/documents/10162/723209fe-d033-bc5e-b307-9066ccf0426d
Hexahydro-4-methylphthalic anhydride |  | 243-072-0 | 19438-60-9 | Respiratory sensitising properties (Article 57(f) - human health) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/81052549-dd49-40b5-9388-b5a72a1a3fb3 | https://echa.europa.eu/documents/10162/bd282c4d-91b6-4537-9f27-b65a5fccbd48#https://echa.europa.eu/documents/10162/670b8e64-298e-4ed7-939f-67da04164a42#https://echa.europa.eu/documents/10162/7e96ce49-1e77-4333-8e0e-35892531aa5c#https://echa.europa.eu/documents/10162/081b759a-1aae-42d9-9f62-4051fb9187bc | https://echa.europa.eu/documents/10162/abbc92cc-984a-16e1-efce-37b665ecb79a | https://echa.europa.eu/documents/10162/723209fe-d033-bc5e-b307-9066ccf0426d
Hexahydro-3-methylphthalic anhydride |  | 260-566-1 | 57110-29-9 | Respiratory sensitising properties (Article 57(f) - human health) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/81052549-dd49-40b5-9388-b5a72a1a3fb3 | https://echa.europa.eu/documents/10162/bd282c4d-91b6-4537-9f27-b65a5fccbd48#https://echa.europa.eu/documents/10162/670b8e64-298e-4ed7-939f-67da04164a42#https://echa.europa.eu/documents/10162/7e96ce49-1e77-4333-8e0e-35892531aa5c#https://echa.europa.eu/documents/10162/081b759a-1aae-42d9-9f62-4051fb9187bc | https://echa.europa.eu/documents/10162/abbc92cc-984a-16e1-efce-37b665ecb79a | https://echa.europa.eu/documents/10162/723209fe-d033-bc5e-b307-9066ccf0426d
Heptacosafluorotetradecanoic acid |  | 206-803-4 | 376-06-7 | vPvB (Article 57e) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/bb2caa51-1bcc-407b-b587-91d4ac23b564 | https://echa.europa.eu/documents/10162/de2ad408-ea06-484e-9015-ac6af581863c | https://echa.europa.eu/documents/10162/08046ee8-0933-a564-0df2-da48cda2bc6b | https://echa.europa.eu/documents/10162/76c2d4f3-1d2e-f263-c819-74f4373a976c
Henicosafluoroundecanoic acid |  | 218-165-4 | 2058-94-8 | vPvB (Article 57e) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/02881bec-75ad-4e83-a9bb-c9f6dc81042e | https://echa.europa.eu/documents/10162/60adb1fb-d001-4a2a-b8bf-a0e05e29d156 | https://echa.europa.eu/documents/10162/9bcff514-6625-baa9-edd9-80ee9e82c86a | https://echa.europa.eu/documents/10162/bd48b5f6-b407-60e1-f06e-76dd605d92ae
Furan |  | 203-727-3 | 110-00-9 | Carcinogenic (Article 57a) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/1986ed9a-a382-4409-8faf-abd0f9af440a | https://echa.europa.eu/documents/10162/f2a41a5f-68db-44e9-a925-2446de4ec9b1 | https://echa.europa.eu/documents/10162/8ab8b38b-26e8-49b7-9f04-b83c457023ec | https://echa.europa.eu/documents/10162/783a7398-0b23-d9ab-4afb-da6a77d2dfcb
Fatty acids, C16-18, lead salts |  | 292-966-7 | 91031-62-8 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/7f59b681-ee89-484b-bec2-3412f3f7d0e5 | https://echa.europa.eu/documents/10162/575f1c14-49f5-4a3d-9368-ebbb989edcf4 | https://echa.europa.eu/documents/10162/13e375d0-709b-4203-b851-7cb5e0a87c80 | https://echa.europa.eu/documents/10162/6e6bc1b6-4695-bfb6-8746-445a90dad5ff
Dioxobis(stearato)trilead |  | 235-702-8 | 12578-12-0 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b3332ac2-958d-4b79-86be-12a13f9cd8c8 | https://echa.europa.eu/documents/10162/2b533763-b36b-481c-a7b2-ff28b898d343 | https://echa.europa.eu/documents/10162/01088b8d-0195-44cf-8760-f1797c18e789 | https://echa.europa.eu/documents/10162/3ce25576-1795-cda5-ea46-acdb9741089a
Dinoseb (6-sec-butyl-2,4-dinitrophenol) |  | 201-861-7 | 88-85-7 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b23de2c1-7142-43ec-9fba-ee150ed98bb1 | https://echa.europa.eu/documents/10162/876747a8-aa98-42f4-bf6c-7ca2be3fed59 | https://echa.europa.eu/documents/10162/34ec55f2-27d2-47d6-94f1-d39acaf3634f | https://echa.europa.eu/documents/10162/f90845a4-dc67-9a63-a3cb-57787b8b97e3
Dimethyl sulphate |  | 201-058-1 | 77-78-1 | Carcinogenic (Article 57a) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/40a80d81-ac48-4ce6-891c-53b7a3779646 | https://echa.europa.eu/documents/10162/5182ed44-1e76-4374-898d-0e8cef94f7cb | https://echa.europa.eu/documents/10162/8fcaf1f0-ea6d-4ec8-a129-f0405def7c33 | https://echa.europa.eu/documents/10162/4e7ad358-f6e0-ea83-0681-94a9e414f29b
Diisopentyl phthalate |  | 210-088-4 | 605-50-5 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/917ec7f1-d611-4dad-8cf7-36c1ca07a400 | https://echa.europa.eu/documents/10162/cf490b69-0c28-453b-a531-e8c2ad32d22c | https://echa.europa.eu/documents/10162/34d906a5-2558-4faa-ba85-65a765aef016 | https://echa.europa.eu/documents/10162/5836c6ce-0043-af59-0755-2f8262791d64
Diethyl sulphate |  | 200-589-6 | 64-67-5 | Carcinogenic (Article 57a)#Mutagenic (Article 57b) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/3f6a4989-b155-4827-abfc-51e3fa7e352d | https://echa.europa.eu/documents/10162/296efee3-75ed-4ead-a79c-28a3f48d2d90 | https://echa.europa.eu/documents/10162/470b7f44-e538-40b0-9110-4575a5726d14 | https://echa.europa.eu/documents/10162/27b4f21d-3bb0-bde6-d9ba-4faadc87aaab
Dibutyltin dichloride (DBTC) |  | 211-670-0 | 683-18-1 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/046189d5-908b-447f-b91b-b3f97c9de5b5 | https://echa.europa.eu/documents/10162/f06ebc42-a92c-43a4-828a-a7c95ea2c961 | https://echa.europa.eu/documents/10162/2cbcb185-4207-1270-7189-6b9009db7b45 | https://echa.europa.eu/documents/10162/f92c8584-d0c2-ebb6-58a6-135b7f3d6f07
Diazene-1,2-dicarboxamide (C,C'-azodi(formamide)) (ADCA) |  | 204-650-8 | 123-77-3 | Respiratory sensitising properties (Article 57(f) - human health) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/4739d799-2888-4031-b3eb-73a9d475347b | https://echa.europa.eu/documents/10162/055d66a6-9924-4386-900c-62c52c42bc62 | https://echa.europa.eu/documents/10162/a8f3a173-0bb2-cb27-d9ae-fafe356c329c | https://echa.europa.eu/documents/10162/b9b09f2a-7230-5ddc-921e-9b0de3c6a956
Cyclohexane-1,2-dicarboxylic anhydride | all possible combinations of the cis- and trans-isomers | - | - | Respiratory sensitising properties (Article 57(f) - human health) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b7a93309-0b81-4dbf-90e0-859bc0f41ce4 | https://echa.europa.eu/documents/10162/542f4e26-c203-4587-a498-fbf33c9edeba#https://echa.europa.eu/documents/10162/d0d33d3f-8361-48a5-b5a9-5db91928b4c5#https://echa.europa.eu/documents/10162/dca0c593-37d7-4257-9ad9-14059e525771 | https://echa.europa.eu/documents/10162/555606e7-97e2-592e-c100-03f5639a82f4 | https://echa.europa.eu/documents/10162/acfa14cb-0ed3-5a82-5366-964eeaab14cb
trans-cyclohexane-1,2-dicarboxylic anhydride |  | 238-009-9 | 14166-21-3 | Respiratory sensitising properties (Article 57(f) - human health) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b7a93309-0b81-4dbf-90e0-859bc0f41ce4 | https://echa.europa.eu/documents/10162/542f4e26-c203-4587-a498-fbf33c9edeba#https://echa.europa.eu/documents/10162/d0d33d3f-8361-48a5-b5a9-5db91928b4c5#https://echa.europa.eu/documents/10162/dca0c593-37d7-4257-9ad9-14059e525771 | https://echa.europa.eu/documents/10162/555606e7-97e2-592e-c100-03f5639a82f4 | https://echa.europa.eu/documents/10162/acfa14cb-0ed3-5a82-5366-964eeaab14cb
Cyclohexane-1,2-dicarboxylic anhydride |  | 201-604-9 | 85-42-7 | Respiratory sensitising properties (Article 57(f) - human health) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b7a93309-0b81-4dbf-90e0-859bc0f41ce4 | https://echa.europa.eu/documents/10162/542f4e26-c203-4587-a498-fbf33c9edeba#https://echa.europa.eu/documents/10162/d0d33d3f-8361-48a5-b5a9-5db91928b4c5#https://echa.europa.eu/documents/10162/dca0c593-37d7-4257-9ad9-14059e525771 | https://echa.europa.eu/documents/10162/555606e7-97e2-592e-c100-03f5639a82f4 | https://echa.europa.eu/documents/10162/acfa14cb-0ed3-5a82-5366-964eeaab14cb
cis-cyclohexane-1,2-dicarboxylic anhydride |  | 236-086-3 | 13149-00-3 | Respiratory sensitising properties (Article 57(f) - human health) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b7a93309-0b81-4dbf-90e0-859bc0f41ce4 | https://echa.europa.eu/documents/10162/542f4e26-c203-4587-a498-fbf33c9edeba#https://echa.europa.eu/documents/10162/d0d33d3f-8361-48a5-b5a9-5db91928b4c5#https://echa.europa.eu/documents/10162/dca0c593-37d7-4257-9ad9-14059e525771 | https://echa.europa.eu/documents/10162/555606e7-97e2-592e-c100-03f5639a82f4 | https://echa.europa.eu/documents/10162/acfa14cb-0ed3-5a82-5366-964eeaab14cb
Bis(pentabromophenyl) ether (decabromodiphenyl ether) (DecaBDE) |  | 214-604-9 | 1163-19-5 | PBT (Article 57d)#vPvB (Article 57e) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/7de8998e-c87a-4fa6-82fa-927fffb0fe79 | https://echa.europa.eu/documents/10162/64b5b9c4-c549-4254-a917-0ff1b7c9fa1b | https://echa.europa.eu/documents/10162/046e853b-03a0-99e4-a682-f5842bef9a1c | https://echa.europa.eu/documents/10162/f919b5a9-8dac-e913-a862-a87cb5e00787
Biphenyl-4-ylamine |  | 202-177-1 | 92-67-1 | Carcinogenic (Article 57a) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/cefb8a85-ed61-4751-a722-2b3510d5f7a0 | https://echa.europa.eu/documents/10162/22458f36-2a5f-4c1c-83b8-b8a9c45bda95 | https://echa.europa.eu/documents/10162/9d518786-9c9a-4d61-b841-45ff84ed54cc | https://echa.europa.eu/documents/10162/1520b492-61af-d235-ffb1-532532fd8c5b
Acetic acid, lead salt, basic |  | 257-175-3 | 51404-69-4 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/a68aa6d7-bc05-4f79-9445-4dd0ecabc1bb | https://echa.europa.eu/documents/10162/5848ac0e-7286-4ec4-9185-49042cf54408 | https://echa.europa.eu/documents/10162/e27fbdcb-b252-4912-8965-a5bb89711203 | https://echa.europa.eu/documents/10162/32f33799-9835-71d0-b0d2-be45c318decc
[Phthalato(2-)]dioxotrilead |  | 273-688-5 | 69011-06-9 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/490c7cfa-ae94-48a0-83c9-99bc3c0a5e13 | https://echa.europa.eu/documents/10162/430ee50c-c98e-4e96-a689-d556fdd06665 | https://echa.europa.eu/documents/10162/e91058f4-7ab4-4648-91e3-3220f7ee6e8c | https://echa.europa.eu/documents/10162/8b848bd8-4151-2f61-c674-7da51d88aa8f
6-methoxy-m-toluidine (p-cresidine) |  | 204-419-1 | 120-71-8 | Carcinogenic (Article 57a) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/251a17a0-cc98-4151-9b3f-3526c667a90e | https://echa.europa.eu/documents/10162/5c6a3acb-99f9-4561-a0a5-47b33f13c1f0 | https://echa.europa.eu/documents/10162/9ee4bdb4-62a8-4ff5-a4f0-f0ced371f88f | https://echa.europa.eu/documents/10162/691261ea-6df2-8037-1230-7bdc04b1494c
4-Nonylphenol, branched and linear | substances with a linear and/or branched alkyl chain with a carbon number of 9 covalently bound in position 4 to phenol, covering also UVCB- and well-defined substances which include any of the individual isomers or a combination thereof | - | - | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
4-(1-ethyl-1-methylhexyl)phenol |  | 257-907-1 | 52427-13-1 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
p-(1,1-dimethylheptyl)phenol |  | 250-339-5 | 30784-30-6 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
p-(1-methyloctyl)phenol |  | 241-427-4 | 17404-66-9 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
p-isononylphenol |  | 247-770-6 | 26543-97-5 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
p-nonylphenol |  | 203-199-4 | 104-40-5 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
4-(1-Ethyl-1,3-dimethylpentyl)phenol |  | - | 186825-36-5 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
4-(1-Ethyl-1,4-dimethylpentyl)phenol |  | - | 142731-63-3 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
Phenol, 4-nonyl-, branched |  | 284-325-5 | 84852-15-3 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
4-(3-ethylheptan-2-yl)phenol |  | - | 186825-39-8 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
Isononylphenol |  | 234-284-4 | 11066-49-2 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
Nonylphenol |  | 246-672-0 | 25154-52-3 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
4-(1,1,5-Trimethylhexyl)phenol |  | - | 521947-27-3 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
Phenol, nonyl-, branched |  | 291-844-0 | 90481-04-2 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/dea74d46-dc8e-4b10-947b-51a19d890153 | https://echa.europa.eu/documents/10162/80751a04-fdb3-4bef-aabe-cc49da2a1148 | https://echa.europa.eu/documents/10162/bc140e0b-b407-fd1c-f750-6d43c99f82a4#https://echa.europa.eu/documents/10162/3b60b4bc-0376-a9ea-bffd-65da8e2c9602 | https://echa.europa.eu/documents/10162/1984ccdb-3321-2890-c598-5735786e5eb2
4-methyl-m-phenylenediamine (toluene-2,4-diamine) |  | 202-453-1 | 95-80-7 | Carcinogenic (Article 57a) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/e7bb9247-f300-4887-9565-574dad453b71 | https://echa.europa.eu/documents/10162/0ffceb1b-180c-4965-a526-0b2d4b73c675 | https://echa.europa.eu/documents/10162/98b66e79-3f52-4792-aa74-1e9a1c674db4 | https://echa.europa.eu/documents/10162/b73979f8-961d-dc60-232e-2a8b4626e547
4-aminoazobenzene |  | 200-453-6 | 60-09-3 | Carcinogenic (Article 57a) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/9cf3cd45-97d2-455a-a58b-86d721fb4f54 | https://echa.europa.eu/documents/10162/4740aca7-2e63-4f48-adf0-8c98bb12244c | https://echa.europa.eu/documents/10162/bd560d2d-2ea1-4e93-950f-cca356500dbb | https://echa.europa.eu/documents/10162/1ec37243-f332-a43a-fc36-9aac9cbd485a
4-(1,1,3,3-tetramethylbutyl)phenol, ethoxylated | covering well-defined substances and UVCB substances, polymers and homologues | - | - | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b823f588-7c35-4bad-9e3c-b639b627df8f | https://echa.europa.eu/documents/10162/f6ae4ec7-0a73-45eb-bafb-43e37bd3b42e | https://echa.europa.eu/documents/10162/430c2613-588f-8b08-8a72-df4013727ef8 | https://echa.europa.eu/documents/10162/e3830756-209a-a21c-95ea-fad990b6748a
2-[2-[4-(1,1,3,3-tetramethylbutyl)phenoxy]ethoxy]ethanol |  | - | 2315-61-9 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b823f588-7c35-4bad-9e3c-b639b627df8f | https://echa.europa.eu/documents/10162/f6ae4ec7-0a73-45eb-bafb-43e37bd3b42e | https://echa.europa.eu/documents/10162/430c2613-588f-8b08-8a72-df4013727ef8 | https://echa.europa.eu/documents/10162/e3830756-209a-a21c-95ea-fad990b6748a
2-[4-(1,1,3,3-tetramethylbutyl)phenoxy]ethanol |  | - | 2315-67-5 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b823f588-7c35-4bad-9e3c-b639b627df8f | https://echa.europa.eu/documents/10162/f6ae4ec7-0a73-45eb-bafb-43e37bd3b42e | https://echa.europa.eu/documents/10162/430c2613-588f-8b08-8a72-df4013727ef8 | https://echa.europa.eu/documents/10162/e3830756-209a-a21c-95ea-fad990b6748a
20-[4-(1,1,3,3-tetramethylbutyl)phenoxy]-3,6,9,12,15,18-hexaoxaicosan-1-ol |  | 219-682-8 | 2497-59-8 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b823f588-7c35-4bad-9e3c-b639b627df8f | https://echa.europa.eu/documents/10162/f6ae4ec7-0a73-45eb-bafb-43e37bd3b42e | https://echa.europa.eu/documents/10162/430c2613-588f-8b08-8a72-df4013727ef8 | https://echa.europa.eu/documents/10162/e3830756-209a-a21c-95ea-fad990b6748a
Polyethylene glycol p-(1,1,3,3-tetramethylbutyl)phenyl ether |  | - | 9002-93-1 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b823f588-7c35-4bad-9e3c-b639b627df8f | https://echa.europa.eu/documents/10162/f6ae4ec7-0a73-45eb-bafb-43e37bd3b42e | https://echa.europa.eu/documents/10162/430c2613-588f-8b08-8a72-df4013727ef8 | https://echa.europa.eu/documents/10162/e3830756-209a-a21c-95ea-fad990b6748a
Poly(oxy-1,2-ethanediyl), α-[(1,1,3,3-tetramethylbutyl)phenyl]-ω-hydroxy- |  | - | 9036-19-5 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/b823f588-7c35-4bad-9e3c-b639b627df8f | https://echa.europa.eu/documents/10162/f6ae4ec7-0a73-45eb-bafb-43e37bd3b42e | https://echa.europa.eu/documents/10162/430c2613-588f-8b08-8a72-df4013727ef8 | https://echa.europa.eu/documents/10162/e3830756-209a-a21c-95ea-fad990b6748a
4,4'-oxydianiline and its salts |  | - | - | Carcinogenic (Article 57a)#Mutagenic (Article 57b) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/7e13f20e-aa4c-4ab9-9b57-f0bd3cf750c9 | https://echa.europa.eu/documents/10162/390a50ab-4461-49fb-87da-b366f77c105b | https://echa.europa.eu/documents/10162/80464fce-e4ae-4e28-b6a0-8e332daaeb95 | https://echa.europa.eu/documents/10162/2b197fb5-9330-3567-b9af-f10f7ef30cc6
4,4'-oxydianiline |  | 202-977-0 | 101-80-4 | Carcinogenic (Article 57a)#Mutagenic (Article 57b) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/7e13f20e-aa4c-4ab9-9b57-f0bd3cf750c9 | https://echa.europa.eu/documents/10162/390a50ab-4461-49fb-87da-b366f77c105b | https://echa.europa.eu/documents/10162/80464fce-e4ae-4e28-b6a0-8e332daaeb95 | https://echa.europa.eu/documents/10162/2b197fb5-9330-3567-b9af-f10f7ef30cc6
4,4'-methylenedi-o-toluidine |  | 212-658-8 | 838-88-0 | Carcinogenic (Article 57a) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/aa8680a2-aecb-4b86-b4e9-e6d6ba41ed42 | https://echa.europa.eu/documents/10162/7fabc585-9a8c-4105-99ad-4083a41b2310 | https://echa.europa.eu/documents/10162/42094854-61af-4a51-abc4-0f85ef9bd51e | https://echa.europa.eu/documents/10162/a16f7d17-5558-227d-65bc-d628d6209328
3-ethyl-2-methyl-2-(3-methylbutyl)-1,3-oxazolidine | ZOLDINE MS-PLUS | 421-150-7 | 143860-04-2 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/2e6fa094-c665-4367-b4d6-d61b841e82c3 | https://echa.europa.eu/documents/10162/c2ba8775-69b2-4dbe-b725-54e8219f10ae | https://echa.europa.eu/documents/10162/782e61c7-07e8-42be-bfb7-99c2bdab913a | https://echa.europa.eu/documents/10162/45e26414-e789-45a9-302a-4ce0ddef9981
1-bromopropane (n-propyl bromide) |  | 203-445-0 | 106-94-5 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/322977d5-5c50-467b-a5aa-14dd621301af | https://echa.europa.eu/documents/10162/0fde3877-5c84-4343-ad20-a3327752a13a | https://echa.europa.eu/documents/10162/cfa3375a-66c2-02a7-43c1-3760fb518448 | https://echa.europa.eu/documents/10162/fb68a467-fa8f-3ea3-3735-3dde15221666
1,2-diethoxyethane |  | 211-076-1 | 629-14-1 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/c5d3b97f-9168-4c49-aa69-9197cd899a8d | https://echa.europa.eu/documents/10162/8393ce5a-9456-4abb-b6ba-008e68941823 | https://echa.europa.eu/documents/10162/8c04401d-d0bb-409c-9250-c6e574f47305 | https://echa.europa.eu/documents/10162/53a23e4f-6934-d9d1-a3cb-3d8bcffbf224
1,2-Benzenedicarboxylic acid, dipentyl ester, branched and linear |  | 284-032-2 | 84777-06-0 | Toxic for reproduction (Article 57c) | 19-Dec-2012 | https://echa.europa.eu/documents/10162/cdf66098-5920-44ea-9d26-aaa390c5d5ef | https://echa.europa.eu/documents/10162/6f9bf9c8-8442-48fe-b38e-4dca50adea4c | https://echa.europa.eu/documents/10162/329c8367-0ec9-426b-bcfe-27ddf27fb777 | https://echa.europa.eu/documents/10162/28bc2e86-e6e2-569b-45ab-724e664a3b24
α,α-Bis[4-(dimethylamino)phenyl]-4 (phenylamino)naphthalene-1-methanol (C.I. Solvent Blue 4) | with ≥ 0.1% of Michler's ketone (EC No. 202-027-5) or Michler's base (EC No. 202-959-2) | 229-851-8 | 6786-83-0 | Carcinogenic (Article 57a) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/6ea9ffc1-a165-4afc-b504-af2bc3694c64 | https://echa.europa.eu/documents/10162/09c5c1e5-3d1f-4c79-8d8a-f142e7dc42ab | https://echa.europa.eu/documents/10162/d3bc3564-1f27-4b6b-b076-695ee0a3ad10 | https://echa.europa.eu/documents/10162/15746c0a-be4d-7014-cf24-9b893fc46146
N,N,N',N'-tetramethyl-4,4'-methylenedianiline (Michler’s base) |  | 202-959-2 | 101-61-1 | Carcinogenic (Article 57a) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/55c7a0fd-13f8-41bd-b1b2-123cdeef4434 | https://echa.europa.eu/documents/10162/bb71dff3-cbb8-42e4-8d18-529e5876b710 | https://echa.europa.eu/documents/10162/14f93f87-28aa-4295-a2a6-6b18e53abbf2 | https://echa.europa.eu/documents/10162/39e9a04b-b21f-83a5-c892-0922c59f6570
Lead(II) bis(methanesulfonate) | - | 401-750-5 | 17570-76-2 | Toxic for reproduction (Article 57c) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/670a5290-68d3-4b88-83e8-6c06a4e430a8 | https://echa.europa.eu/documents/10162/71993287-68f1-43c4-a13e-326ad171b80c | https://echa.europa.eu/documents/10162/ec18fb5f-b968-4e34-a146-e696fccad746 | https://echa.europa.eu/documents/10162/2cba9d62-81b9-7a72-e33f-4d0674fd7aff
Formamide |  | 200-842-0 | 75-12-7 | Toxic for reproduction (Article 57c) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/64c72b55-a1da-47ef-b47d-edad9c9f124e | https://echa.europa.eu/documents/10162/a330365d-7a44-43a1-a896-0fec7f34c4e6 | https://echa.europa.eu/documents/10162/4ff18575-7bcc-4e02-96db-1090a1bdd3a1 | https://echa.europa.eu/documents/10162/ed0578c2-0153-12b5-fe8f-b9e1b2923a40
Diboron trioxide |  | 215-125-8 | 1303-86-2 | Toxic for reproduction (Article 57c) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/c42c7cd8-b8d9-4a53-987e-55f8b377ea03 | https://echa.europa.eu/documents/10162/de509422-112a-4234-8e67-3260d2051364 | https://echa.europa.eu/documents/10162/6b35228c-45c4-4ed8-88b0-823fafb0d795 | https://echa.europa.eu/documents/10162/07417d9a-29c2-61e2-135b-2084d2c8d350
[4-[[4-anilino-1-naphthyl][4-(dimethylamino)phenyl]methylene]cyclohexa-2,5-dien-1-ylidene] dimethylammonium chloride (C.I. Basic Blue 26) | with ≥ 0.1% of Michler's ketone (EC No. 202-027-5) or Michler's base (EC No. 202-959-2) | 219-943-6 | 2580-56-5 | Carcinogenic (Article 57a) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/4575515a-f67b-447e-afb0-9de9bdf2d8b4 | https://echa.europa.eu/documents/10162/e788094b-c091-40fc-af06-0cb993da22ef | https://echa.europa.eu/documents/10162/4005564c-2ff2-4ea5-941d-04525939d70b | https://echa.europa.eu/documents/10162/03e30fbb-e4b7-4ff1-8c54-cfba9e21d505
[4-[4,4'-bis(dimethylamino) benzhydrylidene]cyclohexa-2,5-dien-1-ylidene]dimethylammonium chloride (C.I. Basic Violet 3) | with ≥ 0.1% of Michler's ketone (EC No. 202-027-5) or Michler's base (EC No. 202-959-2) | 208-953-6 | 548-62-9 | Carcinogenic (Article 57a) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/9c2106dd-ab9a-415d-8b47-039c5b5b223a | https://echa.europa.eu/documents/10162/2b47d136-adc9-46de-88c8-4ac990880005 | https://echa.europa.eu/documents/10162/a3290dc9-a3e1-4a9d-8e58-102dca785d19 | https://echa.europa.eu/documents/10162/67b63e21-6302-ef0d-b8b9-fa3c26a9e2bf
4,4'-bis(dimethylamino)benzophenone (Michler’s ketone) |  | 202-027-5 | 90-94-8 | Carcinogenic (Article 57a) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/fb031b0c-5e13-438a-a0ad-334a822c2e04 | https://echa.europa.eu/documents/10162/a32d3dd2-1045-457b-9a6b-988bc641e662 | https://echa.europa.eu/documents/10162/3cbcf0fc-f6bc-4d08-8dcf-b2837657e1c5 | https://echa.europa.eu/documents/10162/f793511d-97cb-70a6-6fa7-ae48395ce858
4,4'-bis(dimethylamino)-4''-(methylamino)trityl alcohol | with ≥ 0.1% of Michler's ketone (EC No. 202-027-5) or Michler's base (EC No. 202-959-2) | 209-218-2 | 561-41-1 | Carcinogenic (Article 57a) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/bcf29e5e-c152-4f88-8df9-1f9515582e3e | https://echa.europa.eu/documents/10162/0bc3c72f-1e50-45e7-98c4-595639fe3010 | https://echa.europa.eu/documents/10162/4300b1a0-79ae-4d90-bcac-07a34691c435 | https://echa.europa.eu/documents/10162/3905164d-a630-516a-720a-ebaf0b47d0c1
1,3,5-tris[(2S and 2R)-2,3-epoxypropyl]-1,3,5-triazine-2,4,6-(1H,3H,5H)-trione (β-TGIC) | - | 423-400-0 | 59653-74-6 | Mutagenic (Article 57b) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/b396da4a-879d-4d97-88c0-f2a97c3160e2 | https://echa.europa.eu/documents/10162/1b1db3da-9d3c-4dd9-aec6-82e5f2bb9681 | https://echa.europa.eu/documents/10162/9c7289c0-b92d-46cb-9007-6a919f003935 | https://echa.europa.eu/documents/10162/94969b35-04d7-5cc9-d1b0-bd120223a3c2
1,3,5-Tris(oxiran-2-ylmethyl)-1,3,5-triazinane-2,4,6-trione (TGIC) |  | 219-514-3 | 2451-62-9 | Mutagenic (Article 57b) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/f4f2cac7-9277-4dd0-bcfc-a0b3134331a3 | https://echa.europa.eu/documents/10162/5e64cc7a-5599-48af-86ea-4cd5d853cb9a | https://echa.europa.eu/documents/10162/b0f7fcea-3dd5-4281-ab06-fd46c48d2ec2 | https://echa.europa.eu/documents/10162/41e3ed1d-b112-e10c-31e1-11fd05c61743
1,2-bis(2-methoxyethoxy)ethane (TEGDME; triglyme) |  | 203-977-3 | 112-49-2 | Toxic for reproduction (Article 57c) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/1797f7da-12f2-4db5-95ee-d8b2f119255b | https://echa.europa.eu/documents/10162/077454da-571b-43dd-98c6-ce04427310e9 | https://echa.europa.eu/documents/10162/04a4b3ea-09d6-439a-bf05-079784355127 | https://echa.europa.eu/documents/10162/08fead86-7723-14f9-4c68-fad56f67e321
1, 2-dimethoxyethane; ethylene glycol dimethyl ether (EGDME) |  | 203-794-9 | 110-71-4 | Toxic for reproduction (Article 57c) | 18-Jun-2012 | https://echa.europa.eu/documents/10162/8baf4265-833f-467d-bb3d-cf3dfe92be88 | https://echa.europa.eu/documents/10162/964d2219-db3d-4fd4-8fb4-5264b7cea5ab | https://echa.europa.eu/documents/10162/466dce51-6adf-4936-874b-2c1fe8c9de5e | https://echa.europa.eu/documents/10162/1d689ed5-8d41-62e9-16c9-6d687bf28377
Zirconia Aluminosilicate Refractory Ceramic Fibres | are fibres covered by index number 650-017-00-8 in Annex VI, part 3, table 3.1 of Regulation (EC) No 1272/2008 of the European Parliament and of the Council of 16 December 2008 on classification, labelling and packaging of substances and mixtures, and fulfil the three following conditions: a) oxides of aluminium, silicon and zirconium are the main components present (in the fibres) within variable concentration ranges b) fibres have a length weighted geometric mean diameter less two standard geometric errors of 6 or less micrometres (µm). c) alkaline oxide and alkali earth oxide (Na2O+K2O+CaO+MgO+BaO) content less or equal to 18% by weight | - | - | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/0d43d978-af14-4acf-9613-95d028fef958#https://echa.europa.eu/documents/10162/46da04ed-c0ac-482e-aa63-c75553312e84 | https://echa.europa.eu/documents/10162/a9f47855-ffbe-4312-84f9-8d74aa409f00 | https://echa.europa.eu/documents/10162/8190e8e4-e315-4bcb-91f5-1548fbe46bc1 | https://echa.europa.eu/documents/10162/93002806-25a6-a48a-adf8-1626177e6722
Refractories, fibers, aluminosilicate |  | - | 142844-00-6 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/0d43d978-af14-4acf-9613-95d028fef958#https://echa.europa.eu/documents/10162/46da04ed-c0ac-482e-aa63-c75553312e84 | https://echa.europa.eu/documents/10162/a9f47855-ffbe-4312-84f9-8d74aa409f00 | https://echa.europa.eu/documents/10162/8190e8e4-e315-4bcb-91f5-1548fbe46bc1 | https://echa.europa.eu/documents/10162/93002806-25a6-a48a-adf8-1626177e6722
Trilead diarsenate |  | 222-979-5 | 3687-31-8 | Carcinogenic (Article 57a)#Toxic for reproduction (Article 57c) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/f94e0b50-062a-4150-8652-f5d3ac218fcd | https://echa.europa.eu/documents/10162/18ce9105-2407-45b5-852b-5329e6042900 | https://echa.europa.eu/documents/10162/69699398-ddf0-4188-988f-a2464175a019 | https://echa.europa.eu/documents/10162/7fe7b854-e3c4-0da8-4611-ba59657eb361
Potassium hydroxyoctaoxodizincatedichromate |  | 234-329-8 | 11103-86-9 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/cbc1bcf1-2910-4813-a818-9f2069f69b3c | https://echa.europa.eu/documents/10162/2869cfe1-049b-43b2-a1b0-c8503d2c2586 | https://echa.europa.eu/documents/10162/fb533ed0-2765-4ad3-acf7-79d9ca8a5b02 | https://echa.europa.eu/documents/10162/82f1b187-86e9-a2c1-05ce-37037c0d42c8
Phenolphthalein |  | 201-004-7 | 77-09-8 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/37893eb4-a79b-4982-a3f0-34438dc72904 | https://echa.europa.eu/documents/10162/2cc034c5-ecbf-4c6c-8de4-9085c3e4f2a0 | https://echa.europa.eu/documents/10162/1f29b212-4a70-40c0-94de-8dfce7791aa9 | https://echa.europa.eu/documents/10162/e77f71f2-2587-e6a6-3e18-41c44ad031cf
Pentazinc chromate octahydroxide |  | 256-418-0 | 49663-84-5 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/756bf74c-275b-4779-8d92-83fc3ac27318 | https://echa.europa.eu/documents/10162/3acce581-104b-4f80-8b0b-94728b9d501c | https://echa.europa.eu/documents/10162/ce28c0be-379f-461e-9197-3af36d5c7f53 | https://echa.europa.eu/documents/10162/15a09341-5a63-825f-9677-995a7685a951
N,N-dimethylacetamide |  | 204-826-4 | 127-19-5 | Toxic for reproduction (Article 57c) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/65231f68-fe7a-4712-9f3a-99a9fdfb7dea | https://echa.europa.eu/documents/10162/f3d28fdc-113a-454a-9de2-e5fa2bb7786d | https://echa.europa.eu/documents/10162/42751bc8-cad5-4035-82ea-a91c0efe5768 | https://echa.europa.eu/documents/10162/2349dc58-0d76-f548-d77f-a2d40eab4bb3
Lead styphnate |  | 239-290-0 | 15245-44-0 | Toxic for reproduction (Article 57c) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/9f6e771e-a50c-4079-9a1a-1d2441ed0033 | https://echa.europa.eu/documents/10162/940e8b74-d5c3-446e-a041-b86297a880c6 | https://echa.europa.eu/documents/10162/fbb1d8a9-4e09-4c36-8e79-c46f9a73bb7c | https://echa.europa.eu/documents/10162/1518f0d7-21c6-797c-1e0f-99db6a1c573c
Lead dipicrate |  | 229-335-2 | 6477-64-1 | Toxic for reproduction (Article 57c) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/3a01a0ae-69e4-4b0c-8d87-e5c42e8586a3 | https://echa.europa.eu/documents/10162/0a46d3e7-e3be-46bc-886b-d04e944ad448 | https://echa.europa.eu/documents/10162/8a7c6915-0578-4c50-8fea-5208e237e0a7 | https://echa.europa.eu/documents/10162/047d6d28-1493-63ea-5c06-69f94dda35dd
Lead diazide, Lead azide |  | 236-542-1 | 13424-46-9 | Toxic for reproduction (Article 57c) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/defeb72b-e314-4707-b01d-dd25464eb5f1 | https://echa.europa.eu/documents/10162/cacccb51-ab67-4ee5-803f-f147d01feddc | https://echa.europa.eu/documents/10162/9890a34b-0a57-4286-a2ab-b226ba3dfef3 | https://echa.europa.eu/documents/10162/72716893-7cd5-259b-b02e-27a8251b3d16
Formaldehyde, oligomeric reaction products with aniline |  | 500-036-1 | 25214-70-4 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/f5ad72aa-57b0-456b-acca-080ea744a5d1 | https://echa.europa.eu/documents/10162/3ccc2258-f2d9-4330-b218-b781178c00a2 | https://echa.europa.eu/documents/10162/9cd084c9-8c83-4cf4-81c9-f43d5c1f1fe8 | https://echa.europa.eu/documents/10162/9ccd6638-ecf1-3cd8-1f64-f7b0c6e931a4
Dichromium tris(chromate) |  | 246-356-2 | 24613-89-6 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/b5600741-d5d7-413c-9755-625a596d1ee1 | https://echa.europa.eu/documents/10162/3d86d90e-9593-4987-9856-639068f57402 | https://echa.europa.eu/documents/10162/6426d317-2c05-40ba-90ae-0fbd67774f53 | https://echa.europa.eu/documents/10162/2e3dc8f8-6055-183f-fc46-f86b85b089b6
Calcium arsenate |  | 231-904-5 | 7778-44-1 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/cb4e87e7-ebdc-4a02-8333-7b30084a9be4 | https://echa.europa.eu/documents/10162/00c3ba78-f33e-461f-9fdd-4c365f2b5b23 | https://echa.europa.eu/documents/10162/ccc629ba-7feb-462d-bce8-86ac200de8f5 | https://echa.europa.eu/documents/10162/79ea9462-05c3-e0bd-c4ca-5fe289ab6a86
Bis(2-methoxyethyl) phthalate |  | 204-212-6 | 117-82-8 | Toxic for reproduction (Article 57c) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/18fc6f65-fb18-4151-b737-45ba0aefb181 | https://echa.europa.eu/documents/10162/5d0a79ed-4557-4a6c-8017-fda04a4eb326 | https://echa.europa.eu/documents/10162/bee2930b-03bb-4293-bda1-89381f8da2a6 | https://echa.europa.eu/documents/10162/d4c509f3-22a8-ae10-60b3-84c2b574e4e2
Bis(2-methoxyethyl) ether |  | 203-924-4 | 111-96-6 | Toxic for reproduction (Article 57c) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/ee023359-daa8-43a2-8c82-242f0a7588f7 | https://echa.europa.eu/documents/10162/169db6b2-a42d-46f3-8daf-a7d8385e84b4 | https://echa.europa.eu/documents/10162/2e6c10fa-03d5-47b2-9337-e047743554b9 | https://echa.europa.eu/documents/10162/da73be84-553c-2cdb-6c20-6f329649af7c
Arsenic acid |  | 231-901-9 | 7778-39-4 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/3646fac6-947b-41ea-b3bf-e231511be65c | https://echa.europa.eu/documents/10162/c61d2694-b4f4-41c0-93da-68e959341b9c | https://echa.europa.eu/documents/10162/bc77d03b-2f02-4084-8fab-2b8d586f342a | https://echa.europa.eu/documents/10162/704cd872-1329-10b8-7326-3d93eea4c147
Aluminosilicate Refractory Ceramic Fibres | are fibres covered by index number 650-017-00-8 in Annex VI, part 3, table 3.1 of Regulation (EC) No 1272/2008 of the European Parliament and of the Council of 16 December 2008 on classification, labelling and packaging of substances and mixtures, and fulfil the three following conditions: a) oxides of aluminium and silicon are the main components present (in the fibres) within variable concentration ranges b) fibres have a length weighted geometric mean diameter less two standard geometric errors of 6 or less micrometres (µm) c) alkaline oxide and alkali earth oxide (Na2O+K2O+CaO+MgO+BaO) content less or equal to 18% by weight | - | - | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/df7f314c-bda4-496c-8d86-04fe2421e157#https://echa.europa.eu/documents/10162/9c0854f5-3f46-4dc5-9776-72ac20f908e2 | https://echa.europa.eu/documents/10162/71a6c340-4c9c-41d9-8c26-6e8d3236e456 | https://echa.europa.eu/documents/10162/611cb779-7564-4a89-b381-c5a126b5a04b | https://echa.europa.eu/documents/10162/c7882d5b-b6f3-94ca-c81a-756cf2d5c130
Refractories, fibers, aluminosilicate |  | - | 142844-00-6 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/df7f314c-bda4-496c-8d86-04fe2421e157#https://echa.europa.eu/documents/10162/9c0854f5-3f46-4dc5-9776-72ac20f908e2 | https://echa.europa.eu/documents/10162/71a6c340-4c9c-41d9-8c26-6e8d3236e456 | https://echa.europa.eu/documents/10162/611cb779-7564-4a89-b381-c5a126b5a04b | https://echa.europa.eu/documents/10162/c7882d5b-b6f3-94ca-c81a-756cf2d5c130
4-(1,1,3,3-tetramethylbutyl)phenol |  | 205-426-2 | 140-66-9 | Endocrine disrupting properties (Article 57(f) - environment) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/189c55b4-d54b-4271-9e50-ce798faf9f4c | https://echa.europa.eu/documents/10162/4c196bd6-2a91-4cb1-82ae-c65cb8a637ac | https://echa.europa.eu/documents/10162/4c6cccfd-d366-4a00-87e5-65aa77181fb6 | https://echa.europa.eu/documents/10162/a08e8cda-06ee-de39-c975-5f36973dbb63
2-Methoxyaniline, o-Anisidine |  | 201-963-1 | 90-04-0 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/30d99b89-85a5-45d1-8701-1ad0680afb1e | https://echa.europa.eu/documents/10162/80f2ea14-b438-4ae6-bbcd-b6b214c7ae04 | https://echa.europa.eu/documents/10162/faf9ec9a-af59-4c1d-b944-cc831b3b52b3 | https://echa.europa.eu/documents/10162/1b56118d-d4af-2d06-d15c-6e2354ff4990
2,2'-dichloro-4,4'-methylenedianiline |  | 202-918-9 | 101-14-4 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/0a89c9f0-1da4-406b-8285-9db907cdb66a | https://echa.europa.eu/documents/10162/40909414-f320-4c4d-8519-3b9895c794e1 | https://echa.europa.eu/documents/10162/6f0cee92-fe4f-492d-b8e1-78ac46ee1a85 | https://echa.europa.eu/documents/10162/883f8b33-697f-5708-dbec-fa3e4979fa57
1,2-dichloroethane |  | 203-458-1 | 107-06-2 | Carcinogenic (Article 57a) | 19-Dec-2011 | https://echa.europa.eu/documents/10162/55390ac1-1267-4141-af33-67c787b50a0d | https://echa.europa.eu/documents/10162/9e0c9155-9ec9-489a-976d-4312ab5838cf | https://echa.europa.eu/documents/10162/a40ec08f-918f-4f60-b1f4-43a9517f6433 | https://echa.europa.eu/documents/10162/77829f20-a0f1-adb8-fb45-4450e9db64b7
Strontium chromate |  | 232-142-6 | 7789-06-2 | Carcinogenic (Article 57a) | 20-Jun-2011 | https://echa.europa.eu/documents/10162/8003d372-68f9-4a5a-9094-4e6d0b58169c | https://echa.europa.eu/documents/10162/dbf5f0a3-13f7-4dcc-9638-be11aa9796bc | https://echa.europa.eu/documents/10162/b404f6d6-e6a2-412a-8362-542d297fd1bd | https://echa.europa.eu/documents/10162/98b171ab-1ab9-d334-ccea-76179abd24f6
Hydrazine |  | 206-114-9 | 302-01-2, 7803-57-8 | Carcinogenic (Article 57a) | 20-Jun-2011 | https://echa.europa.eu/documents/10162/c5b972a9-f57f-4fd5-8177-04b4e46c5e93 | https://echa.europa.eu/documents/10162/8c489ab1-cc43-4655-9fdf-9ba8faffec98 | https://echa.europa.eu/documents/10162/dcbad586-92e2-496c-bf17-f2d4fd7a7e5e | https://echa.europa.eu/documents/10162/8077c3d0-b547-5cec-dd34-0804abf35d93
2-ethoxyethyl acetate |  | 203-839-2 | 111-15-9 | Toxic for reproduction (Article 57c) | 20-Jun-2011 | https://echa.europa.eu/documents/10162/4b69d66f-b3f9-4b3d-a1fb-a90c4825f7d5 | https://echa.europa.eu/documents/10162/a4d1618e-1de0-45ee-95d4-7c0b97f8f4bb | https://echa.europa.eu/documents/10162/e3566164-121d-4007-9f6c-eeb1bbe67c14 | https://echa.europa.eu/documents/10162/4143edd3-cb57-c9f6-f0d4-2c859a736be8
1-Methyl-2-pyrrolidone (NMP) |  | 212-828-1 | 872-50-4 | Toxic for reproduction (Article 57c) | 20-Jun-2011 | https://echa.europa.eu/documents/10162/2d6cdfd7-e924-45fa-85d4-4450a7b4bf5a | https://echa.europa.eu/documents/10162/2a739583-e50d-458f-860b-2170d391621d | https://echa.europa.eu/documents/10162/1c4e3474-34ee-4c15-aaef-dafd1cb47779 | https://echa.europa.eu/documents/10162/e8ed4407-d09b-02bd-441e-f02950db4f74
1,2-Benzenedicarboxylic acid, di-C7-11-branched and linear alkyl esters |  | 271-084-6 | 68515-42-4 | Toxic for reproduction (Article 57c) | 20-Jun-2011 | https://echa.europa.eu/documents/10162/e0602e69-f758-4eef-88c6-6e9c7a0359e6 | https://echa.europa.eu/documents/10162/607fe319-a2a5-412f-9411-7fe86a8dd261 | https://echa.europa.eu/documents/10162/9064c470-918c-4459-b1de-f6112da0a8ec | https://echa.europa.eu/documents/10162/0ee9eaa7-282e-aaf9-03d7-58a70984249e
1,2-Benzenedicarboxylic acid, di-C6-8-branched alkyl esters, C7-rich |  | 276-158-1 | 71888-89-6 | Toxic for reproduction (Article 57c) | 20-Jun-2011 | https://echa.europa.eu/documents/10162/1d278ff8-5bc6-4fa4-8b6a-6847121fcc9d | https://echa.europa.eu/documents/10162/23661523-674a-4366-bf5d-8597e97f09a5 | https://echa.europa.eu/documents/10162/4ecec0c5-de0a-403e-8bad-bb2439aed1cd | https://echa.europa.eu/documents/10162/caf060a0-d277-1665-a24d-335d971a255a
1,2,3-trichloropropane |  | 202-486-1 | 96-18-4 | Carcinogenic (Article 57a)#Toxic for reproduction (Article 57c) | 20-Jun-2011 | https://echa.europa.eu/documents/10162/6b5d8cc7-e37e-46e1-9d91-0f8d529ca833 | https://echa.europa.eu/documents/10162/5c763e23-3df9-4568-990f-df7d38977711 | https://echa.europa.eu/documents/10162/e8d4bcbd-81c1-4a59-a97d-d1cb30c29ec8 | https://echa.europa.eu/documents/10162/16196c31-ff09-f53e-9f78-aef429e56548
Cobalt(II) sulphate |  | 233-334-2 | 10124-43-3 | Carcinogenic (Article 57a)#Toxic for reproduction (Article 57c) | 15-Dec-2010 | https://echa.europa.eu/documents/10162/034150e5-fa99-4659-a208-3492fed19012 | https://echa.europa.eu/documents/10162/6156179b-31de-4d5a-8bde-4c596e132251 | https://echa.europa.eu/documents/10162/5fa87d07-2872-4502-b07c-4186797aa442 | https://echa.europa.eu/documents/10162/2609fc0f-7c31-ebb9-50cd-293fa9f76238
Cobalt(II) dinitrate |  | 233-402-1 | 10141-05-6 | Carcinogenic (Article 57a)#Toxic for reproduction (Article 57c) | 15-Dec-2010 | https://echa.europa.eu/documents/10162/93d2146d-6f41-43f1-9e4c-6c65ee52bfee | https://echa.europa.eu/documents/10162/b0db8920-3429-46d7-9788-9f327712301d | https://echa.europa.eu/documents/10162/a6ff4bdb-6aa3-48ff-b9b4-6f086d8f35fa | https://echa.europa.eu/documents/10162/48eb838f-79f0-3322-7ccc-c2a198040538
Cobalt(II) diacetate |  | 200-755-8 | 71-48-7 | Carcinogenic (Article 57a)#Toxic for reproduction (Article 57c) | 15-Dec-2010 | https://echa.europa.eu/documents/10162/454374bb-5c3a-4716-bce2-2bf2879536d3 | https://echa.europa.eu/documents/10162/b7623ce3-4a1d-4b72-8bd3-b40872dde269 | https://echa.europa.eu/documents/10162/f68e76b3-751d-48b4-a225-4a23c0ee6249 | https://echa.europa.eu/documents/10162/725dc1e5-4900-4e07-e15c-c41633ddd11f
Cobalt(II) carbonate |  | 208-169-4 | 513-79-1 | Carcinogenic (Article 57a)#Toxic for reproduction (Article 57c) | 15-Dec-2010 | https://echa.europa.eu/documents/10162/75006cd4-1dc6-47de-bc84-dd92e07201a9 | https://echa.europa.eu/documents/10162/8724ec6d-5db0-4768-a9a9-43174499cde5 | https://echa.europa.eu/documents/10162/efdc02e9-f1e6-47c7-872b-7d6285457495 | https://echa.europa.eu/documents/10162/2a97b8e1-274e-0df6-24de-1791868ff37c
Chromium trioxide |  | 215-607-8 | 1333-82-0 | Carcinogenic (Article 57a)#Mutagenic (Article 57b) | 15-Dec-2010 | https://echa.europa.eu/documents/10162/6b11ec66-9d90-400a-a61a-90de9a0fd8b1 | https://echa.europa.eu/documents/10162/26e69088-ff48-41b2-86be-ad93559d4c1d | https://echa.europa.eu/documents/10162/6b7798d0-6e23-4e72-a1ea-ec989ef1e9b6 | https://echa.europa.eu/documents/10162/412bacac-471b-5982-5db4-dbe5b67f947b
Acids generated from chromium trioxide and their oligomers |  | - | - | Carcinogenic (Article 57a) | 15-Dec-2010 | https://echa.europa.eu/documents/10162/7ec3a9eb-a9f6-4e82-9e7d-99794b3dc95e | https://echa.europa.eu/documents/10162/e49af511-8109-4671-87c9-d1929bd1f9df#https://echa.europa.eu/documents/10162/b3342013-b216-403f-b9f6-ea2d704d45e3 | https://echa.europa.eu/documents/10162/b4e7126f-3045-4159-8142-958061cf13ac | https://echa.europa.eu/documents/10162/bf1c6776-60e5-cd31-7fc2-8a073e592fcb
Oligomers of chromic acid and dichromic acid |  | - | - | Carcinogenic (Article 57a) | 15-Dec-2010 | https://echa.europa.eu/documents/10162/7ec3a9eb-a9f6-4e82-9e7d-99794b3dc95e | https://echa.europa.eu/documents/10162/e49af511-8109-4671-87c9-d1929bd1f9df#https://echa.europa.eu/documents/10162/b3342013-b216-403f-b9f6-ea2d704d45e3 | https://echa.europa.eu/documents/10162/b4e7126f-3045-4159-8142-958061cf13ac | https://echa.europa.eu/documents/10162/bf1c6776-60e5-cd31-7fc2-8a073e592fcb
Dichromic acid |  | 236-881-5 | 13530-68-2 | Carcinogenic (Article 57a) | 15-Dec-2010 | https://echa.europa.eu/documents/10162/7ec3a9eb-a9f6-4e82-9e7d-99794b3dc95e | https://echa.europa.eu/documents/10162/e49af511-8109-4671-87c9-d1929bd1f9df#https://echa.europa.eu/documents/10162/b3342013-b216-403f-b9f6-ea2d704d45e3 | https://echa.europa.eu/documents/10162/b4e7126f-3045-4159-8142-958061cf13ac | https://echa.europa.eu/documents/10162/bf1c6776-60e5-cd31-7fc2-8a073e592fcb
Chromic acid |  | 231-801-5 | 7738-94-5 | Carcinogenic (Article 57a) | 15-Dec-2010 | https://echa.europa.eu/documents/10162/7ec3a9eb-a9f6-4e82-9e7d-99794b3dc95e | https://echa.europa.eu/documents/10162/e49af511-8109-4671-87c9-d1929bd1f9df#https://echa.europa.eu/documents/10162/b3342013-b216-403f-b9f6-ea2d704d45e3 | https://echa.europa.eu/documents/10162/b4e7126f-3045-4159-8142-958061cf13ac | https://echa.europa.eu/documents/10162/bf1c6776-60e5-cd31-7fc2-8a073e592fcb
2-methoxyethanol |  | 203-713-7 | 109-86-4 | Toxic for reproduction (Article 57c) | 15-Dec-2010 | https://echa.europa.eu/documents/10162/d0434782-81cb-43af-ba42-88c301aeb105 | https://echa.europa.eu/documents/10162/2c32ce81-9fe4-4bec-ab99-cefb0db64f70 | https://echa.europa.eu/documents/10162/de7fb120-e51b-408c-9753-5c605be53de9 | https://echa.europa.eu/documents/10162/97777174-b86f-1e5e-fb0a-92a7ab4e1fda
2-ethoxyethanol |  | 203-804-1 | 110-80-5 | Toxic for reproduction (Article 57c) | 15-Dec-2010 | https://echa.europa.eu/documents/10162/f25b7ab7-c339-4b4a-900b-7a2d38c32c1f | https://echa.europa.eu/documents/10162/16b140dd-f4a7-4b4c-a7c4-696ff16785ea | https://echa.europa.eu/documents/10162/b8b0b5ed-43be-4c24-8d81-96b799916955 | https://echa.europa.eu/documents/10162/7c99a5c1-5dad-1d72-f660-e4c451739bbf
Trichloroethylene |  | 201-167-4 | 79-01-6 | Carcinogenic (Article 57a) | 18-Jun-2010 | https://echa.europa.eu/documents/10162/d54e4381-eebe-41dd-9eb5-5794d383f043 | https://echa.europa.eu/documents/10162/bf5436ce-44ed-49c8-b7e6-51804d3cb0a5 | https://echa.europa.eu/documents/10162/8bbf5e70-f0a4-46c0-84db-f2340c74d64e | https://echa.europa.eu/documents/10162/50be92c9-8f8f-8cb9-3b52-12d018a788b7
Tetraboron disodium heptaoxide, hydrate |  | 235-541-3 | 12267-73-1 | Toxic for reproduction (Article 57c) | 18-Jun-2010 | https://echa.europa.eu/documents/10162/4ed94bb5-533a-4a8a-823f-6348859263d1 | https://echa.europa.eu/documents/10162/b743c746-3dd3-4ddb-83f9-8446d5615222 | https://echa.europa.eu/documents/10162/a049f842-ad66-4b06-b1ec-39fd6c6f47ba | https://echa.europa.eu/documents/10162/e893fee3-07e0-8f9b-a3be-29c7c75008f2
Sodium chromate |  | 231-889-5 | 7775-11-3 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Toxic for reproduction (Article 57c) | 18-Jun-2010 | https://echa.europa.eu/documents/10162/f7a917fb-3073-4f18-93db-d49e3587df84 | https://echa.europa.eu/documents/10162/604acc0d-b03f-4c88-b9e1-574a0709b463 | https://echa.europa.eu/documents/10162/c5d6f8de-bbe1-4b31-b7c4-d8e5b8d65517 | https://echa.europa.eu/documents/10162/7bfc757c-3719-2959-1ffd-f2bdb7160259
Potassium dichromate |  | 231-906-6 | 7778-50-9 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Toxic for reproduction (Article 57c) | 18-Jun-2010 | https://echa.europa.eu/documents/10162/0f342468-8b83-48af-b1ce-c46eb9e011bf | https://echa.europa.eu/documents/10162/cb89dff1-a881-4c96-98aa-058ddaf85532 | https://echa.europa.eu/documents/10162/3ac167c4-9f8e-4375-92f6-285a2911cee0 | https://echa.europa.eu/documents/10162/a1656dde-4912-2cfb-8605-690a46b7b790
Potassium chromate |  | 232-140-5 | 7789-00-6 | Carcinogenic (Article 57a)#Mutagenic (Article 57b) | 18-Jun-2010 | https://echa.europa.eu/documents/10162/4116dc63-d4ff-4465-b598-a81622c5f58b | https://echa.europa.eu/documents/10162/6ca9ac30-7102-4c5a-9eb8-bbcefc64c411 | https://echa.europa.eu/documents/10162/e6da86f5-0988-4fb2-9dcf-84087571c5af | https://echa.europa.eu/documents/10162/3384d4c1-6650-fae8-5f86-22be13577d99
Disodium tetraborate, anhydrous |  | 215-540-4 | 12179-04-3, 1303-96-4, 1330-43-4 | Toxic for reproduction (Article 57c) | 18-Jun-2010 | https://echa.europa.eu/documents/10162/4912ba63-c1c1-4698-b1a0-3b58f6024cd5 | https://echa.europa.eu/documents/10162/423223b2-23c6-4203-a1c5-c1e4cf47e854 | https://echa.europa.eu/documents/10162/04e29ef1-f56d-4926-96f6-c2865c5e10a7 | https://echa.europa.eu/documents/10162/9004f5d0-c545-42ce-8839-e07b8f938e7a
Boric acid | EC No. 233-139-2 and EC No. 234-343-4 | - | - | Toxic for reproduction (Article 57c) | 18-Jun-2010 | https://echa.europa.eu/documents/10162/6877f24b-c82b-4021-b05b-a25af94185c0 | https://echa.europa.eu/documents/10162/cad21145-7d6a-4c1f-b640-80cfe475b45c#https://echa.europa.eu/documents/10162/910d71c0-785a-4b54-997a-6b4396389739 | https://echa.europa.eu/documents/10162/d51fd473-40ec-4831-bc2d-6f53bdf9cbbe | https://echa.europa.eu/documents/10162/2991b041-b25e-7f6a-0900-2ebef8c9e030
Boric acid |  | 233-139-2 | 10043-35-3 | Toxic for reproduction (Article 57c) | 18-Jun-2010 | https://echa.europa.eu/documents/10162/6877f24b-c82b-4021-b05b-a25af94185c0 | https://echa.europa.eu/documents/10162/cad21145-7d6a-4c1f-b640-80cfe475b45c#https://echa.europa.eu/documents/10162/910d71c0-785a-4b54-997a-6b4396389739 | https://echa.europa.eu/documents/10162/d51fd473-40ec-4831-bc2d-6f53bdf9cbbe | https://echa.europa.eu/documents/10162/2991b041-b25e-7f6a-0900-2ebef8c9e030
Boric acid, crude natural |  | 234-343-4 | 11113-50-1 | Toxic for reproduction (Article 57c) | 18-Jun-2010 | https://echa.europa.eu/documents/10162/6877f24b-c82b-4021-b05b-a25af94185c0 | https://echa.europa.eu/documents/10162/cad21145-7d6a-4c1f-b640-80cfe475b45c#https://echa.europa.eu/documents/10162/910d71c0-785a-4b54-997a-6b4396389739 | https://echa.europa.eu/documents/10162/d51fd473-40ec-4831-bc2d-6f53bdf9cbbe | https://echa.europa.eu/documents/10162/2991b041-b25e-7f6a-0900-2ebef8c9e030
Ammonium dichromate |  | 232-143-1 | 7789-09-5 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Toxic for reproduction (Article 57c) | 18-Jun-2010 | https://echa.europa.eu/documents/10162/ff5cd363-1d18-4b42-a208-b0aa4034348e | https://echa.europa.eu/documents/10162/e3fb1a7c-4bd5-4540-9997-fc9f040c1a2f | https://echa.europa.eu/documents/10162/b09ce051-1fe5-425a-bdd0-0f6eb628230b | https://echa.europa.eu/documents/10162/875e5e2d-586f-5c6a-5440-b40ed7024aee
Acrylamide |  | 201-173-7 | 79-06-1 | Carcinogenic (Article 57a)#Mutagenic (Article 57b) | 30-Mar-2010 | https://echa.europa.eu/documents/10162/b280c0f2-f957-44d9-9584-097444464b22 | https://echa.europa.eu/documents/10162/13141b6e-f285-44bf-9868-d5e089875101 | https://echa.europa.eu/documents/10162/7f1a698d-9910-4e8a-a6a7-0806427abd26 | https://echa.europa.eu/documents/10162/4c9a4bdf-23eb-bd23-7a56-9d34b5641826
Tris(2-chloroethyl) phosphate |  | 204-118-5 | 115-96-8 | Toxic for reproduction (Article 57c) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/78d8ecfd-5e83-4299-9f16-15681ee11bbb | https://echa.europa.eu/documents/10162/d36f5718-6e5c-4fca-9e31-efed9c50d0d0 | https://echa.europa.eu/documents/10162/6d09755f-7fcb-4a00-b7ce-91ab45a2e5af | https://echa.europa.eu/documents/10162/da069526-8af2-c6c3-3d6d-2c06b7446f82
Pitch, coal tar, high-temp. | - | 266-028-2 | 65996-93-2 | Carcinogenic (Article 57a)#PBT (Article 57d)#vPvB (Article 57e) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/30488018-7d14-42b1-90eb-8235974bf023 | https://echa.europa.eu/documents/10162/d3d19672-64d5-47f4-aa47-66729f6990ec | https://echa.europa.eu/documents/10162/73d246d4-8c2a-4150-b656-c15948bf0e77 | https://echa.europa.eu/documents/10162/9762a7b3-6123-562f-0bdf-2213e07a427d
Lead sulfochromate yellow (C.I. Pigment Yellow 34) | - | 215-693-7 | 1344-37-2 | Carcinogenic (Article 57a)#Toxic for reproduction (Article 57c) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/7d4b0549-413a-4642-ac69-069566c2f624 | https://echa.europa.eu/documents/10162/cbb77a48-22b8-4888-b485-8e85731648d2 | https://echa.europa.eu/documents/10162/cd44c55f-3473-4c73-a356-ce732b2f3d87 | https://echa.europa.eu/documents/10162/8dbe3010-814b-d92b-8bb1-353fa4c32f9f
Lead chromate molybdate sulphate red (C.I. Pigment Red 104) | - | 235-759-9 | 12656-85-8 | Carcinogenic (Article 57a)#Toxic for reproduction (Article 57c) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/9a10159e-0846-488b-8fdb-5162e2bf97f4 | https://echa.europa.eu/documents/10162/efa8a662-a744-45f7-a507-14a3dcfc3d15 | https://echa.europa.eu/documents/10162/624c2151-d7f2-47d7-ab26-7f4996e81e36 | https://echa.europa.eu/documents/10162/d46d12dd-a1a3-e556-549a-e81aac3b5905
Lead chromate |  | 231-846-0 | 7758-97-6 | Carcinogenic (Article 57a)#Toxic for reproduction (Article 57c) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/1f928dd4-69c2-494d-8d94-08699be302e5 | https://echa.europa.eu/documents/10162/df3d989f-4229-497e-942c-44763337dcf2 | https://echa.europa.eu/documents/10162/25e3a7a4-41a4-4f72-8d0f-2ab06a6bf51a
Diisobutyl phthalate |  | 201-553-2 | 84-69-5 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - human health) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/8f861ec5-40ca-43d6-be8a-7bc22b96f84f#https://echa.europa.eu/documents/10162/a956b752-1a1b-1316-6ed7-a01d09cd52cd#https://echa.europa.eu/documents/10162/22021d20-ffe6-5b92-8961-4420d5788c06 | https://echa.europa.eu/documents/10162/5f086082-c721-442d-8b7b-a6eb49ffb3ab | https://echa.europa.eu/documents/10162/96caf4ad-1095-4929-a1b9-84aeea1aff00#https://echa.europa.eu/documents/10162/872ff9cf-52a5-49f4-1e0a-da6ce3e9e074 | https://echa.europa.eu/documents/10162/566bf9dd-444f-e07f-877f-cc0f53e596b7
Anthracene oil, anthracene-low | - | 292-604-8 | 90640-82-7 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#PBT (Article 57d)#vPvB (Article 57e) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/709cf01f-1445-4606-a0f2-1940a468dab6 | https://echa.europa.eu/documents/10162/ce2619c5-654b-4934-a835-a3e2b00d4e13 | https://echa.europa.eu/documents/10162/3e5e6a56-dd4f-4951-b26e-985b7f692624 | https://echa.europa.eu/documents/10162/621cceb7-32a6-81b3-a5e6-d31f01cf4ce4 | Does not meet the criteria for identification as a carcinogen if it contains &lt; 0.005 % (w/w) benzo[a]pyrene (EINECS No 200-028-5) and &lt; 0.1 % w/w benzene (EINECS No 200-753-7). Does not meet the criteria for identification as a mutagen if it contains &lt; 0.1 % w/w benzene (EINECS No 200-753-7).
Anthracene oil, anthracene paste, distn. lights | - | 295-278-5 | 91995-17-4 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#PBT (Article 57d)#vPvB (Article 57e) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/960a181f-7094-4be7-b02f-c49513a0442e | https://echa.europa.eu/documents/10162/2a5afd0a-9ce1-4b37-acc9-942143c01c7a | https://echa.europa.eu/documents/10162/f0ff9790-013c-49b3-ae93-ac11d2177bf9 | https://echa.europa.eu/documents/10162/3879120b-4739-bfb8-40dc-ed001dbc2f09 | Does not meet the criteria for identification as a carcinogen if it contains &lt; 0.005 % (w/w) benzo[a]pyrene (EINECS No 200-028-5) and &lt; 0.1 % w/w benzene (EINECS No 200-753-7). Does not meet the criteria for identification as a mutagen if it contains &lt; 0.1 % w/w benzene (EINECS No 200-753-7).
Anthracene oil, anthracene paste, anthracene fraction | - | 295-275-9 | 91995-15-2 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#PBT (Article 57d)#vPvB (Article 57e) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/b5013ad5-8980-478b-aa32-4859860f9aef | https://echa.europa.eu/documents/10162/4efcaa35-8b9f-4b19-9f81-9f1fb98533b2 | https://echa.europa.eu/documents/10162/8de3cd61-aae1-4864-b694-8928dcdea7fd | https://echa.europa.eu/documents/10162/3c6d1b64-fe59-72a7-699d-1735d3c4ec48 | Does not meet the criteria for identification as a carcinogen if it contains &lt; 0.005 % (w/w) benzo[a]pyrene (EINECS No 200-028-5) and &lt; 0.1 % w/w benzene (EINECS No 200-753-7). Does not meet the criteria for identification as a mutagen if it contains &lt; 0.1 % w/w benzene (EINECS No 200-753-7).
Anthracene oil, anthracene paste | - | 292-603-2 | 90640-81-6 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#PBT (Article 57d)#vPvB (Article 57e) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/df5aecc8-d928-4a70-9bf0-90ca5953bc44 | https://echa.europa.eu/documents/10162/ad049eea-4aeb-4741-9574-ae6c4bd02b66 | https://echa.europa.eu/documents/10162/1995aafa-663a-43a0-ae2a-285c842d7f53 | https://echa.europa.eu/documents/10162/c517a94d-eed3-8371-3308-508060b19c7f | Does not meet the criteria for identification as a carcinogen if it contains &lt; 0.005 % (w/w) benzo[a]pyrene (EINECS No 200-028-5) and &lt; 0.1 % w/w benzene (EINECS No 200-753-7). Does not meet the criteria for identification as a mutagen if it contains &lt; 0.1 % w/w benzene (EINECS No 200-753-7).
Anthracene oil | - | 292-602-7 | 90640-80-5 | Carcinogenic (Article 57a)#PBT (Article 57d)#vPvB (Article 57e) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/f7cfbd7b-2a5a-49aa-9d16-61689ad3a8e3 | https://echa.europa.eu/documents/10162/91cfe38e-ed0d-4c28-b970-bded2289de26 | https://echa.europa.eu/documents/10162/ef47ccc5-0786-4bf5-b16c-2f13c6c49385 | https://echa.europa.eu/documents/10162/70eb45c5-9fec-0871-e281-e1adeace762b | Does not meet the criteria for identification as a carcinogen if it contains &lt; 0.005 % (w/w) benzo[a]pyrene (EINECS No 200-028-5)
2,4-dinitrotoluene |  | 204-450-0 | 121-14-2 | Carcinogenic (Article 57a) | 13-Jan-2010 | https://echa.europa.eu/documents/10162/7700e235-3ba5-44d4-9ab9-5ace916ca7b3 | https://echa.europa.eu/documents/10162/a8b8d40f-1cea-4daf-9f49-af3a2b2d50d9 | https://echa.europa.eu/documents/10162/b5db559d-a89d-4fc4-a273-dba250fb4201 | https://echa.europa.eu/documents/10162/2bca3e81-17de-68e1-47ad-f34f0e3dd1de
Triethyl arsenate | - | 427-700-2 | 15606-95-8 | Carcinogenic (Article 57a) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/6d48ec99-f364-44de-92a5-62d505ea3724 | https://echa.europa.eu/documents/10162/e7756763-29b1-4135-8aa6-2b20d8e672a1 | https://echa.europa.eu/documents/10162/d3a393bf-e221-4fa1-8d81-082c2ac3cf8f
Sodium dichromate |  | 234-190-3 | 10588-01-9, 7789-12-0 | Carcinogenic (Article 57a)#Mutagenic (Article 57b)#Toxic for reproduction (Article 57c) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/0594aba7-e06e-45bd-807d-3eb9bd74024c | https://echa.europa.eu/documents/10162/03566bd3-3a60-40b1-8a7d-86ea5737a0f6 | https://echa.europa.eu/documents/10162/52c64496-821b-4a74-bc1a-d9fc2040b651 | https://echa.europa.eu/documents/10162/23a67aca-ee8c-be65-16b8-cf4b1ea48c67
Lead hydrogen arsenate |  | 232-064-2 | 7784-40-9 | Carcinogenic (Article 57a)#Toxic for reproduction (Article 57c) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/ebdb28e4-ea63-4f11-8b2e-c299ad5ebd07 | https://echa.europa.eu/documents/10162/f1ac3df6-c52a-4870-847b-1d6a74e1b034 | https://echa.europa.eu/documents/10162/ee0a9b5c-6f5e-43d7-8ef5-76b068bb4732 | https://echa.europa.eu/documents/10162/7b55c9ab-6322-4ebe-2deb-6914e6e0ec87
Hexabromocyclododecane (HBCDD) | and all major diastereoisomers identified | - | - | PBT (Article 57d) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/471aceac-4e5e-4c53-a4b2-23159a290893 | https://echa.europa.eu/documents/10162/df4ad28a-899d-4e76-8529-eda8bc6591ae#https://echa.europa.eu/documents/10162/938b85d2-d87c-480a-b697-bad2e5eced3d#https://echa.europa.eu/documents/10162/0412300e-9170-4eb7-87c4-9bd3bad4ca42#https://echa.europa.eu/documents/10162/7b768673-ea38-4822-9712-45b7dca0c2d8#https://echa.europa.eu/documents/10162/63d654f6-5400-449d-81bb-1aa6518d7ca3 | https://echa.europa.eu/documents/10162/1fc837ef-f922-476e-8e00-825ab60213c0 | https://echa.europa.eu/documents/10162/e109c7be-6fd0-e12e-4c92-7e12d56fead8
1,2,5,6,9,10-hexabromocyclododecane |  | 221-695-9 | 3194-55-6 | PBT (Article 57d) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/471aceac-4e5e-4c53-a4b2-23159a290893 | https://echa.europa.eu/documents/10162/df4ad28a-899d-4e76-8529-eda8bc6591ae#https://echa.europa.eu/documents/10162/938b85d2-d87c-480a-b697-bad2e5eced3d#https://echa.europa.eu/documents/10162/0412300e-9170-4eb7-87c4-9bd3bad4ca42#https://echa.europa.eu/documents/10162/7b768673-ea38-4822-9712-45b7dca0c2d8#https://echa.europa.eu/documents/10162/63d654f6-5400-449d-81bb-1aa6518d7ca3 | https://echa.europa.eu/documents/10162/1fc837ef-f922-476e-8e00-825ab60213c0 | https://echa.europa.eu/documents/10162/e109c7be-6fd0-e12e-4c92-7e12d56fead8
Hexabromocyclododecane |  | 247-148-4 | 25637-99-4 | PBT (Article 57d) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/471aceac-4e5e-4c53-a4b2-23159a290893 | https://echa.europa.eu/documents/10162/df4ad28a-899d-4e76-8529-eda8bc6591ae#https://echa.europa.eu/documents/10162/938b85d2-d87c-480a-b697-bad2e5eced3d#https://echa.europa.eu/documents/10162/0412300e-9170-4eb7-87c4-9bd3bad4ca42#https://echa.europa.eu/documents/10162/7b768673-ea38-4822-9712-45b7dca0c2d8#https://echa.europa.eu/documents/10162/63d654f6-5400-449d-81bb-1aa6518d7ca3 | https://echa.europa.eu/documents/10162/1fc837ef-f922-476e-8e00-825ab60213c0 | https://echa.europa.eu/documents/10162/e109c7be-6fd0-e12e-4c92-7e12d56fead8
alpha-hexabromocyclododecane |  | - | 134237-50-6 | PBT (Article 57d) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/471aceac-4e5e-4c53-a4b2-23159a290893 | https://echa.europa.eu/documents/10162/df4ad28a-899d-4e76-8529-eda8bc6591ae#https://echa.europa.eu/documents/10162/938b85d2-d87c-480a-b697-bad2e5eced3d#https://echa.europa.eu/documents/10162/0412300e-9170-4eb7-87c4-9bd3bad4ca42#https://echa.europa.eu/documents/10162/7b768673-ea38-4822-9712-45b7dca0c2d8#https://echa.europa.eu/documents/10162/63d654f6-5400-449d-81bb-1aa6518d7ca3 | https://echa.europa.eu/documents/10162/1fc837ef-f922-476e-8e00-825ab60213c0 | https://echa.europa.eu/documents/10162/e109c7be-6fd0-e12e-4c92-7e12d56fead8
beta-hexabromocyclododecane |  | - | 134237-51-7 | PBT (Article 57d) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/471aceac-4e5e-4c53-a4b2-23159a290893 | https://echa.europa.eu/documents/10162/df4ad28a-899d-4e76-8529-eda8bc6591ae#https://echa.europa.eu/documents/10162/938b85d2-d87c-480a-b697-bad2e5eced3d#https://echa.europa.eu/documents/10162/0412300e-9170-4eb7-87c4-9bd3bad4ca42#https://echa.europa.eu/documents/10162/7b768673-ea38-4822-9712-45b7dca0c2d8#https://echa.europa.eu/documents/10162/63d654f6-5400-449d-81bb-1aa6518d7ca3 | https://echa.europa.eu/documents/10162/1fc837ef-f922-476e-8e00-825ab60213c0 | https://echa.europa.eu/documents/10162/e109c7be-6fd0-e12e-4c92-7e12d56fead8
gamma-hexabromocyclododecane |  | - | 134237-52-8 | PBT (Article 57d) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/471aceac-4e5e-4c53-a4b2-23159a290893 | https://echa.europa.eu/documents/10162/df4ad28a-899d-4e76-8529-eda8bc6591ae#https://echa.europa.eu/documents/10162/938b85d2-d87c-480a-b697-bad2e5eced3d#https://echa.europa.eu/documents/10162/0412300e-9170-4eb7-87c4-9bd3bad4ca42#https://echa.europa.eu/documents/10162/7b768673-ea38-4822-9712-45b7dca0c2d8#https://echa.europa.eu/documents/10162/63d654f6-5400-449d-81bb-1aa6518d7ca3 | https://echa.europa.eu/documents/10162/1fc837ef-f922-476e-8e00-825ab60213c0 | https://echa.europa.eu/documents/10162/e109c7be-6fd0-e12e-4c92-7e12d56fead8
Dibutyl phthalate (DBP) |  | 201-557-4 | 84-74-2 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/d33fd29e-5170-440f-8c16-10ecd87cf472#https://echa.europa.eu/documents/10162/2f10864f-e748-0e86-9bc4-2b3964fbdace#https://echa.europa.eu/documents/10162/326a8103-bc47-fb36-0871-22b46ec9d66f#https://echa.europa.eu/documents/10162/94f8f2db-076b-b817-31ae-c14f487642e9 | https://echa.europa.eu/documents/10162/3686ef70-c4fc-2824-7fe5-23a1ddb5b06c | https://echa.europa.eu/documents/10162/f36205a8-df0a-4e94-8763-2a6a02f0228c#https://echa.europa.eu/documents/10162/6b0b59e9-8ba8-03de-e8e9-735414f0611d#https://echa.europa.eu/documents/10162/4cd3c3e7-10bd-5ebe-3492-3696c8a43a9b | https://echa.europa.eu/documents/10162/173eff20-d119-4ed5-258a-1e613480c3ce#https://echa.europa.eu/documents/10162/f305eb04-5468-e615-a78d-b6a8aaf20e8e
Diarsenic trioxide |  | 215-481-4 | 1327-53-3 | Carcinogenic (Article 57a) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/d502da18-827f-408d-aa24-9fe0c7e6ed1b | https://echa.europa.eu/documents/10162/03f099fc-0057-49d9-a816-24c44d69b5c1 | https://echa.europa.eu/documents/10162/8730b910-c1d4-4b87-b3e3-a94ce47fbc9b | https://echa.europa.eu/documents/10162/7dc14675-1df3-e89a-e05d-798e5f51749e
Diarsenic pentaoxide |  | 215-116-9 | 1303-28-2 | Carcinogenic (Article 57a) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/7b11c857-a72d-4a8b-9efb-68d99ead7902 | https://echa.europa.eu/documents/10162/76037d34-1e8c-405e-b594-8fe837e6095e | https://echa.europa.eu/documents/10162/91206b72-787f-4a2d-aa80-9d6f1f22fc55 | https://echa.europa.eu/documents/10162/d1ff5c51-11a5-7aca-a48e-2393f45ce653
Cobalt dichloride |  | 231-589-4 | 7646-79-9 | Carcinogenic (Article 57a)#Toxic for reproduction (Article 57c) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/5633d67c-a61b-4bab-b6c5-c4a61cc99e00#https://echa.europa.eu/documents/10162/39173d0d-3258-4f23-ab18-a586bcb39567 | https://echa.europa.eu/documents/10162/383cf518-9e20-495c-bbcf-8875b9373714 | https://echa.europa.eu/documents/10162/7d541979-6f03-421b-91bb-039bfb326de1#https://echa.europa.eu/documents/10162/fdbfa509-c244-4537-85eb-07994f393f23 | https://echa.europa.eu/documents/10162/1c469e55-cee2-7841-84f3-2d6abc372ed8
Bis(tributyltin) oxide (TBTO) |  | 200-268-0 | 56-35-9 | PBT (Article 57d) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/1b18d017-1056-4c92-aeeb-f190ad674174 | https://echa.europa.eu/documents/10162/e7faa0f6-cb40-4930-8931-537d261e43ff | https://echa.europa.eu/documents/10162/52f3fc94-c78f-436f-98ca-e0f845f37a9a | https://echa.europa.eu/documents/10162/6e66a086-935f-ab28-012c-262f5df9ef3b
Bis (2-ethylhexyl)phthalate (DEHP) |  | 204-211-0 | 117-81-7 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - environment)#Endocrine disrupting properties (Article 57(f) - human health) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/c2ecc989-445d-40b9-a054-28671849b092#https://echa.europa.eu/documents/10162/30b654ce-1de3-487a-8696-e05617c3173b#https://echa.europa.eu/documents/10162/3b0d2893-b8db-86b9-6db0-6e06dc9aa10e#https://echa.europa.eu/documents/10162/88c20879-606b-03a6-11e4-9edb90e7e615 | https://echa.europa.eu/documents/10162/331e209a-627f-4533-8690-fd33b74f305d | https://echa.europa.eu/documents/10162/9e33725c-05fc-41af-b06d-bb969c91b0ab#https://echa.europa.eu/documents/10162/fa429d23-21e7-4764-b223-6c8c98f8a01c#https://echa.europa.eu/documents/10162/3f5d91bc-0a63-04f8-8f62-ccc9c8265e7d | https://echa.europa.eu/documents/10162/9cc2ee2d-89f9-39c8-718e-c961a65d3fb9
Benzyl butyl phthalate (BBP) |  | 201-622-7 | 85-68-7 | Toxic for reproduction (Article 57c)#Endocrine disrupting properties (Article 57(f) - human health) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/7cad31ee-4818-4166-9212-b39ca33240bb#https://echa.europa.eu/documents/10162/dd4a532e-e277-5b52-9ba2-a94be60b7f17#https://echa.europa.eu/documents/10162/cfb492f3-424f-712c-130b-01c6e2d5074a | https://echa.europa.eu/documents/10162/20b09a46-f289-4e35-92d2-9cd039e71ba5 | https://echa.europa.eu/documents/10162/b4024445-00ce-4849-9ab0-69aed88c51f2#https://echa.europa.eu/documents/10162/076491e2-fd0e-50be-61b4-908c663d0245 | https://echa.europa.eu/documents/10162/77871a51-ef10-f53b-a903-a1f37266a674
Anthracene |  | 204-371-1 | 120-12-7 | PBT (Article 57d) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/143cd93d-ee58-40c1-b0f1-3713cbb11535 | https://echa.europa.eu/documents/10162/c5f3267f-609d-4ccb-8645-5b7e02d3946c | https://echa.europa.eu/documents/10162/f7c1321a-6709-40d6-b683-1fb870fb0ac4 | https://echa.europa.eu/documents/10162/79e1b8f9-85e4-63d0-b3b2-51da407473b7
Alkanes, C10-13, chloro (Short Chain Chlorinated Paraffins) |  | 287-476-5 | 85535-84-8 | PBT (Article 57d)#vPvB (Article 57e) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/e9713dc8-1855-4dab-9635-72ca8be244ae | https://echa.europa.eu/documents/10162/c0c51b06-7e2e-4934-b46f-0ba82015047c | https://echa.europa.eu/documents/10162/2edcfedb-ec53-4754-8598-e787a8ff7a58 | https://echa.europa.eu/documents/10162/8eada19b-b66d-d5bf-d47c-76e934b88015
5-tert-butyl-2,4,6-trinitro-m-xylene (Musk xylene) |  | 201-329-4 | 81-15-2 | vPvB (Article 57e) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/e3747ed6-99b4-43d1-92ae-9c837ded241d | https://echa.europa.eu/documents/10162/e8716cf9-db7f-4073-88fa-18dd4480c61f | https://echa.europa.eu/documents/10162/909dd42e-2554-4f59-911a-729a2da1d529 | https://echa.europa.eu/documents/10162/c85e5b93-3e53-a492-9e5e-677066967516
4,4'- Diaminodiphenylmethane (MDA) |  | 202-974-4 | 101-77-9 | Carcinogenic (Article 57a) | 28-Oct-2008 | https://echa.europa.eu/documents/10162/68f50f10-f55f-40be-9d1b-4fbbfdec7f96 | https://echa.europa.eu/documents/10162/1c864bd5-399d-49a2-880b-a8ad54d458b5 | https://echa.europa.eu/documents/10162/d36424e7-b12d-4dd8-832e-6d7e3e283fc3 | https://echa.europa.eu/documents/10162/f243739c-1e71-3e19-9fa0-f2f15f861d23

## REACH Annex XVII
Substances restricted under REACH
The table below is the Annex XVII to REACH and includes all the restrictions adopted in the framework of REACH and the previous legislation, Directive 76/769/EEC.
https://echa.europa.eu/substances-restricted-under-reach?p_p_id=disslists_WAR_disslistsportlet&p_p_lifecycle=0&p_p_state=normal&p_p_mode=view&_disslists_WAR_disslistsportlet_cur=1&_disslists_WAR_disslistsportlet_orderByCol=prc_entry_no&_disslists_WAR_disslistsportlet_substance_identifier_field_key=&_disslists_WAR_disslistsportlet_orderByType=asc&_disslists_WAR_disslistsportlet_doSearch=&_disslists_WAR_disslistsportlet_prc_entry_no=&_disslists_WAR_disslistsportlet_deltaParamValue=50&_disslists_WAR_disslistsportlet_resetCur=false&_disslists_WAR_disslistsportlet_delta=50
Last updated 10 December 2025. Database contains 78 unique substances/entries.
Substance name | Description | EC No. | CAS No. | Entry no. | Conditions | Appendices | Standards | History | Q&As | Remarks | First published
Polychlorinated terphenyls (PCTs) |  | - | - | 01 | https://echa.europa.eu/documents/10162/3fd87368-345a-4b56-b545-53b2b8d4aa8c |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1976.262.01.0201.01.ENG&toc=OJ:L:1976:262:TOC">D 76/769/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1982.350.01.0034.01.ENG&toc=OJ:L:1982:350:TOC">D 82/828/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1985.269.01.0056.01.ENG&toc=OJ:L:1985:269:TOC">D 85/467</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 12-Jan-2016
Terphenyl, chlorinated |  | 262-968-2 | 61788-33-8 | 01 | https://echa.europa.eu/documents/10162/3fd87368-345a-4b56-b545-53b2b8d4aa8c |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1976.262.01.0201.01.ENG&toc=OJ:L:1976:262:TOC">D 76/769/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1982.350.01.0034.01.ENG&toc=OJ:L:1982:350:TOC">D 82/828/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1985.269.01.0056.01.ENG&toc=OJ:L:1985:269:TOC">D 85/467</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 12-Jan-2016
Chloroethene | Vinyl Chloride | 200-831-0 | 75-01-4 | 02 | https://echa.europa.eu/documents/10162/a7ed88f5-3b13-46ef-8fb7-dbb08087cfeb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1976.262.01.0201.01.ENG&toc=OJ:L:1976:262:TOC">D 76/769/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 15-Jan-2016
Liquid substances or mixtures which are regarded as dangerous in accordance with Directive 1999/45/EC or are fulfilling the criteria for any of the following hazard classes or categories set out in Annex I to Regulation (EC) No 1272/2008 (See group members): |  | - | - | 03 | https://echa.europa.eu/documents/10162/6e698812-f218-5f39-aff3-e46a1555a67d |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT:18429&cs=1A5FE3B096F5CE8B96EF0786DA3EF0594">EN 14059:2002</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1979.197.01.0037.01.ENG&toc=OJ:L:1979:197:TOC">D 79/663/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.315.01.0013.01.ENG&toc=OJ:L:1997:315:TOC">D 97/64/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.138.01.0008.01.ENG&toc=OJ:L:2009:138:TOC">D 2009/424/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2010.086.01.0007.01.ENG&toc=OJ:L:2010:086:TOC">R 276/2010</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1068">entry 1068</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1521">entry 1521</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1799">entry 1799</a> |  | 15-Jan-2016
Hazard class 4.1 |  | - | - | 03 | https://echa.europa.eu/documents/10162/6e698812-f218-5f39-aff3-e46a1555a67d |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT:18429&cs=1A5FE3B096F5CE8B96EF0786DA3EF0594">EN 14059:2002</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1979.197.01.0037.01.ENG&toc=OJ:L:1979:197:TOC">D 79/663/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.315.01.0013.01.ENG&toc=OJ:L:1997:315:TOC">D 97/64/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.138.01.0008.01.ENG&toc=OJ:L:2009:138:TOC">D 2009/424/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2010.086.01.0007.01.ENG&toc=OJ:L:2010:086:TOC">R 276/2010</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1068">entry 1068</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1521">entry 1521</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1799">entry 1799</a> |  | 15-Jan-2016
Hazard classes 3.1 to 3.6, 3.7 adverse effects on sexual function and fertility or on development, 3.8 effects other than narcotic effects, 3.9 and 3.10 |  | - | - | 03 | https://echa.europa.eu/documents/10162/6e698812-f218-5f39-aff3-e46a1555a67d |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT:18429&cs=1A5FE3B096F5CE8B96EF0786DA3EF0594">EN 14059:2002</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1979.197.01.0037.01.ENG&toc=OJ:L:1979:197:TOC">D 79/663/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.315.01.0013.01.ENG&toc=OJ:L:1997:315:TOC">D 97/64/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.138.01.0008.01.ENG&toc=OJ:L:2009:138:TOC">D 2009/424/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2010.086.01.0007.01.ENG&toc=OJ:L:2010:086:TOC">R 276/2010</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1068">entry 1068</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1521">entry 1521</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1799">entry 1799</a> |  | 15-Jan-2016
Hazard classes 2.1 to 2.4, 2.6 and 2.7, 2.8 types A and B, 2.9, 2.10, 2.12, 2.13 categories 1 and 2, 2.14 categories 1 and 2, 2.15 types A to F |  | - | - | 03 | https://echa.europa.eu/documents/10162/6e698812-f218-5f39-aff3-e46a1555a67d |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT:18429&cs=1A5FE3B096F5CE8B96EF0786DA3EF0594">EN 14059:2002</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1979.197.01.0037.01.ENG&toc=OJ:L:1979:197:TOC">D 79/663/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.315.01.0013.01.ENG&toc=OJ:L:1997:315:TOC">D 97/64/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.138.01.0008.01.ENG&toc=OJ:L:2009:138:TOC">D 2009/424/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2010.086.01.0007.01.ENG&toc=OJ:L:2010:086:TOC">R 276/2010</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1068">entry 1068</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1521">entry 1521</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1799">entry 1799</a> |  | 15-Jan-2016
Hazard class 5.1. |  | - | - | 03 | https://echa.europa.eu/documents/10162/6e698812-f218-5f39-aff3-e46a1555a67d |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT:18429&cs=1A5FE3B096F5CE8B96EF0786DA3EF0594">EN 14059:2002</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1979.197.01.0037.01.ENG&toc=OJ:L:1979:197:TOC">D 79/663/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.315.01.0013.01.ENG&toc=OJ:L:1997:315:TOC">D 97/64/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.138.01.0008.01.ENG&toc=OJ:L:2009:138:TOC">D 2009/424/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2010.086.01.0007.01.ENG&toc=OJ:L:2010:086:TOC">R 276/2010</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1068">entry 1068</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1521">entry 1521</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1799">entry 1799</a> |  | 15-Jan-2016
Tris (2,3 dibromopropyl) phosphate |  | 204-799-9 | 126-72-7 | 04 | https://echa.europa.eu/documents/10162/c05269fe-1ecd-4186-9534-aaec4e15e0f1 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1979.197.01.0037.01.ENG&toc=OJ:L:1979:197:TOC">D 79/663/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Benzene |  | 200-753-7 | 71-43-2 | 05 | https://echa.europa.eu/documents/10162/7c8cf4ac-baf9-a05a-2cc7-c9bca4a9d5b7 |  |  | <a href="https://eur-lex.europa.eu/LexUriServ/LexUriServ.do?uri=OJ:L:1982:339:0055:0056:EN:PDF">D 82/806/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.233.01.0002.01.ENG&toc=OJ:L:2015:233:TOC">R 1494/2015</a> |  |  | 13-Jan-2016
Asbestos fibres |  | - | - | 06 | https://echa.europa.eu/documents/10162/ea9dc42d-7656-8afd-09e4-d8b41fae2c9c | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=575">Appendix 7</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.263.01.0033.01.ENG&toc=OJ:L:1983:263:TOC">D 83/478/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1985.375.01.0001.01.ENG&toc=OJ:L:1985:375:TOC">D 85/610/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.363.01.0036.01.ENG&toc=OJ:L:1991:363:TOC">D 91/659/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.207.01.0018.01.ENG&toc=OJ:L:1999:207:TOC">D 1999/77/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.165.01.0004.01.ENG&toc=OJ:L:2016:165:TOC">R 2016/1005</a> |  |  | 13-Jan-2016
Actinolite |  | - | 77536-66-4 | 06 | https://echa.europa.eu/documents/10162/ea9dc42d-7656-8afd-09e4-d8b41fae2c9c | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=575">Appendix 7</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.263.01.0033.01.ENG&toc=OJ:L:1983:263:TOC">D 83/478/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1985.375.01.0001.01.ENG&toc=OJ:L:1985:375:TOC">D 85/610/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.363.01.0036.01.ENG&toc=OJ:L:1991:363:TOC">D 91/659/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.207.01.0018.01.ENG&toc=OJ:L:1999:207:TOC">D 1999/77/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.165.01.0004.01.ENG&toc=OJ:L:2016:165:TOC">R 2016/1005</a> |  |  | 13-Jan-2016
Tremolite |  | - | 77536-68-6 | 06 | https://echa.europa.eu/documents/10162/ea9dc42d-7656-8afd-09e4-d8b41fae2c9c | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=575">Appendix 7</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.263.01.0033.01.ENG&toc=OJ:L:1983:263:TOC">D 83/478/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1985.375.01.0001.01.ENG&toc=OJ:L:1985:375:TOC">D 85/610/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.363.01.0036.01.ENG&toc=OJ:L:1991:363:TOC">D 91/659/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.207.01.0018.01.ENG&toc=OJ:L:1999:207:TOC">D 1999/77/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.165.01.0004.01.ENG&toc=OJ:L:2016:165:TOC">R 2016/1005</a> |  |  | 13-Jan-2016
Anthophyllite |  | - | 77536-67-5 | 06 | https://echa.europa.eu/documents/10162/ea9dc42d-7656-8afd-09e4-d8b41fae2c9c | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=575">Appendix 7</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.263.01.0033.01.ENG&toc=OJ:L:1983:263:TOC">D 83/478/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1985.375.01.0001.01.ENG&toc=OJ:L:1985:375:TOC">D 85/610/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.363.01.0036.01.ENG&toc=OJ:L:1991:363:TOC">D 91/659/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.207.01.0018.01.ENG&toc=OJ:L:1999:207:TOC">D 1999/77/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.165.01.0004.01.ENG&toc=OJ:L:2016:165:TOC">R 2016/1005</a> |  |  | 13-Jan-2016
Chrysotile asbestos |  | - | 132207-32-0 | 06 | https://echa.europa.eu/documents/10162/ea9dc42d-7656-8afd-09e4-d8b41fae2c9c | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=575">Appendix 7</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.263.01.0033.01.ENG&toc=OJ:L:1983:263:TOC">D 83/478/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1985.375.01.0001.01.ENG&toc=OJ:L:1985:375:TOC">D 85/610/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.363.01.0036.01.ENG&toc=OJ:L:1991:363:TOC">D 91/659/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.207.01.0018.01.ENG&toc=OJ:L:1999:207:TOC">D 1999/77/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.165.01.0004.01.ENG&toc=OJ:L:2016:165:TOC">R 2016/1005</a> |  |  | 13-Jan-2016
Asbestos, chrysotile |  | 601-650-3 | 12001-29-5 | 06 | https://echa.europa.eu/documents/10162/ea9dc42d-7656-8afd-09e4-d8b41fae2c9c | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=575">Appendix 7</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.263.01.0033.01.ENG&toc=OJ:L:1983:263:TOC">D 83/478/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1985.375.01.0001.01.ENG&toc=OJ:L:1985:375:TOC">D 85/610/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.363.01.0036.01.ENG&toc=OJ:L:1991:363:TOC">D 91/659/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.207.01.0018.01.ENG&toc=OJ:L:1999:207:TOC">D 1999/77/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.165.01.0004.01.ENG&toc=OJ:L:2016:165:TOC">R 2016/1005</a> |  |  | 13-Jan-2016
Crocidolite |  | - | 12001-28-4 | 06 | https://echa.europa.eu/documents/10162/ea9dc42d-7656-8afd-09e4-d8b41fae2c9c | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=575">Appendix 7</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.263.01.0033.01.ENG&toc=OJ:L:1983:263:TOC">D 83/478/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1985.375.01.0001.01.ENG&toc=OJ:L:1985:375:TOC">D 85/610/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.363.01.0036.01.ENG&toc=OJ:L:1991:363:TOC">D 91/659/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.207.01.0018.01.ENG&toc=OJ:L:1999:207:TOC">D 1999/77/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.165.01.0004.01.ENG&toc=OJ:L:2016:165:TOC">R 2016/1005</a> |  |  | 13-Jan-2016
Amosite |  | - | 12172-73-5 | 06 | https://echa.europa.eu/documents/10162/ea9dc42d-7656-8afd-09e4-d8b41fae2c9c | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=575">Appendix 7</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.263.01.0033.01.ENG&toc=OJ:L:1983:263:TOC">D 83/478/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1985.375.01.0001.01.ENG&toc=OJ:L:1985:375:TOC">D 85/610/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.363.01.0036.01.ENG&toc=OJ:L:1991:363:TOC">D 91/659/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.207.01.0018.01.ENG&toc=OJ:L:1999:207:TOC">D 1999/77/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.165.01.0004.01.ENG&toc=OJ:L:2016:165:TOC">R 2016/1005</a> |  |  | 13-Jan-2016
Tris(aziridinyl)phosphinoxide |  | 208-892-5 | 545-55-1 | 07 | https://echa.europa.eu/documents/10162/48077533-21cf-40dc-a8a8-d32408b8cda4 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Polybromobiphenyls, Polybrominatedbiphenyls (PBB) |  | - | - | 08 | https://echa.europa.eu/documents/10162/bdd717aa-7466-40ce-8a46-08c83ecc3aeb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1522">entry 1522</a> |  | 13-Jan-2016
Hexabromo-1,1'-biphenyl |  | 252-994-2 | 36355-01-8 | 08 | https://echa.europa.eu/documents/10162/bdd717aa-7466-40ce-8a46-08c83ecc3aeb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1522">entry 1522</a> |  | 13-Jan-2016
Nonabromo-1,1'-biphenyl |  | 248-637-5 | 27753-52-2 | 08 | https://echa.europa.eu/documents/10162/bdd717aa-7466-40ce-8a46-08c83ecc3aeb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1522">entry 1522</a> |  | 13-Jan-2016
Tetrabromo(tetrabromophenyl)benzene |  | 248-696-7 | 27858-07-7 | 08 | https://echa.europa.eu/documents/10162/bdd717aa-7466-40ce-8a46-08c83ecc3aeb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1522">entry 1522</a> |  | 13-Jan-2016
Polybromobiphenyls, Polybrominatedbiphenyls (PBB) |  | - | 59536-65-1 | 08 | https://echa.europa.eu/documents/10162/bdd717aa-7466-40ce-8a46-08c83ecc3aeb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1522">entry 1522</a> |  | 13-Jan-2016
4,4'-dibromobiphenyl |  | 202-198-6 | 92-86-4 | 08 | https://echa.europa.eu/documents/10162/bdd717aa-7466-40ce-8a46-08c83ecc3aeb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1522">entry 1522</a> |  | 13-Jan-2016
3-bromobiphenyl |  | 218-304-9 | 2113-57-7 | 08 | https://echa.europa.eu/documents/10162/bdd717aa-7466-40ce-8a46-08c83ecc3aeb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1522">entry 1522</a> |  | 13-Jan-2016
Decabromo-1,1'-biphenyl |  | 237-137-2 | 13654-09-6 | 08 | https://echa.europa.eu/documents/10162/bdd717aa-7466-40ce-8a46-08c83ecc3aeb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1522">entry 1522</a> |  | 13-Jan-2016
4-bromobiphenyl |  | 202-176-6 | 92-66-0 | 08 | https://echa.europa.eu/documents/10162/bdd717aa-7466-40ce-8a46-08c83ecc3aeb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1522">entry 1522</a> |  | 13-Jan-2016
Polybrominated biphenyls (PBB) except hexabromo-biphenyl |  | - | - | 08 | https://echa.europa.eu/documents/10162/bdd717aa-7466-40ce-8a46-08c83ecc3aeb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1522">entry 1522</a> |  | 13-Jan-2016
Entry 9 |  | - | - | 09 | https://echa.europa.eu/documents/10162/bc246292-b01b-4203-90b7-a4ba81a6780b |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Soap bark powder (Quillaja saponaria) and its derivatives containing saponines | Extractives and their physically modified derivatives such as tinctures, concretes, absolutes, essential oils, oleoresins, terpenes, terpene-free fractions, distillates, residues, etc., obtained from Quillaja saponaria, Rosaceae. | 273-620-4 | 68990-67-0 | 09 | https://echa.europa.eu/documents/10162/bc246292-b01b-4203-90b7-a4ba81a6780b |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Powder of the roots of Helleborus viridis and Helleborus niger |  | - | - | 09 | https://echa.europa.eu/documents/10162/bc246292-b01b-4203-90b7-a4ba81a6780b |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
o-Nitrobenzaldehyde |  | 209-025-3 | 552-89-6 | 09 | https://echa.europa.eu/documents/10162/bc246292-b01b-4203-90b7-a4ba81a6780b |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Wood powder |  | - | - | 09 | https://echa.europa.eu/documents/10162/bc246292-b01b-4203-90b7-a4ba81a6780b |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Benzidine and / or its derivatives |  | - | - | 09 | https://echa.europa.eu/documents/10162/bc246292-b01b-4203-90b7-a4ba81a6780b |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Powder of the roots of Veratrum album and Veratrum nigrum |  | - | - | 09 | https://echa.europa.eu/documents/10162/bc246292-b01b-4203-90b7-a4ba81a6780b |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Entry 10 |  | - | - | 10 | https://echa.europa.eu/documents/10162/1c454ceb-4a7e-4d11-bde6-59bb4bfe7bb5 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Ammonium hydrogen sulphide |  | 235-184-3 | 12124-99-1 | 10 | https://echa.europa.eu/documents/10162/1c454ceb-4a7e-4d11-bde6-59bb4bfe7bb5 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Ammonium polysulphide |  | 232-989-1 | 9080-17-5 | 10 | https://echa.europa.eu/documents/10162/1c454ceb-4a7e-4d11-bde6-59bb4bfe7bb5 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Ammonium sulphide |  | 235-223-4 | 12135-76-1 | 10 | https://echa.europa.eu/documents/10162/1c454ceb-4a7e-4d11-bde6-59bb4bfe7bb5 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Volatile esters of bromoacetic acids |  | - | - | 11 | https://echa.europa.eu/documents/10162/b515d01c-ab18-44a6-9ada-72529f1573ed |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Methyl bromoacetate |  | 202-499-2 | 96-32-2 | 11 | https://echa.europa.eu/documents/10162/b515d01c-ab18-44a6-9ada-72529f1573ed |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Butyl bromoacetate |  | 242-729-9 | 18991-98-5 | 11 | https://echa.europa.eu/documents/10162/b515d01c-ab18-44a6-9ada-72529f1573ed |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Propyl bromoacetate |  | - | 35223-80-4 | 11 | https://echa.europa.eu/documents/10162/b515d01c-ab18-44a6-9ada-72529f1573ed |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Ethyl bromoacetate |  | 203-290-9 | 105-36-2 | 11 | https://echa.europa.eu/documents/10162/b515d01c-ab18-44a6-9ada-72529f1573ed |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1983.147.01.0009.01.ENG&toc=OJ:L:1983:147:TOC">D 83/264/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
2-naphthylamine and its salts |  | - | - | 12 | https://echa.europa.eu/documents/10162/1435a321-63c2-4baf-b325-1fb4bf29d950 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Salts of 2-naphthylamine |  | - | - | 12 | https://echa.europa.eu/documents/10162/1435a321-63c2-4baf-b325-1fb4bf29d950 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
2-naphthylammonium chloride |  | 210-313-6 | 612-52-2 | 12 | https://echa.europa.eu/documents/10162/1435a321-63c2-4baf-b325-1fb4bf29d950 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
2-naphthylammonium acetate |  | 209-030-0 | 553-00-4 | 12 | https://echa.europa.eu/documents/10162/1435a321-63c2-4baf-b325-1fb4bf29d950 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
2-naphthylamine |  | 202-080-4 | 91-59-8 | 12 | https://echa.europa.eu/documents/10162/1435a321-63c2-4baf-b325-1fb4bf29d950 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Benzidine and its salts |  | - | - | 13 | https://echa.europa.eu/documents/10162/3b95527c-4d0e-54f8-b1e4-7277bf458209 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Benzidine |  | 202-199-1 | 92-87-5 | 13 | https://echa.europa.eu/documents/10162/3b95527c-4d0e-54f8-b1e4-7277bf458209 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Salts of benzidine |  | - | - | 13 | https://echa.europa.eu/documents/10162/3b95527c-4d0e-54f8-b1e4-7277bf458209 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
4-Nitrobiphenyl |  | 202-204-7 | 92-93-3 | 14 | https://echa.europa.eu/documents/10162/4abe7020-a86a-409c-87ac-22ae2c89b500 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
4-Aminobiphenyl xenylamine and its salts |  | - | - | 15 | https://echa.europa.eu/documents/10162/0e1cbf82-fe57-4c3b-97a4-0917ad8f212c |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
4-Aminobiphenyl xenylamine |  | 202-177-1 | 92-67-1 | 15 | https://echa.europa.eu/documents/10162/0e1cbf82-fe57-4c3b-97a4-0917ad8f212c |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Lead carbonates |  | - | - | 16 | https://echa.europa.eu/documents/10162/22dd9386-7fac-4e8d-953a-ef3c71025ad4 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1178">entry 1178</a> |  | 13-Jan-2016
Neutral anhydrous carbonate (PbCO3 ) |  | 209-943-4 | 598-63-0 | 16 | https://echa.europa.eu/documents/10162/22dd9386-7fac-4e8d-953a-ef3c71025ad4 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1178">entry 1178</a> |  | 13-Jan-2016
Trilead-bis(carbonate)-dihydroxide 2PbCO3-Pb(OH)2 |  | 215-290-6 | 1319-46-6 | 16 | https://echa.europa.eu/documents/10162/22dd9386-7fac-4e8d-953a-ef3c71025ad4 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1178">entry 1178</a> |  | 13-Jan-2016
Lead sulphates |  | - | - | 17 | https://echa.europa.eu/documents/10162/ffd7653b-98cc-4bcc-9085-616559280314 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1178">entry 1178</a> |  | 15-Jan-2016
Lead sulphate PbSO4 |  | 231-198-9 | 7446-14-2 | 17 | https://echa.europa.eu/documents/10162/ffd7653b-98cc-4bcc-9085-616559280314 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1178">entry 1178</a> |  | 15-Jan-2016
Sulphuric acid, lead salt Pbx SO4 |  | 239-831-0 | 15739-80-7 | 17 | https://echa.europa.eu/documents/10162/ffd7653b-98cc-4bcc-9085-616559280314 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1178">entry 1178</a> |  | 15-Jan-2016
Mercury compounds |  | - | - | 18 | https://echa.europa.eu/documents/10162/5a7222b0-9d3a-4a90-9e55-258149e92b1a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Mercury |  | 231-106-7 | 7439-97-6 | 18a | https://echa.europa.eu/documents/10162/dbcaaec7-bd5b-4a7d-b164-23fa97950a86 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2007.257.01.0013.01.ENG&toc=OJ:L:2007:257:TOC">D 2007/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.253.01.0001.01.ENG&toc=OJ:L:2012:253:TOC">R 847/2012</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/658">entry 658</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/659">entry 659</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1179">entry 1179</a> |  | 13-Jan-2016
Arsenic compounds |  | - | - | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Calcium arsenate |  | 231-904-5 | 7778-44-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trisilver arsenite |  | 232-048-5 | 7784-08-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic tribromide |  | 232-057-4 | 7784-33-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic trichloride |  | 232-059-5 | 7784-34-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trifluoroarsine |  | 232-060-0 | 7784-35-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Pentafluoroarsorane |  | 232-061-6 | 7784-36-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Mercury hydrogenarsenate |  | 232-062-1 | 7784-37-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Manganese hydrogenarsenate |  | 232-063-7 | 7784-38-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Lead hydrogen arsenate |  | 232-064-2 | 7784-40-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Potassium dihydrogenarsenate |  | 232-065-8 | 7784-41-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsine |  | 232-066-3 | 7784-42-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diammonium hydrogenarsenate |  | 232-067-9 | 7784-44-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic triiodide |  | 232-068-4 | 7784-45-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium dioxoarsenate |  | 232-070-5 | 7784-46-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Pentahydroxyarsorane |  | 232-096-7 | 7786-36-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Dimethylarsinic acid |  | 200-883-4 | 75-60-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Phenylarsonic acid |  | 202-631-9 | 98-05-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-hydroxyphenylarsonic acid |  | 202-641-3 | 98-14-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsanilic acid |  | 202-674-3 | 98-50-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Nitarsone |  | 202-695-8 | 98-72-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Roxarsone |  | 204-453-7 | 121-19-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium dimethylarsinate |  | 204-708-2 | 124-65-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium hydrogen 4-aminophenylarsonate |  | 204-869-9 | 127-85-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
6,6'-dihydroxy-3,3'-diarsene-1,2-diyldianilinium dichloride |  | 205-386-6 | 139-93-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Oxophenarsine |  | 206-178-8 | 306-12-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Phenylarsine oxide |  | 211-275-3 | 637-03-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
2,7-(bis(2-arsonophenylazo))-1,8-dihydroxynaphthalene-3,6-disulphonic acid |  | 216-788-6 | 1668-00-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Gallium arsenide |  | 215-114-8 | 1303-00-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Indium arsenide |  | 215-115-3 | 1303-11-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diarsenic pentaoxide |  | 215-116-9 | 1303-28-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
arsenic sulfide |  | 215-117-4 | 1303-33-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diarsenic triselenide |  | 215-119-5 | 1303-36-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diarsenic trioxide |  | 215-481-4 | 1327-53-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Methyltriphenylarsonium iodide |  | 216-108-8 | 1499-33-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tritylium hexafluoroarsenate |  | 207-111-5 | 437-15-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Neoarsphenamine |  | 207-273-7 | 457-60-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Oxophenarsine hydrochloride |  | 208-682-3 | 538-03-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tris[(8α,9R)-6'-methoxycinchonan-9-ol] bis(arsenate) |  | 208-971-4 | 549-59-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Triphenylarsine |  | 210-032-9 | 603-32-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sulfarsphenamine |  | 210-564-1 | 618-82-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Triethyl arsenite |  | 221-543-1 | 3141-12-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disodium 3-[(o-arsonophenyl)azo]-4,5-dihydroxynaphthalene-2,7-disulphonate |  | 222-600-3 | 3547-38-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trilead diarsenate |  | 222-979-5 | 3687-31-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disodium 4-[(o-arsonophenyl)azo]-3-hydroxynaphthalene-2,7-disulphonate |  | 222-993-1 | 3688-92-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
potassium dimethylarsinate |  | 881-273-0 | 21416-85-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
thallium arsenic selenide |  | - | 16142-89-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Slimes and Sludges, copper-lead ore roasting off gas scrubbing, arsenic-contg. | The product obtained by the purification of copper-lead ore concentrate roasting offgas. Composed primarily of arsenic oxide (As2O3). | 310-063-9 | 102110-62-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tetraphenylarsonium Chloride Hydrochloride Hydrate |  | 621-302-4 | 123334-18-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium hexafluoroarsenate(V) |  | 624-772-9 | 12005-86-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Silver hexafluoroarsenate |  | 624-776-0 | 12005-82-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic(II) sulfide |  | 627-949-9 | 12044-79-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diphenyldiarsenic acid |  | 224-845-1 | 4519-32-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
2-nitrophenylarsonic acid |  | 226-485-0 | 5410-29-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
2-aminophenylarsonic acid |  | 218-064-5 | 2045-00-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Lead arsenite |  | 233-083-9 | 10031-13-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Iron arsenate |  | 233-274-7 | 10102-49-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Iron bis(arsenate) |  | 233-275-2 | 10102-50-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid, magnesium salt |  | 233-285-7 | 10103-50-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid, copper salt |  | 233-286-2 | 10103-61-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid, calcium salt |  | 233-287-8 | 10103-62-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Potassium arsonate |  | 233-337-9 | 10124-50-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Flue dust, arsenic-contg. | Formed when arsenic and metal oxide particles are driven off during the roasting and converting of copper concentrates and matte in the production of anode copper. | 232-434-3 | 8028-73-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tribarium diarsenate |  | 236-762-8 | 13477-04-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trinickel bis(arsenate) |  | 236-771-7 | 13477-70-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trilithium arsenate |  | 236-773-8 | 13478-14-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trisilver arsenate |  | 236-841-7 | 13510-44-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Copper diarsenite |  | 240-574-1 | 16509-22-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Potassium hexafluoroarsenate |  | 241-102-7 | 17029-22-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Hydrogen hexafluoroarsenate |  | 241-128-9 | 17068-85-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Strychnine arsenate |  | 233-970-0 | 10476-82-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tricopper arsenide |  | 234-472-6 | 12005-75-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Dysprosium arsenide |  | 234-473-1 | 12005-81-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diiron arsenide |  | 234-474-7 | 12005-88-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Gadolinium arsenide |  | 234-475-2 | 12005-89-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Holmium arsenide |  | 234-476-8 | 12005-92-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Lutetium arsenide |  | 234-477-3 | 12005-94-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Manganese arsenide |  | 234-478-9 | 12005-95-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Terbium arsenide |  | 234-479-4 | 12006-08-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Thallium arsenide |  | 234-481-5 | 12006-09-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Thulium arsenide |  | 234-482-0 | 12006-10-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Ytterbium arsenide |  | 234-483-6 | 12006-12-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tricadmium diarsenide |  | 234-484-1 | 12006-15-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Iron diarsenide |  | 234-485-7 | 12006-21-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trizinc diarsenide |  | 234-486-2 | 12006-40-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Dichromium arsenide |  | 235-499-6 | 12254-85-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Erbium arsenide |  | 235-501-5 | 12254-88-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Lanthanum arsenide |  | 235-502-0 | 12255-04-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Niobium arsenide |  | 235-503-6 | 12255-08-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Neodymium arsenide |  | 235-504-1 | 12255-09-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Triantimony arsenide |  | 235-505-7 | 12255-36-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Samarium arsenide |  | 235-506-2 | 12255-39-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Yttrium arsenide |  | 235-507-8 | 12255-48-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tribarium diarsenide |  | 235-508-3 | 12255-50-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tricalcium diarsenide |  | 235-509-9 | 12255-53-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Germanium arsenide |  | 235-547-6 | 12271-72-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trisilver arsenide |  | 235-652-7 | 12417-99-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic sulfide |  | 235-720-6 | 12612-21-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Iron arsenide |  | 234-947-8 | 12044-16-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Digallium arsenide phosphide |  | 234-948-3 | 12044-20-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tripotassium arsenide |  | 234-949-9 | 12044-21-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trilithium arsenide |  | 234-950-4 | 12044-22-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trisodium arsenide |  | 234-952-5 | 12044-25-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Praseodymium arsenide |  | 234-953-0 | 12044-28-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trimagnesium diarsenide |  | 234-954-6 | 12044-49-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diarsenic tritelluride |  | 234-955-1 | 12044-54-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Zinc diarsenide |  | 234-956-7 | 12044-55-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Nickel diarsenide |  | 235-103-1 | 12068-61-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid, sodium salt |  | 231-547-5 | 7631-89-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid |  | 231-901-9 | 7778-39-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disodium hydrogenarsenate |  | 231-902-4 | 7778-43-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Antimony arsenic oxide |  | 264-904-9 | 64475-90-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic bromide |  | 265-296-8 | 64973-06-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Cobalt arsenide |  | 265-784-0 | 65453-05-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tris(pentane-2,4-dionato-O,O')silicon hexafluoroarsenate |  | 266-621-6 | 67251-38-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Thallium triarsenide |  | 281-902-3 | 84057-85-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
2,6-dimethyl-4-(1-naphthyl)pyrylium hexafluoroarsenate |  | 282-682-1 | 84282-36-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
2,6-dimethyl-4-phenylpyrylium hexafluoroarsenate |  | 282-700-8 | 84304-15-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-cyclohexyl-2,6-dimethylpyrylium hexafluoroarsenate |  | 282-701-3 | 84304-16-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Silicic acid (H4SiO4), zinc salt (1:2), arsenic and manganese-doped |  | 271-895-5 | 68611-46-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Bis(pentane-2,4-dionato-O,O')boron(1+) hexafluoroarsenate(1-) |  | 272-591-5 | 68892-01-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Antimony oxide (Sb2O3), mixed with arsenic oxide (As2O3) |  | 273-156-2 | 68951-38-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Lead alloy, base, dross | A scum formed on the surface of molten lead-base alloys. Includes those cases in which aluminum is used to remove arsenic, nickel and antimony. | 273-700-9 | 69011-59-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Lead, antimonial, dross | A scum formed on the surface of antimonial lead. Consists primarily of sodium arsenate and sodium antimonate with some lead oxide and free caustic soda. | 273-795-7 | 69029-51-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Flue dust, lead-refining | By-product of refining lead ores obtained from baghouse and electro-static precipitator and as slurry from scrubbers. | 273-809-1 | 69029-67-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disilver arsenide |  | 274-573-2 | 70333-07-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium metaarsenate |  | 239-171-3 | 15120-17-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Ammonium dihydrogenarsenate |  | 236-667-1 | 13462-93-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Potassium arsenite |  | 236-680-2 | 13464-35-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trisodium arsenite |  | 236-681-8 | 13464-37-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trisodium arsenate |  | 236-682-3 | 13464-38-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Zinc arsenate |  | 236-683-9 | 13464-44-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tristrontium diarsenate |  | 236-684-4 | 13464-68-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Cobalt arsenide |  | 248-168-6 | 27016-73-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Nickel arsenide |  | 248-169-1 | 27016-75-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tricalcium diarsenite |  | 248-266-9 | 27152-57-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
3-methyl-4-(pyrrolidin-1-yl)benzenediazonium hexafluoroarsenate |  | 248-532-4 | 27569-09-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Antimony arsenate |  | 249-347-1 | 28980-47-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tristrontium diarsenide |  | 254-407-5 | 39297-24-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid, copper(2+) salt |  | 249-916-4 | 29871-13-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Lithium hexafluoroarsenate |  | 249-963-0 | 29935-35-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Ammonium copper arsenate |  | 251-151-6 | 32680-29-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Europium arsenide |  | 251-206-4 | 32775-46-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
N-(p-arsenosophenyl)-1,3,5-triazine-2,4,6-triamine |  | 244-612-8 | 21840-08-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Aluminium arsenide |  | 245-255-0 | 22831-42-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Triammonium arsenate |  | 246-428-3 | 24719-13-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tricobalt diarsenate |  | 246-429-9 | 24719-19-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Calcium arsenite |  | 258-147-3 | 52740-16-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Triphenylsulphonium hexafluoroarsenate(1-) |  | 261-009-5 | 57900-42-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Zirconium arsenide |  | 262-524-8 | 60909-47-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trimanganese arsenide |  | 262-667-6 | 61219-26-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disodium 3,6-bis[(o-arsonophenyl)azo]-4,5-dihydroxynaphthalene-2,7-disulphonate |  | 263-516-7 | 62337-00-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diphenyliodonium hexafluoroarsenate |  | 263-638-0 | 62613-15-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-(ethylamino)-2-methylbenzenediazonium hexafluoroarsenate |  | 264-026-6 | 63217-32-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-(diethylamino)-2-ethoxybenzenediazonium hexafluoroarsenate |  | 264-027-1 | 63217-33-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
2,4,6,8,9,10-Hexaoxa-1,3,5,7-tetraarsatricyclo[3.3.1.13,7]decane |  | - | 12505-67-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
tris(trimethylsilyl)arsine |  | 821-113-9 | 17729-30-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
ARSENIC(V) OXIDE HYDRATE |  | 629-626-8 | 12044-50-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), monocesium salt |  | 677-899-7 | 16331-85-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium arsenate dibasic heptahydrate |  | 677-900-0 | 10048-95-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium cacodylate trihydrate |  | 682-793-9 | 6131-99-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
3-[(dimethylcarbamothioyl)thio]-N,N,6-trimethyl-1,5-bis(thioxo)-2,4-dithia-6-aza-3-arsaheptan-1-amine; asomate |  | 803-606-0 | 3586-60-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsonic acid, [4-[[4-(dimethylamino)phenyl]azo]phenyl]-, monohydrochloride |  | 620-678-7 | 73688-85-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), copper(2+) salt (2:3), tetrahydrate |  | - | 13478-34-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenic acid |  | - | 10102-53-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), monosodium salt |  | - | 10103-60-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diarsenic acid, sodium salt |  | - | 13464-42-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), zinc salt (1:1) |  | - | 13464-33-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, zinc salt |  | - | 10326-24-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tris[(8α)-6'-methoxycinchonan-9(R)-ol] arsenite |  | 303-002-2 | 94138-87-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Gallium zinc triarsenide |  | 308-577-3 | 98106-56-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Vanadium(4+) diarsenate (1:1) |  | 308-917-0 | 99035-51-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Strychnidin-10-one, arsenite (1:1) |  | 309-388-9 | 100258-44-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Slimes and Sludges, copper electrolytic refining, decopperized, arsenic-rich | Product obtained by centrifuging the slime discharged at the bases of cells for decopperization of electrolytic copper solutions. Composed primarily of a copper powder rich in arsenic. | 309-772-6 | 100995-81-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), magnesium salt, manganese-doped |  | 310-019-9 | 102110-21-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Bromo(hydroxytetraphenylarsoranato)magnesium |  | 264-254-6 | 63468-73-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trisodium 3-[(o-arsonatophenyl)azo]-4,5-dihydroxynaphthalene-2,7-disulphonate |  | 266-069-6 | 66019-20-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tetraphenylarsonium chloride, compound with hydrochloric acid (1:1) |  | 277-208-5 | 73003-83-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Strychnidin-10-one, compd. with methylarsonate (1:1) |  | 279-614-8 | 80879-64-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Methyl 17α-hydroxyyohimban-16α-carboxylate, methylarsonate (1:1) |  | 279-628-4 | 80925-03-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disodium 3,6-bis[[2-[(dihydroxyarsino)oxy]phenyl]azo]-4,5-dihydroxynaphthalene-2,7-disulphonate |  | 282-453-6 | 84215-47-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disodium 3-[[2-[(dihydroxyarsino)oxy]phenyl]azo]-4,5-dihydroxynaphthalene-2,7-disulphonate |  | 282-454-1 | 84215-48-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Dibismuth tris(methylarsonate) |  | 286-401-3 | 85237-42-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disodium hydrogen 2-[[7-[(2-arsonophenyl)azo]-1,8-dihydroxy-3,6-disulphonato-2-naphthyl]azo]benzoate |  | 238-715-7 | 14674-83-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
2,2'-[ethylenebis(oxy)]bis[1,3,2-dioxarsolane] |  | 238-911-2 | 14849-23-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Chlorotris(triphenylarsine)rhodium |  | 239-051-0 | 14973-92-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Proustite (Ag3(AsS3)) |  | 239-179-7 | 15122-58-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
o-phenylenebis(dimethylarsine) |  | 236-227-9 | 13246-32-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Melarsonyl potassium |  | 236-405-6 | 13355-00-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Thiobis[methylarsine], anhydrosulphide |  | 236-436-5 | 13367-92-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tetraphenylarsonium (hydrogen dichloride) |  | 244-144-4 | 21006-73-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Methylenebis(diphenylarsine) |  | 244-639-5 | 21892-63-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diphenylarsinecarbonitrile |  | 245-716-6 | 23525-22-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Bis[(2-diphenylarsinoethyl)phenyl]phosphine |  | 245-756-4 | 23582-05-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
[2-(diphenylarsino)ethyl]diphenylphosphine |  | 245-758-5 | 23582-06-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trisodium 4-[(o-arsonophenyl)azo]-3-oxidonaphthalene-2,7-disulphonate |  | 258-693-2 | 53669-45-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Acetarsol sodium |  | 259-714-8 | 55588-51-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
3-formamido-4-hydroxyphenylarsonic acid |  | 262-085-2 | 60154-16-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Calcium arsenate (CaHAsO4) (6CI,7CI) |  | 672-782-7 | 15195-00-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Aluminium gallium arsenide |  | 680-872-2 | 37382-15-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsine, dichloromethyl- |  | 695-189-5 | 593-89-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic oxide |  | - | 12777-38-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), ammonium salt (2:1) |  | - | 127795-79-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Nickel arsenide |  | - | 12795-30-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Rammelsbergite (NiAs2) |  | - | 1303-22-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Zinc arsenenate oxide (Zn5(AsO3)4O3), tetrahydrate |  | - | 1303-39-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diarsenic acid |  | - | 13453-15-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), copper(2+) salt (1:1) |  | - | 13464-31-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid, calcium salt (1:2) |  | - | 13464-39-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsonic acid, disodium salt |  | - | 13466-06-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Boron arsenate |  | - | 13510-31-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Zinc arsenate, octahydrate |  | - | 13510-72-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic chloride (AsCl) |  | - | 17522-78-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
(2-Chloroethenyl)arsonous dichloride |  | - | 541-25-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenous acid, gallium salt (1:1) |  | - | 24343-41-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Manganese arsenate |  | - | 27526-45-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), cobalt(2+) salt |  | - | 29871-10-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
arsenic acid (H3AsO4), calcium salt(2:3), dihydrate |  | - | 30621-31-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenous acid, tricopper(1+) salt |  | - | 33382-64-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsonic acid, methyl-, iron salt (9CI) |  | - | 33972-75-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, nickel(2+) salt |  | - | 33992-49-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic chloride |  | - | 37226-49-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Iron arsenide |  | - | 39310-41-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic chloride (AsCl2) |  | - | 41996-37-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenous acid, lead(2+) salt (1:1) |  | - | 100822-74-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
(4-aminophenyl)arsonic acid, compound with piperazine (1:1) |  | 303-952-8 | 94232-26-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Leucomycin V, 9-O-[5-(dimethylamino)tetrahydro-6-methyl-2H-pyran-2-yl]-, [9(5S,6R)]-, [1,2-ethanediylbis(imino-4,1-phenylene)]bis[arsonate] (1:1) (salt) |  | 299-051-1 | 93841-79-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
(diphenylarsino)dimethylgallium |  | 302-587-1 | 94113-53-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Copper acetoarsenite |  | 601-658-7 | 12002-03-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium hydrogen allylarsonate |  | 304-710-4 | 94278-22-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disodium p-tolylarsonate |  | 304-956-2 | 94313-58-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
2-[[7-[(2-arsonophenyl)azo]-1,8-dihydroxy-3,6-disulpho-2-naphthyl]azo]benzoic acid |  | 223-216-9 | 3772-44-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
p-tolylarsonic acid |  | 223-588-2 | 3969-54-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Ethylenebis(diphenylarsine) |  | 224-629-7 | 4431-24-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
10-[(dimethylthiocarbamoyl)thio]-5,10-dihydrophenarsazine |  | 225-368-1 | 4808-24-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-(2-chloroacetamido)phenylarsonic acid |  | 226-569-7 | 5425-62-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-chlorophenylarsonic acid |  | 226-622-4 | 5440-04-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Calcium bis(dimethylarsinate) |  | 227-311-6 | 5785-43-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
[4-[(4,6-diamino-1,3,5-triazin-2-yl)amino]phenyl]arsonic acid |  | 227-365-0 | 5806-89-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium hydrogen (3-acetamido-4-hydroxyphenyl)arsonate |  | 227-573-1 | 5892-48-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Calcium bis(methylarsonate) |  | 227-598-8 | 5902-95-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Iron tris(dimethylarsinate) |  | 227-759-2 | 5968-84-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
3-amino-4-hydroxyphenylarsonic acid |  | 218-494-3 | 2163-77-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium methylarsonate |  | 218-495-9 | 2163-80-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Copper arsonate |  | 233-644-8 | 10290-12-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Strychnine dimethylarsinate |  | 233-973-7 | 10476-87-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium 4-(glycolloylamino)phenylarsonate |  | 231-680-9 | 7681-83-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diphenoxarsin-10-yl oxide |  | 200-377-3 | 58-36-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Acetarsol |  | 202-582-3 | 97-44-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Propylarsonic acid |  | 203-482-2 | 107-34-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Glycobiarsol |  | 204-143-1 | 116-49-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsthinol |  | 204-361-7 | 119-96-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Carbarsone |  | 204-484-6 | 121-59-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Methylarsonic acid |  | 204-705-6 | 124-58-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disodium arsonoacetate |  | 204-806-5 | 126-82-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-[(2-arsonophenyl)azo]-3-hydroxynaphthalene-2,7-disulphonic acid |  | 205-058-2 | 132-33-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium hydrogen [4-[(hydroxyacetyl)amino]phenyl]arsonate |  | 205-415-2 | 140-45-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disodium methylarsonate |  | 205-620-7 | 144-21-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-(glycolloylamino)phenylarsonic acid |  | 205-643-2 | 144-87-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Iododimethylarsine |  | 211-630-2 | 676-75-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Dichloro(phenyl)arsine |  | 211-791-9 | 696-28-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Chlorodiphenylarsine |  | 211-921-4 | 712-48-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diphenylarsine |  | 212-589-3 | 829-83-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Methyl 1,2,5,6-tetrahydro-1-methylnicotinate, mono[(3-acetamido-4-hydroxyphenyl)arsonate] |  | 212-983-5 | 900-77-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Methyldiphenylarsine |  | 213-414-3 | 945-48-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Triphenylarsine oxide |  | 214-571-0 | 1153-05-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
o-phenylenediarsonic acid |  | 217-154-1 | 1758-48-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium p-[[4-[3-(2-arsono-4-nitrophenyl)triazen-1-yl]phenyl]azo]benzenesulphonate |  | 217-197-6 | 1772-02-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tetramethyldiarsine |  | 207-440-4 | 471-35-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Melarsoprol |  | 207-793-4 | 494-79-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tetraphenylarsonium bromide |  | 208-069-0 | 507-27-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tetraphenylarsonium chloride |  | 208-070-6 | 507-28-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Difetarsone disodium |  | 208-209-0 | 515-76-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trisodium 2-(-hydrogen arsonatophenylazo)-1,8-dihydroxynaphthalene-3,6-disulphonate |  | 208-285-5 | 520-10-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Acetarsol-diethylamine (1:1) |  | 208-597-1 | 534-33-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tryparsamide |  | 209-070-9 | 554-72-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
10-chloro-5,10-dihydrophenarsazine |  | 209-433-1 | 578-94-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium hydrogen [4-(acetamido)phenyl]arsonate |  | 209-558-1 | 585-54-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Allylarsonic acid |  | 209-678-4 | 590-34-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Methyloxoarsine |  | 209-799-2 | 593-58-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Trimethylarsine |  | 209-815-8 | 593-88-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Dichloro(ethyl)arsine |  | 209-919-3 | 598-14-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Triethylarsine |  | 210-526-4 | 617-75-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-acetamidophenylarsonic acid |  | 210-541-6 | 618-22-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-(4-dimethylaminophenylazo)phenylarsonic acid |  | 210-750-2 | 622-68-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Diiodo(phenyl)arsine |  | 228-965-5 | 6380-34-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium p-arsonobenzenesulphonate |  | 229-632-7 | 6634-88-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-(4-aminophenylazo)phenylarsonic acid |  | 230-176-6 | 6966-64-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
10-chloro-10H-phenoxarsine |  | 220-684-6 | 2865-70-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tri-p-tolylarsine |  | 220-776-6 | 2896-10-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Disodium [4-[(4,6-diamino-1,3,5-triazin-2-yl)amino]phenyl]arsonate |  | 222-749-4 | 3599-28-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Difetarsone |  | 222-866-0 | 3639-19-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
4-Aminophenylarsenoxide |  | - | 1122-90-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Magnesium arsenate, decahydrate |  | - | 117746-50-6 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Manganese arsenide (Mn2As) |  | - | 12005-96-8 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tin arsenide (SnAs) |  | - | 12044-32-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Cobalt arsenide (CoAs2) |  | - | 12044-42-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Platinum arsenide (PtAs2) |  | - | 12044-52-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Cobalt arsenide (CoAs3) |  | - | 12256-04-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic sulfide (As2S4) |  | - | 12344-68-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenopyrite, cobaltoan |  | - | 12414-94-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenous acid, antimony(3+) salt (1:1) |  | - | 12523-20-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Silicic acid (H4SiO4), magnesium manganese(2+) zinc salt, arsenic and lead-doped |  | - | 68784-76-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Silicic acid (H4SiO4), tetraethyl ester, polymer with arsenic oxide (As2O3) |  | - | 68957-75-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, silver(1+) salt |  | - | 69507-43-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Benzenesulfonic acid, 4-arsenoso- |  | - | 71130-51-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, lithium salt |  | - | 72845-34-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, neodymium(3+) salt |  | - | 76871-63-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, lutetium(3+) salt |  | - | 76871-65-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, cadmium neodymium(3+) salt (5:1:1) |  | - | 76871-66-4 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), copper(2+) salt (2:3) |  | - | 76407-89-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, cesium salt |  | - | 82005-78-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, mercury(2+) salt |  | - | 82980-40-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, antimony(3+) salt |  | - | 83877-96-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, cadmium salt |  | - | 84953-43-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, cobalt(2+) salt |  | - | 85561-28-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, manganese(2+) salt |  | - | 85561-29-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, rubidium salt |  | - | 85906-44-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), calcium salt (7:10) |  | - | 85949-61-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, praseodymium(3+) salt |  | - | 86859-92-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, cerium(3+) salt |  | - | 86859-93-0 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), copper(2+) salt (1:1), sesquihydrate |  | - | 88442-64-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), copper(2+) salt (4:5) |  | - | 89054-01-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), calcium salt (4:5) |  | - | 89067-81-2 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenous acid, titanium salt |  | - | 92278-94-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsonic acid, tin salt |  | - | 93080-00-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, ammonium copper salt |  | - | 94854-78-1 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenenous acid, ammonium zinc salt |  | - | 94854-80-5 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Arsenic acid (H3AsO4), monopotassium salt, 1/5hydrate |  | - | 164170-84-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tripotassium arsenate 3/2hydrate |  | - | 62702-52-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Sodium oxidoarsonous acid |  | - | 60189-99-3 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
(4-((2-Amino-2-oxoethyl)amino)phenyl)arsonic acid |  | - | 618-25-7 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Tin arsenide (Sn4As3) |  | - | 12397-66-9 | 19 | https://echa.europa.eu/documents/10162/a798c758-371f-41e5-a38d-5f8dc9ba739d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.004.01.0009.01.ENG&toc=OJ:L:2003:004:TOC">D 2003/2/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.384.01.0094.01.ENG&toc=OJ:L:2006:384:TOC">D 2006/139/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/660">entry 660</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/661">entry 661</a> |  | 13-Jan-2016
Organostannic compounds |  | - | - | 20 | https://echa.europa.eu/documents/10162/0979d035-4a80-0a4e-ddcc-03d31d5467d6 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2002.183.01.0058.01.ENG&toc=OJ:L:2002:183:TOC">D 2002/62/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.138.01.0011.01.ENG&toc=OJ:L:2009:138:TOC">D 2009/425/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2010.086.01.0007.01.ENG&toc=OJ:L:2010:086:TOC">R 276/2010</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/157">entry 157</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1069">entry 1069</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1308">entry 1308</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1309">entry 1309</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1801">entry 1801</a> |  | 13-Jan-2016
Di-μ-oxo-di-n-butylstanniohydroxyborane / Dibutyltin hydrogen borate C8H19BO3Sn (DBB) | DBB | 401-040-5 | 75113-37-0 | 21 | https://echa.europa.eu/documents/10162/ce4f0c81-1ba1-4b5d-8d6e-6e1e05a36902 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1989.398.01.0019.01.ENG&toc=OJ:L:1989:398:TOC">D 89/677/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Cadmium and its compounds |  | - | - | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
cadmium zinc telluride |  | - | 303114-50-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Leach residues, cadmium-contg. flue dust | The substance formed during oxidative leaching of cadmium containing flue dust. Consists primarily of cadmium, lead and zinc compounds with chlorine, oxygen and sulfur and contains other nonferrous metal compounds. | 305-417-4 | 94551-70-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium isohexadecanoate |  | 306-072-2 | 95892-12-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium diisobutyl dimaleate |  | 306-446-5 | 97259-82-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmate(1-), triiodo-, potassium |  | - | 14429-88-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Zircon, cadmium orange |  | 309-029-6 | 99749-34-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium chloride phosphate (Cd5Cl(PO4)3), manganese-doped |  | 309-489-8 | 100402-53-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Flue dust, copper-lead blast furnace, cadmium-indium-enriched | A cadmium-indium-enriched product obtained from the recirculation of copper-lead blast furnace flue dusts. Composed primarily of cadmium, indium and lead. | 309-645-5 | 100656-55-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Dodecanoic acid, cadmium salt, basic |  | 309-789-9 | 101012-89-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Octadecanoic acid, cadmium salt, basic |  | 309-794-6 | 101012-93-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Octadecanoic acid, 12-hydroxy-, cadmium salt, basic |  | 309-795-1 | 101012-94-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium oxide (CdO), solid soln. with calcium oxide and titanium oxide (TiO2), praseodymium-doped |  | 309-896-0 | 101356-99-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenide (CdSe), solid soln. with cadmium sulfide, zinc selenide and zinc sulfide, aluminum and copper-doped |  | 309-897-6 | 101357-00-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenide (CdSe), solid soln. with cadmium sulfide, zinc selenide and zinc sulfide, copper and manganese-doped |  | 309-898-1 | 101357-01-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenide (CdSe), solid soln. with cadmium sulfide, zinc selenide and zinc sulfide, europium-doped |  | 309-899-7 | 101357-02-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenide (CdSe), solid soln. with cadmium sulfide, zinc selenide and zinc sulfide, gold and manganese-doped |  | 309-900-0 | 101357-03-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenide (CdSe), solid soln. with cadmium sulfide, zinc selenide and zinc sulfide, manganese and silver-doped |  | 309-901-6 | 101357-04-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium oxide (CdO), solid soln. with magnesium oxide, tungsten oxide (WO3) and zinc oxide |  | 310-029-3 | 102110-30-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Silicic acid, zirconium salt, cadmium pigment-encapsulated |  | 310-077-5 | 102184-95-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium (II) chloride monohydrate |  | 621-024-3 | 35658-65-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium compounds |  | - | - | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Sulfuric acid, cadmium salt (1:1), hydrate |  | 626-360-4 | 15244-35-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium perchlorate hexahydrate |  | 629-339-8 | 10326-28-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium chloride hydrate |  | 629-592-4 | 654054-66-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bromide tetrahydrate |  | 629-623-1 | 13464-92-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
cadmium nitrate tetrahydrate |  | 638-751-7 | 10022-68-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
cadmium chloride (CdCl2), hydrate (2:5) |  | 640-998-0 | 7790-78-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium zinc lithopone yellow | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77205:1. | 292-385-9 | 90604-89-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium lithopone yellow | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77199:1. | 292-386-4 | 90604-90-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Leach residues, cadmium cake | Residues obtained by cementation of cadmium by iron dust out of cadmium sulfate solutions. Composed primarily of metallic cadmium and zinc. | 293-309-7 | 91053-44-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Leach residues, zinc ore-calcine, cadmium-copper ppt. | Insoluble material precipitated by hydrolysis during hydrometallurgical treatment of crude zinc sulfate solution. Consists primarily of cadmium, cobalt, copper, lead, manganese, nickel, thallium, tin and zinc. | 293-311-8 | 91053-46-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, castor-oil, hydrogenated, cadmium salts |  | 294-296-0 | 91697-35-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, C8-10-branched, cadmium salts |  | 296-092-7 | 92257-06-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Leach residues, zinc refining flue dust, cadmium-thallium ppt. | Sponge produced by leaching and precipitating cadmium and thallium fumes and flue dusts from lead/zinc smelting operations. | 296-097-4 | 92257-11-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, C9-13-neo-, cadmium salts |  | 296-441-3 | 92704-12-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, olive-oil, cadmium salts |  | 296-445-5 | 92704-15-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, peanut-oil, cadmium salts |  | 296-449-7 | 92704-19-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, rape-oil, cadmium salts |  | 296-454-4 | 92704-24-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, C14-18 and C18-unsatd., branched and linear, hydrogenated, cadmium salts |  | 296-564-2 | 92797-28-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Nonanoic acid, branched, cadmium salt |  | 297-692-1 | 93686-40-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Carbonic acid, cadmium salt |  | 298-586-8 | 93820-02-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium di(acetate), dihydrate |  | 611-525-5 | 5743-04-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Sulfuric acid, cadmium salt, hydrate (3:3:8) |  | 616-572-5 | 7790-84-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Barium cadmium tetrastearate |  | 214-740-9 | 1191-79-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium oxide |  | 215-146-2 | 1306-19-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulphide |  | 215-147-8 | 1306-23-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenide |  | 215-148-3 | 1306-24-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium telluride |  | 215-149-9 | 1306-25-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium dioleate |  | 233-954-3 | 10468-30-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenide sulphide |  | 234-342-9 | 11112-63-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Tricadmium diarsenide |  | 234-484-1 | 12006-15-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium titanium trioxide |  | 234-593-4 | 12014-14-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Tricadmium diphosphide |  | 234-595-5 | 12014-28-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Antimony, compound with cadmium (2:3) |  | 234-596-0 | 12014-29-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium zirconium trioxide |  | 235-251-7 | 12139-23-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Pentacadmium chloridetriphosphate |  | 235-353-1 | 12185-64-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Dicadmium niobate |  | 235-357-3 | 12187-14-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Dicadmium selenide sulphide |  | 235-392-4 | 12214-12-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium ditantalum hexaoxide |  | 235-561-2 | 12292-07-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium zinc sulphide |  | 235-672-6 | 12442-27-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenide sulfide |  | 235-724-8 | 12626-36-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfoselenide orange | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77202. | 235-758-3 | 12656-57-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium |  | 231-152-8 | 7440-43-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium fluoride |  | 232-222-0 | 7790-79-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium iodide |  | 232-223-6 | 7790-80-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium iodate |  | 232-224-1 | 7790-81-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium dinitrite |  | 232-225-7 | 7790-83-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium wolframate |  | 232-226-2 | 7790-85-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium zinc sulfide yellow | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77205. | 232-466-8 | 8048-07-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Tris(ethylenediamine)cadmium dihydroxide |  | 238-945-8 | 14874-24-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium diicosanoate |  | 238-994-5 | 14923-81-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis(piperidine-1-carbodithioate) |  | 239-025-9 | 14949-59-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Bis(dimethyldithiocarbamato-S,S')cadmium |  | 239-026-4 | 14949-60-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bromide |  | 232-165-1 | 7789-42-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Lauric acid, barium cadmium salt |  | 239-371-0 | 15337-60-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Disodium tetrakis(cyano-C)cadmate(2-) |  | 239-765-2 | 15682-87-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Dipotassium [[N,N'-ethylenebis[N-(carboxymethyl)glycinato]](4-)-N,N',O,O',ON,ON']cadmate(2-) |  | 239-801-7 | 15708-29-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium acrylate |  | 239-835-2 | 15743-19-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium tellurium trioxide |  | 239-963-9 | 15851-44-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium tellurium tetraoxide |  | 239-973-3 | 15852-14-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium dilactate |  | 240-181-5 | 16039-55-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium divanadium hexoxide |  | 240-203-3 | 16056-72-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
5-oxo-L-proline, cadmium salt |  | 240-269-3 | 16105-06-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Dimethylcadmium |  | 208-055-4 | 506-82-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium cyanide |  | 208-829-1 | 542-83-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium di(acetate) |  | 208-853-2 | 543-90-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium carbonate |  | 208-168-9 | 513-78-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium di(octanoate) |  | 218-585-8 | 2191-10-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium distearate |  | 218-743-6 | 2223-93-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium 4-(1,1-dimethylethyl)benzoate |  | 224-022-7 | 4167-05-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium cinnamate |  | 224-509-4 | 4390-97-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium diformate |  | 224-729-0 | 4464-23-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sebacate |  | 224-754-7 | 4476-04-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium nonan-1-oate |  | 225-839-1 | 5112-16-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium dipalmitate |  | 229-199-4 | 6427-86-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium dianthranilate |  | 230-343-3 | 7058-55-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium p-toluate |  | 219-345-5 | 2420-97-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis(2-ethylhexanoate) |  | 219-346-0 | 2420-98-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium dilaurate |  | 220-017-9 | 2605-44-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium didecanoate |  | 220-650-0 | 2847-16-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis[benzoate] |  | 221-187-7 | 3026-22-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium chloride |  | 233-296-7 | 10108-64-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulphate |  | 233-331-6 | 10124-36-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium myristate |  | 233-489-6 | 10196-67-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium nitrate |  | 233-710-6 | 10325-94-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium succinate |  | 205-446-1 | 141-00-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium oxalate |  | 212-408-8 | 814-88-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium dithiocyanate |  | 212-738-2 | 865-38-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium [R-(R*,R*)]-tartrate |  | 251-827-0 | 34100-40-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium didocosanoate |  | 251-927-4 | 34303-23-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Disodium [[N,N'-ethylenebis[N-(carboxymethyl)glycinato]](4-)-N,N',O,O',ON,ON']cadmate(2-) |  | 252-736-9 | 35803-35-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium 3,5,5-trimethylhexanoate |  | 252-918-8 | 36211-44-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium(2+) (R)-12-hydroxyoctadecanoate |  | 253-979-3 | 38517-19-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Bis(ethylenediamine)cadmium(2+) bis[dicyanoaurate(1-)] |  | 242-708-4 | 18974-20-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium diphenolate |  | 242-727-8 | 18991-05-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis(dipentyldithiocarbamate) |  | 242-747-7 | 19010-65-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium disalicylate |  | 242-749-8 | 19010-79-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium hydroxide |  | 244-168-5 | 21041-95-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Bis(5-oxo-L-prolinato-N1,O2)cadmium |  | 288-974-5 | 85958-86-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Bis(5-oxo-DL-prolinato-N1,O2)cadmium |  | 289-081-3 | 85994-31-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Benzenesulfonic acid, mono-C10-13-alkyl derivs., cadmium salts |  | 290-645-6 | 90194-35-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Benzoic acid, cadmium salt, basic |  | 290-764-3 | 90218-85-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Decanoic acid, branched, cadmium salts |  | 291-155-5 | 90342-19-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), solid soln. with zinc sulfide, aluminum and cobalt and copper and silver-doped |  | 272-220-7 | 68784-10-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Barium cadmium calcium chloride fluoride phosphate, antimony and manganese-doped |  | 272-257-9 | 68784-55-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, tall-oil, cadmium salts |  | 272-499-5 | 68855-80-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, C8-18 and C18-unsatd., cadmium salts |  | 272-529-7 | 68876-84-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), aluminum and copper-doped |  | 272-539-1 | 68876-98-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), aluminum and silver-doped |  | 272-540-7 | 68876-99-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), copper chloride-doped |  | 272-541-2 | 68877-00-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), silver chloride-doped |  | 272-542-8 | 68877-01-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), copper and lead-doped |  | 272-581-0 | 68891-87-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, tallow, hydrogenated, cadmium salts |  | 273-203-7 | 68953-39-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Resin acids and Rosin acids, cadmium salts |  | 273-320-3 | 68956-81-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Hydrogen [4-[(5-chloro-4-methyl-2-sulphophenyl)azo]-3-hydroxynaphthalene-2-carboxylato(3-)]cadmate(1-) |  | 273-461-0 | 68966-97-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium, dross | A scum formed on the surface of molten cadmium. | 273-707-7 | 69011-69-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Wastewater, cadmium sulfate electrolytic, acid | Electrolytic solution from electrolysis of cadmium sulfate consisting primarily of cadmium sulfate and sulfuric acid. | 273-721-3 | 69012-21-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Flue dust, cadmium-refining | By-product of refining of cadmium consisting primarily of oxides and chlorides of cadmium, lead, arsenic and zinc. | 273-754-3 | 69012-57-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Calcines, cadmium residue | Product of the roasting of cadmium-enriched lead smelting dusts to remove cadmium. Consists primarily of oxides and sulfates of lead and zinc. | 273-806-5 | 69029-63-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Leach residues, cadmium-refining | Product of leaching calcine and sump tank mud from lead ore refining with sulfuric acid. Consists primarily of lead sulfate and cadmium arsenate. | 273-811-2 | 69029-70-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Residues, cadmium-refining | Product from the washing of sweeps and cleanings from a cadmium plant. Consists primarily of metallic cadmium and iron. | 273-819-6 | 69029-77-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Slimes and Sludges, cadmium-refining, oxidized | Product of adding oxidizer to solution in the cadmium plant. Consists primarily of hydroxides of cadmium, thallium and indium and cadmium arsenate. | 273-831-1 | 69029-90-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Slimes and Sludges, cadmium sump tank | Product of adding sodium carbonate to solutions in the cadmium plant. Consists primarily of cadmium carbonate with lesser amounts of carbonates and hydroxides of other nonferrous metals. | 273-832-7 | 69029-91-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium(2+) 12-hydroxyoctadecanoate |  | 273-881-4 | 69121-20-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium potassium 1-(hydroxyethylidene)bisphosphonate(1:2:1) |  | 273-906-9 | 69190-99-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, C12-18, barium cadmium salts |  | 274-304-9 | 70084-75-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Hexanoic acid, 2-ethyl-, cadmium salt, basic |  | 291-438-3 | 90411-62-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Propanoic acid, cadmium salt, basic |  | 292-013-5 | 90529-78-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Bis(2-ethylhexyl mercaptoacetato -O',S)cadmium |  | 299-281-2 | 93858-50-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis(o-nonylphenolate) |  | 299-701-4 | 93894-07-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis(p-nonylphenolate) |  | 299-703-5 | 93894-08-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis[p-(1,1,3,3-tetramethylbutyl)phenolate] |  | 299-704-0 | 93894-09-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium (Z)-hexadec-9-enoate |  | 299-705-6 | 93894-10-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium isodecanoate |  | 300-973-4 | 93965-24-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis(isoundecanoate) |  | 300-980-2 | 93965-30-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium dimethylhexanoate |  | 301-320-6 | 93983-65-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium tetrapentyl bis(phosphate) |  | 303-977-4 | 94232-49-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium isooctyl phthalate (1:2:2) |  | 304-193-5 | 94247-16-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium (1-ethylhexyl) phthalate (1:2:2) |  | 304-482-6 | 94275-93-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium octyl phthalate (1:2:2) |  | 304-483-1 | 94275-94-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium methacrylate |  | 246-183-2 | 24345-60-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Potassium [N,N-bis(carboxymethyl)glycinato(3-)-N,O,O',O'']cadmate(1-) |  | 256-488-2 | 49784-42-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Bis[N,N-bis(carboxymethyl)glycinato(3-)]tricadmium |  | 256-679-0 | 50648-02-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Boric acid, cadmium salt |  | 257-067-6 | 51222-60-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium o-toluate |  | 257-860-7 | 52337-78-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Bis(propane-1,2-diyldiamine-N,N')cadmium(2+) bis[bis(cyano-C)aurate(1-)] |  | 267-692-6 | 67906-19-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium dilinoleate |  | 267-845-7 | 67939-62-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Tetrapotassium [[[nitrilotris(methylene)]tris[phosphonato]](6-)-N,O,O'',O'''']cadmate(6-) |  | 268-020-4 | 67989-93-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium m-toluate |  | 268-458-6 | 68092-45-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, C10-18, cadmium salts |  | 268-620-6 | 68131-58-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, C12-18, cadmium salts |  | 268-621-1 | 68131-59-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Benzyltriphenylphosphonium tetrachlorocadmate |  | 269-289-0 | 68214-25-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), solid soln. with zinc sulfide, copper and lead-doped |  | 269-773-1 | 68332-81-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, C14-18, cadmium salts |  | 270-067-0 | 68409-82-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium, benzoate p-tert-butylbenzoate complexes |  | 270-824-5 | 68478-53-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Pyrochlore, bismuth cadmium ruthenium | An inorganic pigment that is the reaction product of high temperature calcination in which bismuth oxide, cadmium oxide, and ruthenium oxide in varying amounts are homogeneously and ionically interdiffused to form a crystalline matrix of pyrochlore. | 270-855-4 | 68479-13-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis(4-cyclohexylbutyrate) |  | 259-767-7 | 55700-14-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium divalerate |  | 260-498-2 | 56982-42-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Pentapotassium hydrogen [[[ethylenebis[nitrilobis(methylene)]]tetrakis[phosphonato]](8-)]cadmate(6-) |  | 269-685-3 | 68309-98-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfoselenide red | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77202. | 261-218-1 | 58339-34-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Naphthenic acids, cadmium salts |  | 263-053-0 | 61789-34-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium neodecanoate |  | 263-352-6 | 61951-96-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis(heptadecanoate) |  | 263-434-1 | 62149-56-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium pentadecanoate |  | 264-124-9 | 63400-09-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
(S)-dichloro[2-[[(2,3-dihydroxypropoxy)hydroxyphosphinyl]oxy]triethylmethylammoniumato]cadmium |  | 265-010-1 | 64681-08-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenide (CdSe), solid soln. with cadmium sulfide |  | 275-290-7 | 71243-75-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
(R)-12-hydroxyoleic acid, barium cadmium salt |  | 275-370-1 | 71411-66-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Tetra-μ-chlorodichlorobis[2-[[(2,3-dihydroxypropoxy)hydroxyphosphinyl]oxy]triethylmethylammoniumato]tricadmium, stereoisomer |  | 276-100-5 | 71861-27-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, coco, cadmium salts |  | 276-952-8 | 72869-63-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Zircon, cadmium yellow |  | 277-135-9 | 72968-34-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium isononanoate |  | 283-660-4 | 84696-56-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium isooctadecanoate |  | 284-428-5 | 84878-36-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium tert-decanoate |  | 284-429-0 | 84878-37-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis(nonylphenolate) |  | 284-441-6 | 84878-48-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis(octylphenolate) |  | 284-444-2 | 84878-51-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Flue dust, lead-manufg., cadmium-rich | Residue obtained in the metallurgical treatment of lead concentrate in a lead blast furnace. The substance is composed of cadmium oxides, lead oxides, and impurities containing compounds of arsenic, chlorine, indium and tellurium. | 285-554-3 | 85117-02-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Waste solids, cadmium-electrolysis, thallium-rich | Residue obtained in the electrolysis of cadmium, composed primarily of thallium chromate. Other non-ferrous metals or metal compounds may also be present. | 285-572-1 | 85117-20-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Fatty acids, C9-11-branched, cadmium salts |  | 287-817-8 | 85586-15-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium propionate |  | 241-066-2 | 16986-83-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium hexafluorosilicate(2-) |  | 241-084-0 | 17010-21-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Tricadmium bis(phosphate) |  | 236-764-9 | 13477-17-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium silicate |  | 236-765-4 | 13477-19-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulphite |  | 236-767-5 | 13477-23-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Diboron tricadmium hexaoxide |  | 237-225-0 | 13701-66-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Dicadmium hexakis(cyano-C)ferrate(4-) |  | 237-341-1 | 13755-33-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenite |  | 237-480-8 | 13814-59-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenate |  | 237-481-3 | 13814-62-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium diricinoleate |  | 237-544-5 | 13832-25-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium orthophosphate |  | 237-581-7 | 13847-17-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium molybdenum tetroxide |  | 237-752-6 | 13972-68-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium disulphamate |  | 237-832-0 | 14017-36-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium hydrogen phosphate |  | 237-920-9 | 14067-62-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bis(diethyldithiocarbamate) |  | 238-113-4 | 14239-68-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium chromate |  | 238-252-0 | 14312-00-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium dipotassium tetracyanide |  | 238-371-8 | 14402-75-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium tetrafluoroborate |  | 238-490-5 | 14486-19-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Bis(dibutyldithiocarbamato-S,S')cadmium |  | 238-609-0 | 14566-86-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Bis(pentane-2,4-dionato-O,O')cadmium |  | 238-730-9 | 14689-45-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium epoxyoctadecanoate |  | 247-560-4 | 26264-48-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium toluate |  | 248-480-2 | 27476-27-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
[[N,N'-ethylenebis[glycinato]](2-)-N,N',O,O']cadmium |  | 249-987-1 | 29977-13-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium isooctanoate |  | 250-118-3 | 30304-32-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium dodecylbenzenesulphonate |  | 250-433-6 | 31017-44-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium (1,1-dimethylethyl)benzoate |  | 250-515-1 | 31215-94-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium peroxide |  | - | 12139-22-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium mercury telluride |  | 249-914-3 | 29870-72-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Selenic acid, cadmium salt (1:1), dihydrate |  | - | 10060-09-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Tetraiodide diammonium cadmium |  | - | 105034-60-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium trichloride monopotassium |  | - | 14429-85-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium potassium hexachloride |  | - | 15276-40-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium bromide rubidium |  | - | 16593-57-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium chloride rubidium |  | - | 18532-58-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sodium tetrachloride |  | - | 21360-94-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cesium sulfate cadmium |  | - | 28041-77-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium chloride potassium hydrate |  | - | 28302-54-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Potassium sulfate cadmium hexahydrate |  | - | 30623-04-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium trichloride 1 sodium |  | - | 55425-74-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfate sodium |  | - | 28038-18-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Potassium sulfate cadmium |  | - | 28038-25-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Monoammonium cadmium triiodide |  | - | 32593-99-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfate caesium hexahydrate |  | - | 34345-39-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Potassium sulfate cadmium dihydrate |  | - | 38386-25-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium chloride magnesium dodecahydrate |  | - | 77289-75-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Benzenediazonium, 4-(phenylamino)-, sulfate (1:1), polymer with formaldehyde, cadmium chloride complexes |  | - | 68441-39-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium borate oxide (Cd3(BO2)4O), manganese-doped |  | - | 68784-58-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Barium cadmium zinc sulfide (Ba2(Cd,Zn)S3), manganese-doped |  | - | 68876-90-4 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium laurate, palmitate, stearate |  | - | 68954-18-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sponge |  | - | 69011-70-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Slimes and Sludges, cadmium electrolytic |  | - | 69029-89-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium, dichlorotetrakis(1H-imidazole-kappaN3)- |  | - | 72275-93-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium zinc sulfide ((Cd,Zn)S), cobalt and copper-doped |  | - | 72869-26-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Barium tetrachlorocadmate(2-) |  | 308-988-8 | 99587-10-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Zircon, cadmium red |  | 691-407-8 | 72828-62-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium, diethyl- |  | 804-213-7 | 592-02-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium acetate hydrate |  | 628-590-0 | 89759-80-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfate, tetrahydrate |  | - | 13477-21-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Phosphoric acid, ammonium cadmium salt (1:1:1) |  | - | 14520-70-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Octadecanoic acid, barium cadmium salt |  | - | 17033-07-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Ammonium cadmium chloride (Ammonium cadmium trichloride) |  | - | 18532-52-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium pyrophosphate |  | - | 19262-93-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfate octahydrate |  | - | 22465-18-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium, chloro(1,10-phenanthroline-2-carboxylato)-, polymers |  | - | 25685-75-0 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Disodium ethylenediaminetetraacetate cadmium |  | - | 30363-28-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium zinc sulfide |  | 234-372-2 | 11129-14-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium mercury sulfide |  | 215-717-6 | 1345-09-1 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Dipotassium tetrachlorocadmate(2-) |  | 243-938-8 | 20648-91-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), solid soln. with zinc sulfide, copper and silver-doped |  | 271-511-6 | 68583-43-7 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), solid soln. with zinc sulfide, nickel and silver-doped |  | 271-512-1 | 68583-44-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), solid soln. with zinc sulfide, silver chloride-doped |  | 271-513-7 | 68583-45-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), solid soln. with zinc sulfide, aluminum and silver-doped |  | 271-538-3 | 68584-41-8 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), solid soln. with zinc sulfide, copper and nickel-doped |  | 271-539-9 | 68584-42-9 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Hydrogen [N,N-bis(carboxymethyl)glycinato(3-)-N,O,O',O'']cadmate |  | 256-489-8 | 49784-44-3 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), solid soln. with zinc sulfide, copper chloride-doped |  | 270-977-8 | 68512-49-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), solid soln. with zinc sulfide, copper and manganese-doped |  | 270-978-3 | 68512-50-5 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium sulfide (CdS), solid soln. with zinc sulfide, aluminum and copper-doped |  | 270-979-9 | 68512-51-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Barium di-μ-chlorotetrachlorodicadmate(2-) |  | 262-618-9 | 61129-40-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Dihydrogen bis[hydroxysuccinato(2-)-O1,O2,O4]cadmate(2-) |  | 275-459-5 | 71436-99-2 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Cadmium selenide sulfide, (Cd2SeS) |  | - | 12213-70-6 | 23 | https://echa.europa.eu/documents/10162/3bfef8a3-8c97-4d85-ae0b-ac6827de49a9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0059.01.ENG&toc=OJ:L:1991:186:TOC">D 91/338/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.142.01.0022.01.ENG&toc=OJ:L:1999:142:TOC">D 1999/51/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.134.01.0002.01.ENG&toc=OJ:L:2011:134:TOC">R 494/2011</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0001.01.ENG&toc=OJ:L:2012:252:TOC">R 835/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.040.01.0005.01.ENG&toc=OJ:L:2016:040:TOC">R 217/2016</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/158">entry 158</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/662">entry 662</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/824">entry 824</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/825">entry 825</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1180">entry 1180</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1310">entry 1310</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1311">entry 1311</a> |  | 13-Jan-2016
Monomethyl — tetrachlorodiphenyl methane Trade name: Ugilec 141 |  | - | 76253-60-6 | 24 | https://echa.europa.eu/documents/10162/1d81010a-17e1-4026-b01c-70f647c7356b |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0064.01.ENG&toc=OJ:L:1991:186:TOC">D 91/339/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Monomethyl-dichloro-diphenyl methane Trade name: Ugilec 121, Ugilec 21 |  | - | - | 25 | https://echa.europa.eu/documents/10162/08960286-ad89-47b1-a538-7f4c8d7d9074 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0064.01.ENG&toc=OJ:L:1991:186:TOC">D 91/339/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Monomethyl-dibromo-diphenyl methane bromobenzylbromotoluene, mixture of isomers | DBBT | 402-210-1 | 99688-47-8 | 26 | https://echa.europa.eu/documents/10162/c2ca4109-be74-0301-9131-d01105d2c2cb |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1991.186.01.0064.01.ENG&toc=OJ:L:1991:186:TOC">D 91/339/EEC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Nickel and its compounds |  | - | - | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel |  | 231-111-4 | 7440-02-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel compounds |  | - | - | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, acetate carbonate C8-10-branched fatty acids C9-11-neofatty acids complexes |  | 291-673-1 | 90459-30-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, borate C8-10-branched carboxylate complexes |  | 291-674-7 | 90459-31-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, C5-C23-branched carboxylate octanoate complexes |  | 291-675-2 | 90459-32-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, isooctanoate naphthenate complexes |  | 291-676-8 | 90459-33-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, acetylacetone 6-methyl-2,4-heptanedione complexes |  | 291-677-3 | 90459-34-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [29H,31H-phthalocyaninato(2-)-N29,N30,N31,N32]-, [[3-[(5-chloro-2,6-difluoro-4-pyrimidinyl)amino]phenyl]amino]sulfonyl sulfo derivs., sodium salts |  | 291-678-9 | 90459-35-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickelate(4-), [bis[[[3-[[4,5-dihydro-3-methyl-5-oxo-1-[4-[[2-(sulfooxy)ethyl]sulfonyl]phenyl]-1H-pyrazol-4-yl]azo]phenyl]amino]sulfonyl]-29H,31H-phthalocyaninedisulfonato(6-)-N29,N30,N31,N32]-, sodium |  | 291-679-4 | 90459-36-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Waste solids, chromium-nickel steel manuf. | By-product from refined steel manufacture containing C, Cr, Fe and Ni (as Cr-Ni-grindings) and Cr2O3, FeO, MnO, NiO and SiO2 (as burned slag). | 293-796-6 | 91082-81-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Waste solids, nickel-manuf. | By-product from nickel production containing NiO, SiO2, CuO, and MnO and sulfur. | 293-799-2 | 91082-84-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Fatty acids, C6-19-branched, nickel salts |  | 294-302-1 | 91697-41-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, C5-23-branched carboxylate naphthenate complexes |  | 295-925-1 | 92200-98-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, C5-25-branched carboxylate naphthenate octanoate complexes |  | 295-926-7 | 92200-99-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, borate neodecanoate complexes |  | 296-343-0 | 92502-55-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, C5-23-branched carboxylate C4-10-fatty acids naphthenate complexes |  | 297-548-8 | 93573-14-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, C4-10 fatty acids naphthenate complexes |  | 297-549-3 | 93573-15-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, C4-10 fatty acids octanoate complexes |  | 297-550-9 | 93573-16-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [29H,31H-phthalocyaninato(2-)-N29,N30,N31,N32]-, chlorosulfonyl derivs., reaction products with 2-[(4-aminophenyl)sulfonyl]ethyl hydrogen sulfate monosodium salt, potassium sodium salts, compds. with pyridine |  | 297-551-4 | 93573-17-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, C5-23-branched carboxylate C4-10 fatty acids complexes |  | 297-774-7 | 93762-59-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, aqua[2-[2-[4,5-dihydro-3-methyl-5-(oxo-κO)-1H-pyrazol-4-yl]diazenyl-κN1]benzoato(2-)-κO]- |  | - | 106316-55-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [29H,31H-phthalocyaninato(2-)-κN29,κN30,κN31,κN32]-, sulfo [[4-[[2-(sulfooxy)ethyl]sulfonyl]phenyl]amino]sulfonyl derivs., potassium sodium salts |  | - | 113894-88-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickelate(6-), [4-[2-[5-[[(3,6-dichloro-4-pyridazinyl)carbonyl]amino]-2-sulfophenyl]diazenyl]-4,5-dihydro-5-oxo-1-[5-[[(C,C,C-trisulfo-29H,31H-phthalocyanin-C-yl)sulfonyl]amino]-2-sulfophenyl]-1H-pyrazole-3-carboxylato(8-)-κN29,κN30,κN31,κN32]-, sodium |  | - | 113894-89-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel disulfide |  | - | 12035-50-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel silicide (NiSi) |  | - | 12035-57-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Polydymite (Ni3S4) |  | - | 12059-21-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel Black |  | - | 12125-56-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel oxide (Ni3O4) |  | - | 12137-09-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [carbonato(2-)]hexahydroxytetra- |  | - | 12334-31-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
phosphoric acid, calcium nickel salt |  | - | 17169-61-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Diphosphoric acid, nickel(2+) salt |  | - | 19372-20-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
cobalt nickel oxide |  | 620-395-9 | 12737-30-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel boride phosphide |  | - | 65229-23-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
lithium nickel dioxide |  | 620-400-4 | 12031-65-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
molybdenum nickel oxide |  | - | 12673-58-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel oxide (Ni2O) |  | - | 12359-17-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel arsenide |  | - | 12795-30-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Rammelsbergite (NiAs2) |  | - | 1303-22-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel chloride(NiCl) |  | - | 13931-83-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Sulfuric acid, nickel(2+) salt (1:1), monohydrate |  | - | 14168-76-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Phosphoric acid, nickel salt |  | - | 14396-43-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, bis(hexahydro-1H-azepine-1-carbodithioato-κS1,κS'1)-, (SP-4-1)- |  | - | 14434-67-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Octadecanoic acid, nickel salt |  | - | 14448-69-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Carbonic acid, nickel(2+) salt (1:1), hexahydrate |  | - | 17030-82-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, tetrakis(triphenyl phosphite-.kappa.p)-, (t-4)- |  | 622-848-6 | 14221-00-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dichlorobis(tributylphosphine)nickel(II) |  | 624-942-2 | 15274-43-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tetrakis(triphenylphosphine)nickel(0) |  | 626-363-0 | 15133-82-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
1000 NN Ferrite |  | 629-315-7 | 12645-50-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(II) sulfate hexa-/ heptahydrate |  | 630-456-1 | 15244-37-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel carbide |  | 640-090-4 | 12710-36-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel carbonate hydroxide tetrahydrate |  | 641-002-7 | 39430-27-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
NICKEL(II) TRIFLUOROMETHANESULFONATE |  | 663-776-5 | 60871-84-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Ethylenediaminetetraacetic Acid Disodium Nickel(II) Salt Hydrate |  | 679-904-8 | 11079-07-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel ammonium sulfate hexahydrate |  | 680-452-9 | 7785-20-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel sulfide (NiS2) |  | 683-050-1 | 12035-51-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, bis[(2-phenyldiazenecarbothioic acid-κS) 2-phenylhydrazidato-κN2]- |  | - | 12213-15-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dihydroxide hydrate |  | - | 36897-37-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel hydride |  | - | 37187-84-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Phosphoric acid,compounds,nickel(2+) zinc salt (2:1:2) tetrahydrate |  | - | 501953-51-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, bis[1-[4-(diethylamino)phenyl]-2-phenyl-1,2-ethenedithiolato(2-)-κS1,κS2]- |  | - | 51449-18-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Antimony nickel titanium oxide |  | - | 54576-53-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [29H,31H-phthalocyaninato(2-)-N29,N30,N31,N32]-, chlorosulfonyl derivs., reaction products with 2-[(4-aminophenyl)sulfonyl]ethyl hydrogen sulfate monosodium salt, potassium salts |  | 305-643-3 | 94891-42-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [29H,31H-phthalocyaninato(2-)-N29,N30,N31,N32]-, chlorosulfonyl derivs., reaction products with 2-[(4-aminophenyl)sulfonyl]ethyl hydrogen sulfate monosodium salt, sodium salts |  | 305-644-9 | 94891-43-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Spinels, cobalt nickel zinc grey |  | 305-835-7 | 95046-47-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tetrasodium [bis[[[4-[[2-(sulphooxy)ethyl]sulphonyl]phenyl]amino]sulphonyl]-29H,31H-phthalocyaninedisulphonato(6)-N29,N30,N31,N32]nickelate(4-) |  | 306-462-2 | 97280-68-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[[N,N',N''-[29H,31H-phthalocyaninetriyltris(sulphonylimino-3,1-phenylene)]tris[3-oxobutyramidato]](2-)-N29,N30,N31,N32]nickel |  | 306-784-3 | 97404-21-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[[N,N',N'',N'''-[29H,31H-phthalocyaninetetrayltetrakis(sulphonylimino-3,1-phenylene)]tetrakis[3-oxobutyramidato]](2-)-N29,N30,N31,N32]nickel |  | 306-785-9 | 97404-22-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Diiron nickel zinc tetraoxide |  | 306-902-3 | 97435-21-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Copper, bis(8-quinolinolato-N1,O8)-, reaction products with C8-10-branched fatty acids, tert-decanoic acid, nickel(2+) diacetate, nickel(2+) carbonate (1:1) and nickel hydroxide (Ni(OH)2) |  | 307-554-5 | 97660-42-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Speiss, lead, nickel-contg. | Product obtained and separated during the melting of nickel and other non-ferrous metals containing raw materials. Consists primarily of antimonides and arsenides of copper and nickel. | 308-765-5 | 98246-91-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Diammonium tetrachloronickelate(2-) |  | 308-989-3 | 99587-11-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Cassiterite, cobalt manganese nickel grey |  | 309-018-6 | 99749-23-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Residues, copper-iron-lead-nickel matte, sulfuric acid-insol. |  | 310-050-8 | 102110-49-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[2,2'-[1,2-phenylenebis(nitrilomethylidyne)]-bis(phenolato)]-N,N',O,O'-nickel(II) | BRUN JAUNE ESTOFIL S-3RL; BRUNO GIALLO ESTOFIL S-3RL | 400-870-5 | 14406-71-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tetrasodium (c-(3-(1-(3-(e-6-dichloro-5-cyanopyrimidin-f-yl(methyl)amino)propyl)-1,6-dihydro-2-hydroxy-4-methyl-6-oxo-3-pyridylazo)-4-sulfonatophenylsulfamoyl)phthalocyanine-a,b,d-trisulfonato(6-))nickelato II, where a is 1 or 2 or 3 or 4,b is 8 or 9 or 10 or 11,c is 15 or 16 or 17 or 18, d is 22 or 23 or 24 or 25 and where e and f together are 2 and 4 or 4and 2 respectively | BRILLANTGREEN K-RWA 6083; BRILLIANT GREEN K-RWA 6083 | 410-160-7 | 148732-74-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Hexasodium (di-[N-(3-(4-[5-(5-amino-3-methyl-1-phenylpyrazol-4-yl-azo)-2,4-disulfo-anilino]-6-chloro-1,3,5-triazin-2-ylamino)phenyl)-sulfamoyl](di-sulfo)-phthalocyaninato)nickel | PROCION EMERALD H-EXL COLOUR; SUBSTANCE S156718 | 417-250-5 | 151436-99-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
dinickel(2+) sodium bis(3-carboxy-5-[(E)-2-(7-oxido-2,6-disulfonaphthalen-1-yl)diazen-1-yl]-1H-1,2,4-triazol-1-ide) |  | 443-510-2 | 738587-10-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, compound with Titanium (1:1) |  | 610-765-8 | 52013-44-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel fluoride (NiF2), tetrahydrate |  | 604-130-4 | 13940-83-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [29H,31H-phthalocyaninato(2-)-[?]N29,[?]N30,[?]N31,[?]N32]-, [[3-[(5-chloro-6-fluoro-4-pyrimidinyl)amino]phenyl]amino]sulfonyl sulfo derivs., sodium salts |  | 603-857-4 | 134737-17-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [6,8,16,18-tetrachloro-1,11-bis(2-furanylmethyl)-1,10,11,20-tetrahydrodibenzo[c,j]dipyrazolo[3,4-f:3',4'-m][1,2,5,8,9,12]hexaazacyclotetradecinato(2-)-N5,N10,N15,N20]- |  | 616-726-1 | 79745-01-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel hydride (NiH) |  | 604-345-3 | 14332-32-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Lithium Nickel Manganese Cobalt Oxide |  | 620-032-4 | 346417-97-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [29H,31H-phthalocyaninato(2-)-[?]N29,[?]N30,[?]N31,[?]N32]-, sulfo [[4-[[2-(sulfooxy)ethyl]sulfonyl]phenyl]amino]sulfonyl derivs., sodium salts |  | 603-528-5 | 131866-99-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Phosphinic acid, nickel(2+) salt, hexahydrate |  | 603-867-9 | 13477-97-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, dichloro[1,1'-(1,3-propanediyl)bis[1,1-diphenylphosphine-[?]P]]- |  | 605-052-3 | 15629-92-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
nickel(2+) disodium 2,2',2'',2'''-(ethane-1,2-diyldinitrilo)tetraacetate |  | 692-065-2 | 15708-55-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Lithium Nickel Cobalt Aluminium Oxide |  | 803-110-4 | 193214-24-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trinickel phosphide |  | 822-331-7 | 12059-19-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Cobalt(2+) dinickel(2+) bis[2-hydroxypropane-1,2,3-tricarboxylate] |  | 303-972-7 | 94232-44-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dicobalt(2+) nickel(2+) bis[2-hydroxypropane-1,2,3-tricarboxylate] |  | 304-013-5 | 94232-84-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel methacrylate |  | 304-466-9 | 94275-78-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Aluminiummagnesiumnickelsiliziumoxide |  | - | 198831-12-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Aluminum boron cobalt lithium nickel oxide |  | - | 207803-51-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Phosphoric acid,compounds,nickel(2+) zinc salt (2:1:2) |  | - | 90053-13-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Fatty acids, (C=3-22), nickel salts, basic |  | - | 91845-72-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel Sulfide (Ni5S3) |  | - | 97599-06-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel Sulfide (Ni5S4) |  | - | 97599-07-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel Sulfide (Ni5S6) |  | - | 97599-08-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel Sulfide (Ni9S5) |  | - | 98443-26-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel sulfide (Ni(HS)2) |  | - | 27911-69-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, tris(neodecanoato-.kappa.O)[.mu.3-[orthoborato(3-)-.kappa.O:.kappa.O':.kappa.O'']]tri- |  | - | 936576-64-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel arsenide sulfide |  | - | 12255-10-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Perchloric acid, nickel salt |  | - | 41743-88-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickelate(Ni4O5)(1-), lithium |  | - | 67163-62-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel subarsenide |  | - | 12255-80-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Sulfuric acid, nickel(2+) salt (1:1), tetrahydrate |  | - | 61092-77-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Sulfuric acid, nickel(2+) salt (1:1), pentahydrate |  | - | 66341-42-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Sulfuric acid, nickel(2+) salt (1:1), dihydrate |  | - | 66634-47-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Godlevskite (Ni7S6) |  | - | 66732-80-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel Sulfide (Ni4S) |  | - | 67272-59-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, bis[[2-(hydroxy-κO)-4-octylphenyl]phenylmethanonato-κO]- |  | - | 68189-15-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [29H,31H-phthalocyaninato(2-)-κN29,κN30,κN31,κN32]-, [(3-aminophenyl)amino]sulfonyl sulfo derivs. |  | - | 68412-19-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, dextrin complexes |  | - | 68412-20-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Sulfuric acid, nickel(2+) salt (1:1), reaction products with nickel and nickel oxide (NiO) |  | - | 68585-48-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
2-ethylhexanoic acid;nickel;2,2,2-trifluoroacetic acid |  | - | 70776-98-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Acetic acid, nickel(2+) salt (2:1), polymer with formaldehyde and 4-(1,1,3,3-tetramethylbutyl)phenol |  | - | 71050-57-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [[2,2'-[methylenebis(thio-κS)]bis[acetato-κO]](2-)]- |  | - | 71215-73-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+), tris(1,2-ethanediamine-κN,κN')-, (OC-6-11)-, salt with dimethylbenzenesulfonic acid (1:2) |  | - | 71215-97-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, bis[N-hydroxy-3-(hydroxyimino-κN)-N'-(2-methoxyphenyl)butanimidamidato-κN']- |  | - | 71605-83-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [N-(4-chlorophenyl)-3-[2-[[1-(4-chlorophenyl)-4,5-dihydro-3-methyl-5-(oxo-κO)-1H-pyrazol-4-yl]methylene]hydrazinylidene-κN2]-α-cyano-1H-isoindole-3-acetamidato(2-)-κN2,κO3]- |  | - | 106887-33-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Sulfuric acid, nickel salt (1:?), reaction products with sulfurized phenol calcium salt (2:1) |  | - | 72162-32-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [N,N',N''-tris[4-(4,5-dihydro-3-methyl-5-oxo-1H-pyrazol-1-yl)phenyl]-29H,31H-phthalocyanine-C,C,C-trisulfonamidato(2-)-κN29,κN30,κN31,κN32]- |  | - | 72252-57-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickelate(6-), [C-[[[3-[2-[4,5-dihydro-3-methyl-5-oxo-1-[3-sulfo-4-[2-[2-sulfo-4-[(2,5,6-trichloro-4-pyrimidinyl)amino]phenyl]ethenyl]phenyl]-1H-pyrazol-4-yl]diazenyl]-4-sulfophenyl]amino]sulfonyl]-29H,31H-phthalocyanine-C,C,C-trisulfonato(8-)-κN29,κN30, |  | - | 72453-55-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Aluminum, triethyl-, reaction products with nickel(2+) 2-ethylhexanoate (1:2) |  | - | 79357-65-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, bis[(cyano-κC)triphenylborato(1-)-κN]bis(hexanedinitrile-κN,κN')- |  | - | 83864-02-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel Sulfide (Ni7S3) |  | - | 84013-51-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Sulfuric acid, nickel(2+) salt (1:1), trihydrate |  | - | 85017-75-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, bis(diisononylcarbamodithioato-S,S')- |  | - | 85298-61-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel Sulfide (Ni3S5), trihydrate |  | - | 86498-45-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel di(acetate) |  | 206-761-7 | 373-02-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trinickel bis[tetracosa-μ-oxododecaoxo[μ12-[phosphato(3-)-O:O:O:O':O':O':O'':O'':O'':O''':O''':O''']]dodecamolybdate(3-)] |  | 235-531-9 | 12263-13-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Antimony, compound with nickel (1:3) |  | 235-676-8 | 12503-49-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trihydrogen hydroxybis[orthosilicato(4-)]trinickelate(3-) |  | 235-688-3 | 12519-85-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[carbonato(2-)]tetrahydroxytrinickel |  | 235-715-9 | 12607-70-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel titanium oxide |  | 235-752-0 | 12653-76-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bismuth, compound with nickel (1:1) |  | 235-773-5 | 12688-64-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(1-nitroso-2-naphtholato)nickel |  | 235-829-9 | 12794-26-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) oleate |  | 235-832-5 | 13001-15-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
2-ethylhexanoic acid, nickel salt |  | 231-480-1 | 7580-31-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dichloride |  | 231-743-0 | 7718-54-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) sulphite |  | 231-827-7 | 7757-95-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel sulphate |  | 232-104-9 | 7786-81-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Antimony nickel titanium oxide yellow | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77788. | 232-353-3 | 8007-18-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel difluoride |  | 233-071-3 | 10028-18-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) selenite |  | 233-263-7 | 10101-96-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trinickel bis(orthophosphate) |  | 233-844-5 | 10381-36-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Aluminium, compound with nickel (1:1) |  | 234-439-6 | 12003-78-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dialuminium nickel tetraoxide |  | 234-454-8 | 12004-35-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel boride |  | 234-493-0 | 12007-00-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dinickel boride |  | 234-494-6 | 12007-01-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trinickel boride |  | 234-495-1 | 12007-02-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dichromium nickel tetraoxide |  | 234-636-7 | 12018-18-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel triuranium decaoxide |  | 239-876-6 | 15780-33-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Formic acid, nickel salt |  | 239-946-6 | 15843-02-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[2-hydroxy-4-(octyloxy)benzophenonato]nickel |  | 239-949-2 | 15843-91-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel tellurium trioxide |  | 239-967-0 | 15851-52-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel tellurium tetraoxide |  | 239-974-9 | 15852-21-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) trifluoroacetate |  | 240-235-8 | 16083-14-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[[2,2'-sulphonylbis[4-(1,1,3,3-tetramethylbutyl)phenolato]](2-)-O1,O1',O2]nickel |  | 240-485-8 | 16432-37-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel sulphide |  | 240-841-2 | 16812-54-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Citric acid , ammonium nickel salt |  | 242-161-1 | 18283-82-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dicarbonylbis(triphenylphosphine)nickel |  | 235-851-9 | 13007-90-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dinitrate |  | 236-068-5 | 13138-45-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, compound with niobium (1:1) |  | 234-807-6 | 12034-55-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dioxide |  | 234-823-3 | 12035-36-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel tin trioxide |  | 234-824-9 | 12035-38-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel titanium trioxide |  | 234-825-4 | 12035-39-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Antimony, compound with nickel (1:1) |  | 234-827-5 | 12035-52-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dinickel phosphide |  | 234-828-0 | 12035-64-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trinickel disulphide |  | 234-829-6 | 12035-72-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dinickel hexacyanoferrate |  | 238-946-3 | 14874-78-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(1,1,1,5,5,5-hexafluoropentane-2,4-dionato-O,O')nickel |  | 239-028-5 | 14949-69-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel acetate |  | 239-086-1 | 14998-37-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel selenate |  | 239-125-2 | 15060-62-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(diisobutyldithiocarbamato)nickel |  | 239-354-8 | 15317-78-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bis(dimethyldithiocarbamate) |  | 239-560-8 | 15521-65-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dichromate |  | 239-646-5 | 15586-38-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Diammonium nickel bis(sulphate) |  | 239-793-5 | 15699-18-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Hexakis(1H-imidazole-N3)nickel(2+) dichloride |  | 239-841-5 | 15751-00-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel oxalate |  | 208-933-7 | 547-67-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dibenzoate |  | 209-046-8 | 553-71-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dicyanide |  | 209-160-8 | 557-19-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(η5-2,4-cyclopentadien-1-yl)nickel |  | 215-039-0 | 1271-28-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(1,5-cyclooctadiene)nickel |  | 215-072-0 | 1295-35-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel monoxide |  | 215-215-7 | 1313-99-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel selenide |  | 215-216-2 | 1314-05-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dinickel trioxide |  | 215-217-8 | 1314-06-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bis(4-cyclohexylbutyrate) |  | 223-463-2 | 3906-55-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bis(2-ethylhexanoate) |  | 224-699-9 | 4454-16-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) octanoate |  | 225-656-7 | 4995-91-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trinickel dicitrate |  | 227-873-2 | 6018-92-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(II) fumarate |  | 228-501-1 | 6283-67-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) stearate |  | 218-744-1 | 2223-95-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(pentane-2,4-dionato-O,O')nickel |  | 221-875-7 | 3264-82-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel carbonate |  | 222-068-2 | 3333-67-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel diformate |  | 222-101-0 | 3349-06-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) propionate |  | 222-102-6 | 3349-08-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dihydroxide |  | 235-008-5 | 12054-48-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dinickel silicide |  | 235-033-1 | 12059-14-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, compound with tin (3:1) |  | 235-034-7 | 12059-23-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel diarsenide |  | 235-103-1 | 12068-61-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel telluride |  | 235-260-6 | 12142-88-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, compound with zirconium (1:2) |  | 235-261-1 | 12142-92-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Diiron nickel tetraoxide |  | 235-335-3 | 12168-54-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Di-μ-carbonylbis(η5-2,4-cyclopentadien-1-yl)dinickel |  | 235-339-5 | 12170-92-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dysprosium, compound with nickel (1:2) |  | 235-341-6 | 12175-27-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Lanthanum, compound with nickel (1:5) |  | 235-372-5 | 12196-72-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel disilicide |  | 235-379-3 | 12201-89-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dibromide |  | 236-665-0 | 13462-88-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel diiodide |  | 236-666-6 | 13462-90-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tetracarbonylnickel |  | 236-669-2 | 13463-39-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trinickel bis(arsenate) |  | 236-771-7 | 13477-70-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(butanedione dioximato)nickel |  | 236-782-7 | 13478-93-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel diperchlorate |  | 237-124-1 | 13637-71-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) palmitate |  | 237-138-8 | 13654-40-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dithiocyanate |  | 237-205-1 | 13689-92-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bis(sulphamidate) |  | 237-396-1 | 13770-89-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dinickel orthosilicate |  | 237-411-1 | 13775-54-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dipotassium bis(sulphate) |  | 237-563-9 | 13842-46-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dipotassium tetrafluoronickelate(2-) |  | 237-595-3 | 13859-60-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tetrakis(trifluorophosphine)nickel |  | 237-597-4 | 13859-65-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[N-(carboxymethyl)glycinato(2-)-N,O,ON]nickel |  | 237-618-7 | 13869-33-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Hexaamminenickel(2+) bis[tetrafluoroborate(1-)] |  | 237-638-6 | 13877-20-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bis(dibutyldithiocarbamate) |  | 237-696-2 | 13927-77-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Disodium tetracyanonickelate(2-) |  | 237-877-6 | 14038-85-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[29H,31H-phthalocyaninato(2-)-N29,N30,N31,N32]nickel |  | 237-893-3 | 14055-02-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(quinolin-8-olato-N1,O8)nickel |  | 237-950-2 | 14100-15-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel tungsten tetraoxide |  | 238-032-4 | 14177-51-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Molybdenum nickel tetraoxide |  | 238-034-5 | 14177-55-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nitric acid, nickel salt |  | 238-076-4 | 14216-75-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dipotassium tetracyanoniccolate |  | 238-082-7 | 14220-17-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(triphenylphosphine)nickel(II) chloride |  | 238-154-8 | 14264-16-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(diethyldithiocarbamato-S,S')nickel |  | 238-157-4 | 14267-17-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel hydrogen phosphate |  | 238-278-2 | 14332-34-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(1-nitrosonaphthalen-2-olato-N1,O2)nickel |  | 238-380-7 | 14406-66-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[bis(2-ethylhexyl)dithiocarbamato-S,S']nickel |  | 238-400-4 | 14428-08-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dinickel diphosphate |  | 238-426-6 | 14448-18-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bis(phosphinate) |  | 238-511-8 | 14507-36-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(butylamine)[[2,2'-thiobis[4-(1,1,3,3-tetramethylbutyl)phenolato]](2-)-O,O',S]nickel |  | 238-523-3 | 14516-71-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(6-methylheptane-2,4-dionato-O,O')nickel |  | 238-536-4 | 14522-99-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dibromate |  | 238-596-1 | 14550-87-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bis(tetrafluoroborate) |  | 238-753-4 | 14708-14-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel chromate |  | 238-766-5 | 14721-18-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[[2,2'-thiobis[3-octylphenolato]](2-)-O,O',S]nickel |  | 251-715-1 | 33882-09-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Hydrogen [N,N-bis(carboxymethyl)glycinato(3-)-N,O,O',O'']nickelate(1-) |  | 252-235-5 | 34831-03-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tetrakis(tritolyl phosphite )nickel |  | 252-777-2 | 35884-66-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bisphosphinate |  | 252-840-4 | 36026-88-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(dipentyldithiocarbamato-S,S')nickel |  | 252-937-1 | 36259-37-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel chloride |  | 253-399-0 | 37211-05-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Silicic acid, nickel salt |  | 253-461-7 | 37321-15-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[1-[4-(dimethylamino)phenyl]-2-phenylethylene-1,2-dithiolato(2-)-S,S']nickel |  | 253-958-9 | 38465-55-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tris(4,7-diphenyl-1,10-phenanthroline-N1,N10)nickel(2+) dinitrate |  | 254-127-3 | 38780-90-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[2-butene-2,3-dithiolato(2-)-S,S']nickel |  | 254-210-4 | 38951-94-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[4,4'-dimethoxy-α,α'-stilbenedithiolato(2-)]nickel |  | 254-212-5 | 38951-97-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dipotassium tris(cyano-C)nickelate(2-) |  | 254-261-2 | 39049-81-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bis(benzenesulphonate) |  | 254-642-3 | 39819-65-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bis(piperidine-1-carbodithioate) |  | 255-387-0 | 41476-75-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bis(dihydrogen phosphate) |  | 242-522-3 | 18718-11-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) hydrogen citrate |  | 242-533-3 | 18721-51-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[[1,1'-[1,2-phenylenebis(nitrilomethylidyne)]bis[2-naphtholato]](2-)-N,N',O,O']nickel |  | 243-820-6 | 20437-10-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Oxalic acid, nickel salt |  | 243-867-2 | 20543-06-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(ethane-1,2-diamine)nickel(2+) sulphate |  | 244-300-1 | 21264-77-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) silicate |  | 244-578-4 | 21784-78-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[μ-[[1,1',1'',1'''-[benzene-1,2,4,5-tetrayltetrakis(nitromethylidyne)]naphth-2-olato](4-)]]dinickel |  | 245-028-6 | 22484-07-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Citric acid, nickel salt |  | 245-119-0 | 22605-92-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Ammonium nickel trichloride |  | 246-378-2 | 24640-21-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dihydrogen [[N,N'-ethylenebis[N-(carboxymethyl)glycinato]](4-)-N,N',O,O',ON,ON']nickelate(2-) |  | 247-019-2 | 25481-21-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel hexafluorosilicate |  | 247-430-7 | 26043-11-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel arsenide |  | 248-169-1 | 27016-75-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[[2,2'-thiobis[4-(1,1,3,3-tetramethylbutyl)phenolato]](2-)-O,O',S]nickel |  | 248-536-6 | 27574-34-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel isooctanoate |  | 248-585-3 | 27637-46-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[29H,31H-phthalocyaninetetrasulphonyl tetrachloridato(2)-N29,N30,N31,N32 ]nickel |  | 249-155-8 | 28680-76-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[stilbene-α,β-dithiolato(2-)]nickel |  | 249-353-4 | 28984-20-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[2,3-bis(hydroxyimino)-N-phenylbutyramidato-N2,N3]nickel |  | 249-503-9 | 29204-84-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(II) isooctanoate |  | 249-555-2 | 29317-63-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel hexafluorozirconate(2-) |  | 250-370-4 | 30868-55-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) diethyl bis[[[3,5-bis(1,1-dimethylethyl)-4-hydroxyphenyl]methyl]phosphonate] |  | 250-401-1 | 30947-30-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel silicate(3:4) |  | 250-788-7 | 31748-25-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[μ-[carbonato(2-)-O:O']]dihydroxydinickel |  | 265-748-4 | 65405-96-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(propan-2-ol)[[2,2'-thiobis[4-(1,1,3,3-tetramethylbutyl)phenolato]](2-)-O,O',S]nickel |  | 267-045-8 | 67763-27-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Potassium [[N,N'-ethylenebis[N-(carboxymethyl)glycinato]](4-)-N,N',O,O',ON,ON']nickelate(1-) |  | 267-686-3 | 67906-12-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) dihydrogen bis[R-(R*,R*)]-tartrate |  | 267-894-4 | 67952-41-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel dichlorate |  | 267-897-0 | 67952-43-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) glycerol phosphate |  | 267-923-0 | 67952-69-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Triammonium hydrogen [[[nitrilotris(methylene)]tris[phosphonato]](6-)]nickelate(4-) |  | 267-961-8 | 67968-22-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Cobalt dimolybdenum nickel octaoxide |  | 268-169-5 | 68016-03-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Diammonium nickel bis(hydrogen citrate) |  | 268-176-3 | 68025-13-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Triammonium [N,N-bis(phosphonomethyl)glycinato(5-)]nickelate(3-) |  | 268-195-7 | 68025-40-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trisodium [N,N-bis(phosphonomethyl)glycinato(5-)]nickelate(3-) |  | 268-196-2 | 68025-41-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tetrasodium [[[nitrilotris(methylene)]tris[phosphonato]](6-)-N,O,O'',O'''nickelate(4-)] |  | 268-296-6 | 68052-00-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Molybdenum nickel hydroxide oxide phosphate |  | 268-585-7 | 68130-36-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[(2-amino-2-oxoethoxy)acetato(2-)]nickel |  | 268-711-0 | 68133-84-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Formic acid, copper nickel salt |  | 268-755-0 | 68134-59-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Cobalt titanite green spinel | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77377. | 269-047-4 | 68186-85-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Cobalt nickel gray periclase | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77332. | 269-051-6 | 68186-89-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel ferrite brown spinel | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77497. | 269-071-5 | 68187-10-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tris(4,7-diphenyl-1,10-phenanthroline-N1,N10)nickel(2+) bis[tetrafluoroborate(1-)] |  | 269-684-8 | 68309-97-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Resin acids and Rosin acids, nickel salts |  | 269-826-9 | 68334-36-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
2,3-dihydroxypropyl (dihydrogen phosphate), nickel salt (1:1) |  | 269-946-1 | 68391-37-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [29H,31H-phthalocyaninato(2-)-N29,N30,N31,N32]-, sulfo [[4-[[2-(sulfooxy)ethyl]sulfonyl]phenyl]amino]sulfonyl derivs. |  | 270-174-2 | 68412-18-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, 5,5'-azobis-2,4,6(1H,3H,5H)-pyrimidinetrione complexes |  | 270-944-8 | 68511-62-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Zinc sulfide (ZnS), nickel and silver-doped |  | 270-961-0 | 68512-22-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Olivine, nickel green | An inorganic pigment that is the reaction product of high temperature calcination in which nickel (II) oxide and silicon oxide in varying amounts are homogeneously and ionically interdiffused to form a crystalline matrix of olivine. Its composition may include any one or a combination of the modifiers alkali or alkaline earth halides. | 271-112-7 | 68515-84-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[2,3-bis(hydroxyimino)-N-(2-methoxyphenyl)butyramidato]nickel |  | 255-924-9 | 42739-61-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[1,3-dihydro-5,6-bis[[(2-hydroxy-1-naphthyl)methylene]amino]-2H-benzimidazol-2-onato(2-)-N5,N6,O5,O6]nickel |  | 255-965-2 | 42844-93-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[[2,2'-(4,8-dichlorobenzo[1,2-d:4,5-d']bisoxazole-2,6-diyl)bis[4,6-dichlorophenolato]](2-)]nickel |  | 256-331-8 | 47726-62-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel acrylate |  | 257-066-0 | 51222-18-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Neodecanoic acid, nickel salt |  | 257-447-1 | 51818-56-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Copper, compound with lanthanum and nickel (4:1:1) |  | 257-510-3 | 51912-52-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[3-[(4-chlorophenyl)azo]quinoline-2,4(1H,3H)-dionato]nickel |  | 257-521-3 | 51931-46-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel [R(R*,R*)]-tartrate |  | 257-610-7 | 52022-10-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[(2-hydroxyethyl)dithiocarbamato-S,S']nickel |  | 257-953-2 | 52486-98-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[bis(2-hydroxyethyl)dithiocarbamato-S,S']nickel |  | 257-954-8 | 52486-99-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) methacrylate |  | 257-963-7 | 52496-91-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel divanadium hexaoxide |  | 257-970-5 | 52502-12-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(diethyldithiocarbamato-S,S')nickel |  | 258-044-3 | 52610-81-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel 3,5-bis(tert-butyl-4-hydroxybenzoate (1:2) |  | 258-051-1 | 52625-25-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[2,4-dihydro-5-methyl-4-(1-oxodecyl)-2-phenyl-3H-pyrazol-3-onato-O,O']nickel |  | 260-258-7 | 56557-00-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Cobalt nickel dioxide |  | 261-346-8 | 58591-45-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) acrylate |  | 262-383-2 | 60700-37-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Hydrogen [3,4-bis[[(2-hydroxy-1-naphthyl)methylene]amino]benzoato(3-)-N3,N4,O3,O4]nickelate(1-) |  | 262-703-0 | 61300-98-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, 3-[(4-chlorophenyl)azo]-4-hydroxy-2(1H)-quinolinone complex | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 12775. | 262-934-7 | 61725-51-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Naphthenic acids, nickel salts |  | 263-000-1 | 61788-71-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(ethylenediamine-N,N')copper tetrakis(cyano-C)nickelate |  | 264-136-4 | 63427-32-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tetrapotassium [[[nitrilotris(methylene)]tris[phosphonato]](6-)]nickelate(4-) |  | 264-338-2 | 63588-33-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tripotassium [N,N-bis(phosphonomethyl)glycinato(5-)]nickelate(3-) |  | 264-360-2 | 63597-34-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Potassium [N,N-bis(carboxymethyl)glycinato(3-)-N,O,O',O'']nickelate(1-) |  | 264-377-5 | 63640-18-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[2,3'-bis[[(2-hydroxyphenyl)methylene]amino]but-2-enedinitrilato(2-)-N2,N3,O2,O3]nickel |  | 265-022-7 | 64696-98-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(5-oxo-DL-prolinato-N1,O2)nickel |  | 285-069-7 | 85026-81-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(2-ethylhexanoato-O)(neodecanoato-O)nickel |  | 285-698-7 | 85135-77-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(isodecanoato-O)(isooctanoato-O)nickel |  | 285-909-2 | 85166-19-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[N-(2,4-dimethoxyphenyl)-2,3-bis(hydroxyimino)butyramidato-N2,N3]nickel |  | 286-563-5 | 85269-39-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, 2,2'-thiobis[4-nonylphenol] complexes |  | 287-356-2 | 85480-75-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(isodecanoato-O)(neodecanoato-O)nickel |  | 287-467-6 | 85508-42-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(II) isodecanoate |  | 287-468-1 | 85508-43-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) neodecanoate |  | 287-469-7 | 85508-44-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(2-ethylhexanoato-O)(isononanoato-O)nickel |  | 287-470-2 | 85508-45-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(isononanoato-O)(isooctanoato-O)nickel |  | 287-471-8 | 85508-46-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(isononanoato-O)(neodecanoato-O)nickel |  | 287-592-6 | 85551-28-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, isodecanoate naphthenate complexes |  | 287-800-5 | 85585-97-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, isononanoate naphthenate complexes |  | 287-801-0 | 85585-98-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, naphthenate neodecanoate complexes |  | 287-802-6 | 85585-99-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(1H-1,2,4-triazole-3-sulphonato-N2,O3)nickel |  | 287-849-2 | 85586-46-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[[3-[1-cyano-2-(methylamino)-2-oxoethylidene]-2,3-dihydro-1H-isoindol-1-ylidene](salicylic)hydrazidato(2-)]nickel |  | 288-967-7 | 85958-80-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Zinc sulfide (ZnS), copper and nickel-doped |  | 271-601-5 | 68585-93-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Resin acids and Rosin acids, calcium nickel salts |  | 271-764-2 | 68607-31-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel barium titanium primrose priderite | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77900. | 271-853-6 | 68610-24-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel niobium titanium yellow rutile | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77895. | 271-892-9 | 68611-43-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickelate(6-), [4-[[5-[[(3,6-dichloro-4-pyridazinyl)carbonyl]amino]-2-sulfophenyl]azo]-4,5-dihydro-5-oxo-1-[5-[[(trisulfo-29H,31H-phthalocyaninyl)sulfonyl]amino]-2-sulfophenyl]-1H-pyrazole-3-carboxylato(8-)-N29,N30,N31,N32]-, hexahydrogen |  | 272-095-9 | 68698-80-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Hexaamminenickel(2+) diformate |  | 272-149-1 | 68758-60-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Zinc sulfide (ZnS), copper and nickel and silver-doped |  | 272-277-8 | 68784-84-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(2-heptadecyl-1H-imidazole-N3)bis(octanoato-O)nickel |  | 272-799-6 | 68912-08-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Pentaammonium hydrogen [[[ethylenebis[nitrilobis(methylene)]]tetrakis[phosphonato]](8-)]nickelate(6-) |  | 273-375-3 | 68958-86-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Hydrogen pentapotassium [[[ethylenebis[nitrilobis(methylene)]]tetrakis[phosphonato]](8-)]nickelate(6-) |  | 273-376-9 | 68958-87-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Hydrogen pentasodium [[[ethylenebis[nitrilobis(methylene)]]tetrakis[phosphonato]](8-)]nickelate(6-) |  | 273-377-4 | 68958-88-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(ethane-1,2-diamine-N,N')nickel(2+) bis[bis(cyano-C)aurate(1-)] |  | 273-379-5 | 68958-89-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel icosatitanium pentatriacontaoxide diwolframate |  | 273-686-4 | 69011-05-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Slags, ferronickel-manufg. | By-product from the production of ferronickel from a complex ore. Consists primarily of oxides of aluminum, iron, magnesium and silicon. | 273-729-7 | 69012-29-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Matte, nickel | Product of blowing smelted nickel ore in a converter to lower the iron content. | 273-749-6 | 69012-50-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(4-benzoyl-2,4-dihydro-5-methyl-2-phenyl-3H-pyrazol-3-onato-O,O')nickel |  | 274-027-3 | 69524-96-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel zirconium trioxide |  | 274-755-1 | 70692-93-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(5-oxo-L-prolinato-N1,O2)nickel |  | 274-912-4 | 70824-02-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(3-amino-4,5,6,7-tetrachloro-1H-isoindol-1-one oximato-N2,O1)nickel |  | 274-916-6 | 70833-37-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trisodium [22-[[[3-[(5-chloro-2,6-difluoro-4-pyrimidinyl)amino]phenyl]amino]sulphonyl]-29H,31H-phthalocyanine-1,8,15-trisulphonato(5-)-N29,N30,N31,N32]nickelate(3-) |  | 275-295-4 | 71243-96-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel iron chromite black spinel | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77504. | 275-738-1 | 71631-15-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Ethyl hydrogen sulphate, nickel(2+) salt |  | 275-897-7 | 71720-48-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel uranyl tetraacetate, of uranium depleted in uranium-235 |  | 275-994-4 | 71767-12-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[μ-(piperazine-N1:N4)]bis[3-[1-[(4,5,6,7-tetrachloro-1-oxo-1H-isoindol-3-yl)hydrazono]ethyl]quinoline-2,4(1H,3H)-dionato(2-)]dinickel |  | 276-168-6 | 71889-22-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(D-gluconato-O1,O2)nickel |  | 276-205-6 | 71957-07-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dihydrogen hexasodium bis[3-[(2-amino-8-hydroxy-6-sulpho-1-naphthyl)azo]-2-hydroxy-5-sulphobenzoato(5-)]nickelate(8-) |  | 276-364-1 | 72139-08-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickelate(6-), [22-[[[3-[[4,5-dihydro-3-methyl-5-oxo-1-[3-sulfo-4-[2-[2-sulfo-4-[(2,5,6-trichloro-4-pyrimidinyl)amino]phenyl]ethenyl]phenyl]-1H-pyrazol-4-yl]azo]-4-sulfophenyl]amino]sulfonyl]-29H,31H-phthalocyanine-1,8,15-trisulfonato(8-)-N29,N30,N31,N32]-, hexasodium, (SP-4-2)- |  | 276-399-2 | 72152-45-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trisodium [[[[3-[(4-amino-6-chloro-1,3,5-triazin-2-yl)amino]phenyl]amino]sulphonyl]tris(aminosulphonyl)-29H,31H-phthalocyaninetrisulphonato(5-)-N29,N30,N31,N32]nickelate(3-) |  | 276-491-2 | 72229-81-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel, [29H,31H-phthalocyaninato(2-)-N29,N30,N31,N32]-, [[3-[(1,3-dioxobutyl)amino]phenyl]amino]sulfonyl derivs. |  | 276-877-0 | 72828-53-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
[N,N',N'',N'''-tetrakis[4-(4,5-dihydro-3-methyl-5-oxo-1H-pyrazol-1-yl)phenyl]-29H,31H-phthalocyaninetetrasulphonamidato(2-)-N29,N30,N31,N32]nickel |  | 277-174-1 | 72986-45-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Antimony oxide (Sb2O3), solid soln. with nickel oxide (NiO) and titanium oxide (TiO2) |  | 277-627-3 | 73892-02-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[N-(2-hydroxyethyl)-N-methylglycinato-N,O,ON]nickel |  | 278-504-7 | 76625-10-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tetrahydrogen [[[(3-amino-4-sulphophenyl)amino]sulphonyl]-29H,31H-phthalocyaninetrisulphonato(6-)-N29,N30,N31,N32]nickelate(4-) |  | 279-060-7 | 79102-62-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis(4-benzoyl-2,4-dihydro-5-methyl-2-phenyl-3H-pyrazol-3-onato-O,O')(2,2,4,4-tetramethyl-7-oxa-3,20-diazadispiro[5.1.11.2]henicosan-21-one-O21)nickel |  | 279-067-5 | 79121-51-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Trisodium [5-[(4,5-dihydro-3-methyl-5-oxo-1-phenyl-1H-pyrazol-4-yl)azo]-4-hydroxy-3-[(2-hydroxy-3-nitro-5-sulphophenyl)azo]naphthalene-2,7-disulphonato(5-)]nickelate(3-) |  | 279-314-7 | 79817-91-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dimethoxy[29H,31H-phthalocyaninato(2-)-N29,N30,N31,N32]nickel |  | 281-282-4 | 83898-70-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Leach residues, nickel-vanadium ore | Residues from basic leaching of nickel-bearing vanadium ores. Composed primarily of silica and insoluble compounds of nickel and vanadium with minor quantities of other metals, such as arsenic, lead, tin and zinc. | 282-214-6 | 84144-92-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Bis[di(3,5,5-trimethylhexyl)dithiocarbamato-S,S']nickel |  | 283-380-2 | 84604-95-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Fatty acids, C8-18 and C18-unsatd., nickel salts |  | 283-972-0 | 84776-45-4 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(isooctanoato-O)(neodecanoato-O)nickel |  | 284-347-5 | 84852-35-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(isodecanoato-O)(isononanoato-O)nickel |  | 284-348-0 | 84852-36-8 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel bis(isononanoate) |  | 284-349-6 | 84852-37-9 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(2-ethylhexanoato-O)(isooctanoato-O)nickel |  | 284-350-1 | 84852-38-0 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(2-ethylhexanoato-O)(isodecanoato-O)nickel |  | 284-351-7 | 84852-39-1 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickelate(6-), [4-[[5-[[(3,6-dichloro-4-pyridazinyl)carbonyl]amino]-2-sulfophenyl]azo]-4,5-dihydro-5-oxo-1-[2-sulfo-5-[[(trisulfo-29H,31H-phthalocyaninyl)sulfonyl]amino]phenyl]-1H-pyrazole-3-carboxylato(8-)-N29,N30,N31,N32]-, hexasodium |  | 299-467-3 | 93891-86-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
(neononanoato-O)(neoundecanoato-O)nickel |  | 300-092-5 | 93920-08-2 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) neoundecanoate |  | 300-093-0 | 93920-09-3 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Nickel(2+) neononanoate |  | 300-094-6 | 93920-10-6 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Tetrasodium [[[(3-amino-4-sulphophenyl)amino]sulphonyl]-29H,31H-phthalocyaninetrisulphonato(6-)-N29,N30,N31,N32]nickelate(4-) |  | 300-365-9 | 93939-76-5 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Dimethylhexanoic acid, nickel salt |  | 301-323-2 | 93983-68-7 | 27 | https://echa.europa.eu/documents/10162/3bbe9024-52a6-8e63-5581-e686331eb459 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:65837,413439&cs=11CE3B335EC0E7D1F395E7774F8D5B9AD">EN 1811:2023</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:65833,413439&cs=1AC0C9913499251FF8308468E90169473">EN 12472:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:35087,6151&cs=190FF52D3B7C2630A24DBFA86AD81474D">EN 16128:2015</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.188.01.0001.01.ENG&toc=OJ:L:1994:188:TOC">D 94/27/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.301.01.0051.01.ENG&toc=OJ:L:2004:301:TOC">D 2004/96/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/663">entry 663</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/935">entry 935</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1911">entry 1911</a> | Titles and references of standards to be used: https://ec.europa.eu/docsroom/documents/45591 | 13-Jan-2016
Substances which are classified as carcinogen category 1A or 1B in Part 3 of Annex VI to Regulation (EC) No 1272/2008 and are listed in Appendix 1 or Appendix 2, respectively. |  | - | - | 28 | https://echa.europa.eu/documents/10162/0645e093-576f-c279-ceb9-4f2d1ec3e3bd | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=307">Foreword</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=310">Appendix 1</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=350">Appendix 2</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=581">Appendix 11</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.068.01.0024.01.ENG&toc=OJ:L:1997:068:TOC">D 97/10/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.333.01.0001.01.ENG&toc=OJ:L:1997:333:TOC">D 97/56/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.166.01.0087.01.ENG&toc=OJ:L:1999:166:TOC">D 1999/43/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0014.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/34/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0026.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/36/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.033.01.0028.01.ENG&toc=OJ:L:2006:033:TOC">D 2005/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.037.01.0001.01.ENG&toc=OJ:L:2012:037:TOC">R 109/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2017.224.01.0110.01.ENG&toc=OJ:L:2017:224:TOC">R 1510/2017</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.114.01.0004.01.ENG&toc=OJ:L:2018:114:TOC">R 2018/675</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.446.01.0034.01.ENG&toc=OJ%3AL%3A2021%3A446%3ATOC">R 2021/2204</a>#<a href="https://eur-lex.europa.eu/eli/reg/2023/1132/oj/eng">R 2023/1132</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX:32025R1731">R 2025/1731</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/156">entry 156</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1523">entry 1523</a> |  | 15-Jan-2016
Carcinogen category 1B (Table 3.1)/ carcinogen category 2 (Table 3.2) listed in Appendix 2 |  | - | - | 28 | https://echa.europa.eu/documents/10162/0645e093-576f-c279-ceb9-4f2d1ec3e3bd | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=307">Foreword</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=310">Appendix 1</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=350">Appendix 2</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=581">Appendix 11</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.068.01.0024.01.ENG&toc=OJ:L:1997:068:TOC">D 97/10/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.333.01.0001.01.ENG&toc=OJ:L:1997:333:TOC">D 97/56/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.166.01.0087.01.ENG&toc=OJ:L:1999:166:TOC">D 1999/43/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0014.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/34/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0026.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/36/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.033.01.0028.01.ENG&toc=OJ:L:2006:033:TOC">D 2005/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.037.01.0001.01.ENG&toc=OJ:L:2012:037:TOC">R 109/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2017.224.01.0110.01.ENG&toc=OJ:L:2017:224:TOC">R 1510/2017</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.114.01.0004.01.ENG&toc=OJ:L:2018:114:TOC">R 2018/675</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.446.01.0034.01.ENG&toc=OJ%3AL%3A2021%3A446%3ATOC">R 2021/2204</a>#<a href="https://eur-lex.europa.eu/eli/reg/2023/1132/oj/eng">R 2023/1132</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX:32025R1731">R 2025/1731</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/156">entry 156</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1523">entry 1523</a> |  | 15-Jan-2016
Carcinogen category 1A (Table 3.1)/ carcinogen category 1 (Table 3.2) listed in Appendix 1 |  | - | - | 28 | https://echa.europa.eu/documents/10162/0645e093-576f-c279-ceb9-4f2d1ec3e3bd | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=307">Foreword</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=310">Appendix 1</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=350">Appendix 2</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=581">Appendix 11</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.068.01.0024.01.ENG&toc=OJ:L:1997:068:TOC">D 97/10/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.333.01.0001.01.ENG&toc=OJ:L:1997:333:TOC">D 97/56/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.166.01.0087.01.ENG&toc=OJ:L:1999:166:TOC">D 1999/43/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0014.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/34/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0026.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/36/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.033.01.0028.01.ENG&toc=OJ:L:2006:033:TOC">D 2005/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.037.01.0001.01.ENG&toc=OJ:L:2012:037:TOC">R 109/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2017.224.01.0110.01.ENG&toc=OJ:L:2017:224:TOC">R 1510/2017</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.114.01.0004.01.ENG&toc=OJ:L:2018:114:TOC">R 2018/675</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.446.01.0034.01.ENG&toc=OJ%3AL%3A2021%3A446%3ATOC">R 2021/2204</a>#<a href="https://eur-lex.europa.eu/eli/reg/2023/1132/oj/eng">R 2023/1132</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX:32025R1731">R 2025/1731</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/156">entry 156</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1523">entry 1523</a> |  | 15-Jan-2016
Substances which are classified as germ cell mutagen category 1A or 1B in Part 3 of Annex VI to Regulation (EC) No 1272/2008 and are listed in Appendix 3 or Appendix 4, respectively. |  | - | - | 29 | https://echa.europa.eu/documents/10162/ec709879-c587-5b0e-6384-89c34b408717 | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=307">Foreword</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=468">Appendix 3</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=469">Appendix 4</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=581">Appendix 11</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.068.01.0024.01.ENG&toc=OJ:L:1997:068:TOC">D 97/10/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.333.01.0001.01.ENG&toc=OJ:L:1997:333:TOC">D 97/56/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.166.01.0087.01.ENG&toc=OJ:L:1999:166:TOC">D 1999/43/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0014.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/34/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0026.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/36/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.033.01.0028.01.ENG&toc=OJ:L:2006:033:TOC">D 2005/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.037.01.0001.01.ENG&toc=OJ:L:2012:037:TOC">R 109/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2017.224.01.0110.01.ENG&toc=OJ:L:2017:224:TOC">R 1510/2017</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.114.01.0004.01.ENG&toc=OJ:L:2018:114:TOC">R 2018/675</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.446.01.0034.01.ENG&toc=OJ%3AL%3A2021%3A446%3ATOC">R 2021/2204</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX:32025R1731">R 2025/1731</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/156">entry 156</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1523">entry 1523</a> |  | 15-Jan-2016
Mutagen category 1B (Table 3.1)/ mutagen category 2 (Table 3.2) listed in Appendix 4 |  | - | - | 29 | https://echa.europa.eu/documents/10162/ec709879-c587-5b0e-6384-89c34b408717 | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=307">Foreword</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=468">Appendix 3</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=469">Appendix 4</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=581">Appendix 11</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.068.01.0024.01.ENG&toc=OJ:L:1997:068:TOC">D 97/10/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.333.01.0001.01.ENG&toc=OJ:L:1997:333:TOC">D 97/56/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.166.01.0087.01.ENG&toc=OJ:L:1999:166:TOC">D 1999/43/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0014.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/34/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0026.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/36/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.033.01.0028.01.ENG&toc=OJ:L:2006:033:TOC">D 2005/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.037.01.0001.01.ENG&toc=OJ:L:2012:037:TOC">R 109/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2017.224.01.0110.01.ENG&toc=OJ:L:2017:224:TOC">R 1510/2017</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.114.01.0004.01.ENG&toc=OJ:L:2018:114:TOC">R 2018/675</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.446.01.0034.01.ENG&toc=OJ%3AL%3A2021%3A446%3ATOC">R 2021/2204</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX:32025R1731">R 2025/1731</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/156">entry 156</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1523">entry 1523</a> |  | 15-Jan-2016
Mutagen category 1A (Table 3.1)/ mutagen category 1 (Table 3.2) listed in Appendix 3 |  | - | - | 29 | https://echa.europa.eu/documents/10162/ec709879-c587-5b0e-6384-89c34b408717 | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=307">Foreword</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=468">Appendix 3</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=469">Appendix 4</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=581">Appendix 11</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.068.01.0024.01.ENG&toc=OJ:L:1997:068:TOC">D 97/10/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.333.01.0001.01.ENG&toc=OJ:L:1997:333:TOC">D 97/56/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.166.01.0087.01.ENG&toc=OJ:L:1999:166:TOC">D 1999/43/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0014.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/34/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0026.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/36/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.033.01.0028.01.ENG&toc=OJ:L:2006:033:TOC">D 2005/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.037.01.0001.01.ENG&toc=OJ:L:2012:037:TOC">R 109/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2017.224.01.0110.01.ENG&toc=OJ:L:2017:224:TOC">R 1510/2017</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.114.01.0004.01.ENG&toc=OJ:L:2018:114:TOC">R 2018/675</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.446.01.0034.01.ENG&toc=OJ%3AL%3A2021%3A446%3ATOC">R 2021/2204</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX:32025R1731">R 2025/1731</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/156">entry 156</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1523">entry 1523</a> |  | 15-Jan-2016
Substances which are classified as reproductive toxicant category 1A or 1B in Part 3 of Annex VI to Regulation (EC) No 1272/2008 and are listed in Appendix 5 or Appendix 6, respectively. |  | - | - | 30 | https://echa.europa.eu/documents/10162/61845f2b-f319-ab2e-24aa-6fc4f8fc150f | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=307">Foreword</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=556">Appendix 5</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=558">Appendix 6</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=581">Appendix 11</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.068.01.0024.01.ENG&toc=OJ:L:1997:068:TOC">D 97/10/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.333.01.0001.01.ENG&toc=OJ:L:1997:333:TOC">D 97/56/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.166.01.0087.01.ENG&toc=OJ:L:1999:166:TOC">D 1999/43/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0014.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/34/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0026.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/36/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.033.01.0028.01.ENG&toc=OJ:L:2006:033:TOC">D 2005/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.037.01.0001.01.ENG&toc=OJ:L:2012:037:TOC">R 109/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2017.224.01.0110.01.ENG&toc=OJ:L:2017:224:TOC">R 1510/2017</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.114.01.0004.01.ENG&toc=OJ:L:2018:114:TOC">R 2018/675</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.446.01.0034.01.ENG&toc=OJ%3AL%3A2021%3A446%3ATOC">R 2021/2204</a>#<a href="https://eur-lex.europa.eu/eli/reg/2023/1132/oj/eng">R 2023/1132</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX:32025R1731">R 2025/1731</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/156">entry 156</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1523">entry 1523</a> |  | 15-Jan-2016
Reproductive toxicant category 1B adverse effects on sexual function and fertility or on development (Table 3.1) or reproductive toxicant category 2 with R60 (May impair fertility) or R61 (May cause harm to the unborn child) (Table 3.2) listed in Appendix 6 |  | - | - | 30 | https://echa.europa.eu/documents/10162/61845f2b-f319-ab2e-24aa-6fc4f8fc150f | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=307">Foreword</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=556">Appendix 5</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=558">Appendix 6</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=581">Appendix 11</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.068.01.0024.01.ENG&toc=OJ:L:1997:068:TOC">D 97/10/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.333.01.0001.01.ENG&toc=OJ:L:1997:333:TOC">D 97/56/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.166.01.0087.01.ENG&toc=OJ:L:1999:166:TOC">D 1999/43/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0014.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/34/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0026.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/36/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.033.01.0028.01.ENG&toc=OJ:L:2006:033:TOC">D 2005/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.037.01.0001.01.ENG&toc=OJ:L:2012:037:TOC">R 109/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2017.224.01.0110.01.ENG&toc=OJ:L:2017:224:TOC">R 1510/2017</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.114.01.0004.01.ENG&toc=OJ:L:2018:114:TOC">R 2018/675</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.446.01.0034.01.ENG&toc=OJ%3AL%3A2021%3A446%3ATOC">R 2021/2204</a>#<a href="https://eur-lex.europa.eu/eli/reg/2023/1132/oj/eng">R 2023/1132</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX:32025R1731">R 2025/1731</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/156">entry 156</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1523">entry 1523</a> |  | 15-Jan-2016
Reproductive toxicant category 1A adverse effects on sexual function and fertility or on development (Table 3.1) or reproductive toxicant category 1 with R60 (May impair fertility) or R61 (May cause harm to the unborn child) (Table 3.2) listed in Appendix 5 |  | - | - | 30 | https://echa.europa.eu/documents/10162/61845f2b-f319-ab2e-24aa-6fc4f8fc150f | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=307">Foreword</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=556">Appendix 5</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=558">Appendix 6</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=581">Appendix 11</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.068.01.0024.01.ENG&toc=OJ:L:1997:068:TOC">D 97/10/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.333.01.0001.01.ENG&toc=OJ:L:1997:333:TOC">D 97/56/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1999.166.01.0087.01.ENG&toc=OJ:L:1999:166:TOC">D 1999/43/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0014.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/34/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.156.01.0026.01.ENG&toc=OJ:L:2003:156:TOC">D 2003/36/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2006.033.01.0028.01.ENG&toc=OJ:L:2006:033:TOC">D 2005/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.037.01.0001.01.ENG&toc=OJ:L:2012:037:TOC">R 109/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2017.224.01.0110.01.ENG&toc=OJ:L:2017:224:TOC">R 1510/2017</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.114.01.0004.01.ENG&toc=OJ:L:2018:114:TOC">R 2018/675</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.446.01.0034.01.ENG&toc=OJ%3AL%3A2021%3A446%3ATOC">R 2021/2204</a>#<a href="https://eur-lex.europa.eu/eli/reg/2023/1132/oj/eng">R 2023/1132</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX:32025R1731">R 2025/1731</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/156">entry 156</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1523">entry 1523</a> |  | 15-Jan-2016
Creosote and Creosote related substances | Tar Oils derived from coal and/or wood by distillation or by other processes | - | - | 31 | https://echa.europa.eu/documents/10162/a27e80a3-3798-3c76-01a0-32357cc09f6f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1070">entry 1070</a> |  | 13-Jan-2016
Creosote oil, high-boiling distillate | The high-boiling distillation fraction obtained from the high temperature carbonization of bituminous coal which is further refined to remove excess crystalline salts. It consists primarily of creosote oil with some of the normal polynuclear aromatic salts, which are components of coal tar distillates, removed. It is crystal free at approximately 5°C (41°F). | 274-565-9 | 70321-79-8 | 31 | https://echa.europa.eu/documents/10162/a27e80a3-3798-3c76-01a0-32357cc09f6f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1070">entry 1070</a> |  | 13-Jan-2016
Distillates (coal tar), upper; heavy anthracene oil | The distillate from coal tar having an approximate distillation range of 220°C to 450°C (428°F to 842°F). Composed primarily of three to four membered condensed ring aromatic hydrocarbons and other hydrocarbons. | 266-026-1 | 65996-91-0 | 31 | https://echa.europa.eu/documents/10162/a27e80a3-3798-3c76-01a0-32357cc09f6f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1070">entry 1070</a> |  | 13-Jan-2016
Anthracene oil | A complex combination of polycyclic aromatic hydrocarbons obtained from coal tar having an approximate distillation range of 300°C to 400°C (572°F to 752°F). Composed primarily of phenanthrene, anthracene and carbazole | 292-602-7 | 90640-80-5 | 31 | https://echa.europa.eu/documents/10162/a27e80a3-3798-3c76-01a0-32357cc09f6f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1070">entry 1070</a> |  | 13-Jan-2016
Distillates (coal tar), naphthalene oils; naphthalene oil | A complex combination of hydrocarbons obtained by the distillation of coal tar. It consists primarily of aromatic and other hydrocarbons, phenolic compounds and aromatic nitrogen compounds and distills in the approximate range of 200°C to 250°C (392°F to 482°F). | 283-484-8 | 84650-04-4 | 31 | https://echa.europa.eu/documents/10162/a27e80a3-3798-3c76-01a0-32357cc09f6f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1070">entry 1070</a> |  | 13-Jan-2016
Low temperature tar oil, alkaline; extract residues (coal), low temperature coal tar alkaline | The residue from low temperature coal tar oils after an alkaline wash, such as aqueous sodium hydroxide, to remove crude coal tar acids. Composed primarily of hydrocarbons and aromatic nitrogen bases. | 310-191-5 | 122384-78-5 | 31 | https://echa.europa.eu/documents/10162/a27e80a3-3798-3c76-01a0-32357cc09f6f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1070">entry 1070</a> |  | 13-Jan-2016
Tar acids, coal, crude; crude phenols | The reaction product obtained by neutralizing coal tar oil alkaline extract with an acidic solution, such as aqueous sulfuric acid, or gaseous carbon dioxide, to obtain the free acids. Composed primarily of tar acids such as phenol, cresols, and xylenols. | 266-019-3 | 65996-85-2 | 31 | https://echa.europa.eu/documents/10162/a27e80a3-3798-3c76-01a0-32357cc09f6f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1070">entry 1070</a> |  | 13-Jan-2016
Creosote oil; wash oil | A complex combination of hydrocarbons obtained by the distillation of coal tar. It consists primarily of aromatic hydrocarbons and may contain appreciable quantities of tar acids and tar bases. It distills at the approximate range of 200°C to 325°C (392°F to 617°F). | 263-047-8 | 61789-28-4 | 31 | https://echa.europa.eu/documents/10162/a27e80a3-3798-3c76-01a0-32357cc09f6f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1070">entry 1070</a> |  | 13-Jan-2016
Creosote, wood | A complex combination of phenols obtained as a distillate from wood tar. | 232-419-1 | 8021-39-4 | 31 | https://echa.europa.eu/documents/10162/a27e80a3-3798-3c76-01a0-32357cc09f6f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1070">entry 1070</a> |  | 13-Jan-2016
Creosote; wash oil | The distillate of coal tar produced by the high temperature carbonization of bituminous coal. It consists primarily of aromatic hydrocarbons, tar acids and tar bases. | 232-287-5 | 8001-58-9 | 31 | https://echa.europa.eu/documents/10162/a27e80a3-3798-3c76-01a0-32357cc09f6f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1070">entry 1070</a> |  | 13-Jan-2016
Creosote oil, acenaphthene fraction; wash oil | A complex combination of hydrocarbons produced by the distillation of coal tar and boiling in the range of approximately 240°C to 280°C (464°F to 536°F). Composed primarily of acenaphthene, naphthalene and alkyl naphthalene. | 292-605-3 | 90640-84-9 | 31 | https://echa.europa.eu/documents/10162/a27e80a3-3798-3c76-01a0-32357cc09f6f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1070">entry 1070</a> |  | 13-Jan-2016
Chloroform |  | 200-663-8 | 67-66-3 | 32 | https://echa.europa.eu/documents/10162/3a704004-7613-4635-8f97-d5e06a3dd860 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.283.01.0041.01.ENG&toc=OJ:L:2001:283:TOC">D 2001/90/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
1,1,2-Trichloroethane |  | 201-166-9 | 79-00-5 | 34 | https://echa.europa.eu/documents/10162/b6152d9d-052a-4f8c-849f-661fa4b39e99 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1996.231.01.0020.01.ENG&toc=OJ:L:1996:231:TOC">D 96/55/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
1,1,2,2-Tetrachloroethane |  | 201-197-8 | 79-34-5 | 35 | https://echa.europa.eu/documents/10162/efce8a7e-d8cb-4ab3-a266-80e92c515780 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1996.231.01.0020.01.ENG&toc=OJ:L:1996:231:TOC">D 96/55/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
1,1,1,2-Tetrachloroethane |  | - | 630-20-6 | 36 | https://echa.europa.eu/documents/10162/0eaea875-9f18-4164-98af-a46ee96b6493 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1996.231.01.0020.01.ENG&toc=OJ:L:1996:231:TOC">D 96/55/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Pentachloroethane |  | 200-925-1 | 76-01-7 | 37 | https://echa.europa.eu/documents/10162/29afaaf7-80f9-41a5-bf70-6bea78773e65 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1996.231.01.0020.01.ENG&toc=OJ:L:1996:231:TOC">D 96/55/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
1,1-Dichloroethene |  | 200-864-0 | 75-35-4 | 38 | https://echa.europa.eu/documents/10162/47239303-dea1-4d12-8a2f-b95bc7c5f1b8 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.365.01.0001.01.ENG&toc=OJ:L:1994:365:TOC">D 94/60/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1996.231.01.0020.01.ENG&toc=OJ:L:1996:231:TOC">D 96/55/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Substances classified as flammable gases category 1 or 2, flammable liquids categories 1, 2 or 3, flammable solids category 1 or 2,substances and mixtures which, in contact with water, emit flammable gases, category 1, 2 or 3, pyrophoric liquids category 1 or pyrophoric solids category 1, regardless of whether they appear in Part 3 of Annex VI to Regulation (EC) No 1272/2008 or not |  | - | - | 40 | https://echa.europa.eu/documents/10162/8551911c-3918-0cc6-c8a7-65910cb3a218 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1994.331.01.0007.01.ENG&toc=OJ:L:1994:331:TOC">D 94/48/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.353.01.0001.01.ENG&toc=OJ:L:2008:353:TOC">R 1272/2008</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/665">entry 665</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1181">entry 1181</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1182">entry 1182</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1800">entry 1800</a> |  | 13-Jan-2016
Hexachloroethane |  | 200-666-4 | 67-72-1 | 41 | https://echa.europa.eu/documents/10162/81999653-d3c7-4f71-88a3-2c9373b12d52 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.1997.116.01.0031.01.ENG&toc=OJ:L:1997:116:TOC">D 97/16/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2001.286.01.0027.01.ENG&toc=OJ:L:2001:286:TOC">D 2001/91/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Azocolourants and Azodyes |  | - | - | 43 | https://echa.europa.eu/documents/10162/bd51087b-9917-b018-69ac-d7594315e2a9 | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=578">Appendix 8</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=579">Appendix 9</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=580">Appendix 10</a> | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:67246,6270&cs=1A7D0EF6F30762E9AD164FEB88559FE50">EN ISO 17234-1:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:32710,6270&cs=1A9508DCF3932AF8C3C73003282DD2328">EN ISO 17234-2:2011</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:41035,6229&cs=12490E181E71D287FCBCD9FE342D23189">EN 14362-1:2017</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:41030,6229&cs=16E75732D3523C88F712B66DCC0D7A89E">EN 14362-3:2017</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2002.243.01.0015.01.ENG&toc=OJ:L:2002:243:TOC">D 2002/61/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32003L0003&qid=1613983904470">D 2003/3/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.057.01.0004.01.ENG&toc=OJ:L:2004:057:TOC">D 2004/21/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/66">entry 666</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1183">entry 1183</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1312">entry 1312</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1313">entry 1313</a> |  | 13-Jan-2016
A mixture of: disodium (6-(4-anisidino)-3-sulfonato-2-(3,5-dinitro-2-oxidophenylazo)-1-naphtholato)(1-(5-chloro-2-oxidophenylazo)-2-naphtholato)chromate(1-); trisodium bis(6-(4-anisidino)-3-sulfonato-2-(3,5-dinitro-2-oxidophenylazo)-1-naphtholato)chromate(1-) | AZUL MARINHO 018112; BLEU MARINE 018112; NAVY 018112; NAVY BLUE 018112 | 405-665-4 | 118685-33-9 | 43 | https://echa.europa.eu/documents/10162/bd51087b-9917-b018-69ac-d7594315e2a9 | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=578">Appendix 8</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=579">Appendix 9</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=580">Appendix 10</a> | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:67246,6270&cs=1A7D0EF6F30762E9AD164FEB88559FE50">EN ISO 17234-1:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:32710,6270&cs=1A9508DCF3932AF8C3C73003282DD2328">EN ISO 17234-2:2011</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:41035,6229&cs=12490E181E71D287FCBCD9FE342D23189">EN 14362-1:2017</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:41030,6229&cs=16E75732D3523C88F712B66DCC0D7A89E">EN 14362-3:2017</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2002.243.01.0015.01.ENG&toc=OJ:L:2002:243:TOC">D 2002/61/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32003L0003&qid=1613983904470">D 2003/3/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.057.01.0004.01.ENG&toc=OJ:L:2004:057:TOC">D 2004/21/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/66">entry 666</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1183">entry 1183</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1312">entry 1312</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1313">entry 1313</a> |  | 13-Jan-2016
4-o-tolylazo-o-toluidine |  | 202-591-2 | 97-56-3 | 43 | https://echa.europa.eu/documents/10162/bd51087b-9917-b018-69ac-d7594315e2a9 | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=578">Appendix 8</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=579">Appendix 9</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=580">Appendix 10</a> | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:67246,6270&cs=1A7D0EF6F30762E9AD164FEB88559FE50">EN ISO 17234-1:2020</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:32710,6270&cs=1A9508DCF3932AF8C3C73003282DD2328">EN ISO 17234-2:2011</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:41035,6229&cs=12490E181E71D287FCBCD9FE342D23189">EN 14362-1:2017</a>#<a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:41030,6229&cs=16E75732D3523C88F712B66DCC0D7A89E">EN 14362-3:2017</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2002.243.01.0015.01.ENG&toc=OJ:L:2002:243:TOC">D 2002/61/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32003L0003&qid=1613983904470">D 2003/3/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2004.057.01.0004.01.ENG&toc=OJ:L:2004:057:TOC">D 2004/21/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/66">entry 666</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1183">entry 1183</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1312">entry 1312</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1313">entry 1313</a> |  | 13-Jan-2016
Diphenylether, octabromo derivative C12H2Br8O |  | - | - | 45 | https://echa.europa.eu/documents/10162/ce525a61-dbc8-4847-965f-db4f6136d5b5 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.042.01.0045.01.ENG&toc=OJ:L:2003:042:TOC">D 2003/11/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Diphenyl ether, octabromo derivative |  | 251-087-9 | 32536-52-0 | 45 | https://echa.europa.eu/documents/10162/ce525a61-dbc8-4847-965f-db4f6136d5b5 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.042.01.0045.01.ENG&toc=OJ:L:2003:042:TOC">D 2003/11/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Entry 46 |  | - | - | 46 | https://echa.europa.eu/documents/10162/e5842a1e-e9f9-6096-2829-72f71c00eaab |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.178.01.0024.01.ENG&toc=OJ:L:2003:178:TOC">D 2003/53/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/667">entry 667</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1113">entry 1113</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1184">entry 1184</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1802">entry 1802</a> |  | 13-Jan-2016
Nonylphenol ethoxylates (C2H4O)nC15H24O |  | - | - | 46 | https://echa.europa.eu/documents/10162/e5842a1e-e9f9-6096-2829-72f71c00eaab |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.178.01.0024.01.ENG&toc=OJ:L:2003:178:TOC">D 2003/53/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/667">entry 667</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1113">entry 1113</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1184">entry 1184</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1802">entry 1802</a> |  | 13-Jan-2016
Nonylphenol C6H4(OH)C9H19 |  | 246-672-0 | 25154-52-3 | 46 | https://echa.europa.eu/documents/10162/e5842a1e-e9f9-6096-2829-72f71c00eaab |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.178.01.0024.01.ENG&toc=OJ:L:2003:178:TOC">D 2003/53/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32020R2096&qid=1611240953841">R 2020/2096</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/667">entry 667</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1113">entry 1113</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1184">entry 1184</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1802">entry 1802</a> |  | 13-Jan-2016
Nonylphenol ethoxylates (C2H4O)nC15H24O |  | - | - | 46a | https://echa.europa.eu/documents/10162/7dcd73a4-e80d-47c5-ba0a-a5f4361bf4b1 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.009.01.0001.01.ENG&toc=OJ:L:2016:009:TOC">R 26/2016</a> |  |  | 27-Jan-2016
[No public or meaningful name is available] |  | - | - | 46a | https://echa.europa.eu/documents/10162/7dcd73a4-e80d-47c5-ba0a-a5f4361bf4b1 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.009.01.0001.01.ENG&toc=OJ:L:2016:009:TOC">R 26/2016</a> |  |  | 27-Jan-2016
4-Nonylphenol, ethoxylated | 1 - 2.5 moles ethoxylated | 500-045-0 | 26027-38-3 | 46a | https://echa.europa.eu/documents/10162/7dcd73a4-e80d-47c5-ba0a-a5f4361bf4b1 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.009.01.0001.01.ENG&toc=OJ:L:2016:009:TOC">R 26/2016</a> |  |  | 27-Jan-2016
Nonylphenol, branched, ethoxylated | 1 - 2.5 moles ethoxylated | 500-209-1 | 68412-54-4 | 46a | https://echa.europa.eu/documents/10162/7dcd73a4-e80d-47c5-ba0a-a5f4361bf4b1 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.009.01.0001.01.ENG&toc=OJ:L:2016:009:TOC">R 26/2016</a> |  |  | 27-Jan-2016
[No public or meaningful name is available] |  | - | - | 46a | https://echa.europa.eu/documents/10162/7dcd73a4-e80d-47c5-ba0a-a5f4361bf4b1 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.009.01.0001.01.ENG&toc=OJ:L:2016:009:TOC">R 26/2016</a> |  |  | 27-Jan-2016
Isononylphenol, ethoxylated |  | 609-346-2 | 37205-87-1 | 46a | https://echa.europa.eu/documents/10162/7dcd73a4-e80d-47c5-ba0a-a5f4361bf4b1 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.009.01.0001.01.ENG&toc=OJ:L:2016:009:TOC">R 26/2016</a> |  |  | 27-Jan-2016
Nonylphenol, ethoxylated |  | 500-024-6 | 9016-45-9 | 46a | https://echa.europa.eu/documents/10162/7dcd73a4-e80d-47c5-ba0a-a5f4361bf4b1 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.009.01.0001.01.ENG&toc=OJ:L:2016:009:TOC">R 26/2016</a> |  |  | 27-Jan-2016
4-Nonylphenol, branched, ethoxylated | 1 - 2.5 moles ethoxylated | 500-315-8 | 127087-87-0 | 46a | https://echa.europa.eu/documents/10162/7dcd73a4-e80d-47c5-ba0a-a5f4361bf4b1 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.009.01.0001.01.ENG&toc=OJ:L:2016:009:TOC">R 26/2016</a> |  |  | 27-Jan-2016
Chromium VI compounds |  | - | - | 47 | https://echa.europa.eu/documents/10162/1f775bd4-b1b0-4847-937f-d6a37e2c0c98 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:40864,6035&cs=19BC92A04D7D087213187D44F697BF65E">EN 196-10:2016</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2003.178.01.0024.01.ENG&toc=OJ:L:2003:178:TOC">D 2003/53/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2014.090.01.0001.01.ENG&toc=OJ:L:2014:090:TOC">R 301/2014</a> |  |  | 13-Jan-2016
Toluene |  | 203-625-9 | 108-88-3 | 48 | https://echa.europa.eu/documents/10162/f28ce048-676d-4d17-b939-0c08e9a0db9f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.309.01.0013.01.ENG&toc=OJ:L:2005:309:TOC">D 2005/59/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/668">entry 668</a> |  | 13-Jan-2016
Trichlorobenzene |  | 204-428-0 | 120-82-1 | 49 | https://echa.europa.eu/documents/10162/fa5ae031-f3bb-41c5-87d2-dde74ccc4d56 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.309.01.0013.01.ENG&toc=OJ:L:2005:309:TOC">D 2005/59/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Polycyclic-aromatic hydrocarbons (PAH) |  | - | - | 50 | https://echa.europa.eu/documents/10162/4f099937-658f-8b86-2f62-5e767fab4d6e |  | <a href="https://publishing.energyinst.org/topics/fuel-quality-and-control/ip-test-methods/ip-346-determination-of-polycyclic-aromatics-in-unused-lubricating-base-oils-and-asphaltene-free-petroleum-fractions-dimethyl-sulphoxide-extraction-refractive-index-method">IP 346</a>#<a href="http://www.iso.org/iso/home/store/catalogue_tc/catalogue_detail.htm?csnumber=57940">ISO 21461:2012</a>#<a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT:71807&cs=13FB6411FF98AE4ECDA0104AF5FF55023">prEN 16143 rev</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.323.01.0051.01.ENG&toc=OJ:L:2005:323:TOC">D 2005/69/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.328.01.0069.01.ENG&toc=OJ:L:2013:328:TOC">R 1272/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.259.01.0001.01.ENG&toc=OJ%3AL%3A2021%3A259%3ATOC">R 2021/1199</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/669">entry 669</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/670">entry 670</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/671">entry 671</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1476">entry 1476</a> |  | 13-Jan-2016
Benzo[a]pyrene (BaP) |  | - | 50-32-8 | 50 | https://echa.europa.eu/documents/10162/4f099937-658f-8b86-2f62-5e767fab4d6e |  | <a href="https://publishing.energyinst.org/topics/fuel-quality-and-control/ip-test-methods/ip-346-determination-of-polycyclic-aromatics-in-unused-lubricating-base-oils-and-asphaltene-free-petroleum-fractions-dimethyl-sulphoxide-extraction-refractive-index-method">IP 346</a>#<a href="http://www.iso.org/iso/home/store/catalogue_tc/catalogue_detail.htm?csnumber=57940">ISO 21461:2012</a>#<a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT:71807&cs=13FB6411FF98AE4ECDA0104AF5FF55023">prEN 16143 rev</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.323.01.0051.01.ENG&toc=OJ:L:2005:323:TOC">D 2005/69/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.328.01.0069.01.ENG&toc=OJ:L:2013:328:TOC">R 1272/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.259.01.0001.01.ENG&toc=OJ%3AL%3A2021%3A259%3ATOC">R 2021/1199</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/669">entry 669</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/670">entry 670</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/671">entry 671</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1476">entry 1476</a> |  | 13-Jan-2016
Dibenzo[a,h]anthracene (DBAhA) |  | - | 53-70-3 | 50 | https://echa.europa.eu/documents/10162/4f099937-658f-8b86-2f62-5e767fab4d6e |  | <a href="https://publishing.energyinst.org/topics/fuel-quality-and-control/ip-test-methods/ip-346-determination-of-polycyclic-aromatics-in-unused-lubricating-base-oils-and-asphaltene-free-petroleum-fractions-dimethyl-sulphoxide-extraction-refractive-index-method">IP 346</a>#<a href="http://www.iso.org/iso/home/store/catalogue_tc/catalogue_detail.htm?csnumber=57940">ISO 21461:2012</a>#<a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT:71807&cs=13FB6411FF98AE4ECDA0104AF5FF55023">prEN 16143 rev</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.323.01.0051.01.ENG&toc=OJ:L:2005:323:TOC">D 2005/69/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.328.01.0069.01.ENG&toc=OJ:L:2013:328:TOC">R 1272/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.259.01.0001.01.ENG&toc=OJ%3AL%3A2021%3A259%3ATOC">R 2021/1199</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/669">entry 669</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/670">entry 670</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/671">entry 671</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1476">entry 1476</a> |  | 13-Jan-2016
Benzo[a]anthracene (BaA) |  | - | 56-55-3 | 50 | https://echa.europa.eu/documents/10162/4f099937-658f-8b86-2f62-5e767fab4d6e |  | <a href="https://publishing.energyinst.org/topics/fuel-quality-and-control/ip-test-methods/ip-346-determination-of-polycyclic-aromatics-in-unused-lubricating-base-oils-and-asphaltene-free-petroleum-fractions-dimethyl-sulphoxide-extraction-refractive-index-method">IP 346</a>#<a href="http://www.iso.org/iso/home/store/catalogue_tc/catalogue_detail.htm?csnumber=57940">ISO 21461:2012</a>#<a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT:71807&cs=13FB6411FF98AE4ECDA0104AF5FF55023">prEN 16143 rev</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.323.01.0051.01.ENG&toc=OJ:L:2005:323:TOC">D 2005/69/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.328.01.0069.01.ENG&toc=OJ:L:2013:328:TOC">R 1272/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.259.01.0001.01.ENG&toc=OJ%3AL%3A2021%3A259%3ATOC">R 2021/1199</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/669">entry 669</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/670">entry 670</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/671">entry 671</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1476">entry 1476</a> |  | 13-Jan-2016
Chrysen (CHR) |  | - | 218-01-9 | 50 | https://echa.europa.eu/documents/10162/4f099937-658f-8b86-2f62-5e767fab4d6e |  | <a href="https://publishing.energyinst.org/topics/fuel-quality-and-control/ip-test-methods/ip-346-determination-of-polycyclic-aromatics-in-unused-lubricating-base-oils-and-asphaltene-free-petroleum-fractions-dimethyl-sulphoxide-extraction-refractive-index-method">IP 346</a>#<a href="http://www.iso.org/iso/home/store/catalogue_tc/catalogue_detail.htm?csnumber=57940">ISO 21461:2012</a>#<a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT:71807&cs=13FB6411FF98AE4ECDA0104AF5FF55023">prEN 16143 rev</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.323.01.0051.01.ENG&toc=OJ:L:2005:323:TOC">D 2005/69/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.328.01.0069.01.ENG&toc=OJ:L:2013:328:TOC">R 1272/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.259.01.0001.01.ENG&toc=OJ%3AL%3A2021%3A259%3ATOC">R 2021/1199</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/669">entry 669</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/670">entry 670</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/671">entry 671</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1476">entry 1476</a> |  | 13-Jan-2016
Benzo[j]fluoranthene (BjFA) |  | - | 205-82-3 | 50 | https://echa.europa.eu/documents/10162/4f099937-658f-8b86-2f62-5e767fab4d6e |  | <a href="https://publishing.energyinst.org/topics/fuel-quality-and-control/ip-test-methods/ip-346-determination-of-polycyclic-aromatics-in-unused-lubricating-base-oils-and-asphaltene-free-petroleum-fractions-dimethyl-sulphoxide-extraction-refractive-index-method">IP 346</a>#<a href="http://www.iso.org/iso/home/store/catalogue_tc/catalogue_detail.htm?csnumber=57940">ISO 21461:2012</a>#<a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT:71807&cs=13FB6411FF98AE4ECDA0104AF5FF55023">prEN 16143 rev</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.323.01.0051.01.ENG&toc=OJ:L:2005:323:TOC">D 2005/69/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.328.01.0069.01.ENG&toc=OJ:L:2013:328:TOC">R 1272/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.259.01.0001.01.ENG&toc=OJ%3AL%3A2021%3A259%3ATOC">R 2021/1199</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/669">entry 669</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/670">entry 670</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/671">entry 671</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1476">entry 1476</a> |  | 13-Jan-2016
Benzo[b]fluoranthene (BbFA) |  | - | 205-99-2 | 50 | https://echa.europa.eu/documents/10162/4f099937-658f-8b86-2f62-5e767fab4d6e |  | <a href="https://publishing.energyinst.org/topics/fuel-quality-and-control/ip-test-methods/ip-346-determination-of-polycyclic-aromatics-in-unused-lubricating-base-oils-and-asphaltene-free-petroleum-fractions-dimethyl-sulphoxide-extraction-refractive-index-method">IP 346</a>#<a href="http://www.iso.org/iso/home/store/catalogue_tc/catalogue_detail.htm?csnumber=57940">ISO 21461:2012</a>#<a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT:71807&cs=13FB6411FF98AE4ECDA0104AF5FF55023">prEN 16143 rev</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.323.01.0051.01.ENG&toc=OJ:L:2005:323:TOC">D 2005/69/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.328.01.0069.01.ENG&toc=OJ:L:2013:328:TOC">R 1272/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.259.01.0001.01.ENG&toc=OJ%3AL%3A2021%3A259%3ATOC">R 2021/1199</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/669">entry 669</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/670">entry 670</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/671">entry 671</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1476">entry 1476</a> |  | 13-Jan-2016
Benzo[k]fluoranthene (BkFA) |  | - | 207-08-9 | 50 | https://echa.europa.eu/documents/10162/4f099937-658f-8b86-2f62-5e767fab4d6e |  | <a href="https://publishing.energyinst.org/topics/fuel-quality-and-control/ip-test-methods/ip-346-determination-of-polycyclic-aromatics-in-unused-lubricating-base-oils-and-asphaltene-free-petroleum-fractions-dimethyl-sulphoxide-extraction-refractive-index-method">IP 346</a>#<a href="http://www.iso.org/iso/home/store/catalogue_tc/catalogue_detail.htm?csnumber=57940">ISO 21461:2012</a>#<a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT:71807&cs=13FB6411FF98AE4ECDA0104AF5FF55023">prEN 16143 rev</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.323.01.0051.01.ENG&toc=OJ:L:2005:323:TOC">D 2005/69/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.328.01.0069.01.ENG&toc=OJ:L:2013:328:TOC">R 1272/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.259.01.0001.01.ENG&toc=OJ%3AL%3A2021%3A259%3ATOC">R 2021/1199</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/669">entry 669</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/670">entry 670</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/671">entry 671</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1476">entry 1476</a> |  | 13-Jan-2016
Benzo[e]pyrene (BeP) |  | - | 192-97-2 | 50 | https://echa.europa.eu/documents/10162/4f099937-658f-8b86-2f62-5e767fab4d6e |  | <a href="https://publishing.energyinst.org/topics/fuel-quality-and-control/ip-test-methods/ip-346-determination-of-polycyclic-aromatics-in-unused-lubricating-base-oils-and-asphaltene-free-petroleum-fractions-dimethyl-sulphoxide-extraction-refractive-index-method">IP 346</a>#<a href="http://www.iso.org/iso/home/store/catalogue_tc/catalogue_detail.htm?csnumber=57940">ISO 21461:2012</a>#<a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT:71807&cs=13FB6411FF98AE4ECDA0104AF5FF55023">prEN 16143 rev</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.323.01.0051.01.ENG&toc=OJ:L:2005:323:TOC">D 2005/69/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.328.01.0069.01.ENG&toc=OJ:L:2013:328:TOC">R 1272/2013</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.259.01.0001.01.ENG&toc=OJ%3AL%3A2021%3A259%3ATOC">R 2021/1199</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/669">entry 669</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/670">entry 670</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/671">entry 671</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1476">entry 1476</a> |  | 13-Jan-2016
Polycyclic aromatic hydrocarbons (PAH) |  | - | - | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Benzo[a]pyrene (BaP) |  | - | 50-32-8 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Chrysen (CHR) |  | - | 218-01-9 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Benzo[a]anthracene (BaA) |  | - | 56-55-3 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Acenaphthene |  | 201-469-6 | 83-32-9 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Phenanthrene |  | 201-581-5 | 85-01-8 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Fluorene |  | 201-695-5 | 86-73-7 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Naphthalene |  | 202-049-5 | 91-20-3 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Anthracene |  | 204-371-1 | 120-12-7 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Pyrene |  | 204-927-3 | 129-00-0 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Benzo[ghi]perylene |  | 205-883-8 | 191-24-2 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Benzo[e]pyrene (BeP) |  | - | 192-97-2 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Indeno[1,2,3-cd]pyrene |  | 205-893-2 | 193-39-5 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Benzo[j]fluoranthene (BjFA) |  | - | 205-82-3 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Benzo[b]fluoranthene (BbFA) |  | - | 205-99-2 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Fluoranthene |  | 205-912-4 | 206-44-0 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Benzo[k]fluoranthene (BkFA) |  | - | 207-08-9 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Acenaphthylene |  | 205-917-1 | 208-96-8 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Dibenzo[a,h]anthracene (DBAhA) |  | - | 53-70-3 | 50a | https://echa.europa.eu/documents/10162/2b2f0a5d-a75d-7aa4-fc01-6208b17a52ce |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202500660">R 2025/660 (Annex XVII entry 50a)</a> |  |  | 03-Jun-2025
Diisobutyl phthalate (DIBP); Dibutyl phthalate (DBP); Benzyl butyl phthalate (BBP); Bis(2-ethylhexyl) phthalate (DEHP) |  | - | - | 51 | https://echa.europa.eu/documents/10162/aaa92146-a005-1dc2-debe-93c80b57c5ee |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.344.01.0040.01.ENG&toc=OJ:L:2005:344:TOC">D 2005/84/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.322.01.0014.01.ENG&toc=OJ:L:2018:322:TOC">R 2018/2005</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/672">entry 672</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/673">entry 673</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/674">entry 674</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/675">entry 675</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/676">entry 676</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/984">entry 984</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/985">entry 985</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1185">entry 1185</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1314">entry 1314</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1524">entry 1524</a> |  | 15-Jan-2016
Bis (2-ethylhexyl) phthalate (DEHP) |  | 204-211-0 | 117-81-7 | 51 | https://echa.europa.eu/documents/10162/aaa92146-a005-1dc2-debe-93c80b57c5ee |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.344.01.0040.01.ENG&toc=OJ:L:2005:344:TOC">D 2005/84/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.322.01.0014.01.ENG&toc=OJ:L:2018:322:TOC">R 2018/2005</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/672">entry 672</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/673">entry 673</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/674">entry 674</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/675">entry 675</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/676">entry 676</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/984">entry 984</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/985">entry 985</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1185">entry 1185</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1314">entry 1314</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1524">entry 1524</a> |  | 15-Jan-2016
Benzyl butyl phthalate (BBP) |  | 201-622-7 | 85-68-7 | 51 | https://echa.europa.eu/documents/10162/aaa92146-a005-1dc2-debe-93c80b57c5ee |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.344.01.0040.01.ENG&toc=OJ:L:2005:344:TOC">D 2005/84/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.322.01.0014.01.ENG&toc=OJ:L:2018:322:TOC">R 2018/2005</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/672">entry 672</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/673">entry 673</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/674">entry 674</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/675">entry 675</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/676">entry 676</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/984">entry 984</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/985">entry 985</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1185">entry 1185</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1314">entry 1314</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1524">entry 1524</a> |  | 15-Jan-2016
Dibutyl phthalate (DBP) |  | 201-557-4 | 84-74-2 | 51 | https://echa.europa.eu/documents/10162/aaa92146-a005-1dc2-debe-93c80b57c5ee |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.344.01.0040.01.ENG&toc=OJ:L:2005:344:TOC">D 2005/84/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.322.01.0014.01.ENG&toc=OJ:L:2018:322:TOC">R 2018/2005</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/672">entry 672</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/673">entry 673</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/674">entry 674</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/675">entry 675</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/676">entry 676</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/984">entry 984</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/985">entry 985</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1185">entry 1185</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1314">entry 1314</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1524">entry 1524</a> |  | 15-Jan-2016
Diisobutyl phthalate |  | 201-553-2 | 84-69-5 | 51 | https://echa.europa.eu/documents/10162/aaa92146-a005-1dc2-debe-93c80b57c5ee |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.344.01.0040.01.ENG&toc=OJ:L:2005:344:TOC">D 2005/84/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.322.01.0014.01.ENG&toc=OJ:L:2018:322:TOC">R 2018/2005</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/672">entry 672</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/673">entry 673</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/674">entry 674</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/675">entry 675</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/676">entry 676</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/984">entry 984</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/985">entry 985</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1185">entry 1185</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1314">entry 1314</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1524">entry 1524</a> |  | 15-Jan-2016
The following phthalates (or other CAS and EC numbers covering the substance) (See group members) [Entry 52] |  | - | - | 52 | https://echa.europa.eu/documents/10162/0923c10f-5495-16ee-5c0f-dd3613fbde6d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.344.01.0040.01.ENG&toc=OJ:L:2005:344:TOC">D 2005/84/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/672">entry 672</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/673">entry 673</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/674">entry 674</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/675">entry 675</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/676">entry 676</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/748">entry 748</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/984">entry 984</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1185">entry 1185</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1314">entry 1314</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1524">entry 1524</a> | Guideline available, see Q&amp;A entry 748 | 15-Jan-2016
Di-n-octyl phthalate (DNOP) |  | 204-214-7 | 117-84-0 | 52 | https://echa.europa.eu/documents/10162/0923c10f-5495-16ee-5c0f-dd3613fbde6d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.344.01.0040.01.ENG&toc=OJ:L:2005:344:TOC">D 2005/84/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/672">entry 672</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/673">entry 673</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/674">entry 674</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/675">entry 675</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/676">entry 676</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/748">entry 748</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/984">entry 984</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1185">entry 1185</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1314">entry 1314</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1524">entry 1524</a> | Guideline available, see Q&amp;A entry 748 | 15-Jan-2016
Di-“isodecyl” phthalate (DIDP) |  | 247-977-1 | 26761-40-0 | 52 | https://echa.europa.eu/documents/10162/0923c10f-5495-16ee-5c0f-dd3613fbde6d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.344.01.0040.01.ENG&toc=OJ:L:2005:344:TOC">D 2005/84/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/672">entry 672</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/673">entry 673</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/674">entry 674</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/675">entry 675</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/676">entry 676</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/748">entry 748</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/984">entry 984</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1185">entry 1185</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1314">entry 1314</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1524">entry 1524</a> | Guideline available, see Q&amp;A entry 748 | 15-Jan-2016
1,2-Benzenedicarboxylic acid, di-C9-11-branched alkyl esters, C10-rich |  | 271-091-4 | 68515-49-1 | 52 | https://echa.europa.eu/documents/10162/0923c10f-5495-16ee-5c0f-dd3613fbde6d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.344.01.0040.01.ENG&toc=OJ:L:2005:344:TOC">D 2005/84/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/672">entry 672</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/673">entry 673</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/674">entry 674</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/675">entry 675</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/676">entry 676</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/748">entry 748</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/984">entry 984</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1185">entry 1185</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1314">entry 1314</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1524">entry 1524</a> | Guideline available, see Q&amp;A entry 748 | 15-Jan-2016
1,2-Benzenedicarboxylic acid, di-C8-10-branched alkyl esters, C9-rich |  | 271-090-9 | 68515-48-0 | 52 | https://echa.europa.eu/documents/10162/0923c10f-5495-16ee-5c0f-dd3613fbde6d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.344.01.0040.01.ENG&toc=OJ:L:2005:344:TOC">D 2005/84/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/672">entry 672</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/673">entry 673</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/674">entry 674</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/675">entry 675</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/676">entry 676</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/748">entry 748</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/984">entry 984</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1185">entry 1185</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1314">entry 1314</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1524">entry 1524</a> | Guideline available, see Q&amp;A entry 748 | 15-Jan-2016
Di-“isononyl” phthalate (DINP) |  | 249-079-5 | 28553-12-0 | 52 | https://echa.europa.eu/documents/10162/0923c10f-5495-16ee-5c0f-dd3613fbde6d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2005.344.01.0040.01.ENG&toc=OJ:L:2005:344:TOC">D 2005/84/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.058.01.0043.01.ENG&toc=OJ:L:2015:058:TOC">R 326/2015</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/672">entry 672</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/673">entry 673</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/674">entry 674</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/675">entry 675</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/676">entry 676</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/748">entry 748</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/984">entry 984</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1185">entry 1185</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1314">entry 1314</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1524">entry 1524</a> | Guideline available, see Q&amp;A entry 748 | 15-Jan-2016
2-(2-methoxyethoxy)ethanol (DEGME) |  | 203-906-6 | 111-77-3 | 54 | https://echa.europa.eu/documents/10162/c0b7baea-d2d6-4fa5-86b8-c322ab573d8f |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.348.01.0108.01.ENG&toc=OJ:L:2008:348:TOC">D 1348/2008/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
2-(2-butoxyethoxy)ethanol (DEGBE) |  | 203-961-6 | 112-34-5 | 55 | https://echa.europa.eu/documents/10162/9e666f9e-efe4-4fc5-ac35-66c557bcc0dc |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.348.01.0108.01.ENG&toc=OJ:L:2008:348:TOC">D 1348/2008/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Methylenediphenyl diisocyanate (MDI) including the following specific isomers (See group members): |  | - | - | 56 | https://echa.europa.eu/documents/10162/a3af0497-4775-1a05-7037-80d653a39272 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:60083,6062&cs=172DF0C3749F218A646C295CFBDCAFD28">EN 14387:2021</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.348.01.0108.01.ENG&toc=OJ:L:2008:348:TOC">D 1348/2008/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/677">entry 677</a> |  | 15-Jan-2016
4,4’-Methylenediphenyl diisocyanate |  | 202-966-0 | 101-68-8 | 56 | https://echa.europa.eu/documents/10162/a3af0497-4775-1a05-7037-80d653a39272 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:60083,6062&cs=172DF0C3749F218A646C295CFBDCAFD28">EN 14387:2021</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.348.01.0108.01.ENG&toc=OJ:L:2008:348:TOC">D 1348/2008/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/677">entry 677</a> |  | 15-Jan-2016
Methylenediphenyl diisocyanate (MDI) |  | 247-714-0 | 26447-40-5 | 56 | https://echa.europa.eu/documents/10162/a3af0497-4775-1a05-7037-80d653a39272 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:60083,6062&cs=172DF0C3749F218A646C295CFBDCAFD28">EN 14387:2021</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.348.01.0108.01.ENG&toc=OJ:L:2008:348:TOC">D 1348/2008/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/677">entry 677</a> |  | 15-Jan-2016
2,4’-Methylenediphenyl diisocyanate |  | 227-534-9 | 5873-54-1 | 56 | https://echa.europa.eu/documents/10162/a3af0497-4775-1a05-7037-80d653a39272 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:60083,6062&cs=172DF0C3749F218A646C295CFBDCAFD28">EN 14387:2021</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.348.01.0108.01.ENG&toc=OJ:L:2008:348:TOC">D 1348/2008/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/677">entry 677</a> |  | 15-Jan-2016
2,2’-Methylenediphenyl diisocyanate |  | 219-799-4 | 2536-05-2 | 56 | https://echa.europa.eu/documents/10162/a3af0497-4775-1a05-7037-80d653a39272 |  | <a href="https://standards.cencenelec.eu/dyn/www/f?p=CEN:110:0::::FSP_PROJECT,FSP_ORG_ID:60083,6062&cs=172DF0C3749F218A646C295CFBDCAFD28">EN 14387:2021</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.348.01.0108.01.ENG&toc=OJ:L:2008:348:TOC">D 1348/2008/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2013.043.01.0024.01.ENG&toc=OJ:L:2013:043:TOC">R 126/2013</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/677">entry 677</a> |  | 15-Jan-2016
Cyclohexane |  | 203-806-2 | 110-82-7 | 57 | https://echa.europa.eu/documents/10162/ca2ec609-0177-402a-911f-6ac524275d6e |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.348.01.0108.01.ENG&toc=OJ:L:2008:348:TOC">D 1348/2008/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a> |  |  | 13-Jan-2016
Ammonium nitrate (AN) |  | 229-347-8 | 6484-52-2 | 58 | https://echa.europa.eu/documents/10162/408d3eb2-6422-7455-a359-7981864b7948 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2008.348.01.0108.01.ENG&toc=OJ:L:2008:348:TOC">D 1348/2008/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.164.01.0007.01.ENG&toc=OJ:L:2009:164:TOC">R 552/2009</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32019R1148&qid=1611240953841">R 2019/1148</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/678">entry 678</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/679">entry 679</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/680">entry 680</a> |  | 13-Jan-2016
Dichloromethane |  | 200-838-9 | 75-09-2 | 59 | https://echa.europa.eu/documents/10162/0ea58491-bb76-4a47-b1d2-36faa1e0f290 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2009.137.01.0003.01.ENG&toc=OJ:L:2009:137:TOC">D 455/2009/EC</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2010.086.01.0007.01.ENG&toc=OJ:L:2010:086:TOC">R 276/2010</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1133">entry 1133</a> |  | 13-Jan-2016
Acrylamide |  | 201-173-7 | 79-06-1 | 60 | https://echa.europa.eu/documents/10162/701f6c15-98ec-4611-ac63-35fcca2e4047 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2011.101.01.0012.01.ENG&toc=OJ:L:2011:101:TOC">R 366/2011</a> |  |  | 13-Jan-2016
Dimethyl fumarate (DMFu) |  | 210-849-0 | 624-49-7 | 61 | https://echa.europa.eu/documents/10162/f9d87b89-5870-483f-bd8a-a5cfe18095cd |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.128.01.0001.01.ENG&toc=OJ:L:2012:128:TOC">R 412/2012</a> |  |  | 13-Jan-2016
Entry 62 |  | - | - | 62 | https://echa.europa.eu/documents/10162/492c2db0-7db6-4d76-8347-fca1b22c0c7d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.253.01.0005.01.ENG&toc=OJ:L:2012:253:TOC">R 848/2012</a> |  |  | 13-Jan-2016
Phenylmercury 2-ethylhexanoate |  | 236-326-7 | 13302-00-6 | 62 | https://echa.europa.eu/documents/10162/492c2db0-7db6-4d76-8347-fca1b22c0c7d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.253.01.0005.01.ENG&toc=OJ:L:2012:253:TOC">R 848/2012</a> |  |  | 13-Jan-2016
Phenylmercury neodecanoate |  | 247-783-7 | 26545-49-3 | 62 | https://echa.europa.eu/documents/10162/492c2db0-7db6-4d76-8347-fca1b22c0c7d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.253.01.0005.01.ENG&toc=OJ:L:2012:253:TOC">R 848/2012</a> |  |  | 13-Jan-2016
Phenylmercury acetate |  | 200-532-5 | 62-38-4 | 62 | https://echa.europa.eu/documents/10162/492c2db0-7db6-4d76-8347-fca1b22c0c7d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.253.01.0005.01.ENG&toc=OJ:L:2012:253:TOC">R 848/2012</a> |  |  | 13-Jan-2016
Phenylmercury propionate |  | 203-094-3 | 103-27-5 | 62 | https://echa.europa.eu/documents/10162/492c2db0-7db6-4d76-8347-fca1b22c0c7d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.253.01.0005.01.ENG&toc=OJ:L:2012:253:TOC">R 848/2012</a> |  |  | 13-Jan-2016
Phenylmercury octanoate |  | - | 13864-38-5 | 62 | https://echa.europa.eu/documents/10162/492c2db0-7db6-4d76-8347-fca1b22c0c7d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.253.01.0005.01.ENG&toc=OJ:L:2012:253:TOC">R 848/2012</a> |  |  | 13-Jan-2016
Lead and its compounds |  | - | - | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead compounds |  | - | - | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead |  | 231-100-4 | 7439-92-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Pyrochlore, antimony lead yellow | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77588. | 232-382-1 | 8012-00-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Resin acids and Rosin acids, lead salts |  | 232-714-5 | 9008-26-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead thiosulphate |  | 236-780-6 | 13478-50-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Diantimony trilead octaoxide |  | 236-845-9 | 13510-89-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead diperchlorate |  | 237-125-7 | 13637-76-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead fumarate |  | 237-220-3 | 13698-55-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
2-ethylhexanoic acid, lead salt |  | 241-076-7 | 16996-40-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(9Z,12Z)-octadeca-9,12-dienoic acid, lead salt |  | 241-077-2 | 16996-51-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(II) maleate |  | 241-430-0 | 17406-54-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Bis(diethyldithiocarbamato-S,S')lead |  | 241-534-6 | 17549-30-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Cyclo-di-μ-oxo(μ-phthalato)trilead |  | 241-894-4 | 17976-43-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead disulphamidate |  | 237-379-9 | 13767-78-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(tetrafluoroborate) |  | 237-486-0 | 13814-96-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead nitrite |  | 237-528-8 | 13826-65-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) tellurium tetraoxide |  | 237-573-3 | 13845-35-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Tetramethyllead |  | 200-897-0 | 75-74-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Tetraethyllead |  | 201-075-4 | 78-00-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead di(acetate) |  | 206-104-4 | 301-04-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(2-ethylhexanoate) |  | 206-107-0 | 301-08-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead diformate |  | 212-371-8 | 811-54-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dipropionate |  | 212-400-4 | 814-70-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead oxalate |  | 212-413-5 | 814-93-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead tartrate |  | 212-426-6 | 815-84-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead malate |  | 212-436-0 | 816-68-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dibutyrate |  | 212-462-2 | 819-73-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) acrylate |  | 212-760-2 | 867-47-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dibenzoate |  | 212-841-2 | 873-54-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Chlorotriethylplumbane |  | 213-925-1 | 1067-14-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) methacrylate |  | 213-949-2 | 1068-61-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead distearate |  | 214-005-2 | 1072-35-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dioleate |  | 214-310-0 | 1120-46-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Chlorotriphenylplumbane |  | 214-572-6 | 1153-06-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Acetoxytriphenylplumbane |  | 214-601-2 | 1162-06-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Ethyltrimethylplumbane |  | 217-169-3 | 1762-26-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Diethyldimethylplumbane |  | 217-170-9 | 1762-27-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Triethylmethylplumbane |  | 217-171-4 | 1762-28-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Tetrabutylplumbane |  | 217-649-2 | 1920-90-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead succinate |  | 214-732-5 | 1191-18-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dioxide |  | 215-174-5 | 1309-60-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Orange lead |  | 215-235-6 | 1314-41-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead sulphide |  | 215-246-6 | 1314-87-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead telluride |  | 215-247-1 | 1314-91-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead monoxide |  | 215-267-0 | 1317-36-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Trilead-bis(carbonate)-dihydroxide 2PbCO3-Pb(OH)2 |  | 215-290-6 | 1319-46-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
2-(2,4,5,7-tetrabromo-3,6-dihydroxyxanthen-9-yl)benzoic acid, lead salt |  | 215-415-4 | 1326-05-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, bis(acetato-O)tetrahydroxytri- |  | 215-630-3 | 1335-32-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead sulfochromate yellow | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77603. | 215-693-7 | 1344-37-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Basic lead chromate orange | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77601. | 215-694-2 | 1344-38-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Chlorotrimethylplumbane |  | 216-187-9 | 1520-78-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Trilead dicitrate |  | 208-141-1 | 512-26-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead tetraacetate |  | 208-908-0 | 546-67-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dicyanide |  | 209-742-1 | 592-05-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dithiocyanate |  | 209-774-6 | 592-87-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Tetraphenyllead |  | 209-871-3 | 595-89-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Acetoxytributylplumbane |  | 219-970-3 | 2587-82-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dipicrate |  | 229-335-2 | 6477-64-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead phthalate |  | 229-927-0 | 6838-85-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Diacetoxydiphenylplumbane |  | 230-052-1 | 6928-68-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Formic acid, lead salt |  | 230-336-5 | 7056-83-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) octanoate |  | 230-784-1 | 7319-86-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Hexaphenyldiplumbane |  | 221-505-4 | 3124-01-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Docosanoic acid, lead salt |  | 221-834-3 | 3249-61-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Trilead diarsenate |  | 222-979-5 | 3687-31-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Acetoxytrimethylplumbane |  | 227-200-2 | 5711-19-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Diphenyllead dichloride |  | 218-325-3 | 2117-69-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Hexaethyldiplumbane |  | 219-216-3 | 2388-00-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead arsenite |  | 233-083-9 | 10031-13-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dibromide |  | 233-084-4 | 10031-22-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dinitrate |  | 233-245-9 | 10099-74-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) silicate |  | 233-246-4 | 10099-76-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead divanadium hexaoxide |  | 233-248-5 | 10099-79-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead diiodide |  | 233-256-9 | 10101-63-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead molybdate |  | 233-459-2 | 10190-55-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead diphosphinate |  | 233-664-7 | 10294-58-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Silicic acid, chromium lead salt |  | 234-347-6 | 11113-70-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dibismuth dilead tetraruthenium tridecaoxide |  | 234-358-6 | 11116-83-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Silicic acid, lead salt |  | 234-363-3 | 11120-22-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dicalcium lead tetraoxide |  | 234-591-3 | 12013-69-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Antimony, compound with lead (1:1) |  | 235-536-6 | 12266-38-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead hydroxide nitrate |  | 235-543-4 | 12268-84-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(maleato)trioxotetralead |  | 235-549-7 | 12275-07-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Potassium pentadecaoxodiplumbatepentaniobate(1-) |  | 235-618-1 | 12372-45-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dihydroxy[styphnato(2-)]dilead |  | 235-642-2 | 12403-82-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead germanate |  | 235-666-3 | 12435-47-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dioxobis(stearato)trilead |  | 235-702-8 | 12578-12-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead titanium zirconium oxide |  | 235-727-4 | 12626-81-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead chromate molybdate sulfate red | This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77605. | 235-759-9 | 12656-85-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead tungsten oxide |  | 235-792-9 | 12737-98-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead oxide sulfate |  | 235-803-7 | 12765-51-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dilead chromate dihydroxide |  | 234-628-3 | 12017-86-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dodecairon lead nonadecaoxide |  | 234-684-9 | 12023-90-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Hafnium lead trioxide |  | 234-735-5 | 12029-23-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead disodium dioxide |  | 234-800-8 | 12034-30-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead diniobium hexaoxide |  | 234-814-4 | 12034-88-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead tin trioxide |  | 234-844-8 | 12036-31-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead oxide sulfate |  | 234-853-7 | 12036-76-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Bismuth, compound with lead (1:1) |  | 234-980-8 | 12048-28-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dilead oxide |  | 235-036-8 | 12059-89-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead titanium trioxide |  | 235-038-9 | 12060-00-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead zirconium trioxide |  | 235-039-4 | 12060-01-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead ditantalum hexaoxide |  | 235-065-6 | 12065-68-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Pentalead tetraoxide sulphate |  | 235-067-7 | 12065-90-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead selenide |  | 235-109-4 | 12069-00-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead disulphide |  | 235-241-2 | 12137-74-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Trilead dioxide phosphonate |  | 235-252-2 | 12141-20-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Tetralead trioxide sulphate |  | 235-380-9 | 12202-17-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead chloride oxide |  | 235-383-5 | 12205-72-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Stearic acid, lead salt |  | 231-068-1 | 7428-48-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead sulphite |  | 231-196-8 | 7446-10-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead sulphate PbSO4 |  | 231-198-9 | 7446-14-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead selenate |  | 231-199-4 | 7446-15-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Trilead bis(orthophosphate) |  | 231-205-5 | 7446-27-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) selenite |  | 231-300-1 | 7488-51-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(4+) stearate |  | 231-740-4 | 7717-46-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dichloride |  | 231-845-5 | 7758-95-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead chromate |  | 231-846-0 | 7758-97-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead tungsten tetraoxide |  | 231-849-7 | 7759-01-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead difluoride |  | 231-998-8 | 7783-46-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(IV) fluoride |  | 232-012-9 | 7783-59-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead hydrogen arsenate |  | 232-064-2 | 7784-40-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Tungstate (WO42-), calcium (1:1), (T-4)-, lead-doped |  | 272-255-8 | 68784-53-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Silicic acid (H6Si2O7), barium magnesium strontium salt, lead-doped |  | 272-269-4 | 68784-74-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Silicic acid (H2Si2O5), barium salt (1:1), lead-doped |  | 272-271-5 | 68784-75-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Cadmium sulfide (CdS), copper and lead-doped |  | 272-581-0 | 68891-87-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(R)-3,5,6-trihydroxy-4,6-bis(3-methylbut-2-enyl)-2-(3-methyl-2-oxobutyl)cyclohexa-2,4-dien-1-one, lead salt |  | 272-654-7 | 68901-11-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
α-d-Glucopyranose, 1-(dihydrogen phosphate), lead salt |  | 272-655-2 | 68901-12-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Acetic acid, reaction products with lead oxide (PbO), silica and sulfuric acid |  | 273-052-7 | 68937-05-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
[phthalato(2-)]dioxotrilead |  | 273-688-5 | 69011-06-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Trilead chromate silicate |  | 273-689-0 | 69011-07-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead alloy, base, dross | A scum formed on the surface of molten lead-base alloys. Includes those cases in which aluminum is used to remove arsenic, nickel and antimony. | 273-700-9 | 69011-59-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead alloy, base, Pb,Sn, dross | Oxides formed during melting, refining, and casting of solders. Major constituents are oxides of tin, lead and antimony; minor constituents are iron, nickel, sulfur, arsenic, copper and silver. | 273-701-4 | 69011-60-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Flue dust, lead-tin alloy-manufg. | Volatilized metals generated during smelting and refining of tin-lead alloys. Consists primarily of oxides of lead and tin. | 273-758-5 | 69012-60-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Flue dust, zinc-refining | By-product of refining of zinc ores consisting primarily of zinc, lead and iron. | 273-760-6 | 69012-63-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Leach residues, zinc ore-calcine, zinc cobalt | Residue from treatment of calcined zinc ore concentrates with antimony trioxide, zinc dust, lead oxide and copper sulfate. Consists primarily of zinc and a composite of metallics: cobalt, copper and lead. | 273-769-5 | 69012-72-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, dross, antimony-rich | A scum or slag formed on the surface of molten lead during the process of removing antimony along with arsenic by oxidation with air. It consists of antimony, arsenic and lead oxides. | 273-791-5 | 69029-45-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, dross, bismuth-rich | A scum formed on the surface of molten lead during the process of removing bismuth by the addition of calcium and magnesium. It consists of lead containing calcium and magnesium bismuthides. | 273-792-0 | 69029-46-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, antimonial, dross | A scum formed on the surface of antimonial lead. Consists primarily of sodium arsenate and sodium antimonate with some lead oxide and free caustic soda. | 273-795-7 | 69029-51-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, dross |  | 273-796-2 | 69029-52-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead oxide (PbO), retort | Oxidation product of metals in lead refinery retort. Consists primarily of oxides of lead, antimony and zinc. | 273-797-8 | 69029-53-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Slags, lead reverbatory smelting | By-product from the smelting of lead ores, scrap lead or lead smelter dross. Consists primarily of oxides and silicates of antimony and lead. | 273-800-2 | 69029-58-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Zinc, desilverizing skims | Crusts formed on the surface of cooling molten lead during the desilverizing of lead. | 273-802-3 | 69029-60-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Bismuth, refinery lead chloride residues | By-product of the refining of bismuth. Principally lead chloride with salts and compounds of various metals. | 273-803-9 | 69029-61-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Flue dust, lead-refining | By-product of refining lead ores obtained from baghouse and electro-static precipitator and as slurry from scrubbers. | 273-809-1 | 69029-67-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Leach residues, lead slag | Residues from leaching slag from a lead refinery with caustic soda and sulfuric acid. Consists primarily of sodium sulfate, tellurium compounds and lead sulfate. | 273-812-8 | 69029-71-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Leach residues, precious metal recovery lead refining | Salt extract from the treatment of speiss refining residue with aqua regia followed by alkali fusion. Consists of precious metals and their salts. | 273-813-3 | 69029-72-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Calcines, lead ore conc. | Sinter formed by heating finely divided lead concentrates and fluxes to a softening condition to agglomerate without fusion. | 273-815-4 | 69029-74-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Residues, lead roaster | Roaster hearth cleanings or barrings from lead refining. Consist primarily of cadmium, thallium and lead and oxides of cadmium, thallium, calcium and arsenic. | 273-820-1 | 69029-78-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Residues, lead smelting | Residues from processing equipment used in a lead refinery. Consist primarily of lead and lead oxide. | 273-821-7 | 69029-79-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Residues, precious metal recovery lead refining | Residues from treating lead refinery ores and residues containing precious metals with sodium boroplumbate followed by thermal fusion. | 273-822-2 | 69029-80-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Slags, tellurium | Product of treating molten lead with sodium salts. Consists primarily of sodium-tellurium salts in various states of oxidation. | 273-828-5 | 69029-86-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Slimes and Sludges, lead refining | Product of neutralizing lead refining wastes to precipitate heavy metals. | 273-833-2 | 69029-93-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, dross, copper-rich | A scum formed on the surface of molten copper. | 273-925-2 | 69227-11-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, C6-19-branched carboxylate naphthenate complexes |  | 274-302-8 | 70084-67-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) nitroresorcinolate |  | 274-521-9 | 70268-38-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, decanoate octanoate complexes |  | 274-558-0 | 70321-55-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, alkyls, manufg. wastes |  | 274-632-2 | 70513-89-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Flue dust, lead blast furnace |  | 274-633-8 | 70514-05-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
2-Propenoic acid, 2-methyl-, lead salt, basic |  | 292-141-1 | 90552-19-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Sulfuric acid, lead(2+) salt, basic |  | 292-204-3 | 90583-07-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Sulfurous acid, lead(2+) salt, basic |  | 292-237-3 | 90583-37-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Tetradecanoic acid, lead salt, basic |  | 292-267-7 | 90583-65-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Citric acid, lead salt |  | 238-432-9 | 14450-60-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead acrylate |  | 238-458-0 | 14466-01-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead metaborate |  | 238-762-3 | 14720-53-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead 2,4,6-trinitro-m-phenylene dioxide |  | 239-290-0 | 15245-44-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Bis(pentane-2,4-dionato-O,O')lead |  | 239-323-9 | 15282-88-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lauric acid, lead salt |  | 239-344-3 | 15306-30-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead oleate |  | 239-378-9 | 15347-55-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead acetate |  | 239-379-4 | 15347-57-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Octanoic acid, lead salt |  | 239-790-9 | 15696-43-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Sulphuric acid, lead salt Pbx SO4 |  | 239-831-0 | 15739-80-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead disalicylate |  | 239-839-4 | 15748-73-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) decanoate |  | 239-867-7 | 15773-52-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dihexanoate |  | 239-868-2 | 15773-53-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dilaurate |  | 239-869-8 | 15773-55-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dipalmitate |  | 239-870-3 | 15773-56-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead hydrogenorthophosphate |  | 239-952-9 | 15845-52-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) tellurium trioxide |  | 239-964-4 | 15851-47-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead silicate |  | 240-047-6 | 15906-71-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead benzoate |  | 240-049-7 | 15907-04-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead phosphite |  | 240-176-8 | 16038-76-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead phthalate |  | 240-321-5 | 16183-12-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Diantimony lead tetroxide |  | 240-499-4 | 16450-50-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) (R)-12-hydroxyoleate |  | 236-010-9 | 13094-04-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) 2,4-dinitroresorcinolate |  | 236-498-3 | 13406-89-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead diazide |  | 236-542-1 | 13424-46-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dilead pyrophosphate |  | 236-629-4 | 13453-66-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead hexafluorosilicate |  | 247-278-1 | 25808-74-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Thiosulphuric acid, lead salt |  | 247-566-7 | 26265-65-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Neodecanoic acid, lead salt |  | 248-369-9 | 27253-28-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Isononanoic acid, lead salt |  | 248-376-7 | 27253-41-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) sebacate |  | 249-660-3 | 29473-77-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead didocosanoate |  | 249-719-3 | 29597-84-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dilead diruthenium hexaoxide |  | 253-389-6 | 37194-88-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dilead dirhodium heptaoxide |  | 253-421-9 | 37240-96-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead isophthalate |  | 254-128-9 | 38787-87-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead chloride silicate |  | 254-443-1 | 39390-00-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(2,4-dihydroxybenzoate) |  | 255-375-5 | 41453-50-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(piperidine-1-carbodithioate) |  | 255-438-7 | 41556-46-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead propionate |  | 255-883-7 | 42558-73-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dimyristate |  | 250-924-5 | 32112-52-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dilinoleate |  | 251-602-7 | 33627-12-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dibromate |  | 251-796-3 | 34018-28-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(II) methylthiolate |  | 252-324-9 | 35029-96-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead cyanamide |  | 252-377-8 | 35112-70-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Orthoboric acid, lead(2+) salt |  | 252-594-8 | 35498-15-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(3,5,5-trimethylhexanoate) |  | 252-746-3 | 35837-70-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(dipentyldithiocarbamate) |  | 253-068-0 | 36501-84-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dilead chromate oxide |  | 242-339-9 | 18454-12-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dilactate |  | 242-666-7 | 18917-82-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(dimethyldithiocarbamate) |  | 242-748-2 | 19010-66-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead maleate |  | 242-832-9 | 19136-34-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead palmitate |  | 243-135-2 | 19528-55-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
5,5,13,13-tetradehydro-4,5-dihydro-4,8,10,15-tetranitro-7,11-metheno-11H,13H-tetrazolo[1,5-c][1,7,3,5,2,6]dioxadiazadiplumbacyclododecine |  | 243-204-7 | 19651-80-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead hydroxide |  | 243-310-3 | 19783-14-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Phosphorodithioate O,O-bis(1,3-dimethylbutyl), lead salt |  | 243-777-3 | 20383-42-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Myristic acid, lead salt |  | 243-788-3 | 20403-41-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Decanoic acid, lead salt |  | 243-789-9 | 20403-42-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead cyanamidate |  | 244-099-0 | 20890-10-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead 2,4-dihydroxybenzoate |  | 244-118-2 | 20936-32-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead silicate |  | 245-090-4 | 22569-74-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Disodium lead N,N'-ethylenebis[N-(carboxylatomethyl)aminoacetate] |  | 245-314-0 | 22904-40-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
3,5,5-trimethylhexanoic acid, lead salt |  | 245-788-9 | 23621-79-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead carbonate |  | 247-054-3 | 25510-11-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead diiodate |  | 247-168-3 | 25659-31-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Sulfuric acid, lead salt, tetrabasic |  | 258-142-6 | 52732-72-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(isononanoate) |  | 258-222-0 | 52847-85-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Phosphonic acid, lead salt, basic |  | 258-787-3 | 53807-64-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
1,3,5-triazine-2,4,6(1H,3H,5H)-trione, lead salt |  | 259-227-0 | 54554-36-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dioxobis(stearato)dilead |  | 260-043-8 | 56189-09-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Cadmium sulfide (CdS), solid soln. with zinc sulfide, copper and lead-doped |  | 269-773-1 | 68332-81-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C8-10-branched, lead salts, basic |  | 270-063-9 | 68409-79-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead oxide (PbO), lead-contg. |  | 270-148-0 | 68411-78-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Oils, menhaden, lead salts |  | 270-321-0 | 68424-76-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Hexanoic acid, dimethyl-, lead(2+) salt, basic |  | 270-498-4 | 68442-95-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Hexanoic acid, 2-ethyl-, lead(2+) salt, basic |  | 271-106-4 | 68515-76-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Hexanoic acid, 3,5,5-trimethyl-, lead(2+) salt, basic |  | 271-108-5 | 68515-77-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isooctanoate naphthenate complexes |  | 271-110-6 | 68515-80-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
[phthalato(2-)]oxodilead |  | 260-590-2 | 57142-78-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(12-hydroxystearate) |  | 261-238-0 | 58405-97-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead 5-nitroterephthalate |  | 262-308-3 | 60580-60-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Naphthenic acids, lead manganese salts |  | 262-983-4 | 61788-52-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, tall-oil, lead manganese salts |  | 262-985-5 | 61788-53-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, tall-oil, lead salts |  | 262-986-0 | 61788-54-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Naphthenic acids, lead salts |  | 263-109-4 | 61790-14-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Sulfurous acid, lead salt, dibasic |  | 263-467-1 | 62229-08-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Bis(o-acetoxybenzoato)lead |  | 263-548-1 | 62451-77-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(4-cyclohexylbutyrate) |  | 263-663-7 | 62637-99-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) heptadecanoate |  | 264-123-3 | 63399-94-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(diisononylnaphthalenesulphonate) |  | 264-316-2 | 63568-30-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Isooctanoic acid, lead salt |  | 264-931-6 | 64504-12-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dibutanolate |  | 265-457-2 | 65119-94-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) 4,6-dinitro-o-cresolate |  | 265-460-9 | 65121-76-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead 12-hydroxyoctadecanoate |  | 265-492-3 | 65127-78-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Bismuth lead ruthenium oxide |  | 265-645-4 | 65229-22-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Frits, chemicals | Frit is a mixture of inorganic chemical substances produced by rapidly quenching a molten, complex combination of materials, confining the chemical substances thus manufactured as nonmigratory components of glassy solid flakes or granules. This category includes all of the chemical substances specified below when they are intentionally manufactured in the production of frit. The primary members of this category are oxides of some or all of the elements listed below. Fluorides of these elements may also be included in combination with these primary substances.@Aluminum@Manganese@Antimony@Molybdenum@Arsenic@Neodymium@Barium@Nickel@Bismuth@Niobium@Boron@Phosphorus@Cadmium@Potassium@Calcium@Silicon@Cerium@Silver@Chromium@Sodium@Cobalt@Strontium@Copper@Tin@Gold@Titanium@Iron@Tungsten@Lanthanum@Vanadium@Lead@Zinc@Lithium@Zirconium@Magnesium | 266-047-6 | 65997-18-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Petrolatum (petroleum), oxidized, lead salt |  | 266-853-8 | 67674-14-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dilead silicate sulphate |  | 266-962-0 | 67711-86-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C12-18, lead salts |  | 268-622-7 | 68131-60-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, 2-ethylhexanoate tall-oil fatty acids complexes |  | 269-090-9 | 68187-37-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Sulphuric acid, barium lead salt |  | 255-896-8 | 42579-89-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead naphthenate |  | 256-790-4 | 50825-29-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
3-(triphenylplumbyl)-1H-pyrazole |  | 256-968-1 | 51105-45-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Nitroresorcinol, lead salt |  | 257-133-4 | 51317-24-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Trinitrophloroglucinol, lead salt |  | 257-135-5 | 51325-28-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Acetic acid, lead salt, basic |  | 257-175-3 | 51404-69-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Octadecanoic acid, lead(2+) salt, tribasic |  | 257-645-8 | 52080-60-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Sulfurous acid, lead salt, basic |  | 257-762-4 | 52231-92-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead methacrylate |  | 258-041-7 | 52609-46-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Slimes and Sludges, lead sinter dust scrubber |  | 274-639-0 | 70514-37-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) isooctadecanoate |  | 274-822-5 | 70727-02-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) neodecanoate |  | 275-836-4 | 71684-29-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(II) fumarate |  | 275-845-3 | 71686-03-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Hydroxy(neodecanoato-O)lead |  | 275-988-1 | 71753-04-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(nonylphenolate) |  | 276-725-3 | 72586-00-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C9-11-branched, lead salts |  | 279-755-5 | 81412-57-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, C5-23-branched carboxylate C4-10-fatty acid naphthenate complexes |  | 280-541-9 | 83711-45-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, C5-23-branched carboxylate naphthenate complexes |  | 280-542-4 | 83711-46-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, C5-23-branched carboxylate naphthenate octanoate complexes |  | 280-544-5 | 83711-47-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(tetracosylbenzenesulphonate) |  | 288-691-7 | 85865-91-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis[didodecylbenzenesulphonate] |  | 288-692-2 | 85865-92-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
1,2-Benzenedicarboxylic acid, lead(2+) salt, basic |  | 290-588-7 | 90193-83-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
2-Butenedioic acid (E)-, lead(2+) salt, basic |  | 290-862-6 | 90268-59-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
2-Butenedioic acid (Z)-, lead(2+) salt, basic |  | 290-869-4 | 90268-66-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, C5-23-branched carboxylate C4-10-fatty acid complexes |  | 281-934-8 | 84066-98-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, C5-23-branched carboxylate octanoate complexes |  | 281-935-3 | 84066-99-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, C4-10-fatty acid naphthenate complexes |  | 281-936-9 | 84067-00-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Matte, lead | Substance resulting from the smelting of lead and its alloys obtained from primary and secondary sources and including recycled plant intermediates. Composed primarily of iron and lead (mainly in sulfide form) and may contain other residual non-ferrous metals and their compounds. | 282-356-9 | 84195-51-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Slimes and Sludges, lead, electrolytic | Substance resulting from the electrolytic refining of lead obtained from primary and secondary sources and including recycled plant intermediates. Composed primarily of lead, antimony and precious metals and may contain other residual non-ferrous metals and their compounds. | 282-365-8 | 84195-60-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Speiss, lead | Substance resulting from the smelting of lead and its alloys obtained from primary and secondary sources and including recycled plant intermediates. Composed primarily of arsenic, lead and iron and may contain other residual non-ferrous metals and their compounds. | 282-366-3 | 84195-61-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(p-octylphenolate) |  | 282-756-3 | 84394-98-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C8-18 and C18-unsatd., lead salts |  | 283-962-6 | 84776-36-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C8-12, lead salts |  | 283-980-4 | 84776-53-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C18-24, lead salts |  | 283-982-5 | 84776-54-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
[μ-(4,6-dinitroresorcinolato(2-)-O1,O3)]dihydroxydilead |  | 284-297-4 | 84837-22-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(II) isodecanoate |  | 284-345-4 | 84852-34-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isononanoate isooctanoate complexes, basic |  | 284-573-4 | 84929-94-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isooctanoate neodecanoate complexes, basic |  | 284-575-5 | 84929-95-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, naphthenate neodecanoate complexes, basic |  | 284-576-0 | 84929-96-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isononanoate naphthenate complexes |  | 284-577-6 | 84929-97-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Benzenesulfonic acid, 4-C10-13-sec-alkyl derivs., lead(2+) salts |  | 284-665-4 | 84961-75-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C8-10-branched, lead salts |  | 285-212-3 | 85049-42-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Flue dust, lead-manufg., cadmium-rich | Residue obtained in the metallurgical treatment of lead concentrate in a lead blast furnace. The substance is composed of cadmium oxides, lead oxides, and impurities containing compounds of arsenic, chlorine, indium and tellurium. | 285-554-3 | 85117-02-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) 4-(1,1-dimethylethyl)benzoate |  | 286-637-7 | 85292-77-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(5-oxo-L-prolinate) |  | 286-935-7 | 85392-77-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(5-oxo-DL-prolinate) |  | 286-936-2 | 85392-78-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead uranate pigment |  | 287-565-9 | 85536-79-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Linseed oil, lead manganese salt |  | 271-394-1 | 68553-17-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Oils, fish, lead salts |  | 271-396-2 | 68553-63-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Spinels, boron calcium lead silicon white | An inorganic pigment that is the reaction product of high temperature calcination in which boron oxide, calcium oxide, lead (II) oxide, and silicon oxide in varying amounts are homogeneously and ionically interdiffused to form a crystalline matrix of spinel. | 271-410-7 | 68555-05-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Spinels, lead silicon tin zinc white | An inorganic pigment that is the reaction product of high temperature calcination in which lead (II) oxide, silicon oxide, tin (IV) oxide, and zinc oxide in varying amounts are homogeneously and ionically interdiffused to form a crystalline matrix of spinel. | 271-412-8 | 68555-07-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Zinc sulfide (ZnS), copper and lead-doped |  | 271-598-0 | 68585-90-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) dodecylphenolate |  | 271-610-4 | 68586-21-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C6-19-branched, lead salts, basic |  | 271-675-9 | 68603-83-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Octanoic acid, branched, lead salts, basic |  | 271-698-4 | 68604-56-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, tallow, hydrogenated, lead salts |  | 271-722-3 | 68605-98-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, ethyl methyl derivs. |  | 271-852-0 | 68610-17-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Acetic acid, lead(2+) salt, dihydrate |  | - | 22723-52-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Phosphonic acid, lead(2+) salt |  | - | 24824-71-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, tetrapropyl- |  | - | 3440-75-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Sulfuric acid, lead(2+) lead(4+) salt (3:1:1) |  | - | 35830-81-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Tin alloy |  | - | 39412-44-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead diundec-10-enoate |  | 303-968-5 | 94232-40-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(isononanoato-O)(isooctanoato-O)lead |  | 304-157-9 | 94246-84-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(isodecanoato-O)(isooctanoato-O)lead |  | 304-158-4 | 94246-85-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(isodecanoato-O)(isononanoato-O)lead |  | 304-160-5 | 94246-86-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(isodecanoato-O)(neodecanoato-O)lead |  | 304-161-0 | 94246-87-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(2-ethylhexanoato-O)(isooctanoato-O)lead |  | 304-164-7 | 94246-90-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(2-ethylhexanoato-O)(isononanoato-O)lead |  | 304-165-2 | 94246-91-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(2-ethylhexanoato-O)(isodecanoato-O)lead |  | 304-166-8 | 94246-92-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(2-ethylhexanoato-O)(neodecanoato-O)lead |  | 304-167-3 | 94246-93-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead icosanoate (1:2) |  | 304-438-6 | 94266-31-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead icosanoate |  | 304-439-1 | 94266-32-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C6-19-branched, lead salts |  | 292-897-2 | 91002-20-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C8-9, lead salts |  | 292-964-6 | 91031-60-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C8-10, lead salts |  | 292-965-1 | 91031-61-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C16-18, lead salts |  | 292-966-7 | 91031-62-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Leach residues, zinc ore, lead-contg. | Insoluble substance obtained during dissolution of zinc ores or concentrate in sulfuric acid for the production of zinc sulfate solutions after physical separation such as flotation and filtration. | 293-314-4 | 91053-49-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Naphthenic acids, lead (2+) salts |  | 293-365-2 | 91078-81-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Isodecanoic acid, lead(2+) salt, basic |  | 293-984-8 | 91671-82-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Isooctanoic acid, lead(2+) salt, basic |  | 293-985-3 | 91671-83-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Isoundecanoic acid, lead(2+) salt, basic |  | 293-986-9 | 91671-84-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, castor-oil, hydrogenated, lead salts |  | 294-297-6 | 91697-36-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Phosphorodithioic acid, mixed O,O-bis(Bu and pentyl) esters, lead(2+) salt |  | 295-087-7 | 91783-10-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, coco, lead salts |  | 295-368-4 | 92044-89-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Naphthenic acids, lead salts, basic |  | 295-449-4 | 92045-67-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, C4-10-fatty acid octanoate complexes |  | 295-918-3 | 92200-92-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C14-26, lead salts |  | 296-911-8 | 93165-26-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Flue gases, lead-zinc blast furnace | The off-gases from the zinc/lead blast furnace. Consists primarily of carbon monoxide, carbon dioxide and nitrogen. | 298-735-7 | 93821-47-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Slimes and Sludges, lead-zinc blast furnace, offgas wet cleaning | Product generated from the wet cleaning of off-gases from the zinc/lead blast furnace. Consists primarily of zinc and lead in metallic or oxide form with varying amounts of other elements and compounds. | 298-758-2 | 93821-70-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Speiss, lead-zinc |  | 298-759-8 | 93821-72-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead 3-(acetamido)phthalate |  | 298-872-2 | 93839-98-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(2-ethylhexanolate) |  | 298-879-0 | 93840-04-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) 4,4'-isopropylidenebisphenolate |  | 299-255-0 | 93858-23-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) (Z)-hexadec-9-enoate |  | 299-256-6 | 93858-24-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) ethylphenyldithiocarbamate |  | 299-550-4 | 93892-65-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) neononanoate |  | 299-747-5 | 93894-48-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) neoundecanoate |  | 299-748-0 | 93894-49-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(neononanoato-O)(neoundecanoato-O)lead |  | 299-764-8 | 93894-64-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Phosphoric acid, mixed Bu and hexyl diesters, lead(2+) salts |  | 300-328-7 | 93925-27-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(isoundecanoate) |  | 300-979-7 | 93965-29-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead bis(tricosanoate) |  | 300-988-6 | 93966-37-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead tetracosanoate |  | 300-989-1 | 93966-38-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead pentadecanoate |  | 301-027-3 | 93966-74-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(II) isooctanoate |  | 301-153-9 | 93981-67-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Hexacosanoic acid, lead salt |  | 301-357-8 | 94006-20-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
[μ-[[5,5'-azobis[1H-tetrazolato]](2-)]]dihydroxydilead |  | 301-376-1 | 94015-57-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Decanoic acid, branched, lead salts |  | 291-160-2 | 90342-24-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dodecanoic acid, lead salt, basic |  | 291-195-3 | 90342-56-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Hexadecanoic acid, lead salt, basic |  | 291-388-2 | 90388-09-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Hexadecanoic acid, lead(2+) salt, basic |  | 291-389-8 | 90388-10-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
9-Hexadecenoic acid, lead(2+) salt, (Z)-, basic |  | 291-395-0 | 90388-15-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Isodecanoic acid, lead salt, basic |  | 291-536-6 | 90431-14-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Isononanoic acid, lead salt, basic |  | 291-543-4 | 90431-21-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Isooctanoic acid, lead salt, basic |  | 291-549-7 | 90431-26-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, C8-10-branched fatty acids C9-11-neofatty acids naphthenate complexes, overbased |  | 291-550-2 | 90431-27-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, C8-10-branched fatty acids C9-11-neofatty acids naphthenate complexes |  | 291-551-8 | 90431-28-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, 2-ethylhexanoate isodecanoate complexes, basic |  | 291-552-3 | 90431-30-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, 2-ethylhexanoate isononanoate complexes, basic |  | 291-553-9 | 90431-31-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, 2-ethylhexanoate isooctanoate complexes, basic |  | 291-554-4 | 90431-32-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, 2-ethylhexanoate naphthenate complexes |  | 291-556-5 | 90431-33-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, 2-ethylhexanoate naphthenate complexes, basic |  | 291-557-0 | 90431-34-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, 2-ethylhexanoate neodecanoate complexes, basic |  | 291-558-6 | 90431-35-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isodecanoate isononanoate complexes, basic |  | 291-559-1 | 90431-36-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isodecanoate isooctanoate complexes, basic |  | 291-560-7 | 90431-37-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isodecanoate naphthenate complexes |  | 291-561-2 | 90431-38-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isodecanoate neodecanoate complexes, basic |  | 291-562-8 | 90431-39-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isononanoate naphthenate complexes, basic |  | 291-563-3 | 90431-40-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isononanoate neodecanoate complexes, basic |  | 291-564-9 | 90431-41-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isooctanoate naphthenate complexes, basic |  | 291-565-4 | 90431-42-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, naphthenate neodecanoate complexes |  | 291-567-5 | 90431-43-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, neononanoate neoundecanoate complexes, basic |  | 291-568-0 | 90431-44-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Neodecanoic acid, lead salt, basic |  | 291-667-9 | 90459-25-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Neononanoic acid, lead salt, basic |  | 291-668-4 | 90459-26-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Neoundecanoic acid, lead salt, basic |  | 291-671-0 | 90459-28-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Octadecanoic acid, lead salt, basic |  | 291-695-1 | 90459-51-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Octadecanoic acid, lead(2+) salt, basic |  | 291-696-7 | 90459-52-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
9-Octadecenoic acid (Z)-, lead salt, basic |  | 291-735-8 | 90459-88-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead C3-13-alkanecarboxylate naphthenate complexes |  | 616-735-0 | 79803-79-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Acetic acid, oleic acid, candelilla wax, tallow, calcium-lead complex salt |  | 614-927-9 | 69103-03-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
lead(2+) phosphonate |  | 603-832-8 | 13453-65-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, tallow, reaction products with lead oxide |  | 305-196-4 | 94349-78-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
(isononanoato-O)(neodecanoato-O)lead |  | 305-385-1 | 94481-58-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, zinc dross |  | 305-409-0 | 94551-60-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Calcines, lead-zinc ore conc. | A thermally agglomerated substance formed by heating a mixture of metal sulfide concentrates, limestone, sand, furnace dross, miscellaneous zinc, lead and copper bearing materials, together with already roasted material to a temperature of 1000°C to 1200°C (538°F to 649°F). | 305-411-1 | 94551-62-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fumes, lead | Residue produced in lead smelting operations from the volatillsation of lead from materials smelted. Consists primarily of chlorides and oxides of antimony, arsenic and zinc. | 305-414-8 | 94551-66-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Leach residues, copper-lead | Substance formed by the oxidative leaching of lead- and copper-containing materials with sulfuric acid. Consists primarily of lead sulfates and other lead compounds with lesser amounts of other non-ferrous compounds. | 305-420-0 | 94551-72-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Matte, copper-lead | Substance produced as a by-product of lead smelting operations. Consists primarily of copper sulfide with significant amounts of lead sulfide and precious metals. | 305-422-1 | 94551-74-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Residues, lead smelting wastewater treatment | Material obtained during waste water treatment in lead production. Consists primarily of lead with additional metal carbonates, hydroxides, sulfides or sulfates. | 305-424-2 | 94551-78-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Residues, lead-zinc smelting wastewater treatment | Product of hydrolysis and neutralisation of zinc/lead smelter waste waters. Consists primarily of cadmium hydroxide, calcium sulfate, calcium sulfite, iron hydroxide, lead hydroxide and zinc hydroxide. | 305-425-8 | 94551-79-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Wastes, lead battery reprocessing | Material obtained during the recycling of exhausted lead storage batteries. Consists primarily of oxides and sulfates of lead and lead alloys. | 305-445-7 | 94551-99-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Waste solids, lead silver anode | The slag or residue obtained when lead/silver anodes used in the electrolytic production of zinc are recast. Fusion of the alloys of lead and silver (manganese may also be present) and simultaneous oxidation occur. | 305-449-9 | 94552-05-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead(2+) isohexadecanoate |  | 306-073-8 | 95892-13-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Slimes and Sludges, lead acetate manuf. | The residue obtained from the treatment of lead with acetic acid. Composed primarily of elemental lead, lead acetate and basic lead carbonate hydroxide. | 306-241-0 | 96690-46-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, bullion |  | 308-011-5 | 97808-88-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead fluoride hydroxide |  | 308-169-5 | 97889-90-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
7-methyloctanoic acid, lead salt |  | 308-331-5 | 97952-39-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Nitric acid, lead(2+) salt, reaction products with sodium tin oxide |  | 308-406-2 | 97953-08-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Speiss, lead, nickel-contg. | Product obtained and separated during the melting of nickel and other non-ferrous metals containing raw materials. Consists primarily of antimonides and arsenides of copper and nickel. | 308-765-5 | 98246-91-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Sulfuric acid, barium salt (1:1), lead-doped |  | 308-972-0 | 99328-54-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Perchloric acid, reaction products with lead oxide (PbO) and triethanolamine |  | 309-027-5 | 99749-31-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Silicic acid (H2SiO3), calcium salt (1:1), lead and manganese-doped |  | 309-535-7 | 100402-96-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, dross, vanadium-zinc-contg. |  | 309-638-7 | 100656-49-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Matte, copper-lead, tellurium-contg. | A matte obtained from reduction of copper electrolysis sludge from which selenium has been removed with coal. Composed primarily of copper, lead and tellurium. | 309-642-9 | 100656-53-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Flue dust, copper-lead blast furnace, cadmium-indium-enriched | A cadmium-indium-enriched product obtained from the recirculation of copper-lead blast furnace flue dusts. Composed primarily of cadmium, indium and lead. | 309-645-5 | 100656-55-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isodecanoate naphthenate complexes, basic |  | 309-793-0 | 101012-92-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, isooctanoate neodecanoate complexes |  | 309-808-0 | 101013-06-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Phosphoric acid, calcium salt (1:1), solid soln. with calcium chloride, calcium fluoride, calcium oxide, phosphorus oxide (P2O5) and strontium oxide, lead and manganese-doped |  | 310-002-6 | 102047-25-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Barium oxide (BaO), solid soln. with calcium oxide, strontium oxide and tungsten oxide (WO3), lead-doped |  | 310-023-0 | 102110-24-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Boric acid (H3BO3), solid soln. with barium oxide, calcium oxide and strontium oxide, lead and manganese-doped |  | 310-025-1 | 102110-26-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Silicic acid, calcium salt, lead and manganese-doped |  | 310-036-1 | 102110-36-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead ores, concs., leached | The residue obtained from leaching of a lead concentrate or ore to remove some of the soluble elements such as arsenic or iron. Composed primarily of lead but also containing other non-ferrous metals and gangue constituents. | 310-049-2 | 102110-48-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Residues, copper-iron-lead-nickel matte, sulfuric acid-insol. |  | 310-050-8 | 102110-49-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Slimes and Sludges, battery scrap, antimony- and lead-rich | A product obtained by the treatment of battery scraps to recover lead. Composed primarily of oxides and sulfates of antimony and lead. | 310-061-8 | 102110-60-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Slimes and Sludges, copper conc. roasting off gas scrubbing, lead-mercury-selenium-contg. | The product obtained by the purification of copper ore concentrate roasting offgas. Composed primarily of lead, mercury and selenium. | 310-062-3 | 102110-61-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Slimes and Sludges, copper-lead ore roasting off gas scrubbing, arsenic-contg. | The product obtained by the purification of copper-lead ore concentrate roasting offgas. Composed primarily of arsenic oxide (As2O3). | 310-063-9 | 102110-62-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
barium calcium cesium lead samarium strontium bromide chloride fluoride iodide europium doped |  | 431-780-4 | 199876-46-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Triethyllead acetate |  | 624-392-3 | 2587-81-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead phthalocyanine |  | 628-401-1 | 15187-16-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
2-hydroxypropane-1,2,3-tricarboxylate;lead(2+);trihydrate |  | 629-552-6 | 6107-83-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
lead(II) perchlorate trihydrate |  | 638-754-3 | 13453-62-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, triphenyl(phenylthio)- (7CI) |  | 654-801-0 | 15590-77-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, acetoxytripropyl- |  | 654-802-6 | 13266-07-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
triethylplumbanecarbonitrile |  | 667-806-8 | 13732-17-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead hexafluoro-silicate |  | 677-594-9 | 1310-03-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Boric acid (HBO2), lead(2+) salt, monohydrate |  | 677-597-5 | 10214-39-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead picrate (dry) |  | 684-862-9 | 25721-38-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
1,3-Benzenediol, 2-nitro-, lead salt, basic |  | 685-313-6 | 68411-33-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane dichlorodiethyl- |  | 621-088-2 | 13231-90-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Chromic acid lead salt with lead molybdate |  | - | 12709-98-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Ethanedioic acid, lead(2+) salt (1:1), trihydrate |  | - | 128226-81-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead hydroxide oxide (Pb2(OH)2O) |  | - | 1311-11-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Carbonic acid, lead salt |  | - | 13427-42-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead hydroxide oxide phosphite (Pb3(OH)O(PO3)), hydrate (2:1) |  | - | 1344-40-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Chloric acid, lead(2+) salt, monohydrate |  | - | 13510-96-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Chromic acid (H2CrO4), lead(2+) potassium salt (2:1:2) |  | - | 13845-31-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, tetrakis(1-methylethyl)- |  | - | 14846-40-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead and Lead compounds (as Pb) |  | - | 14931-82-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Acetic acid, lead(2+) salt, monohydrate |  | - | 15773-48-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Phosphoric acid, lead salt |  | - | 16040-38-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Phosphoric acid, lead(2+) salt (2:1) |  | - | 16180-04-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
198Pb |  | - | 16646-00-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead chromate sulfate (Pb9(CrO4)5(SO4)4) |  | - | 51899-02-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
1,2-Benzenedicarboxylic acid, lead(2+) salt |  | - | 18608-34-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Phenol, tetrapropylene-, lead(2+) salt |  | - | 122332-23-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Fatty acids, C4-20-branched, lead salts |  | - | 125328-49-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, C9-28-neocarboxylate 2-ethylhexanoate complexes, basic |  | - | 125494-56-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
2-Propenoic acid, 2-methyl-, methyl ester, polymer with ethenylbenzene, lead(2+) 2-methyl-2-propenoate (1:2) and alpha-(2-methyl-1-oxo-2-propen-1-yl)-omega-((2-methyl-1-oxo-2-propen-1-yl)oxy)poly(oxy-1,2-ethanediyl) |  | - | 68155-47-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Carboxylic acids, tall-oil, lead salts, basic |  | - | 68603-93-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Castor oil, dehydrated, polymer with rosin, calcium lead zinc salt |  | - | 68604-05-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Acetic acid, reaction products with capric acid, carpylic acid, coconut oil, and hydrogenated tallow fatty acids, calcium, lead salts |  | - | 68784-59-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Silicic acid (H4SiO4), magnesium manganese(2+) zinc salt, arsenic and lead-doped |  | - | 68784-76-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Gum rosin, litharge polymer |  | - | 68952-91-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Basic lead beta-resorcylate |  | - | 68954-05-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Gilsonite, polymer with cyclopentadiene, dicyclopentadiene and linseed oil, lead salts |  | - | 68956-49-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Gilsonite, linseed oil polymer, lead salt |  | - | 68989-89-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Linseed oil, polymer with tung oil, lead salt |  | - | 68990-75-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Antimonial lead |  | - | 69029-50-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Acetic acid, capric acid, caprylic acid, coconut oil, hydrogenated tallow fatty acids, calcium-lead complex salt |  | - | 69103-04-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
9H-Carbazole, 9-(triphenylplumbyl)- |  | - | 56240-91-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
1H-Indole, 1-(triphenylplumbyl)- |  | - | 56240-92-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
1H-Pyrrole, 1-(triphenylplumbyl)- |  | - | 56240-93-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, triethyl-, oleate |  | - | 63916-98-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
1H-Isoindole-1,3(2H)-dione, 4,5,6, 7-tetrachloro-2-(triethylplumbyl)- |  | - | 73928-23-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
1H-1,2,3-Triazole-5-methanol, .alpha.,.alpha.-dimethyl-1-[4-(triphenylplumbyl)butyl]- |  | - | 73826-06-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, [(4-aminobenzoyl)oxy]triethyl- |  | - | 73928-17-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Chromic acid (H2CrO4), lead(2 ) salt |  | - | 15804-54-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbylium, trimethyl- |  | - | 14570-16-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Diethyllead |  | - | 24952-65-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dehydrated castor oil fatty acids, glycerine, C36 fatty acid dimers, litharge polymer |  | - | 70879-91-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Silicic acid (H6Si2O7), barium zinc salt (1:1:2), lead-doped |  | - | 71799-66-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Octadecanoic acid, barium lead salt |  | - | 73105-55-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Carbamodithioic acid, phenylethyl-, lead salt |  | - | 75790-73-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Bismuth alloy, Bi 50,Pb 25,Cd 12,Sn 12 (L-pbbi50Sn12.5Cd12.5) |  | - | 76093-98-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Pigment Lightfast Lead-Molybdate Orange OS (9CI) |  | - | 78690-68-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
2-[carboxylatomethyl(carboxymethyl)amino]acetate;lead(2+) |  | - | 79849-02-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead (II) hydroxide salicylate |  | - | 87903-39-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Chromic acid (H2CrO4), lead(2+) sodium salt (2:1:2) |  | - | 93215-61-3 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, di-mu-hydroxy(2-methyl-4,6-dinitrophenolato-kappaO)(nitrato-kappaO)di- |  | - | 96471-22-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lanthanum lead titanium zirconium oxide |  | - | 1227908-26-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, bis[bis(1-methylethyl)carbamodithioato-S,S']-, (T-4)- |  | - | 30051-53-7 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, bis[bis(2-methylpropyl)carbamodithioato-S,S']-, (T-4)- |  | - | 69090-73-9 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, bis(dipropylcarbamodithioato-S,S')-, (T-4)- |  | - | 70995-63-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead sulfite |  | - | 25666-92-6 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Hexafluorosilicate(2-), lead(2+) (1:1), tetrahydrate |  | - | 83689-82-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead dichlorite |  | - | 13453-57-1 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead chloride (PbCl) |  | - | 13931-84-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Dilead dioxide |  | - | 35229-41-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, tributylchloro- |  | - | 13302-14-2 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, chlorotripropyl- |  | - | 1520-71-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Lead, diacetoxydiethyl- (7CI) |  | - | 15773-47-4 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, (4-bromobutyl)triphenyl- |  | - | 16035-34-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, (4-azidobutyl)triphenyl- |  | - | 16035-39-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
1H-Imidazole, 1-(tributylplumbyl)- |  | - | 16128-42-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, bis(acetyloxy)dibutyl- |  | - | 2587-84-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, tris(acetyloxy)phenyl- |  | - | 3076-54-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, triethyl[[(4-methylphenyl)sulfonyl]oxy]- |  | - | 43135-86-0 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Plumbane, tetrakis(1-methylpropyl)- |  | - | 65151-08-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Resin acids and Rosin acids, calcium salts, polymers with lead resinates and sapond. linseed oil |  | - | 68139-27-5 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
Grounding oil |  | - | 68152-99-8 | 63 | https://echa.europa.eu/documents/10162/64e0e958-99c2-e75e-4fa8-d2b71b18f0b4 |  | <a href="https://standards.cen.eu/dyn/www/f?p=204:110:0::::FSP_PROJECT,FSP_ORG_ID:61592,6229&cs=1C8C4C2CFD96C6D94449C448C73B68528">EN 16711-3:2019</a> | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2012.252.01.0004.01.ENG&toc=OJ:L:2012:252:TOC">R 836/2012</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2015.104.01.0002.01.ENG&toc=OJ:L:2015:104:TOC">R 628/2015</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2021.024.01.0019.01.ENG&toc=OJ%3AL%3A2021%3A024%3ATOC">R 2021/57</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R0923&qid=1629107278018">R 2023/923</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1190">entry 1190</a> | Guideline available, see Q&amp;A entry 1190 | 13-Jan-2016
1,4-Dichlorobenzene |  | 203-400-5 | 106-46-7 | 64 | https://echa.europa.eu/documents/10162/f4a325ee-70ff-4324-ab82-b9cf29c54018 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2014.136.01.0019.01.ENG&toc=OJ:L:2014:136:TOC">R 474/2014</a> |  |  | 13-Jan-2016
Inorganic ammonium salts |  | - | - | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Sulfuric acid, ammonium sodium salt, dihydrate |  | - | 7783-10-0 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Sulfuric acid, ammonium magnesium salt (2:2:1), hexahydrate |  | - | 7785-18-4 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium cerous sulfate tetrahydrate |  | - | 10049-02-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium phosphotungstenate, trihydrate |  | - | 12704-02-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium 12-tungstophosphate |  | - | 1311-90-6 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Cobalt ammonium complex |  | - | 14695-95-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Nitric acid, ammonium cerium(3+) salt (5:2:1) |  | - | 15318-60-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Tungstate(2-), dioxodithioxo-, diammonium (9CI) |  | - | 16150-61-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Triammonium uranyl pentafluoride |  | - | 12062-03-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Carbonic acid, ammoniumplutonium salt (8CI) |  | - | 24917-46-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Uranic acid, diammonium salt, hydrate |  | - | 28347-83-3 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Molybdate(2-),tetrafluorodioxo-, diammonium (8CI,9CI) |  | - | 30291-63-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium Tetrachloroaluminate, NH4AlCl4 |  | 885-776-6 | 7784-14-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium ferric chromate |  | - | 7789-08-4 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium tetracyanoplatinate |  | 209-237-6 | 562-79-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Tetraammonium disodium vanadate |  | 235-012-7 | 12055-09-3 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium hexachloroosmate |  | 235-188-5 | 12125-08-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium trivanadium octaoxide |  | 235-384-0 | 12207-63-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium tetratungsten tridecaoxide |  | 235-638-0 | 12398-61-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Tetraammonium hexamolybdate |  | 235-650-6 | 12411-64-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Tetraammonium cerium tetrakis(sulphate) |  | 231-567-4 | 7637-03-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium selenite |  | 231-983-6 | 7783-19-9 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium selenate |  | 231-985-7 | 7783-21-3 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium diuranium heptaoxide |  | 231-986-2 | 7783-22-4 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Triammonium hexafluoroaluminate |  | 232-052-7 | 7784-19-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Aluminium ammonium bis(sulphate) |  | 232-055-3 | 7784-25-0 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium hydrogenarsenate |  | 232-067-9 | 7784-44-3 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium chromate |  | 232-138-4 | 7788-98-9 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium dichromate |  | 232-143-1 | 7789-09-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium trioxovanadate |  | 232-261-3 | 7803-55-6 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium iron bis(sulphate) |  | 233-151-8 | 10045-89-3 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium iron phosphate |  | 233-255-3 | 10101-60-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium iron bis(sulphate) |  | 233-382-4 | 10138-04-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium wolframate |  | 234-364-9 | 11120-25-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium tetraborate |  | 234-513-8 | 12007-58-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Hexaammonium heptamolybdate |  | 234-722-4 | 12027-67-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium wolframate |  | 234-732-9 | 12028-06-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Tetramanganese nitride |  | 234-785-8 | 12033-07-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium hexachlororhenate |  | 234-991-8 | 12051-87-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium hexafluorosilicate |  | 240-968-3 | 16919-19-0 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium hexafluorozirconate |  | 240-970-4 | 16919-31-6 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium hexachloroplatinate |  | 240-973-0 | 16919-58-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium hexachlorostannate |  | 241-033-2 | 16960-53-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium hexafluorotitanate |  | 241-036-9 | 16962-40-6 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium hexafluorogermanate(4) |  | 241-037-4 | 16962-47-3 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Triammonium heptafluorozirconate(3-) |  | 241-283-2 | 17250-81-6 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium hexabromoplatinate |  | 241-394-6 | 17363-02-9 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Tetraammonium uranyl tricarbonate, of uranium depleted in uranium-235 |  | 241-988-5 | 18077-77-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium molybdate(VI) |  | 236-031-3 | 13106-76-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium tetraoxotellurate |  | 236-622-6 | 13453-06-0 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium dihydrogenarsenate |  | 236-667-1 | 13462-93-6 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium chromium bis(sulphate) |  | 236-923-2 | 13548-43-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium lanthanum pentanitrate |  | 236-968-8 | 13566-21-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium manganese bis(sulphate) |  | 236-969-3 | 13566-22-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium diamminetetrakis(thiocyanato-N)chromate(1-) |  | 237-003-3 | 13573-16-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium copper(2+) disulphate |  | 237-021-1 | 13587-25-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium cobalt bis(sulphate) |  | 237-043-1 | 13596-46-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium perrhenate |  | 237-075-6 | 13598-65-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium gadolinium pentanitrate |  | 237-111-0 | 13628-49-4 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium calcium trinitrate |  | 237-425-8 | 13780-11-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium zinc disulphate |  | 237-483-4 | 13814-87-4 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium yttrium pentanitrate |  | 239-605-1 | 15552-06-4 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium manganese phosphate |  | 239-688-4 | 15609-81-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium neodymium pentanitrate |  | 239-721-2 | 15653-40-4 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium tetrabromopalladate(2-) |  | 239-731-7 | 15661-00-4 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium nickel bis(sulphate) |  | 239-793-5 | 15699-18-0 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Triammonium hexachloroiridate |  | 239-842-0 | 15752-05-3 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Triammonium hexachlororhodate |  | 239-364-2 | 15336-18-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium hexachlororuthenate |  | 242-552-7 | 18746-63-9 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium hexachloropalladate |  | 242-854-9 | 19168-23-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium hexakis(thiocyanato)platinate |  | 243-002-9 | 19372-45-3 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium oxobis[sulphato(2-)-O]titanate(2-) |  | 243-090-9 | 19468-86-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium samarium(3+) disulphate |  | 244-699-2 | 21995-29-9 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium europium(3+) disulphate |  | 244-700-6 | 21995-30-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium gadolinium(3+) disulphate |  | 244-701-1 | 21995-31-3 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium lanthanum(3+) disulphate |  | 244-702-7 | 21995-32-4 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium praseodymium(3+) disulphate |  | 244-703-2 | 21995-33-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium neodymium(3+) disulphate |  | 244-704-8 | 21995-34-6 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium iron tetrachloride |  | 246-234-9 | 24411-12-9 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium hexabromoosmate(2-) |  | 246-340-5 | 24598-62-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium nickel trichloride |  | 246-378-2 | 24640-21-9 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Sulphuric acid, ammonium magnesium salt |  | 244-089-6 | 20861-69-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium aquapentachlororuthenate |  | 247-008-2 | 25461-53-4 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Triammonium diaquaoctachloro-μ-nitridodiruthenate(3-) |  | 248-399-2 | 27316-90-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium bis(cyano-C)aurate |  | 250-464-5 | 31096-40-9 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium tetrachloroaurate |  | 250-476-0 | 31113-23-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Triammonium tris[carbonato(2-)-O]hydroxyzirconate(3-) |  | 251-086-3 | 32535-84-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium pentafluorozirconate(1-) |  | 237-596-9 | 13859-62-6 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium sodium sulphate |  | 237-603-5 | 13863-45-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Triammonium hexacyanoferrate |  | 238-091-6 | 14221-48-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Chromic acid, ammonium salt |  | 238-422-4 | 14445-91-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Tetraammonium hexacyanoferrate |  | 238-476-9 | 14481-29-9 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium cobalt phosphate |  | 238-633-1 | 14590-13-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium cerium(4+) trisulphate |  | 238-682-9 | 14638-69-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium tetrachlorozincate(2-) |  | 238-687-6 | 14639-97-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Triammonium pentachlorozincate(3-) |  | 238-688-1 | 14639-98-6 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium magnesium bis(sulphate) |  | 238-782-2 | 14727-95-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Beryllium diammonium tetrafluoride |  | 238-948-4 | 14874-86-3 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium scandium(3+) disulphate |  | 239-145-1 | 15091-94-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Nitric acid, ammonium calcium salt |  | 239-289-5 | 15245-12-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium tetrachloropalladate |  | 237-498-6 | 13820-40-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium tetrachloroplatinate |  | 237-499-1 | 13820-41-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Diammonium pentachloronitrosylruthenate |  | 237-504-7 | 13820-58-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium tetrafluoroborate |  | 237-531-4 | 13826-83-0 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Sulphuric acid, ammonium cerium salt |  | 237-557-6 | 13840-04-5 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium trifluorohydroxyborate(1-) |  | 239-326-5 | 15283-48-4 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium Iron(II) Sulfate dodecahydrate |  | 616-517-5 | 7783-83-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Vanadate(3-), hexafluoro-, triammonium salt |  | 621-084-0 | 13815-31-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium tetrathiotungstate [(NH4)2WS4] |  | 628-727-4 | 13862-78-7 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Ammonium chromic sulfate dodecahydrate |  | 680-414-1 | 10022-47-6 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Nickel ammonium sulfate hexahydrate |  | 680-452-9 | 7785-20-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
azanium;lanthanum(3+);tetranitrate |  | - | 31178-09-3 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Aluminate(2-), difluoro[phosphato(3-)-κO]-, ammonium hydrogen (1:1:1) |  | - | 11095-65-1 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Cryptohalite |  | - | 1309-32-6 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Phosphoric acid, ammonium cadmium salt (1:1:1) |  | - | 14520-70-8 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Selenious acid, ammonium salt (2:1) |  | - | 25425-97-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
Sulfuric acid, ammonium magnesium salt (3:2:2) |  | - | 27733-50-2 | 65 | https://echa.europa.eu/documents/10162/9d7e5f15-403b-f28b-8165-6a497f5e593d |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.166.01.0001.01.ENG&toc=OJ:L:2016:166:TOC">R 2016/1017</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1803">entry 1803</a> |  | 04-Aug-2016
4,4'-isopropylidenediphenol |  | 201-245-8 | 80-05-7 | 66 | https://echa.europa.eu/documents/10162/1451b2a8-117e-e179-cfcd-c57214bb4adf |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2016.337.01.0003.01.ENG&toc=OJ:L:2016:337:TOC">R 2016/2235</a> |  |  | 11-Jan-2017
C9-C14 linear and/or branched perfluorocarboxylic acids (C9-C14 PFCAs), their salts and C9-C14 PFCAs-related substances, perfluorononan-1-oic acid (PFNA); nonadecafluorodecanoic acid (PFDA); henicosafluoroundecanoic acid (PFUnDA); tricosafluorododecanoic acid (PFDoDA); pentacosafluorotridecanoic acid (PFTrDA); heptacosafluorotetradecanoic acid (PFTDA); including their salts and precursors |  | - | - | 68 | https://echa.europa.eu/documents/10162/f9e7b269-87cd-fc26-1a8e-b8c8b6e40c08 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297">R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297R%2801%29&qid=1629107278018">Corrigendum to R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/LV/TXT/PDF/?uri=OJ:L:2022:089:FULL&from=EN">Corrigendum to R 2021/1297 (Latvian)</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1906">entry 1906</a> |  | 14-Aug-2017
Pentacosafluorotridecanoic acid |  | 276-745-2 | 72629-94-8 | 68 | https://echa.europa.eu/documents/10162/f9e7b269-87cd-fc26-1a8e-b8c8b6e40c08 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297">R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297R%2801%29&qid=1629107278018">Corrigendum to R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/LV/TXT/PDF/?uri=OJ:L:2022:089:FULL&from=EN">Corrigendum to R 2021/1297 (Latvian)</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1906">entry 1906</a> |  | 14-Aug-2017
Henicosafluoroundecanoic acid |  | 218-165-4 | 2058-94-8 | 68 | https://echa.europa.eu/documents/10162/f9e7b269-87cd-fc26-1a8e-b8c8b6e40c08 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297">R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297R%2801%29&qid=1629107278018">Corrigendum to R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/LV/TXT/PDF/?uri=OJ:L:2022:089:FULL&from=EN">Corrigendum to R 2021/1297 (Latvian)</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1906">entry 1906</a> |  | 14-Aug-2017
Heptacosafluorotetradecanoic acid |  | 206-803-4 | 376-06-7 | 68 | https://echa.europa.eu/documents/10162/f9e7b269-87cd-fc26-1a8e-b8c8b6e40c08 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297">R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297R%2801%29&qid=1629107278018">Corrigendum to R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/LV/TXT/PDF/?uri=OJ:L:2022:089:FULL&from=EN">Corrigendum to R 2021/1297 (Latvian)</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1906">entry 1906</a> |  | 14-Aug-2017
Nonadecafluorodecanoic acid |  | 206-400-3 | 335-76-2 | 68 | https://echa.europa.eu/documents/10162/f9e7b269-87cd-fc26-1a8e-b8c8b6e40c08 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297">R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297R%2801%29&qid=1629107278018">Corrigendum to R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/LV/TXT/PDF/?uri=OJ:L:2022:089:FULL&from=EN">Corrigendum to R 2021/1297 (Latvian)</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1906">entry 1906</a> |  | 14-Aug-2017
Perfluorononan-1-oic acid |  | 206-801-3 | 375-95-1 | 68 | https://echa.europa.eu/documents/10162/f9e7b269-87cd-fc26-1a8e-b8c8b6e40c08 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297">R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297R%2801%29&qid=1629107278018">Corrigendum to R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/LV/TXT/PDF/?uri=OJ:L:2022:089:FULL&from=EN">Corrigendum to R 2021/1297 (Latvian)</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1906">entry 1906</a> |  | 14-Aug-2017
Tricosafluorododecanoic acid |  | 206-203-2 | 307-55-1 | 68 | https://echa.europa.eu/documents/10162/f9e7b269-87cd-fc26-1a8e-b8c8b6e40c08 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297">R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R1297R%2801%29&qid=1629107278018">Corrigendum to R 2021/1297</a>#<a href="https://eur-lex.europa.eu/legal-content/LV/TXT/PDF/?uri=OJ:L:2022:089:FULL&from=EN">Corrigendum to R 2021/1297 (Latvian)</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1906">entry 1906</a> |  | 14-Aug-2017
Methanol |  | 200-659-6 | 67-56-1 | 69 | https://echa.europa.eu/documents/10162/3de90b7d-d9d7-5396-0c3e-4a7ec86577c7 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.099.01.0007.01.ENG&toc=OJ:L:2018:099:TOC">R 2018/589</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.102.01.0099.01.ENG&toc=OJ:L:2018:102:TOC">Corrigendum to R 2018/589</a> |  |  | 20-Apr-2018
Octamethylcyclotetrasiloxane (D4); Decamethylcyclopentasiloxane (D5); dodecamethylcyclohexasiloxane (D6) |  | - | - | 70 | https://echa.europa.eu/documents/10162/0ac1f453-ad41-4010-e837-a68273b896ca |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.006.01.0045.01.ENG&toc=OJ:L:2018:006:TOC">R 2018/35</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202401328&qid=1715919472757">R 2024/1328</a> |  |  | 22-Jan-2018
Dodecamethylcyclohexasiloxane |  | 208-762-8 | 540-97-6 | 70 | https://echa.europa.eu/documents/10162/0ac1f453-ad41-4010-e837-a68273b896ca |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.006.01.0045.01.ENG&toc=OJ:L:2018:006:TOC">R 2018/35</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202401328&qid=1715919472757">R 2024/1328</a> |  |  | 22-Jan-2018
Octamethylcyclotetrasiloxane |  | 209-136-7 | 556-67-2 | 70 | https://echa.europa.eu/documents/10162/0ac1f453-ad41-4010-e837-a68273b896ca |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.006.01.0045.01.ENG&toc=OJ:L:2018:006:TOC">R 2018/35</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202401328&qid=1715919472757">R 2024/1328</a> |  |  | 22-Jan-2018
Decamethylcyclopentasiloxane |  | 208-764-9 | 541-02-6 | 70 | https://echa.europa.eu/documents/10162/0ac1f453-ad41-4010-e837-a68273b896ca |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.006.01.0045.01.ENG&toc=OJ:L:2018:006:TOC">R 2018/35</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202401328&qid=1715919472757">R 2024/1328</a> |  |  | 22-Jan-2018
1-methyl-2-pyrrolidone |  | 212-828-1 | 872-50-4 | 71 | https://echa.europa.eu/documents/10162/5161f2bf-503c-2107-ab54-c087017d11c9 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.099.01.0003.01.ENG&toc=OJ:L:2018:099:TOC">R 2018/588</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1603">entry 1603</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1828">entry 1828</a> | Restriction guideline: How to comply with REACH Restriction 71, guideline for users of NMP (1-methyl-2-pyrrolidone). The document is published at https://echa.europa.eu/documents/10162/17233/entry_71_how_to_comply_en.pdf/c6e09198-c0b1-44e3-abae-6b3d0bc909a8 Explanatory note: How to comply with REACH Restriction 71, guideline for users of NMP (1-methyl-2-pyrrolidone), Explanatory Note on Biological Monitoring. The document is published at https://echa.europa.eu/documents/10162/17233/entry_71_exp_note_biomonitoring_en.pdf/25c684f2-243e-4321-e7c0-ba7074838ab6 | 20-Apr-2018
The following substances which are classified as carcinogenic, mutagenic or toxic for reproduction, category 1A or 1B (See group members) [Entry 72] |  | - | - | 72 | https://echa.europa.eu/documents/10162/8db10905-d535-0a04-0af5-7628a210dc28 | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=582">Appendix 12</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2018.256.01.0001.01.ENG&toc=OJ:L:2018:256:TOC">R 2018/1513</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1805">entry 1805</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1806">entry 1806</a> |  | 13-Dec-2018
(3,3,4,4,5,5,6,6,7,7,8,8,8-tridecafluorooctyl)silanetriol and any of its mono-, di- or tri-O-(alkyl) derivatives |  | - | - | 73 | https://echa.europa.eu/documents/10162/2906fed0-12ab-9593-23a1-72efd611fe70 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2019.154.01.0037.01.ENG&toc=OJ:L:2019:154:TOC">R 2019/957</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1907">entry 1907</a> |  | 24-Jun-2019
Trimethoxy(3,3,4,4,5,5,6,6,7,7,8,8,8-tridecafluorooctyl)silane |  | 288-657-1 | 85857-16-5 | 73 | https://echa.europa.eu/documents/10162/2906fed0-12ab-9593-23a1-72efd611fe70 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2019.154.01.0037.01.ENG&toc=OJ:L:2019:154:TOC">R 2019/957</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1907">entry 1907</a> |  | 24-Jun-2019
Triethoxy(3,3,4,4,5,5,6,6,7,7,8,8,8-tridecafluorooctyl)silane |  | 257-473-3 | 51851-37-7 | 73 | https://echa.europa.eu/documents/10162/2906fed0-12ab-9593-23a1-72efd611fe70 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv:OJ.L_.2019.154.01.0037.01.ENG&toc=OJ:L:2019:154:TOC">R 2019/957</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1907">entry 1907</a> |  | 24-Jun-2019
Diisocyanates |  | - | - | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
4-methyl-m-phenylene diisocyanate |  | 209-544-5 | 584-84-9 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
Hexamethylene diisocyanate |  | 212-485-8 | 822-06-0 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
1,3-bis(isocyanatomethyl)benzene |  | 222-852-4 | 3634-83-1 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
2,4’-Methylenediphenyl diisocyanate |  | 227-534-9 | 5873-54-1 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
3-isocyanatomethyl-3,5,5-trimethylcyclohexyl isocyanate |  | 223-861-6 | 4098-71-9 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
4,4'-methylenedicyclohexyl diisocyanate |  | 225-863-2 | 5124-30-1 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
m-tolylidene diisocyanate |  | 247-722-4 | 26471-62-5 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
2,2’-Methylenediphenyl diisocyanate |  | 219-799-4 | 2536-05-2 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
1,3-bis(1-isocyanato-1-methylethyl)benzene |  | 220-474-4 | 2778-42-9 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
1,5-naphthylene diisocyanate |  | 221-641-4 | 3173-72-6 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
2-methyl-m-phenylene diisocyanate |  | 202-039-0 | 91-08-7 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
3,3'-dimethylbiphenyl-4,4'-diyl diisocyanate |  | 202-112-7 | 91-97-4 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
4,4’-Methylenediphenyl diisocyanate |  | 202-966-0 | 101-68-8 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
2,4,6-triisopropyl-m-phenylene diisocyanate |  | 218-485-4 | 2162-73-4 | 74 | https://echa.europa.eu/documents/10162/503ac424-3bcb-137b-9247-09e41eb6dd5a |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/ALL/?uri=uriserv:OJ.L_.2020.252.01.0024.01.ENG">R 2020/1149</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1908">entry 1908</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1916">entry 1916</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1917">entry 1917</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1918">entry 1918</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1919">entry 1919</a> |  | 10-Aug-2020
Substances in tattoo inks and permanent make up |  | - | - | 75 | https://echa.europa.eu/documents/10162/0fa98a4c-ff76-6d0b-d48a-8b94ccac9bae | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=585">Appendix 13</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=uriserv%3AOJ.L_.2020.423.01.0006.01.ENG&toc=OJ%3AL%3A2020%3A423%3ATOC">R 2020/2081</a> | <a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1909">entry 1909</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1910">entry 1910</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1912">entry 1912</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1913">entry 1913</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1914">entry 1914</a>#<a href="https://echa.europa.eu/support/qas-support/browse/-/qa/70Qx/view/ids/1915">entry 1915</a> |  | 13-Jan-2021
N,N-dimethylformamide |  | 200-679-5 | 68-12-2 | 76 | https://echa.europa.eu/documents/10162/3dfadfba-a9bd-9b66-4868-153dd9034b7b |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32021R2030&qid=1629107278018">R 2021/2030</a> |  |  | 30-Nov-2021
formaldehyde and formaldehyde releasers |  | - | - | 77 | https://echa.europa.eu/documents/10162/1046d94e-f527-0bbb-7435-ad9eb040bea7 | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=589">Appendix 14</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R1464&qid=1629107278018">R 2023/1464</a> |  |  | 27-Jul-2023
Formaldehyde |  | 200-001-8 | 50-00-0 | 77 | https://echa.europa.eu/documents/10162/1046d94e-f527-0bbb-7435-ad9eb040bea7 | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=589">Appendix 14</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R1464&qid=1629107278018">R 2023/1464</a> |  |  | 27-Jul-2023
synthetic polymer microparticles |  | - | - | 78 | https://echa.europa.eu/documents/10162/5469015c-525b-ec20-3bae-cbbc2bac7fb6 | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=591">Appendix 15</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:02006R1907-20250623&from=EN:#page=597">Appendix 16</a> |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AJOL_2023_238_R_0003&qid=1629107278018">R 2023/2055</a>#<a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32023R2055R%2801%29">Corrigendum to R 2023/2055</a> |  |  | 04-Oct-2023
undecafluorohexanoic acid (PFHxA), its salts and related substances |  | - | - | 79 | https://echa.europa.eu/documents/10162/f1451352-6834-634d-79a9-b43880dae258 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202402462">R 2024/2462</a> |  |  | 30-Sep-2024
undecafluorohexanoic acid and its ammonium salt |  | - | - | 79 | https://echa.europa.eu/documents/10162/f1451352-6834-634d-79a9-b43880dae258 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202402462">R 2024/2462</a> |  |  | 30-Sep-2024
Sodium undecafluorohexanoate |  | 220-881-7 | 2923-26-4 | 79 | https://echa.europa.eu/documents/10162/f1451352-6834-634d-79a9-b43880dae258 |  |  | <a href="https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ%3AL_202402462">R 2024/2462</a> |  |  | 30-Sep-2024
N,N-dimethylacetamide (DMAC) |  | 204-826-4 | 127-19-5 | 80 | https://echa.europa.eu/documents/10162/78d2f2df-f3f7-77f5-3b34-a82f6663a96a |  |  | <a href="https://eur-lex.europa.eu/eli/reg/2025/1090/oj">R 2025/1090 (Annex VII entries 80 and 81))</a> |  |  | 16-Jun-2025
1-ethylpyrrolidin-2-one (NEP) |  | 220-250-6 | 2687-91-4 | 81 | https://echa.europa.eu/documents/10162/68f4621e-0d2e-6e1a-74bf-672a0bc361f2 |  |  | <a href="https://eur-lex.europa.eu/eli/reg/2025/1090/oj">R 2025/1090 (Annex VII entries 80 and 81)</a> |  |  | 16-Jun-2025
Per- and polyfluoroalkyl substances (PFAS) |  | - | - | 82 | https://echa.europa.eu/documents/10162/e8c76d1e-5204-4e99-ea16-3dffb8f1ca8b |  |  | <a href="https://eur-lex.europa.eu/eli/reg/2025/1988/oj">R 2025/1988</a> |  |  | 17-Oct-2025

## Aromatic Amines (Ent.43)
REACH Annex XVII — Entry 43 — Restricted Aromatic Amines (azo dye cleavage)
Source: REACH Regulation (EC) No 1907/2006, Annex XVII, Entry 43 ¦ Reference standard: EN ISO 14362-1:2017
Version note: Verify against current ECHA Annex XVII table at echa.europa.eu/substances-restricted-under-reach
# | Substance name | CAS number | EC number | IARC Group | Notes
1 | 4-Aminodiphenyl | 92-67-1 | 202-177-1 | Group 1 | Proven human carcinogen
2 | Benzidine | 92-87-5 | 202-199-1 | Group 1 | Proven human carcinogen
3 | 4-Chloro-o-toluidine | 95-69-2 | 202-441-6 | Group 2A | Probably carcinogenic
4 | 2-Naphthylamine | 91-59-8 | 202-080-4 | Group 1 | Proven human carcinogen
5 | o-Aminoazotoluene | 97-56-3 | 202-591-2 | Group 2B | Possibly carcinogenic
6 | 5-Nitro-o-toluidine | 99-55-8 | 202-765-8 | Group 2B | Possibly carcinogenic
7 | 4-Chloroaniline | 106-47-8 | 203-401-0 | Group 2A | Probably carcinogenic
8 | 2,4-Diaminoanisole | 615-05-4 | 210-406-1 | Group 2B | Possibly carcinogenic
9 | 4,4'-Methylenedianiline (MDA) | 101-77-9 | 202-974-4 | Group 2B | Possibly carcinogenic
10 | 3,3'-Dichlorobenzidine | 91-94-1 | 202-109-0 | Group 2A | Probably carcinogenic
11 | 3,3'-Dimethoxybenzidine | 119-90-4 | 204-355-4 | Group 2B | Possibly carcinogenic
12 | 3,3'-Dimethylbenzidine | 119-93-7 | 204-358-0 | Group 2B | Possibly carcinogenic
13 | 4,4'-Methylenedi-o-toluidine | 838-88-0 | 212-658-8 | Group 2B | Possibly carcinogenic
14 | 4,4'-Oxydianiline | 101-80-4 | 202-977-0 | Group 2B | Possibly carcinogenic
15 | 4,4'-Thiodianiline | 139-65-1 | 205-370-9 | Group 2B | Possibly carcinogenic
16 | o-Toluidine | 95-53-4 | 202-429-0 | Group 1 | Proven human carcinogen (IARC Vol. 115, 2018)
17 | 2,4-Diaminotoluene | 95-80-7 | 202-453-1 | Group 2B | Possibly carcinogenic
18 | 2,4,5-Trimethylaniline | 137-17-7 | 205-282-0 | Group 3 | Not classifiable
19 | o-Anisidine | 90-04-0 | 201-963-1 | Group 2B | Possibly carcinogenic
20 | 4,4'-Methylenebis(2-chloroaniline) (MOCA) | 101-14-4 | 202-918-9 | Group 1 | Proven human carcinogen
21 | p-Cresidine | 120-71-8 | 204-419-1 | Group 2B | Possibly carcinogenic
22 | 4-Aminoazobenzene | 60-09-3 | 200-453-6 | Group 2B | Possibly carcinogenic

## Azo - Restricted Amines (JTF)
Annex 3 List of dyes that generate specific aromatic amines (Informative)
Note: The following list is made with referring to publicly available information on detection of aromatic amines. However, it is considered possible that other dyes than those mentioned here generate specific aromatic amines. The name of dyes (color index generic name) means a constituent of dyes available in the market and does not indicate a product name. (It is not recommended to use this material as a restricted substance list.) ・ More than one number is given in “CAS no.” column, as the numbers used are different depending on the documents. ・ The “generated primary aromatic amines” column is added in this list (updated on October 1, 2015) to indicate the information on the generated primary aromatic amines that was obtained with mainly referring to color index. ・ The “verification” column indicates the status of verification whether any aromatic amine is generated or not with referring to structural formulas in color index. “Verified” means completion of verification and “not verified” means no verification is made, because of no mentioning of structural formulas and others. ・ “Added” in the “remarks” column means a dye added this time in the list that has been publicly available. ・ “Not applicable” in the “remarks” column means a dye that has the information on detection of aromatic amine in the name of such dye but does not generate any aromatic amine with applying structural formula. The possible reasons include the case that such dye is mixed with other dyes but sold in the same dye name or that any impure substance (unreacted starting materials, intermediate reactants, etc.) is contained.
October 1, 2015
Dye name (color index generic name) | C.I. constitution no. | CAS no. | Generated primary aromatic amines | 2012 version published by JTF | Eco label | SCCNFP | Verification | Remarks | Colonne1
Acid Black 29 |  | 12217-14-0 |  | 1 | 1 | Table 1 | Not verified
Acid Black 94 | 30336 | 6358-80-1 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Acid Black 131 |  | 12219-01-1 |  | 1 | 1 | Table 3 | Not verified
Acid Black 132 |  | 12219-02-2 |  | 1 | 1 | Table 3 | Not verified
Acid Black 209 |  | 72827-68-0 |  | 1 | 1 | Table 1 | Not verified
Acid Black 232 | 30334 |  | Benzidine, IARC Group 1 |  | 1 | Table 1 | Verified | Added
Acid Brown 415 |  | 97199-27-4 |  | 1 | 1 | Table 3 | Not verified
Acid Orange 17 | 16020 | 52749-23-2 | 2,4-Xylidine, IARC Group 3 or 2,6-Xylidine, IARC Group 2 |  | 1 | Table 5 | Verified | Added
Acid Orange 24 | 20170 | 1320-07-6 | 2,4-Xylidine, IARC Group 3 or 2,6-Xylidine, IARC Group 2 | 1 | 1 | Table 5 | Verified
Acid Orange 45 | 22195 | 2429-80-3 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Acid Red 4 | 14710 | 5858-39-9 | o-Anisidine, IARC Group 2B | 1 | 1 | Table 3 | Verified
Acid Red 5 | 14905 | 5858-63-9 | o-Anisidine, IARC Group 2B | 1 | 1 | Table 3 | Verified
Acid Red 8 | 14900 | 4787-93-3 | 2,4-Xylidine, IARC Group 3 |  | 1 | Table 5 | Verified | Added
Acid Red 24 | 16140 | 5858-30-0 | o-Toluidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Acid Red 26 | 16150 | 3761-53-3 | 2,4-Xylidine, IARC Group 3 | 1 | 1 | Table 5 | Verified
Acid Red 26:1 | 16151 | 8004-46-4 | 2,4-Xylidine, IARC Group 3 and o-Anisidine, IARC Group |  | 1 | Table 5 | Verified | Added
Acid Red 26:2 | 16152 | 8004-47-5 | 2,4-Xylidine, IARC Group 3 or 2,6-Xylidine, IARC Group 2 |  | 1 | Table 5 | Verified | Added
Acid Red 35 | 18065 | 6441-93-6 | o-Toluidine, IARC Group 1 |  | 1 | Table 1 | Verified | Added
Acid Red 73 | 27290 | 5413-75-2 | 4-Aminoazobenzene, IARC Group 2B | 1 | 1 | Table 3 | Verified
Acid Red 85 | 22245 | 3567-65-5 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Acid Red 104 | 26420 | 2230325 | o-Aminoazotoluene, IARC Group 2B |  | 1 | Table 1 | Verified
Acid Red 114 | 23635 | 6459-94-5 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Acid Red 115 | 27200 | 6226-80-8 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Acid Red 116 | 26660 | 6245-62-1 | o-Aminoazotoluene, IARC Group 2B | 1 | 1 | Table 3 | Verified
Acid Red 119:1 |  | 90880-75-4 |  |  | 1 | Table 1 | Not verified | Added
Acid Red 128 | 24125 | 6548-30-7 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Acid Red 135 | 14695 | 5858-37-7 | 2,4-Xylidine, IARC Group 3 |  | 1 | Table 5 | Verified | Added
Acid Red 148 | 26665 | 6300-53-4 | o-Aminoazotoluene, IARC Group 2B | 1 | 1 | Table 1 | Verified
Acid Red 150 | 27190 | 6226-78-4 | o-Aminoazotoluene, IARC Group 2B | 1 | 1 | Table 3 | Verified
Acid Red 158 | 20530 | 8004-55-5, 12239-89-3 |  | 1 | 1 | Table 1 | Verified | Not applicable
Acid Red 167 |  | 619901-41-5 |  | 1 | 1 | Table 1 | Not verified
Acid Red 170 | 27210 | 6226-81-9 | 2,4-Xylidine, IARC Group 3 |  | 1 | Table 5 | Verified | Added
Acid Red 264 | 18133 | 6505-96-0 | o-Anisidine, IARC Group 2B | 1 | 1 | Table 3 | Verified
Acid Red 265 | 18129 | 6358-43-6 | o-Toluidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Acid Red 323 | 22238 | 6358-34-5 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Acid Red 420 |  |  |  | 1 | 1 | Table 3 | Not verified
Acid Violet 12 | 18075 | 6625-46-3 | o-Anisidine, IARC Group 2B | 1 | 1 | Table 3 | Verified
Basic Brown 2 | 21030 | 6358-83-4 | No.8 2,4-Diaminoanisole, IARC Group 2B | 1 |  | Table 2 | Verified
Basic Brown 4 | 21010 | 5421-66-9, 8005-78-5 | Toluylendiamine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Basic Orange 1 | 11320 | 4438-16-8 |  | 1 |  |  | Verified | Not Synthetic applicable toluylendiamine, IARC Group 2B
Basic Red 42 |  | 12221-66-8 |  | 1 | 1 | Table 3 | Not verified
Basic Red 76 | 12245 | 68391-30-0 | o-Anisidine, IARC Group 2B |  | 1 | Table 3 | Verified | Added
Basic Red 111 | 284240 | 113741-92-7, 118658-98-3 | 4-Aminoazobenzene, IARC Group 2B | 1 | 1 | Table 3 | Verified
Basic Red 114 |  |  |  |  | 1 | Table 3 | Not verified | Added
Basic Yellow 82 |  | 71872-38-3 |  |  | 1 | Table 3 | Not verified | Added
Basic Yellow 103 |  |  |  |  | 1 | Table 1 | Not verified | Added
Direct Black 4 | 30245 | 2429-83-6, 25156-49-4 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Black 9 | 31565 |  |  | 1 |  |  |  | Not Synthetic applicable toluylendiamine, IARC Group 2B
Direct Black 15 | 22620 | 6426-75-1 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Black 29 | 22580 | 3626-23-1 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Black 38 | 30235 | 1937-37-7 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Black 91 | 30400 | 6739-62-4 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 |  |  | Verified
Direct Black 114 |  | 61703-05-7 |  | 1 |  |  | Not verified
Direct Black 154 | 303865 | 37372-50-2, 54804-85-2 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 1 | 24410 | 2610-05-1, 3814-14-3 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 2 | 22590 | 2429-73-4 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Blue 3 | 23705 | 2429-72-3 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 6 | 22610 | 2602-46, 2602-46-2 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Blue 8 | 24140 | 2429-71-2 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 9 | 24155 | 6428-98-4 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 10 | 24340 | 4198-19-0 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 14 | 23850 | 72-57-1 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 15 | 24400 | 2429-74-5 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 16 | 22475 | 6426-66-0 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Blue 19 | 22485 | 6426-68-2 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Blue 21 | 23710 | 1651144 | 3,3’-Dimethylbenzidine, IARC Group 2B |  | 1 | Table 1 | Verified | Added
Direct Blue 22 | 34280 | 2586-57-4 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 25 | 23790 | 2150-54-1, 25180-27-2 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 26 | 31930 | 7082-31-7 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 |  | Table 2 | Verified
Direct Blue 35 | 24145 | 6473-33-2 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 48 | 22565 | 6459-89-8 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Blue 49 | 22540 | 6426-73-9 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Blue 53 | 23860 | 314-13-6 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 |  | Table 2 | Verified
Direct Blue 58 | 22490 | 6426-69-3 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Blue 64 | 22595 | 6426-74-0 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Blue 76 | 24411 | 16143-79-6 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 |  | Verified
Direct Blue 80 | 24315 | 12222-00-3 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 |  |  | Verified
Direct Blue 90 | 231560 | 12217-56-0 |  | 1 |  |  | Verified | Not applicable
Direct Blue 98 | 23155 | 1737161 |  | 1 |  |  | Verified | Not applicable
Direct Blue 100 |  | 61814-75-3 |  | 1 |  |  | Not verified
Direct Blue 116 | 27980 | 6227-23-2 |  |  | 1 | Table 5 | Verified | Added Not applicable
Direct Blue 151 | 24175 | 6449-35-0, 110735-25-6 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 156 |  | 61724-73-0 |  | 1 |  |  | Not verified
Direct Blue 160 |  | 12222-02-5 |  | 1 | 1 | Table 1 | Not verified
Direct Blue 173 |  | 12235-72-2 |  | 1 | 1 | Table 1 | Not verified
Direct Blue 177 | 22625 | 6426-76-2 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Blue 191 |  | 61373-80-6
Direct Blue 192 |  | 159202-76-3 |  | 1 | 1 | Table 1 | Not verified
Direct Blue 201 | 51300 | 60800-55-7 |  | 1 | 1 | Table 1 |  | Not applicable
Direct Blue 215 | 24415 | 6771-80-5, 6771-80-8 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 218 | 24401 | 28407-37-6 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 |  |  | Verified
Direct Blue 224 |  | 71838-46-5 |  | 1 |  |  | Not verified
Direct Blue 230 | 22455 | 6527-65-7 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Blue 295 | 23820 | 6420-22-0 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Blue 306 | 24203 |  | 3,3’-Dimethoxybenzidine, IARC Group 2B |  | 1 | Table 1 | Verified | Added
Direct Brown 1 | 30045 | 3811-71-0 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 1:2 | 30110 | 2586-58-5 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 2 | 22311 | 2429-82-5, 25255-06-5 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 6 | 30140 | 2893-80-3, 25180-39-6 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 25 | 36030 | 33363-87-0 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 27 | 31725 | 6360-29-8, 6360-59-8 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 31 | 35660 | 2429-81-4, 25180-41-0 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 33 | 35520 | 1324-87-4 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 51 | 31710 | 4623-91-0 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 56 | 22040 | 6486-31-3 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Brown 58 | 22340 | 6426-59-1 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Brown 59 | 22345 | 6247-51-4 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 60 | 22325 | 6426-57-9 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Brown 74 | 36300 | 8014-91-3 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 79 | 30050 | 6483-77-8 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 86 | 22030 | 6486-30-2 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Brown 95 | 30145 | 16071-86-6 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 101 | 31740 | 3626-29-7 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 111 |  | 12222-20-7 |  | 1 |  |  | Not verified
Direct Brown 154 | 30120 | 6360-54-9 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Brown 165 | 22045 | 6486-32-4 | Benzidine, IARC Group 1 | 1 |  |  | Verified
Direct Brown 200 |  | 12222-26-3 |  | 1 |  |  | Not verified
Direct Brown 222 | 30368 | 64743-15-3 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Brown 223 |  | 76930-14-8 |  |  | 1 | Table 1 | Not verified | Added Not applicable
Direct Brown 230 |  |  |  | 1 |  |  | Not verified
Direct Green 1 | 30280 | 3626-28-6 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Green 6 | 30295 | 889614 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Green 8 | 30315 | 5422-17-3, 25180-47-6 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Green 8:1 |  | 76012-70-9 |  | 1 |  | Table 1 | Not verified
Direct Green 21 |  | 8003-52-9 | Benzidine, IARC Group 1 |  |  | Table 2 | Verified
Direct Green 21:1 | [5 empty] | Table 2 | Not verified | Added
Direct Green 60 | 22315 | 6426-56-8 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Green 85 | 30387 | 72390-60-4 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Orange 1 | 22430 | 6459-87-6 | Benzidine, IARC Group 1 |  |  | Table 2 | Verified
Direct Orange 2 | 22380 | 8005-97-8 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Orange 6 | 23375 | 6637-88-3 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Orange 7 | 23380 | 2868-76-0 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Orange 8 | 22130 | 2429-79-0, 64083-59-6 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Orange 10 | 23370 | 6405-94-3 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Orange 25 | 22135 | 6486-43-7 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Orange 33 | 22385 |  | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Orange 72 | 29058 | 12217-64-0 |  | 1 |  |  | Verified | Not applicable
Direct Orange 101 | 22190 | 6528-39-8 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Orange 108 | 29173 | 6358-79-8 | o-Toluidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Red 1 | 22310 | 2429-84-7, 25188-24-3 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Red 2 | 23500 | 992-59-6 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Red 7 | 24100 | 2868-75-9 | 3,3’-Dimethoxybenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Red 10 | 22145 | 2429-70-1, 25188-29-8 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Red 13 | 22155 | 1937-35-5, 25188-30-1 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Red 17 | 22150 | 317583 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Red 18 | 22280 | 6548-26-1 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Red 21 | 23560 | 1645-78-9, 6406-01-5 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Red 22 | 23565 | 6448-80-2 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Red 24 | 29185 | 6420-44-6 | o-Anisidine, IARC Group 2B | 1 | 1 | Table 3 | Verified
Direct Red 26 | 29190 | 3687-80-7 | o-Anisidine, IARC Group 2B | 1 | 1 | Table 3 | Verified
Direct Red 28 | 22120 | 573-58-0 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Red 29 | 22305 | 6426-54-6 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Red 33 | 22306 | 6253-15-2 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Red 37 | 22240 | 3530-19-6 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Red 39 | 23630 | 6358-29-8 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Red 42 | 22180 | 6548-39-6 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Red 43 | 22205 | 6486-50-6 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Red 44 | 22500 | 6548-29-4 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Red 46 | 23050 | 2302-97-8 | No.10 3,3’-Dichlorobenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Red 52 | 22290 | 6797-93-9 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Red 59 | 22420 | 6655-94-3 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Red 60 | 22200 | 6486-49-3 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Red 62 | 29175 | 6420-43-5 | o-Toluidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Red 67 | 23505 | 6598-56-7 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Red 72 | 29200 | 8005-64-9 | o-Toluidine, IARC Group 1 and o-Anisidine, IARC Group | 1 | 1 | Table 3 | Verified
Direct Red 74 | 22170 | 8003-75-6 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Red 88 | 22360 | 6459-86-5 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Red 126 | 17785 | 6369-36-4 | 2,4-Xylidine, IARC Group 3 |  | 1 | Table 5 | Verified | Added
Direct Red 168 | 19575 | 6404-53-1 | 2,4-Xylidine, IARC Group 3 |  | 1 | Table 5 | Verified | Added
Direct Red 216 | 17815 | 8004-49-7 | 2,4-Xylidine, IARC Group 3 or 2,6-Xylidine, IARC Group 2 |  | 1 | Table 5 | Verified | Added
Direct Red 264 | 29187 |  | 2,4-Xylidine, IARC Group 3 |  | 1 | Table 5 | Verified | Added
Direct Violet 1 | 22570 | 2586-60-9, 25188-44-7 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Violet 3 | 22445 | 6507-83-1 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Violet 4 | 22555 | 6472-95-3 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Violet 9 | 27885 | 6227-14-1 |  | 1 |  |  | Verified | Not applicable
Direct Violet 12 | 22550 | 2429-75-6 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Violet 13 | 24080 | 13478-92-7 | 3,3’-Dimethoxybenzidine, IARC Group 2B |  | 1 | Table 1 | Verified | Added
Direct Violet 14 | 29105 | 6420-38-8 | 2,4-Xylidine, IARC Group 3 |  | 1 | Table 5 | Verified | Added
Direct Violet 17 | 22465 | 6426-65-9 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Violet 21 | 23520 | 6470-45-7 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Violet 22 | 22480 | 6426-67-1, 25329-82-2 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Violet 27 | 22460 | 6426-64-8 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Violet 36 | 22470 | 6472-94-2 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Violet 38 | 22630 | 6426-77-3 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Violet 42 | 22450 | 6459-88-7 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Violet 43 | 22440 | 6426-63-7 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Violet 45 | 22510 | 6426-72-8 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Violet 85 | 22520 | 6507-84-2 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Violet 88 | 22046 | 6358-33-4 | Benzidine, IARC Group 1 | 1 |  |  | Verified
Direct Yellow 1 | 22250 | 6472-91-9 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Yellow 20 | 22410 | 6426-62-6 | Benzidine, IARC Group 1 | 1 |  | Table 2 | Verified
Direct Yellow 24 | 22010 | 6486-29-9 | Benzidine, IARC Group 1 | 1 | 1 | Table 1 | Verified
Direct Yellow 48 | 23660 | 6459-97-8 | 3,3’-Dimethylbenzidine, IARC Group 2B | 1 | 1 | Table 1 | Verified
Direct Yellow 68 |  | 61725-07-3 |  | 1 |  |  | Not verified
Direct Yellow 95 |  | 61703-10-4 |  | 1 |  |  | Not verified
Disperse Orange 60 |  | 12270-44-9 |  |  | 1 | Table 1 | Not verified | Added
Disperse Orange 149 |  | 85136-74-4, 85136-74-9, 151126-94-2 |  | 1 | 1 | Table 3 | Not verified
Disperse Red 151 | 26130 | 61968-47-6 | 4-Aminoazobenzene, IARC Group 2B | 1 | 1 | Table 3 | Verified
Disperse Red 221 |  | 64426-35-3 |  |  | 1 | Table 1 | Verified | Not applicable
Disperse Yellow 7 | 26090 | 6300-37-4 | 4-Aminoazobenzene, IARC Group 2B | 1 | 1 | Table 3 | Verified
Disperse Yellow 23 | 26070 | 6250-22-3 | 4-Aminoazobenzene, IARC Group 2B | 1 | 1 | Table 3 | Verified
Disperse Yellow 56 | 261550 | 54077-16-6 | 4-Aminoazobenzene, IARC Group 2B | 1 | 1 | Table 3 | Verified
Disperse Yellow 218 |  | 83929-90-2 |  |  | 1 | Table 1 | Not verified | Added
Solvent Orange 7 | 12140 | 3118-98-2, 3118-98-6 | 2,4-Xylidine, IARC Group 3 | 1 |  | Table 5
Solvent Red 19 | 26050 | 6368-72-5 | 4-Aminoazobenzene, IARC Group 2B | 1 |  | Table 3 | Verified
Solvent Red 23 | 26100 | 85-86-9 | 4-Aminoazobenzene, IARC Group 2B | 1 |  | Table 3 | Verified
Mordant Yellow 16 | 25100 | 8003-87-0 | 4,4’-Thioianiline, IARC Group 2B |  |  | Table 1 | Verified | Added
Mordant Red 57 |  | 12262-42-9 |  |  |  | Table 1 | Not verified | Added
Referable information (1) Names of dyes are found in the 2014 version of the EU Eco Label Textile Restricted Substance List. [As of October 1, 2015] http://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:32014D0350&from=EN (2) A wide range of names of azo colorants and their generated primary aromatic amines can be found in "The safety review of the use of certain azo-dyes in cosmetic products by the Scientific Committee on Cosmetic Products and Non-food Products Intended for Consumers, adopted by the SCCNFP during the 19th plenary meeting of 27 February, 2002". http://ec.europa.eu/health/ph_risk/committees/sccp/documents/out155_en.pdf (3) Information on azo pigments is publicly disclosed by ETAD. http://www.etad.com/lang-en/publications/information-notices.html?start=5
(4) (Informative) Names of dyes that generate primary aromatic amines are found in the 2014 controlled substances list of Levi Strauss & Co. http://levistrauss.com/wp-content/uploads/2014/09/July-2014-RSL-English.pdf (5) (Informative) Names of dyes that generate primary aromatic amines are found in the 2014 Azo Dyes and Pigments: Banned Arylamines of Marks and Spencer http://levistrauss.com/wp-content/uploads/2014/09/July-2014-RSL-English.pdf

## IARC
IARC Monographs
Monographs.iarc.who.int/list-of-classifications/
Last table update: 2026-03-27 10:08 (CET)
CAS No. | Agent | Group | Volume | Volume publication year | Evaluation year | Additional information
| Aloe vera, whole leaf extract | 2B | 108 | 2016 | 2013
| Clonorchis sinensis (infection with) | 1 | 61, 100B | 2012 | 2009
| Fusarium graminearum, F. culmorum, and F. crookwellense, toxins derived from (zearalenone, deoxynivalenol, nivalenol, and fusarenone X) | 3 | Sup 7, 56 | 1993 | 1992
| Fusarium sporotrichioides, toxins derived from (T-2 toxin) | 3 | 56 | 1993 | 1992
| Helicobacter pylori (infection with) | 1 | 61, 100B | 2012 | 2009
| Microcystis extracts | 3 | 94 | 2010 | 2006
| Opisthorchis felineus (infection with) | 3 | 61 | 1994 | 1994
| Opisthorchis viverrini (infection with) | 1 | 61, 100B | 2012 | 2009
| Schistosoma haematobium (infection with) | 1 | 61, 100B | 2012 | 2009
| Schistosoma japonicum (infection with) | 2B | 61 | 1994 | 1994
| Schistosoma mansoni (infection with) | 3 | 61 | 1994 | 1994
| Acheson process, occupational exposure associated with | 1 | 111 | 2017 | 2014
| Acid mists, strong inorganic | 1 | 54, 100F | 2012 | 2009
| Acrylic fibres | 3 | 19, Sup 7 | 1987 | 1987
| Acrylonitrile-butadiene-styrene copolymers | 3 | 19, Sup 7 | 1987 | 1987
| Alcoholic beverages | 1 | 44, 96, 100E | 2012 | 2009 | NB There is "evidence suggesting lack of carcinogenicity" in humans of alcoholic beverages for non-Hodgkin lymphoma and kidney cancer.
Alpha particles (see Radionuclides)
| Aluminium production | 1 | 34, Sup 7, 92, 100F | 2012 | 2009
| Anaesthetics, volatile | 3 | 11, Sup 7 | 1987 | 1987
| Androgenic (anabolic) steroids | 2A | Sup 7 | 1987 | 1987
| Areca nut | 1 | 85, 100E | 2012 | 2009
| Arecoline | 2B | 128 | 2021 online | 2020
| Art glass, glass containers and pressed ware (manufacture of) | 2A | 58 | 1993 | 1993
| Auramine production | 1 | Sup 7, 99, 100F | 2012 | 2009
| Automotive gasoline (see also Gasoline engine exhaust, and Aviation gasoline) | 1 | 45, 138 | In prep. | 2025
| BK polyomavirus (BKV) | 2B | 104 | 2014 | 2012
| Benzidine, dyes metabolized to | 1 | 99, 100F | 2012 | 2009 | NB Overall evaluation upgraded to Group 1 based on mechanistic and other relevant data
Beta particles (see Radionuclides)
| Betel quid with tobacco | 1 | Sup 7, 85, 100E | 2012 | 2009
| Betel quid without tobacco | 1 | Sup 7, 85, 100E | 2012 | 2009
| Biomass fuel (primarily wood), indoor emissions from household combustion of | 2A | 95 | 2010 | 2006
| Bisulfites | 3 | 54 | 1992 | 1991
| Bitumens, occupational exposure to hard bitumens and their emissions during mastic asphalt work | 2B | 103 | 2013 | 2011
| Boot and shoe manufacture and repair (see Leather dust, Benzene) |  | 25, Sup 7 | 1987 | 1987
| Bracken fern | 2B | 40, Sup 7 | 1987 | 1987
| Calcium carbide production | 3 | 92 | 2010 | 2005
| Carbon electrode manufacture | 2A | 92 | 2010 | 2005
| Carpentry and joinery | 2B | 25, Sup 7 | 1987 | 1987
| Ceramic implants | 3 | 74 | 1999 | 1999
| Chimney sweeping (see Soot) |  | 92 | 2010 | 2005
| Chlorinated drinking-water | 3 | 52 | 1991 | 1990
| Chlorinated paraffins of average carbon chain length C12 and average degree of chlorination approximately 60% | 2B | 48 | 1990 | 1989
Chlorophenols (see Polychlorophenols)
| Chlorophenoxy herbicides | 2B | 41, Sup 7 | 1987 | 1987
| Coal dust | 3 | 68 | 1997 | 1996
| Coal gasification | 1 | Sup 7, 92, 100F | 2012 | 2009
| Coal, indoor emissions from household combustion of | 1 | 95, 100E | 2012 | 2009
Cobalt and cobalt compounds, see Soluble cobalt(II) salts, Cobalt metal (without tungsten carbide or other metal alloys), Cobalt(II) oxide, Cobalt(II,III) oxide, Cobalt(II) sulfide, and Other cobalt(II) compounds (not including Soluble cobalt(II) salts, Cobalt(II) oxide, Cobalt(II,III) oxide, or Cobalt(II) sulfide)
Cobalt metal without tungsten carbide, see Cobalt metal (without tungsten carbide or other metal alloys)
Cobalt sulfate and other soluble cobalt(II) salts, see Soluble cobalt(II) salts
| Coffee, drinking | 3 | 51, 116 | 2018 online | 2016 | NB There is "evidence suggesting lack of carcinogenicity" in humans of coffee drinking for cancers of the pancreas, liver, female breast, uterine endometrium, and prostate. Inverse associations with coffee drinking have been observed with cancers of the liver and uterine endometrium.
| Coke production | 1 | Sup 7, 92, 100F | 2012 | 2009
Continuous glass filament (see Glass filament)
| Dental materials | 3 | 74 | 1999 | 1999
Depleted uranium (see Radionuclides, alpha-particle-emitting, internally deposited)
Diesel engine exhaust (see Engine exhaust, diesel)
| Diesel fuel, marine | 2B | 45 | 1989 | 1988 | NB Overall evaluation upgraded to Group 2B with supporting evidence from other relevant data
| Diesel fuels, distillate (light) | 3 | 45 | 1989 | 1988
| Dry cleaning (occupational exposures in) | 2B | 63 | 1995 | 1995
Dyes metabolized to benzidine (see Benzidine, dyes metabolized to)
| Electric fields, extremely low-frequency | 3 | 80 | 2002 | 2001
| Electric fields, static | 3 | 80 | 2002 | 2001
| Engine exhaust, diesel | 1 | 46, 105 | 2014 | 2012
| Engine exhaust, gasoline | 2B | 46, 105 | 2014 | 2012
| Epstein-Barr virus | 1 | 70, 100B | 2012 | 2009
| Estrogen therapy, postmenopausal | 1 | 72, 100A | 2012 | 2008 | NB There is "evidence suggesting lack of carcinogenicity" for estrogen-only menopausal therapy in humans and cancer of the colorectum. An inverse association has been observed between estrogen-only menopausal therapy and cancer of the colorectum.
| Estrogen-progestogen menopausal therapy (combined) | 1 | 72, 91, 100A | 2012 | 2008
| Estrogen-progestogen oral contraceptives (combined) | 1 | 72, 91, 100A | 2012 | 2008 | NB There is "evidence suggesting lack of carcinogenicity" for combined estrogen-progestogen oral contraceptives in humans and cancers of the endometrium, ovary, and colorectum. An inverse association has been observed between combined estrogen-progestogen oral contraceptives and cancers of the endometrium, ovary, and colorectum.
| Firefighter (occupational exposure as a) | 1 | 98, 132 | 2023 online | 2022
| Fission products, including strontium-90 | 1 | 100D | 2012 | 2009
| Flat-glass and specialty glass (manufacture of) | 3 | 58 | 1993 | 1993
| Fluorescent lighting | 3 | 55 | 1992 | 1992
| Fluoro-edenite fibrous amphibole | 1 | 111 | 2017 | 2014
Foreign bodies (see Ceramic implants, Dental materials, Implanted foreign bodies, Metallic implants, Organic polymeric materials, Orthopaedic implants, Polymeric implants, Silicone breast implants)
| Frying, emissions from high-temperature | 2A | 95 | 2010 | 2006
| Fuel oils, distillate (light) | 3 | 45 | 1989 | 1988
| Fuel oils, residual (heavy) | 2B | 45 | 1989 | 1988
| Furniture and cabinet making (see Wood dust) |  | 25, Sup 7 | 1987 | 1987
Gamma-Radiation (see X- and Gamma-Radiation)
| Aviation gasoline (see also Automotive gasoline, and Gasoline engine exhaust) | 2B | 45 | 1989 | 1988 | NB Overall evaluation upgraded to Group 2B with supporting evidence from other relevant data
Gasoline engine exhaust (see Engine exhaust, gasoline)
| Glass filament, continuous | 3 | 43, 81 | 2002 | 2001
| Goldenseal root powder | 2B | 108 | 2016 | 2013
| Haematite mining (underground) | 1 | 1, Sup 7, 100D | 2012 | 2009
| Hair colouring products (personal use of) | 3 | 57, 99 | 2010 | 2008
| Hairdresser or barber (occupational exposure as a) | 2A | 57, 99 | 2010 | 2008
| Hepatitis B virus (chronic infection with) | 1 | 59, 100B | 2012 | 2009
| Hepatitis C virus (chronic infection with) | 1 | 59, 100B | 2012 | 2009
| Hepatitis D virus (HDV) | 1 | 59, 139 | In prep. | 2025
| Hexachlorocyclohexanes | 2B | 20, Sup 7 | 1987 | 1987
High-temperature frying (see Frying)
Household combustion of biomass fuel (see Biomass fuel, indoor emissions from household combustion of)
Household combustion of coal (see Coal, indoor emissions from household combustion)
| Human cytomegalovirus (HCMV) | 2B | 139 | In prep. | 2025
| Human T-cell lymphotropic virus type I | 1 | 67, 100B | 2012 | 2009
| Human T-cell lymphotropic virus type II | 3 | 67 | 1996 | 1996
Human herpesvirus type 4 (see Epstein-Barr virus)
Human herpesvirus type 8 (see Kaposi sarcoma herpesvirus)
| Human immunodeficiency virus type 1 (infection with) | 1 | 67, 100B | 2012 | 2009
| Human immunodeficiency virus type 2 (infection with) | 2B | 67 | 1996 | 1996
| Human papillomavirus (HPV) genus beta (except types 5 and 8) and genus gamma | 3 | 90, 100B | 2012 | 2009
| Human papillomavirus (HPV) type 16 | 1 | 64, 90, 100B | 2012 | 2009 | NB The HPV types that have been classified as carcinogenic to humans can differ by an order of magnitude in risk for cervical cancer
| Human papillomavirus (HPV) type 18 | 1 | 64, 90, 100B | 2012 | 2009 | NB The HPV types that have been classified as carcinogenic to humans can differ by an order of magnitude in risk for cervical cancer
| Human papillomavirus (HPV) type 33 | 1 | 64, 90, 100B | 2012 | 2009 | NB The HPV types that have been classified as carcinogenic to humans can differ by an order of magnitude in risk for cervical cancer
| Human papillomavirus (HPV) type 68 | 2A | 100B | 2012 | 2009
| Human papillomavirus (HPV) types 31, 35, 39, 45, 51, 52, 56, 58, 59 | 1 | 64, 90, 100B | 2012 | 2009 | NB The HPV types that have been classified as carcinogenic to humans can differ by an order of magnitude in risk for cervical cancer
| Human papillomavirus (HPV) types 26, 53, 66, 67, 70, 73, 82 | 2B | 100B | 2012 | 2009
| Human papillomavirus (HPV) types 30, 34, 69, 85, 97 | 2B | 100B | 2012 | 2009 | NB Classified by phylogenetic analogy to the HPV genus alpha types classified in Group 1
| Human papillomavirus (HPV) types 5 and 8 (in patients with epidermodysplasia verruciformis) | 2B | 100B | 2012 | 2009
| Human papillomavirus (HPV) types 6 and 11 | 3 | 90, 100B | 2012 | 2009
| Hypochlorite salts | 3 | 52 | 1991 | 1990
| Implanted foreign bodies of metallic chromium or titanium and of cobalt-based, chromium-based, and titanium-based alloys, and stainless steel | 3 | 74 | 1999 | 1999
| Implanted foreign bodies of metallic cobalt, metallic nickel and an alloy powder containing 66-67% nickel, 13-16% chromium, and 7% iron | 2B | 74 | 1999 | 1999
| Insulation glass wool | 3 | 43, 81 | 2002 | 2001
Involuntary smoking (see Tobacco smoke, secondhand)
| Ionizing radiation (all types) | 1 | 100D | 2012 | 2009
| Iron and steel founding (occupational exposure during) | 1 | 34, Sup 7, 100F | 2012 | 2009
| Isopropyl alcohol manufacture using strong acids | 1 | Sup 7, 100F | 2012 | 2009
| Isopropyl oils | 3 | 15, Sup 7, 71 | 1999 | 1998
| JC polyomavirus (JCV) | 2B | 104 | 2014 | 2012
| Jet fuel | 3 | 45 | 1989 | 1988
| Kaposi sarcoma herpesvirus | 1 | 70, 100B | 2012 | 2009
| Lead compounds, inorganic | 2A | Sup 7, 87 | 2006 | 2004
| Lead compounds, organic | 3 | 23, Sup 7, 87 | 2006 | 2004 | NB Organic lead compounds are metabolized at least in part, to ionic lead both in humans and animals. To the extent that ionic lead, generated from organic lead, is present in the body, it will be expected to exert the toxicities associated with inorganic lead
| Leather dust | 1 | 100C | 2012 | 2009
| Leather goods manufacture (see also Leather dust) | 3 | 25, Sup 7 | 1987 | 1987
| Leather tanning and processing | 3 | 25, Sup 7 | 1987 | 1987
| Lumber and sawmill industries (including logging) | 3 | 25, Sup 7 | 1987 | 1987
| MOPP and other combined chemotherapy including alkylating agents | 1 | Sup 7, 100A | 2012 | 2008
| Madder root (Rubia tinctorum) | 3 | 82 | 2002 | 2002
| Magenta production | 1 | Sup 7, 57, 99, 100F | 2012 | 2009
| Magnetic fields, extremely low-frequency | 2B | 80 | 2002 | 2001
| Magnetic fields, static | 3 | 80 | 2002 | 2001
| Malaria (caused by infection with Plasmodium falciparum in holoendemic areas) | 2A | 104 | 2014 | 2012
Mate, hot (see Very hot beverages)
| Mate, not very hot (drinking) | 3 | 51, 116 | 2018 online | 2016
| Merkel cell polyomavirus (MCPyV) | 1 | 104, 139 | In prep. | 2025
| Metabisulfites | 3 | 54 | 1992 | 1991
| Metallic implants prepared as thin smooth films | 2B | 74 | 1999 | 1999
| Methylmercury compounds | 2B | 58 | 1993 | 1993 | NB Evaluated as a group
| Mineral oils, highly-refined | 3 | 33, Sup 7 | 1987 | 2009
| Mineral oils, untreated or mildly treated | 1 | 33, Sup 7, 100F | 2012 | 2009
| Modacrylic fibres | 3 | 19, Sup 7 | 1987 | 1987
| Neutron radiation | 1 | 75, 100D | 2012 | 2009 | NB Overall evaluation upgraded to Group 1 with supporting evidence from other relevant data
| Nickel compounds | 1 | Sup 7, 49, 100C | 2012 | 2009
| Nickel refining (see Nickel compounds) |  | 11 | 1976 | 1976
| Night shift work | 2A | 98, 124 | 2020 online | 2019 | NB Volume 98 evaluated shiftwork that involves circadian disruption
| Nitrate or nitrite (ingested) under conditions that result in endogenous nitrosation | 2A | 94 | 2010 | 2006
| Non-arsenical insecticides (occupational exposures in spraying and application of) | 2A | 53 | 1991 | 1990
Oestrogen (see Estrogen)
| Opium consumption | 1 | 126 | 2021 online | 2020
Oral contraceptives, combined estrogen-progestogen (see Estrogen-progestogen oral contraceptives)
| Organic polymeric materials | 3 | 74 | 1999 | 1999
| Orthopaedic implants of complex composition and cardiac pacemakers | 3 | 74 | 1999 | 1999
| Other cobalt(II) compounds (not including Soluble cobalt(II) salts, Cobalt(II) oxide, Cobalt(II,III) oxide, and Cobalt(II) sulfide) | 3 | 52, 86, 131 | 2023 online | 2022
| Outdoor air pollution | 1 | 109 | 2016 | 2013
| Outdoor air pollution, particulate matter in | 1 | 109 | 2016 | 2013
Oxymetholone (see Androgenic (anabolic) steroids)
| Paint manufacture (occupational exposure in) | 3 | 47 | 1989 | 2009
| Painter (occupational exposure as a) | 1 | 47, 98, 100F | 2012 | 2009
Particulate matter in outdoor air pollution (see Outdoor air pollution, particulate matter in)
| Paving and roofing with coal-tar pitch (see Coal-tar pitch) |  | 35, Sup 7, 92, 100F
| Pentavalent antimony | 3 | 131 | 2023 online | 2022
| Petroleum refining (occupational exposures in) | 2A | 45 | 1989 | 1988
| Petroleum solvents | 3 | 47 | 1989 | 1988
| Phenacetin, analgesic mixtures containing | 1 | Sup 7, 100A | 2012 | 2008
| Pickled vegetables (traditional Asian) | 2B | 56 | 1993 | 1992
| Polychlorinated biphenyls, dioxin-like, with a Toxicity Equivalency Factor (TEF) according to WHO (PCBs 77, 81, 105, 114, 118, 123, 126, 156, 157, 167, 169, 189) | 1 | 107 | 2016 | 2013 | NB Overall evaluation upgraded to Group 1 with strong supporting evidence from other relevant data
| Polychlorinated dibenzo-para-dioxins (other than 2,3,7,8-tetrachlorodibenzo-para-dioxin) | 3 | 69 | 1997 | 1997
| Polychlorinated dibenzofurans (see 2,3,4,7,8-Pentachlorodibenzofuran) | 3 | 69 | 1997 | 1997
| Polychlorophenols and their sodium salts (mixed exposures) (see Pentachlorophenol; 2,4,6-Trichlorophenol) | 2B | 53, 71 | 1999 | 1998
| Polymeric implant prepared as thin smooth films (with the exception of poly-glycolic acid) | 2B | 74 | 1999 | 1999
| Printing inks | 3 | 65 | 1996 | 1995
| Printing processes (occupational exposures in) | 2B | 65 | 1996 | 1995
| Processed meat (consumption of) | 1 | 114 | 2018 | 2015
| Proflavine salts | 3 | 24, Sup 7 | 1987 | 1987
| Progestins (see also Progestogen-only contraceptives) | 2B | Sup 7 | 1987 | 1987
| Progestogen-only contraceptives | 2B | 72 | 1999 | 1998
| Pulp and paper manufacture | 3 | 25, Sup 7 | 1987 | 1987
| Radiofrequency electromagnetic fields | 2B | 102 | 2013 | 2011
| Radioiodines, including iodine-131 | 1 | 78, 100D | 2012 | 2009
| Radionuclides, alpha-particle-emitting, internally deposited | 1 | 78, 100D | 2012 | 2009 | NB Specific radionuclides for which there is sufficient or limited evidence in humans are also listed individually as Group 1 agents
| Radionuclides, beta-particle-emitting, internally deposited | 1 | 78, 100D | 2012 | 2009 | NB Specific radionuclides for which there is sufficient evidence in humans are also listed individually as Group 1 agents
| Red meat (consumption of) | 2A | 114 | 2018 | 2015
| Refractory ceramic fibres | 2B | 43, 81 | 2002 | 2001
| Rock (stone) wool | 3 | 43, 81 | 2002 | 2001
| Rubber manufacturing industry | 1 | 28, Sup 7, 100F | 2012 | 2009
| SV40 polyomavirus | 3 | 104 | 2014 | 2012
| Salted fish, Chinese-style | 1 | 56, 100E | 2012 | 2009
| Silicone breast implants | 3 | 74 | 1999 | 1999 | NB There is "evidence suggesting lack of carcinogenicity" in humans of silicone breast implants for female breast carcinoma.
| Slag wool | 3 | 43, 81 | 2002 | 2001
| Solar radiation (see Ultraviolet radiation (wavelengths 100–400 nm, encompassing UVA, UVB, and UVC); Ultraviolet radiation from welding; and Ultraviolet-emitting tanning devices) | 1 | 55, 100D | 2012 | 2009
| Soluble cobalt(II) salts | 2A | 52, 86, 131 | 2023 online | 2022
| Soot (as found in occupational exposure of chimney sweeps) | 1 | 35, Sup 7, 92, 100F | 2012 | 2009
| Special-purpose fibres such as E-glass and '475' glass fibres | 2B | 81 | 2002 | 2001
| Sulfites | 3 | 54 | 1992 | 1991
Sunlamps and sunbeds (see Ultraviolet-emitting tanning devices)
Surgical implants (see Ceramic implants, Dental materials, Implanted foreign bodies, Metallic implants, Organic polymeric materials, Orthopaedic implants, Polymeric implants, Silicone breast implants)
| Tea | 3 | 51 | 1991 | 1990
Testosterone (see Androgenic (anabolic) steroids)
| Tetrakis(hydroxymethyl)phosphonium salts | 3 | 48, 71 | 1999 | 1998
| Textile manufacturing industry (work in) | 2B | 48 | 1990 | 1989
| Tobacco smoke, secondhand | 1 | 83, 100E | 2012 | 2009
| Tobacco smoking | 1 | 83, 100E | 2012 | 2009 | NB There is "evidence suggesting lack of carcinogenicity" in humans of tobacco smoking for cancers of the endometrium (postmenopausal) and thyroid gland.
| Tobacco, smokeless | 1 | Sup 7, 89, 100E | 2012 | 2009
Toxins derived from certain Fusarium species (see Fusarium)
| Trivalent antimony | 2A | 47, 131 | 2023 online | 2022
| Ultraviolet radiation from welding (see Ultraviolet radiation (wavelengths 100–400 nm, encompassing UVA, UVB, and UVC); Ultraviolet-emitting tanning devices; and Solar radiation) | 1 | 100D*, 118** | 2018 online | 2017 | *Volume 100D concluded that there is sufficient evidence for ocular melanoma in welders; **Volume 118 concluded that ultraviolet emissions from welding are carcinogenic to humans (Group 1; there is sufficient evidence in humans for the carcinogenicity of ultraviolet radiation from welding)
| Ultraviolet radiation (wavelengths 100–400 nm, encompassing UVA, UVB, and UVC) (see Ultraviolet radiation from welding; Ultraviolet-emitting tanning devices; and Solar radiation) | 1 | 55, 100D | 2012 | 2017 | NB Overall evaluation upgraded to Group 1 based on mechanistic and other relevant data
| Ultraviolet-emitting tanning devices (see Ultraviolet radiation (wavelengths 100–400 nm, encompassing UVA, UVB, and UVC); Ultraviolet radiation from welding; and Solar radiation) | 1 | 100D | 2012 | 2009
Urethane (see Ethyl carbamate)
| Very hot beverages at above 65 °C (drinking) | 2A | 116 | 2018 online | 2016
| Weapons-grade tungsten (with nickel and cobalt) alloy | 2B | 131 | 2023 online | 2022
| Welding fumes | 1 | 49, 118 | 2018 online | 2017
| Wood dust | 1 | 62, 100C | 2012 | 2009
Wood smoke (see Biomass fuel, indoor emissions from household combustion)
| X- and Gamma-Radiation | 1 | 75, 100D | 2012 | 2009
100-00-5 | 4-Chloronitrobenzene | 2B | 65, 123 | 2020 online | 2018
100-17-4 | para-Nitroanisole | 2B | 123 | 2020 online | 2018
100-40-3 | 4-Vinylcyclohexene | 2B | Sup 7, 60 | 1994 | 1994
100-41-4 | Ethylbenzene | 2B | 77 | 2000 | 2000
100-42-5 | Styrene | 2A | 60, 82, 121 | 2019 | 2018
100-75-4 | N-Nitrosopiperidine | 2B | 17, Sup 7 | 1987 | 1987
10043-66-0 | Iodine-131 (see Radioiodines)
10043-92-2 | Radon-222 and its decay products | 1 | 43, 78, 100D | 2012 | 2009
10048-13-2 | Sterigmatocystin | 2B | 10, Sup 7 | 1987 | 1987
10048-32-5 | Parasorbic acid | 3 | 10, Sup 7 | 1987 | 1987
10098-97-2 | Strontium-90 (see Fission products)
101-14-4 | 4,4'-Methylenebis(2-chloroaniline) (MOCA) | 1 | Sup 7, 57, 99, 100F | 2012 | 2009 | NB Overall evaluation upgraded to Group 1 based on mechanistic and other relevant data
101-21-3 | Chloropropham | 3 | 12, Sup 7 | 1987 | 1987
101-25-7 | Dinitrosopentamethylenetetramine | 3 | 11, Sup 7 | 1987 | 1987
101-61-1 | Michler’s base [4,4´-methylenebis(N,N-dimethyl)­benzenamine] | 2B | 27, Sup 7, 99 | 2010 | 2008
101-68-8 | 4,4'-Methylenediphenyl diisocyanate | 3 | 19, Sup 7, 71 | 1999 | 1998
101-77-9 | 4,4'-Methylenedianiline | 2B | 39, Sup 7 | 1987 | 1987
101-80-4 | 4,4'-Diaminodiphenyl ether | 2B | 29, Sup 7 | 1987 | 1987
101-90-6 | Diglycidyl resorcinol ether | 2B | 36, Sup 7, 71 | 1999 | 1998
101043-37-2 | Microcystin-LR | 2B | 94 | 2010 | 2006
102-50-1 | meta-Cresidine | 3 | 27, Sup 7 | 1987 | 1987
102-71-6 | Triethanolamine | 3 | 77 | 2000 | 2000
103-03-7 | Phenicarbazide | 3 | 12, Sup 7 | 1987 | 1987
103-11-7 | 2-Ethylhexyl acrylate | 2B | 60, 122 | 2019 | 2018
103-23-1 | Di(2-ethylhexyl) adipate | 3 | Sup 7, 77 | 2000 | 2000
103-33-3 | Azobenzene | 3 | 8, Sup 7 | 1987 | 1987
103-90-2 | Acetaminophen (see Paracetamol)
103-90-2 | Paracetamol (Acetaminophen) | 3 | 50, 73 | 1999 | 1998
10380-28-6 | Copper 8-hydroxyquinoline | 3 | 15, Sup 7 | 1987 | 1987
104-94-9 | para-Anisidine | 3 | 27, Sup 7 | 1987 | 1987
104987-11-3 | Tacrolimus | 1 | 137 | 2026 online | 2024
105-11-3 | para-Benzoquinone dioxime | 3 | 29, Sup 7, 71 | 1999 | 1998
105-55-5 | N,N'-Diethylthiourea | 3 | 79 | 2001 | 2000
105-60-2 | Caprolactam | 3 | 39, Sup 7, 71 | 1999 | 1998 | Moved to Group 3 following 2019 update to the IARC Monographs Preamble
105-74-8 | Lauroyl peroxide | 3 | 36, Sup 7, 71 | 1999 | 1998
10540-29-1 | Tamoxifen | 1 | 66, 100A | 2012 | 2008 | NB There is "evidence suggesting lack of carcinogenicity" for tamoxifen in humans and cancer of the female breast. An inverse association has been observed between tamoxifen and cancer of the female breast.
105650-23-5 | PhIP (2-Amino-1-methyl-6-phenylimidazo[4,5-b]pyridine) | 2B | 56 | 1993 | 1992
105735-71-5 | 3,7-Dinitrofluoranthene | 2B | 46, 65, 105 | 2014 | 2012
10595-95-6 | N-Nitrosomethylethylamine | 2B | 17, Sup 7 | 1987 | 1987
10599-90-3 | Chloramine | 3 | 84 | 2004 | 2002
106-46-7 | para-Dichlorobenzene | 2B | Sup 7, 73 | 1999 | 1998
106-47-8 | para-Chloroaniline | 2B | 57 | 1993 | 1992
106-50-3 | para-Phenylenediamine | 3 | 16, Sup 7 | 1987 | 1987
106-51-4 | para-Quinone | 3 | 15, Sup 7, 71 | 1999 | 1998
106-87-6 | 4-Vinylcyclohexene diepoxide | 2B | Sup 7, 60 | 1994 | 1994
106-88-7 | 1,2-Epoxybutane | 2B | 47, 71 | 1999 | 1998 | NB Overall evaluation upgraded to Group 2B with supporting evidence from other relevant data
106-89-8 | Epichlorohydrin | 2A | 11, Sup 7, 71 | 1999 | 1998 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
106-91-2 | Glycidyl methacrylate | 2A | 125 | 2020 online | 2019
106-93-4 | Ethylene dibromide | 2A | 15, Sup 7, 71 | 1999 | 1998 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
106-94-5 | 1-Bromopropane | 2B | 115 | 2018 | 2016
106-99-0 | 1,3-Butadiene | 1 | Sup 7, 54, 71, 97, 100F | 2012 | 2009
107-02-8 | Acrolein | 2A | 63, Sup 7, 128 | 2021 online | 2020
107-05-1 | Allyl chloride | 3 | 36, Sup 7, 71, 125 | 2020 online | 2019
107-06-2 | 1,2-Dichloroethane | 2B | 20, Sup 7, 71 | 1999 | 1998
107-13-1 | Acrylonitrile | 1 | 19, Sup 7, 71, 136 | 2025 online | 2024
107-14-2 | Chloroacetonitrile | 3 | 52, 71 | 1999 | 1998
107-30-2 | Chloromethyl methyl ether (see Bis(chloromethyl)ether; chloromethyl methyl ether)
1071-83-6 | Glyphosate | 2A | 112 | 2017 | 2015
1072-52-2 | 2-(1-Aziridinyl)ethanol | 3 | 9, Sup 7 | 1987 | 1987
108-05-4 | Vinyl acetate | 2B | Sup 7, 63 | 1995 | 1995
108-10-1 | Methyl isobutyl ketone | 2B | 101 | 2013 | 2011
108-20-3 | Diisopropyl ether | 3 | 138 | In prep. | 2025
108-30-5 | Succinic anhydride | 3 | 15, Sup 7 | 1987 | 1987
108-45-2 | meta-Phenylenediamine | 3 | 16, Sup 7 | 1987 | 1987
108-46-3 | Resorcinol | 3 | 15, Sup 7, 71 | 1999 | 1998
108-60-1 | Bis(2-chloro-1-methylethyl)ether | 3 | 41, Sup 7, 71 | 1999 | 1998
108-78-1 | Melamine | 2B | Sup 7, 73, 119 | 2019 | 2017
108-88-3 | Toluene | 3 | 47, 71 | 1999 | 1998
108-94-1 | Cyclohexanone | 3 | 47, 71 | 1999 | 1998
108-95-2 | Phenol | 3 | 47, 71 | 1999 | 1998
108-99-6 | β-Picoline | 3 | 122 | 2019 | 2018
109-70-6 | 1-Bromo-3-chloropropane | 2B | 125 | 2020 online | 2019
109-99-9 | Tetrahydrofuran | 2B | 119 | 2019 | 2017
110-00-9 | Furan | 2B | 63 | 1995 | 1995
110-57-6 | trans-1,4-Dichlorobutene | 3 | 15, Sup 7, 71 | 1999 | 1998
110-86-1 | Pyridine | 2B | 77, 119 | 2019 | 2017
110-91-8 | Morpholine | 3 | 47, 71 | 1999 | 1998
11056-06-7 | Bleomycins | 2B | 26, Sup 7 | 1987 | 1987 | NB Overall evaluation upgraded to Group 2B with supporting evidence from other relevant data
111-42-2 | Diethanolamine | 2B | 77, 101 | 2013 | 2011
111-44-4 | Bis(2-chloroethyl)ether | 3 | 9, Sup 7, 71 | 1999 | 1998
111-76-2 | 2-Butoxyethanol | 3 | 88 | 2006 | 2004
111025-46-8 | Pioglitazone | 2A | 108 | 2016 | 2013
111189-32-3 | Naphtho[1,2-b]fluoranthene | 3 | 92 | 2010 | 2005
1116-54-7 | N-Nitrosodiethanolamine | 2B | 17, Sup 7, 77 | 2000 | 2000
1120-71-4 | 1,3-Propane sultone | 2A | 4, Sup 7, 71, 110 | 2017 | 2014 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
1143-38-0 | Dithranol | 3 | 13, Sup 7 | 1987 | 1987
115-02-6 | Azaserine | 2B | 10, Sup 7 | 1987 | 1987
115-07-1 | Propylene | 3 | Sup 7, 60 | 1994 | 1994
115-28-6 | Chlorendic acid | 2B | 48 | 1990 | 1989
115-32-2 | Dicofol | 3 | 30, Sup 7 | 1987 | 1987
115-96-8 | Tris(2-chloroethyl) phosphate | 3 | 48, 71 | 1999 | 1998
116-06-3 | Aldicarb | 3 | 53 | 1991 | 1990
116-14-3 | Tetrafluoroethylene | 2A | 19, Sup 7, 71, 110 | 2017 | 2014 | NB Overall evaluation upgraded to Group 2A on the basis of sufficient evidence in experimental animals with a striking and atypical pattern of tumours
1163-19-5 | Decabromodiphenyl oxide | 3 | 48, 71 | 1999 | 1998
116355-83-0 | Fusarium moniliforme, toxins derived from (fumonisin B1, fumonisin B2, and fusarin C) | 2B | 56 | 1993 | 1992
116355-83-0 | Fumonisin B1 | 2B | 82 | 2002 | 2002
117-10-2 | Dantron (Chrysazin; 1,8-Dihydroxyanthraquinone) | 2B | 50 | 1990 | 1989
117-39-5 | Quercetin | 3 | Sup 7, 73 | 1999 | 1998
117-79-3 | 2-Aminoanthraquinone | 3 | 27, Sup 7 | 1987 | 1987
117-81-7 | Bis(2-ethylhexyl) phthalate (see Di(2-ethylhexyl) phthalate)
117-81-7 | Di(2-ethylhexyl)phthalate | 2B | Sup 7, 77, 101 | 2013 | 2011
118-74-1 | Hexachlorobenzene | 2B | Sup 7, 79 | 2001 | 2000
118-92-3 | Anthranilic acid | 3 | 16, Sup 7 | 1987 | 1987
118-96-7 | 2,4,6-Trinitrotoluene | 3 | 65 | 1996 | 1995
118399-22-7 | Nodularins | 3 | 94 | 2010 | 2006
119-34-6 | 4-Amino-2-nitrophenol | 3 | 16, Sup 7 | 1987 | 1987
119-61-9 | Benzophenone | 2B | 101 | 2013 | 2011
119-90-4 | 3,3'-Dimethoxybenzidine (ortho-Dianisidine) | 2B | 4, Sup 7 | 1987 | 1987
119-93-7 | 3,3'-Dimethylbenzidine (ortho-Tolidine) | 2B | 1, Sup 7 | 1987 | 1987
120-12-7 | Anthracene | 2B | 92, Sup 7, 133 | 2024 online | 2023
120-58-1 | Isosafrole | 3 | 10, Sup 7 | 1987 | 1987
120-71-8 | para-Cresidine | 2B | 27, Sup 7 | 1987 | 1987
120-80-9 | Catechol | 2B | 15, Sup 7, 71 | 1999 | 1998
12001-79-5 | Vitamin K substances | 3 | 76 | 2000 | 1999
12070-12-1, 7440-48-4 | Tungsten carbide with cobalt metal (see Cobalt metal with tungsten carbide)
121-14-2 | 2,4-Dinitrotoluene | 2B | 65 | 1996 | 1995
121-66-4 | 2-Amino-5-nitrothiazole | 3 | 31, Sup 7 | 1987 | 1987
121-69-7 | N,N-Dimethylaniline | 3 | 57 | 1993 | 1992
121-73-3 | 3-Chloronitrobenzene | 3 | 65 | 1996 | 1995
121-75-5 | Malathion | 2A | 30, Sup 7, 112 | 2017 | 2015
121-88-0 | 2-Amino-5-nitrophenol | 3 | 57 | 1993 | 1992
12122-67-7 | Zineb | 3 | 12, Sup 7 | 1987 | 1987
12174-11-7 | Attapulgite (see Palygorskite)
12174-11-7 | Palygorskite (Attapulgite) (long fibres, > 5 micrometres) | 2B | 68 | 1997 | 1996
12174-11-7 | Palygorskite (Attapulgite)(short fibres, < 5 micrometres) | 3 | 68 | 1997 | 1996
12192-57-3 | Aurothioglucose | 3 | 13, Sup 7 | 1987 | 1987
122-34-9 | Simazine | 3 | 53, 73 | 1999 | 1998
122-39-4 | Diphenylamine | 2B | 130 | 2022 online | 2021
122-42-9 | Propham | 3 | 12, Sup 7 | 1987 | 1987
122-60-1 | Phenyl glycidyl ether | 2B | 47, 71 | 1999 | 1998
122-66-7 | 1,2-Diphenylhydrazine | 2B | 130 | 2022 online | 2021
122320-73-4 | Rosiglitazone | 3 | 108 | 2016 | 2013
123-31-9 | Hydroquinone | 3 | 15, Sup 7, 71 | 1999 | 1998
123-33-1 | Maleic hydrazide | 3 | 4, Sup 7 | 1987 | 1987
123-35-3 | β-Myrcene | 2B | 119 | 2019 | 2017
123-72-8 | Butyraldehyde | 2B | 141 | In prep. | 2026
123-91-1 | 1,4-Dioxane | 2B | 11, Sup 7, 71 | 1999 | 1998
124-48-1 | Chlorodibromomethane | 3 | 52, 71 | 1999 | 1998
124-58-3 | Methylarsonic acid | 2B | 100C | 2012 | 2009
124-58-3 | Monomethylarsonic acid (see Methylarsonic acid)
12427-38-2 | Maneb | 3 | 12, Sup 7 | 1987 | 1987
125-33-7 | Primidone | 2B | 108 | 2016 | 2013
126-07-8 | Griseofulvin | 2B | Sup 7, 79 | 2001 | 2000
126-72-7 | Tris(2,3-dibromopropyl) phosphate | 2A | 20, Sup 7, 71 | 1999 | 1998 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
126-85-2 | Nitrogen mustard N-oxide | 2B | 9, Sup 7 | 1987 | 1987
126-99-8 | Chloroprene | 2B | Sup 7, 71 | 1999 | 1998
12663-46-6 | Cyclochlorotine | 3 | 10, Sup 7 | 1987 | 1987
127-07-1 | Hydroxyurea | 3 | 76 | 2000 | 1999
127-18-4 | Tetrachloroethylene (Perchloroethylene) | 2A | Sup 7, 63, 106 | 2014 | 2012
127-19-5 | N,N-Dimethylacetamide | 2B | 123 | 2020 online | 2018
127-69-5 | Sulfafurazole (Sulfisoxazole) | 3 | 24, Sup 7 | 1987 | 1987
128-37-0 | Butylated hydroxytoluene (BHT) | 3 | 40, Sup 7 | 1987 | 1987
128-66-5 | Vat Yellow 4 | 3 | 48 | 1990 | 1989
129-00-0 | Pyrene | 3 | Sup 7, 92 | 2010 | 2005
129-15-7 | 2-Methyl-1-nitroanthraquinone (uncertain purity) | 2B | 27, Sup 7 | 1987 | 1987
129-17-9 | Blue VRS | 3 | 16, Sup 7 | 1987 | 1987
129-20-4 | Oxyphenbutazone | 3 | 13, Sup 7 | 1987 | 1987
129-43-1 | 1-Hydroxyanthraquinone | 2B | 82 | 2002 | 2002
129-73-7 | Leucomalachite green (see also Malachite green) | 2B | 129 | 2022 online | 2021
13010-47-4 | 1-(2-Chloroethyl)-3-cyclohexyl-1-nitrosourea (CCNU) (Lomustine) | 2A | 26, Sup 7 | 1987 | 1987 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
1303-00-0 | Gallium arsenide (see Arsenic and inorganic arsenic compounds) |  | 86, 100C | 2012 | 2009
13045-94-8 | Medphalan | 3 | 9, Sup 7 | 1987 | 1987
1307-96-6 | Cobalt(II) oxide | 2B | 52, 131 | 2023 online | 2022
1308-06-1 | Cobalt(II,III) oxide | 3 | 52, 131 | 2023 online | 2022
1309-37-1 | Ferric oxide | 3 | 1, Sup 7 | 1987 | 1987
1309-64-4 | Antimony trioxide, see Trivalent antimony
131-79-3 | Yellow OB | 3 | 8, Sup 7 | 1987 | 1987
1313-27-5 | Molybdenum trioxide | 2B | 118 | 2018 online | 2017
1314-62-1 | Vanadium pentoxide | 2B | 86 | 2006 | 2003
1317-42-6 | Cobalt(II) sulfide | 3 | 52, 131 | 2023 online | 2022
1317-60-8 | Haematite | 3 | 1, Sup 7 | 1987 | 1987
1318-02-1 | Zeolites other than erionite (clinoptilolite, phillipsite,mordenite, non-fibrous Japanese zeolite, synthetic zeolites) | 3 | 68 | 1997 | 1996
132-27-4 | Sodium ortho-phenylphenate | 2B | Sup 7, 73 | 1999 | 1998
132-65-0 | Dibenzothiophene | 3 | 103 | 2013 | 2011
13233-32-4 | Radium-224 and its decay products | 1 | 78, 100D | 2012 | 2009
13256-22-9 | N-Nitrososarcosine | 2B | 17, Sup 7 | 1987 | 1987
13292-46-1 | Rifampicin | 3 | 24, Sup 7 | 1987 | 1987
133-06-2 | Captan | 3 | 30, Sup 7 | 1987 | 1987
1330-20-7 | Xylenes | 3 | 47, 71 | 1999 | 1998
1332-21-4, 12172-73-5, 12001-29-5, 12001-28-4 | Asbestos (all forms, including actinolite, amosite, anthophyllite, chrysotile, crocidolite, tremolite) | 1 | 14, Sup 7, 100C | 2012 | 2009 | NB Mineral substances (e.g. talc or vermiculite) that contain asbestos should also be regarded as carcinogenic to humans NB2 These CAS Nos are for substances that CAS does not treat in its regular CA index
1333-86-4 | Carbon black | 2B | Sup 7, 65, 93 | 2010 | 2006
1336-36-3 | Polychlorinated biphenyls | 1 | 18, Sup 7, 107 | 2016 | 2013
1338-16-5 | Iron sorbitol-citric acid complex | 3 | 2, Sup 7 | 1987 | 1987
134-29-2 | ortho-Anisidine hydrochloride (see also ortho-Anisidine) | 2A | 127 | 2021 online | 2020
134-32-7 | 1-Naphthylamine | 3 | 4, Sup 7 | 1987 | 1987
1345-04-6 | Antimony trisulfide, see Trivalent antimony
13463-67-7 | Titanium dioxide | 2B | 47, 93 | 2010 | 2006
13483-18-6 | 1,2-Bis(chloromethoxy)ethane | 3 | 15, Sup 7, 71 | 1999 | 1998
135-20-6 | Cupferron | 2B | 127 | 2021 online | 2020
135-88-6 | N-Phenyl-2-naphthylamine | 3 | 16, Sup 7 | 1987 | 1987
136-40-3 | Phenazopyridine hydrochloride | 2B | 24, Sup 7 | 1987 | 1987
13674-84-5, 76025-08-6, 76649-15-5, 6145-73-9 | Tris(chloropropyl) phosphate (TCPP) | 2A | 141 | In prep. | 2026
137-17-7 | 2,4,5-Trimethylaniline | 3 | 27, Sup 7 | 1987 | 1987
137-26-8 | Thiram | 3 | Sup 7, 53 | 1991 | 1990
137-30-4 | Ziram | 3 | Sup 7, 53 | 1991 | 1990
137234-62-9 | Voriconazole | 1 | 137 | 2026 online | 2024
138-59-0 | Shikimic acid | 3 | 40, Sup 7 | 1987 | 1987
139-05-9 | Cyclamates (sodium cyclamate) | 3 | Sup 7, 73 | 1999 | 1998
139-13-9 | Nitrilotriacetic acid and its salts | 2B | 48, 73 | 1999 | 1998 | NB Evaluated as a group
139-65-1 | 4,4'-Thiodianiline | 2B | 27, Sup 7 | 1987 | 1987
139-94-6 | Nithiazide | 3 | 31, Sup 7 | 1987 | 1987
13909-09-6 | 1-(2-Chloroethyl)-3-(4-methylcyclohexyl)-1-nitrosourea (Methyl-CCNU) (Semustine) | 1 | Sup 7, 100A | 2012 | 2008
13982-63-3 | Radium-226 and its decay products | 1 | 78, 100D | 2012 | 2009
13983-17-0 | Wollastonite | 3 | Sup 7, 68 | 1997 | 1996
140-11-4 | Benzyl acetate | 3 | 40, Sup 7, 71 | 1999 | 1998
140-56-7 | para-Dimethylaminoazobenzenediazo sodium sulfonate | 3 | 8, Sup 7 | 1987 | 1987
140-57-8 | Aramite® | 2B | 5, Sup 7 | 1987 | 1987
140-88-5 | Ethyl acrylate | 2B | 39, Sup 7, 71, 122 | 2019 | 2018
1401-55-4 | Tannic acid and tannins | 3 | 10, Sup 7 | 1987 | 1987
1402-68-2 | Aflatoxins | 1 | Sup 7, 56, 82, 100F | 2012 | 2009
14047-09-7 | 3,3′,4,4′-Tetrachloroazobenzene | 2A | 117 | 2019 | 2016 | NB Overall evaluation upgraded to Group 2A
141-32-2 | n-Butyl acrylate | 3 | 39, Sup 7, 71 | 1999 | 1998
141-37-7 | 3,4-Epoxy-6-methylcyclohexylmethyl-3,4-epoxy-6-methylcyclo-hexanecarboxylate | 3 | 11, Sup 7, 71 | 1999 | 1998
141-90-2 | Thiouracil | 2B | Sup 7, 79 | 2001 | 2000
142-04-1 | Aniline hydrochloride (see also Aniline) | 2A | 127 | 2021 online | 2020
142-83-6 | 2,4-Hexadienal | 2B | 101 | 2013 | 2011
143-50-0 | Chlordecone (Kepone) | 2B | 20, Sup 7 | 1987 | 1987
143-67-9 | Vinblastine sulfate | 3 | 26, Sup 7 | 1987 | 1987
144-34-3 | Methyl selenac | 3 | 12, Sup 7 | 1987 | 1987
14484-64-1 | Ferbam | 3 | 12, Sup 7 | 1987 | 1987
14596-37-3 | Phosphorus-32, as phosphate | 1 | 78, 100D | 2012 | 2009
1464-53-5 | 1,2:3,4-Diepoxybutane (see 1,3-Butadiene) |  | 11, Sup 7
148-18-5 | Sodium diethyldithiocarbamate | 3 | 12, Sup 7 | 1987 | 1987
148-24-3 | 8-Hydroxyquinoline | 3 | 13, Sup 7 | 1987 | 1987
148-82-3 | Melphalan | 1 | 9, Sup 7, 100A | 2012 | 2008
14807-96-6 | Talc | 2A | 42, Sup 7, 93, 136 | 2025 online | 2024 | The term "Talc" includes "Talc containing asbestiform fibres other than asbestos" and "Talc not containing asbestiform fibres", which were evaluated separately in Volume 93. For "Talc containing asbestos", see "Asbestos".
14807-96-6 | Talc containing asbestos (see Asbestos) |  | 93, Sup 7, 100C | 2012 | 2009
14807-96-6 | Talc-based body powder (perineal use of) (see Talc) |  | 93 | 2010 | 2006
14808-60-7 | Silica dust, crystalline, in the form of quartz or cristobalite | 1 | Sup 7, 68, 100C | 2012 | 2009
149-29-1 | Patulin | 3 | 40, Sup 7 | 1987 | 1987
149-30-4 | 2-Mercaptobenzothiazole | 2A | 115 | 2018 | 2016
14901-08-7 | Cycasin | 2B | 10, Sup 7 | 1987 | 1987
150-13-0 | para-Aminobenzoic acid | 3 | 16, Sup 7 | 1987 | 1987
150-68-5 | Monuron | 3 | Sup 7, 53 | 1991 | 1990
150-69-6 | Dulcin | 3 | 12, Sup 7 | 1987 | 1987
15086-94-9 | Eosin | 3 | 15, Sup 7 | 1987 | 1987
151-56-4 | Aziridine | 2B | 9, Sup 7, 71 | 1999 | 1998 | NB Overall evaluation upgraded to Group 2B with supporting evidence from other relevant data
15262-20-1 | Radium-228 and its decay products | 1 | 78, 100D | 2012 | 2009
154-93-8 | Bischloroethyl nitrosourea (BCNU) (Carmustine) | 2A | 26, Sup 7 | 1987 | 1987
15501-74-3 | Sepiolite | 3 | Sup 7, 68 | 1997 | 1996
15503-86-3 | Isatidine | 3 | 10, Sup 7 | 1987 | 1987
156-10-5 | para-Nitrosodiphenylamine | 3 | 27, Sup 7 | 1987 | 1987
156-51-4 | Phenelzine sulfate | 3 | 24, Sup 7 | 1987 | 1987
15625-89-5 | Trimethylolpropane triacrylate, technical grade | 2B | 122 | 2019 | 2018
15663-27-1 | Cisplatin | 2A | 26, Sup 7 | 1987 | 1987 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
15721-02-5 | 2,2',5,5'-Tetrachlorobenzidine | 3 | 27, Sup 7 | 1987 | 1987
1582-09-8 | Trifluralin | 3 | 53 | 1991 | 1990
15972-60-8 | Alachlor | 2A | 140 | In prep. | 2025
16065-83-1 | Chromium (III) compounds | 3 | 49 | 1990 | 1989
16071-86-6 | CI Direct Brown 95 (see Benzidine, dyes metabolized to)
1615-80-1 | 1,2-Diethylhydrazine | 2B | 4, Sup 7, 71 | 1999 | 1998
1634-04-4 | Methyl tert-butyl ether | 2B | 73, 138 | In prep. | 2025
16543-55-8, 64091-91-4 | N'-Nitrosonornicotine (NNN) and 4-(N-Nitrosomethylamino)-1-(3-pyridyl)-1-butanone (NNK) | 1 | Sup 7, 89, 100E | 2012 | 2009 | NB Overall evaluation upgraded to Group 1 based on mechanistic and other relevant data
16568-02-8 | Gyromitrin | 3 | 31, Sup 7 | 1987 | 1987
1675-54-3 | Bisphenol A diglycidyl ether (Araldite) | 3 | 47, 71 | 1999 | 1998
1689-82-3 | 4-Hydroxyazobenzene | 3 | 8, Sup 7 | 1987 | 1987
1694-09-3 | Benzyl violet 4B | 2B | 16, Sup 7 | 1987 | 1987
16984-48-8 | Fluorides (inorganic, used in drinking-water) | 3 | 27, Sup 7 | 1987 | 1987
1705-85-7 | 6-Methylchrysene | 3 | Sup 7, 92 | 2010 | 2005
1706-01-0 | 3-Methylfluoranthene | 3 | Sup 7, 92 | 2010 | 2005
17117-34-9 | 3-Nitrobenzanthrone | 2B | 105 | 2014 | 2012 | NB Overall evaluation upgraded to Group 2B with supporting evidence from other relevant data
1746-01-6 | 2,3,7,8-Tetrachlorodibenzo-para-dioxin | 1 | Sup 7, 69, 100F | 2012 | 2009
1763-23-1 | Perfluorooctanesulfonic acid (PFOS) | 2B | 135 | 2025 online | 2023
1836-75-5 | Nitrofen (technical-grade) | 2B | 30, Sup 7 | 1987 | 1987
18540-29-9 | Chromium (VI) compounds | 1 | Sup 7, 49, 100C | 2012 | 2009
18883-66-4 | Streptozotocin | 2B | 17, Sup 7 | 1987 | 1987
189-55-9 | Dibenzo[a,i]pyrene | 2B | 92 | 2010 | 2005
189-64-0 | Dibenzo[a,h]pyrene | 2B | Sup 7, 92 | 2010 | 2005
1897-45-6 | Chlorothalonil | 2B | Sup 7, 73 | 1999 | 1998
191-07-1 | Coronene | 3 | 32, Sup 7, 92 | 2010 | 2005
191-24-2 | Benzo[ghi]perylene | 3 | 92, Sup 7 | 2010 | 2005
191-26-4 | Anthanthrene | 3 | 92, Sup 7 | 2010 | 2005
191-30-0 | Dibenzo[a,l]pyrene | 2A | Sup 7, 92 | 2010 | 2005 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
1912-24-9 | Atrazine | 2A | 53, 73, 140 | In prep. | 2025
1918-02-1 | Picloram | 3 | 53 | 1991 | 1990
192-47-2 | Dibenzo[h,rst]pentaphene | 3 | Sup 7, 92 | 2010 | 2005
192-51-8 | Dibenzo[e,l]pyrene | 3 | 92 | 2010 | 2005
192-65-4 | Dibenzo[a,e]pyrene | 3 | Sup 7, 92 | 2010 | 2005
192-97-2 | Benzo[e]pyrene | 3 | 92, Sup 7 | 2010 | 2005
193-09-9 | Naphtho[2,3-e]pyrene | 3 | 92 | 2010 | 2005
193-39-5 | Indeno[1,2,3-cd]pyrene | 2B | Sup 7, 92 | 2010 | 2005
1936-15-8 | CI Orange G | 3 | 8, Sup 7 | 1987 | 1987
1936-15-8 | Orange G (see CI Orange G)
1937-37-7 | CI Direct Black 38 (see Benzidine, dyes metabolized to)
194-59-2 | 7H-Dibenzo[c,g]carbazole | 2B | 32, Sup 7, 103 | 2013 | 2011
195-19-7 | Benzo[c]phenanthrene | 2B | 92, Sup 7 | 2010 | 2005 | NB Overall evaluation upgraded to Group 2B with supporting evidence from other relevant data
1954-28-5 | Triethylene glycol diglycidyl ether | 3 | 11, Sup 7, 71 | 1999 | 1998
196-78-1 | Benzo[g]chrysene | 3 | 92 | 2010 | 2005
198-55-0 | Perylene | 3 | Sup 7, 92 | 2010 | 2005
20073-24-9 | 3-Carbethoxypsoralen | 3 | 40, Sup 7 | 1987 | 1987
202-33-5 | Benz[j]aceanthrylene | 2B | 92 | 2010 | 2005 | NB Overall evaluation upgraded to Group 2B with supporting mechanistic and other relevant data
202-94-8 | 11H-Benz[bc]aceanthrylene | 3 | 92 | 2010 | 2005
202-98-2 | 4H-Cyclopenta[def]chrysene | 3 | 92 | 2010 | 2005
20268-51-3 | 7-Nitrobenz[a]anthracene | 3 | 46 | 1989 | 1988
203-12-3 | Benzo[ghi]fluoranthene | 3 | 92, Sup 7 | 2010 | 2005
203-20-3 | Naphtho[2,1-a]fluoranthene | 3 | 92 | 2010 | 2005
203-33-8 | Benzo[a]fluoranthene | 3 | Sup 7, 92 | 2010 | 2005
205-12-9 | Benzo[c]fluorene | 3 | 92, Sup 7 | 2010 | 2005
205-82-3 | Benzo[j]fluoranthene | 2B | 92 | 2010 | 2005
205-99-2 | Benzo[b]fluoranthene | 2B | 92 | 2010 | 2005
20589-63-3 | 3-Nitroperylene | 3 | 46 | 1989 | 1988
206-44-0 | Fluoranthene | 3 | Sup 7, 92 | 2010 | 2005
2068-78-2 | Vincristine sulfate | 3 | 26, Sup 7 | 1987 | 1987
207-08-9 | Benzo[k]fluoranthene | 2B | 92 | 2010 | 2005
207-83-0 | 13H-Dibenzo[a,g]fluorene | 3 | 92 | 2010 | 2005
20830-75-5 | Digoxin | 2B | 108 | 2016 | 2013
20830-81-3 | Daunomycin (Daunorubicin) | 2B | 10, Sup 7 | 1987 | 1987
20941-65-5 | Ethyl tellurac | 3 | 12, Sup 7 | 1987 | 1987
211-91-6 | Benz[l]aceanthrylene | 3 | 92 | 2010 | 2005
21259-20-1 | T2-Trichothecene | 3 | 31, Sup 7 | 1987 | 1987
213-46-7 | Picene | 3 | 92 | 2010 | 2005
214-17-5 | Benzo[b]chrysene | 3 | 92 | 2010 | 2005
215-58-7 | Dibenz[a,c]anthracene | 3 | Sup 7, 92 | 2010 | 2005
2164-17-2 | Fluometuron | 3 | 30, Sup 7 | 1987 | 1987
2168-68-5 | Bis(1-aziridinyl)morpholinophosphine sulfide | 3 | 9, Sup 7 | 1987 | 1987
217-59-4 | Triphenylene | 3 | Sup 7, 92 | 2010 | 2005
218-01-9 | Chrysene | 2B | 92 | 2010 | 2005
21884-44-6 | Luteoskyrin | 3 | 10, Sup 7 | 1987 | 1987
22248-79-9 | Tetrachlorvinphos | 2B | 30, Sup 7, 112 | 2017 | 2015
22349-59-3 | 1,4-Dimethylphenanthrene | 3 | Sup 7, 92 | 2010 | 2005
22398-80-7 | Indium phosphide | 2A | 86 | 2006 | 2003 | NB Overall evaluation upgraded to Group 2A
224-41-9 | Dibenz[a,j]anthracene | 3 | Sup 7, 92 | 2010 | 2005
224-42-0 | Dibenz[a,j]acridine | 2A | 32, Sup 7, 103 | 2013 | 2011 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
224-53-3 | Dibenz[c,h]acridine | 2B | 103 | 2013 | 2011 | NB Overall evaluation upgraded to Group 2B with supporting evidence from other relevant data
2243-62-1 | 1,5-Naphthalenediamine | 3 | 27, Sup 7 | 1987 | 1987
225-11-6 | Benz[a]acridine | 3 | 32, Sup 7, 103 | 2013 | 2011
225-51-4 | Benz[c]acridine | 3 | 32, Sup 7, 103 | 2013 | 2011
22506-53-2 | 3,9-Dinitrofluoranthene | 2B | 46, 65, 105 | 2014 | 2012
22571-95-5 | Symphytine | 3 | 31, Sup 7 | 1987 | 1987
226-36-8 | Dibenz[a,h]acridine | 2B | 32, Sup 7, 103 | 2013 | 2011
22839-47-0 | Aspartame | 2B | 134 | 2024 online | 2023
22966-79-6 | Oestradiol mustard | 3 | 9, Sup 7 | 1987 | 1987
22975-76-4 | 4,4'-Dimethylangelicin plus ultraviolet A radiation | 3 | Sup 7 | 1987 | 1987
2303-16-4 | Diallate | 3 | 30, Sup 7 | 1987 | 1987
2318-18-5 | Senkirkine | 3 | 31, Sup 7 | 1987 | 1987
23214-92-8 | Adriamycin (Doxorubicin) | 2A | 10, Sup 7 | 1987 | 1987 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
23246-96-0 | Riddelliine | 2B | 10, Sup 7, 82 | 2002 | 2002
23255-93-8 | Hycanthone mesylate | 3 | 13, Sup 7 | 1987 | 1987
2353-45-9 | Fast Green FCF | 3 | 16, Sup 7 | 1987 | 1987
23537-16-8 | Rugulosin | 3 | 40, Sup 7 | 1987 | 1987
23746-34-1 | Potassium bis(2-hydroxyethyl)dithiocarbamate | 3 | 12, Sup 7 | 1987 | 1987
238-84-6 | Benzo[a]fluorene | 3 | Sup 7, 92 | 2010 | 2005
2385-85-5 | Mirex | 2B | 20, Sup 7 | 1987 | 1987
2386-90-5 | Bis(2,3-epoxycyclopentyl)ether | 3 | 47, 71 | 1999 | 1998
239-35-0 | Benzo[b]naphtho[2,1-d]thiophene | 3 | 103 | 2013 | 2011
2425-06-1 | Captafol | 2A | 53 | 1991 | 1990 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
2425-85-6 | CI Pigment Red 3 | 3 | 57 | 1993 | 1992
2426-08-6 | 1-Butyl glycidyl ether | 2B | 125 | 2020 online | 2019
2429-74-5 | CI Direct Blue 15 | 2B | 57 | 1993 | 1992
243-17-4 | Benzo[b]fluorene | 3 | 92, Sup 7 | 2010 | 2005
2432-99-7 | 11-Aminoundecanoic acid | 3 | 39, Sup 7 | 1987 | 1987
24560-98-3 | cis-9,10-Epoxystearic acid | 3 | 11, Sup 7, 71 | 1999 | 1998
2475-45-8 | Disperse Blue 1 | 2B | 48 | 1990 | 1989
24938-64-5 | para-Aramid fibrils | 3 | 68 | 1997 | 1996
25013-15-4 | Vinyl toluene | 3 | 60 | 1994 | 1994
25013-16-5 | Butylated hydroxyanisole (BHA) | 2B | 40, Sup 7 | 1987 | 1987
25038-54-4 | Nylon 6 | 3 | 19, Sup 7 | 1987 | 1987
25732-74-5 | Acepyrene (3,4-dihydrocyclopenta[cd]pyrene) | 3 | 92 | 2010 | 2005
25812-30-0 | Gemfibrozil | 3 | 66 | 1996 | 1996
25962-77-0 | trans-2-[(Dimethylamino)methylimino]-5-[2-(5-nitro-2-furyl)-vinyl]-1,3,4-oxadiazole | 2B | 7, Sup 7 | 1987 | 1987
2602-46-2 | CI Direct Blue 6 (see Benzidine, dyes metabolized to)
26148-68-5 | A-alpha-C (2-Amino-9H-pyrido[2,3-b]indole) | 2B | 40, Sup 7 | 1987 | 1987
262-12-4 | Dibenzo-para-dioxin | 3 | 69 | 1997 | 1997
26308-28-1 | Ripazepam | 3 | 66 | 1996 | 1996
2646-17-5 | Oil Orange SS | 2B | 8, Sup 7 | 1987 | 1987
26471-62-5 | Toluene diisocyanates | 2B | 39, Sup 7, 71 | 1999 | 1998
26782-43-4 | Hydroxysenkirkine | 3 | 10, Sup 7 | 1987 | 1987
271-89-6 | Benzofuran | 2B | 63 | 1995 | 1995
27208-37-3 | Cyclopenta[cd]pyrene | 2A | Sup 7, 92 | 2010 | 2005 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
2757-90-6 | Agaritine | 3 | 31, Sup 7 | 1987 | 1987
2783-94-0 | Sunset Yellow FCF | 3 | 8, Sup 7 | 1987 | 1987
2784-94-3 | HC Blue No. 1 | 2B | 57 | 1993 | 1992
2832-40-8 | Disperse Yellow 3 | 3 | 48 | 1990 | 1989
2835-39-4 | Allyl isovalerate | 3 | 36, Sup 7, 71 | 1999 | 1998
28407-37-6 | CI Direct Blue 218 | 2B | 129 | 2022 online | 2021
28434-86-8 | 3,3'-Dichloro-4,4'-diaminodiphenyl ether | 2B | 16, Sup 7 | 1987 | 1987
2871-01-4 | HC Red No. 3 | 3 | 57 | 1993 | 1992
29069-24-7 | Prednimustine | 3 | 50 | 1990 | 1989
29291-35-8 | N-Nitrosofolic acid | 3 | 17, Sup 7 | 1987 | 1987
2955-38-6 | Prazepam | 3 | 66 | 1996 | 1996
2973-10-6 | Diisopropyl sulfate | 2B | 54, 71 | 1999 | 1998
29767-20-2 | Teniposide | 2A | 76 | 2000 | 1999 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
298-00-0 | Methyl parathion | 3 | 30, Sup 7 | 1987 | 1987
298-81-7 | Methoxsalen (8-methoxypsoralen) plus ultraviolet A radiation | 1 | 24, Sup 7, 100A | 2012 | 2008
299-75-2 | Treosulfan | 1 | 26, Sup 7, 100A | 2012 | 2008
29975-16-4 | Estazolam | 3 | 66 | 1996 | 1996
3018-12-0 | Dichloroacetonitrile | 3 | 52, 71 | 1999 | 1998
302-01-2 | Hydrazine | 2A | 4, Sup 7, 71, 115 | 2018 | 2016
302-17-0 | Chloral hydrate | 2A | 63, 84, 106 | 2014 | 2012
303-34-4 | Lasiocarpine | 2B | 10, Sup 7 | 1987 | 1987
303-47-9 | Ochratoxin A | 2B | Sup 7, 56 | 1993 | 1992
30310-80-6 | N-Nitrosohydroxyproline | 3 | 17, Sup 7 | 1987 | 1987
305-03-3 | Chlorambucil | 1 | 26, Sup 7, 100A | 2012 | 2008
30516-87-1 | Zidovudine (AZT) | 2B | 76 | 2000 | 1999
3068-88-0 | beta-Butyrolactone | 2B | 11, Sup 7, 71 | 1999 | 1998
308068-56-6 | Carbon nanotubes, multiwalled MWCNT-7 | 2B | 111 | 2017 | 2014
308068-56-6 | Carbon nanotubes, multiwalled, other than MWCNT-7 | 3 | 111 | 2017 | 2014
308068-56-6 | Carbon nanotubes, single-walled | 3 | 111 | 2017 | 2014
308068-56-6 | Multiwalled carbon nanotubes MWCNT-7 (see Carbon nanotubes, multiwalled MWCNT-7)
308068-56-6 | Multiwalled carbon nanotubes other than MWCNT-7 (see Carbon nanotubes, multiwalled, other than MWCNT-7)
308068-56-6 | Single-walled carbon nanotubes (see Carbon nanotubes, single-walled)
308076-74-6 | Silicon carbide, fibrous | 2B | 111 | 2017 | 2014
309-00-2 | Aldrin (see Dieldrin, and aldrin metabolized to dieldrin)
3118-97-6 | Sudan II | 3 | 8, Sup 7 | 1987 | 1987
313-67-7 | Aristolochic acid | 1 | 82, 100A | 2012 | 2008 | NB Overall evaluation upgraded to Group 1 based on mechanistic and other relevant data
313-67-7 | Aristolochic acid, plants containing | 1 | 82, 100A | 2012 | 2008
314-13-6 | Evans blue | 3 | 8, Sup 7 | 1987 | 1987
315-18-4 | Zectran | 3 | 12, Sup 7 | 1987 | 1987
315-22-0 | Monocrotaline | 2B | 10, Sup 7 | 1987 | 1987
3173-72-6 | 1,5-Naphthalene diisocyanate | 3 | 19, Sup 7, 71 | 1999 | 1998
320-67-2 | Azacitidine | 2A | 50 | 1990 | 1989 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
3252-43-5 | Dibromoacetonitrile | 2B | 52, 71, 101 | 2013 | 2011
3296-90-0 | 2,2-Bis(bromomethyl)propane-1,3-diol | 2B | 77 | 2000 | 2000
331-39-5 | Caffeic acid | 2B | 56 | 1993 | 1992
33229-34-4 | HC Blue No. 2 | 3 | 57 | 1993 | 1992
333-41-5 | Diazinon | 2A | 112 | 2017 | 2015 | NB Overall evaluation upgraded to Group 2A based on mechanistic evidence
334-88-3 | Diazomethane | 3 | 7, Sup 7 | 1987 | 1987
33419-42-0 | Etoposide | 1 | 76, 100A | 2012 | 2008 | NB Overall evaluation upgraded to Group 1 based on mechanistic and other relevant data
33419-42-0, 15663-27-1, 11056-06-7 | Etoposide in combination with cisplatin and bleomycin | 1 | 76, 100A | 2012 | 2008
335-67-1 | Perfluorooctanoic acid (PFOA) | 1 | 110, 135 | 2025 online | 2023
3351-28-8 | 1-Methylchrysene | 3 | Sup 7, 92 | 2010 | 2005
3351-30-2 | 4-Methylchrysene | 3 | Sup 7, 92 | 2010 | 2005
3351-31-3 | 3-Methylchrysene | 3 | Sup 7, 92 | 2010 | 2005
3351-32-4 | 2-Methylchrysene | 3 | Sup 7, 92 | 2010 | 2005
33543-31-6 | 2-Methylfluoranthene | 3 | Sup 7, 92 | 2010 | 2005
3564-09-8 | Ponceau 3R | 2B | 8, Sup 7 | 1987 | 1987
3567-69-9 | Carmoisine | 3 | 8, Sup 7 | 1987 | 1987
3570-75-0 | 2-(2-Formylhydrazino)-4-(5-nitro-2-furyl)thiazole | 2B | 7, Sup 7 | 1987 | 1987
366-70-1 | Procarbazine hydrochloride | 2A | 26, Sup 7 | 1987 | 1987 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
3688-53-7 | AF-2 [2-(2-Furyl)-3-(5-nitro-2-furyl)acrylamide] | 2B | 31, Sup 7 | 1987 | 1987
3697-24-3 | 5-Methylchrysene | 2B | Sup 7, 92 | 2010 | 2005
37319-17-8 | Pentosan polysulfate sodium | 2B | 108 | 2016 | 2013
3761-53-3 | Ponceau MX | 2B | 8, Sup 7 | 1987 | 1987
37620-20-5 | N'-Nitrosoanabasine (NAB) | 3 | 37, Sup 7, 89 | 2007 | 2004
3771-19-5 | Nafenopin | 2B | 24, Sup 7 | 1987 | 1987
3778-73-2 | Isophosphamide | 3 | 26, Sup 7 | 1987 | 1987
3795-88-8 | 5-(Morpholinomethyl)-3-[(5-nitrofurfurylidene)­amino]-2-oxazolidinone | 2B | 7, Sup 7 | 1987 | 1987
3844-45-9 | Brilliant Blue FCF, disodium salt | 3 | 16, Sup 7 | 1987 | 1987
38571-73-2 | 1,2,3-Tris(chloromethoxy)propane | 3 | 15, Sup 7, 71 | 1999 | 1998
3902-71-4 | 4,5',8-Trimethylpsoralen | 3 | 40, Sup 7 | 1987 | 1987
396-01-0 | Triamterene | 2B | 108 | 2016 | 2013
4063-41-6 | 4,5'-Dimethylangelicin plus ultraviolet A radiation | 3 | Sup 7 | 1987 | 1987
40762-15-0 | Doxefazepam | 3 | 66 | 1996 | 1996
409-21-2 | Silicon carbide whiskers | 2A | 111 | 2017 | 2014
4170-30-3 | Crotonaldehyde | 2B | 63, 128 | 2021 online | 2020
420-12-2 | Ethylene sulfide | 3 | 11, Sup 7 | 1987 | 1987
42397-64-8 | 1,6-Dinitropyrene | 2B | 46, 105 | 2014 | 2012
42397-65-9 | 1,8-Dinitropyrene | 2B | Sup 7, 46, 105 | 2014 | 2012
4342-03-4 | Dacarbazine | 2B | 26, Sup 7 | 1987 | 1987
439-14-5 | Diazepam | 3 | Sup 7, 66 | 1996 | 1996
443-48-1 | Metronidazole | 2B | 13, Sup 7 | 1987 | 1987
446-86-6 | Azathioprine | 1 | 26, Sup 7, 100A | 2012 | 2008
4548-53-2 | Ponceau SX | 3 | 8, Sup 7 | 1987 | 1987
4549-40-0 | N-Nitrosomethylvinylamine | 2B | 17, Sup 7 | 1987 | 1987
4657-93-6 | 5-Aminoacenaphthene | 3 | 16, Sup 7 | 1987 | 1987
4680-78-8 | Guinea Green B | 3 | 16, Sup 7 | 1987 | 1987
480-54-6 | Retrorsine | 3 | 10, Sup 7 | 1987 | 1987
480-81-9 | Seneciphylline | 3 | 10, Sup 7 | 1987 | 1987
484-20-8 | 5-Methoxypsoralen | 2A | 40, Sup 7 | 1987 | 1987 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
492-17-1 | 2,4'-Diphenyldiamine | 3 | 16, Sup 7 | 1987 | 1987
492-80-8 | Auramine | 2B | 1, Sup 7, 99, 100F | 2012 | 2009
493-52-7 | Methyl red | 3 | 8, Sup 7 | 1987 | 1987
494-03-1 | N,N-Bis(2-chloroethyl)-2-naphthylamine (see Chlornaphazine)
494-03-1 | Chlornaphazine | 1 | 4, Sup 7, 100A | 2012 | 2008
494-38-2 | Acridine orange | 3 | 16, Sup 7 | 1987 | 1987
50-00-0 | Formaldehyde | 1 | Sup 7, 62, 88, 100F | 2012 | 2009
50-06-6 | Phenobarbital | 2B | Sup 7, 79 | 2001 | 2000
50-07-7 | Mitomycin C | 2B | 10, Sup 7 | 1987 | 1987
50-18-0, 6055-19-2 | Cyclophosphamide | 1 | 26, Sup 7, 100A | 2012 | 2008
50-29-3 | DDT (4,4'-dichlorodiphenyltrichloroethane) | 2A | Sup 7, 53, 113 | 2018 | 2015
50-32-8 | Benzo[a]pyrene | 1 | Sup 7, 92, 100F | 2012 | 2009 | NB Overall evaluation upgraded to Group 1 based on mechanistic and other relevant data
50-33-9 | Phenylbutazone | 3 | 13, Sup 7 | 1987 | 1987
50-41-9 | Clomiphene citrate | 3 | 21, Sup 7 | 1987 | 1987
50-44-2 | 6-Mercaptopurine | 3 | 26, Sup 7 | 1987 | 1987
50-55-5 | Reserpine | 3 | 24, Sup 7 | 1987 | 1987
50-76-0 | Actinomycin D | 3 | 10, Sup 7 | 1987 | 1987
501-30-4 | Kojic acid | 3 | 79 | 2001 | 2000
50471-44-8 | Vinclozolin | 2B | 140 | In prep. | 2025
505-60-2 | Mustard gas (see Sulfur mustard)
505-60-2 | Sulfur mustard | 1 | 9, Sup 7, 100F | 2012 | 2009
509-14-8 | Tetranitromethane | 2B | 65 | 1996 | 1995
50926-11-9 | Indium tin oxide | 2B | 118 | 2018 online | 2017
51-02-5 | Pronetalol hydrochloride | 3 | 13, Sup 7 | 1987 | 1987
51-03-6 | Piperonyl butoxide | 3 | 30, Sup 7 | 1987 | 1987
51-18-3 | 2,4,6-Tris(1-aziridinyl)-s-triazine | 3 | 9, Sup 7 | 1987 | 1987
51-21-8 | 5-Fluorouracil | 3 | 26, Sup 7 | 1987 | 1987
51-52-5 | Propylthiouracil | 2B | Sup 7, 79 | 2001 | 2000
51-75-2 | Nitrogen mustard | 2A | 9, Sup 7 | 1987 | 1987
51-79-6 | Ethyl carbamate (Urethane) | 2A | 7, Sup 7, 96 | 2010 | 2007
510-15-6 | Chlorobenzilate | 3 | 30, Sup 7 | 1987 | 1987
51264-14-3 | Amsacrine | 2B | 76 | 2000 | 1999
513-37-1 | 1-Chloro-2-methylpropene | 2B | 63 | 1995 | 1995
5131-60-2 | 4-Chloro-meta-phenylenediamine | 3 | 27, Sup 7 | 1987 | 1987
5141-20-8 | Light Green SF | 3 | 16, Sup 7 | 1987 | 1987
51481-61-9 | Cimetidine | 3 | 50 | 1990 | 1989
5160-02-1 | D & C Red No. 9 | 3 | Sup 7, 57 | 1993 | 1992
51630-58-1 | Fenvalerate | 3 | 53 | 1991 | 1990
518-75-2 | Citrinin | 3 | 40, Sup 7 | 1987 | 1987
52-01-7 | Spironolactone | 3 | Sup 7, 79 | 2001 | 2000
52-24-4 | Thiotepa | 1 | Sup 7, 50, 100A | 2012 | 2008
52-46-0 | Apholate | 3 | 9, Sup 7 | 1987 | 1987
52-68-6 | Trichlorfon | 3 | 30, Sup 7 | 1987 | 1987
520-18-3 | Kaempferol | 3 | 31, Sup 7 | 1987 | 1987
523-44-4 | CI Acid Orange 20 | 3 | 8, Sup 7 | 1987 | 1987
523-44-4 | Orange I (see CI Acid Orange 20)
523-50-2 | Angelicin plus ultraviolet A radiation | 3 | 40, Sup 7 | 1987 | 1987
52645-53-1 | Permethrin | 3 | 53 | 1991 | 1990
52918-63-5 | Deltamethrin | 3 | 53 | 1991 | 1990
53-03-2 | Prednisone | 3 | 26, Sup 7 | 1987 | 1987
53-70-3 | Dibenz[a,h]anthracene | 2A | Sup 7, 92 | 2010 | 2005 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
5307-14-2 | 1,4-Diamino-2-nitrobenzene | 3 | Sup 7, 57 | 1993 | 1992
531-76-0 | Merphalan | 2B | 9, Sup 7 | 1987 | 1987
531-82-8 | N-[4-(5-Nitro-2-furyl)-2-thiazolyl]acetamide | 2B | 7, Sup 7 | 1987 | 1987
532-82-1 | Chrysoidine | 3 | 8, Sup 7 | 1987 | 1987
536-33-4 | Ethionamide | 3 | 13, Sup 7 | 1987 | 1987
5385-75-1 | Dibenzo[a,e]fluoranthene | 3 | Sup 7, 92 | 2010 | 2005
53973-98-1 | Carrageenan, degraded (Poligeenan) | 2B | 31, Sup 7 | 1987 | 1987
53973-98-1 | Poligeenan (see Carrageenan, degraded)
54-05-7 | Chloroquine | 3 | 13, Sup 7 | 1987 | 1987
54-31-9 | Furosemide (Frusemide) | 3 | 50 | 1990 | 1989
54-85-3 | Isonicotinic acid hydrazide (Isoniazid) | 3 | 4, Sup 7 | 1987 | 1987
540-73-8 | 1,2-Dimethylhydrazine | 2A | 4, Sup 7, 71 | 1999 | 1998 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
541-73-1 | meta-Dichlorobenzene | 3 | 73 | 1999 | 1998
542-56-3 | Isobutyl nitrite | 2B | 122 | 2019 | 2018
542-75-6 | 1,3-Dichloropropene (technical-grade) | 2B | 41, Sup 7, 71 | 1999 | 1998
542-78-9 | Malonaldehyde | 3 | 36, Sup 7, 71 | 1999 | 1998
542-88-1, 107-30-2 | Bis(chloromethyl)ether; chloromethyl methyl ether (technical-grade) | 1 | 4, Sup 7, 100F | 2012 | 2009
5431-33-4 | Glycidyl oleate | 3 | 11, Sup 7 | 1987 | 1987
545-06-2 | Trichloroacetonitrile | 3 | 52, 71 | 1999 | 1998
545-55-1 | Tris(1-aziridinyl)phosphine oxide | 3 | 9, Sup 7 | 1987 | 1987
5456-28-0 | Ethyl selenac | 3 | 12, Sup 7 | 1987 | 1987
54749-90-5 | Chlorozotocin | 2A | 50 | 1990 | 1989 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
548-62-9 | Gentian violet (see also Leucogentian violet) | 2B | 129 | 2022 online | 2021
55-18-5 | N-Nitrosodiethylamine | 2A | 17, Sup 7 | 1987 | 1987 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
55-98-1 | 1,4-Butanediol dimethanesulfonate (see Busulfan)
55-98-1 | Busulfan | 1 | 4, Sup 7, 100A | 2012 | 2008
55-98-1 | Myleran (see Busulfan)
551-74-6 | Mannomustine dihydrochloride | 3 | 9, Sup 7 | 1987 | 1987
5522-43-0 | 1-Nitropyrene | 2A | Sup 7, 46, 105 | 2014 | 2012 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
555-84-0 | 1-[(5-Nitrofurfurylidene)amino]-2-imidazolidinone | 2B | 7, Sup 7 | 1987 | 1987
55557-01-2 | N-Nitrosoguvacine | 3 | Sup 7, 85 | 2004 | 2003
55557-02-3 | N-Nitrosoguvacoline | 3 | Sup 7, 85 | 2004 | 2003
556-52-5 | Glycidol | 2A | 77 | 2000 | 2000 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
5589-96-8 | Bromochloroacetic acid | 2B | 101 | 2013 | 2011
56-04-2 | Methylthiouracil | 2B | Sup 7, 79 | 2001 | 2000
56-23-5 | Carbon tetrachloride | 2B | 20, Sup 7, 71 | 1999 | 1998
56-25-7 | Cantharidin | 3 | 10, Sup 7 | 1987 | 1987
56-38-2 | Parathion | 2B | 30, Sup 7, 112 | 2017 | 2015
56-53-1 | Diethylstilbestrol | 1 | 21, Sup 7, 100A | 2012 | 2008
56-55-3 | Benz[a]anthracene | 2B | 92, Sup 7 | 2010 | 2005
56-75-7 | Chloramphenicol | 2A | Sup 7, 50 | 1990 | 1989 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
562-10-7 | Doxylamine succinate | 3 | 79 | 2001 | 2000
563-41-7 | Semicarbazide hydrochloride | 3 | 12, Sup 7 | 1987 | 1987
563-47-3 | 3-Chloro-2-methylpropene, technical grade | 2B | 63, 115 | 2018 | 2016
56894-91-8 | 1,4-Bis(chloromethoxymethyl)benzene | 3 | 15, Sup 7, 71 | 1999 | 1998
569-61-9 | CI Basic Red 9 | 2B | 57, 99 | 2010 | 2008
569-64-2, 14426-28-9, 2437-29-8 | Malachite green (see also Leucomalachite green) | 3 | 129 | 2022 online | 2021
57-06-7 | Allyl isothiocyanate | 3 | 73, Sup 7 | 1999 | 1998
57-14-7 | 1,1-Dimethylhydrazine | 2B | 4, Sup 7, 71 | 1999 | 1998
57-39-6 | Tris(2-methyl-1-aziridinyl)phosphine oxide | 3 | 9, Sup 7 | 1987 | 1987
57-41-0 | Phenytoin | 2B | Sup 7, 66 | 1996 | 1996
57-57-8 | beta-Propiolactone | 2B | 4, Sup 7, 71 | 1999 | 1998
57-68-1 | Sulfamethazine | 3 | 79 | 2001 | 2000 | NB Overall evaluation downgraded to Group 3 with supporting evidence from other relevant data
57-74-9 | Chlordane | 2B | Sup 7, 53, 79 | 2001 | 2000
57-88-5 | Cholesterol | 3 | 31, Sup 7 | 1987 | 1987
57018-52-7 | 1-tert-Butoxypropan-2-ol | 2B | 88, 119 | 2019 | 2017
57117-31-4 | 2,3,4,7,8-Pentachlorodibenzofuran | 1 | 100F | 2012 | 2009 | NB Overall evaluation upgraded to Group 1 based on mechanistic and other relevant data
57465-28-8 | 3,3’,4,4’,5-Pentachlorobiphenyl (PCB-126) | 1 | 100F | 2012 | 2009 | See Polychlorinated biphenyls, dioxin-like, with a TEF according to WHO
57835-92-4 | 4-Nitropyrene | 2B | 46, 105 | 2014 | 2012
58-08-2 | Caffeine | 3 | 51 | 1991 | 1990
58-14-0 | Pyrimethamine | 3 | 13, Sup 7 | 1987 | 1987
58-55-9 | Theophylline | 3 | 51 | 1991 | 1990
58-89-9 | Lindane (see also Hexachlorocyclohexanes) | 1 | 113 | 2018 | 2015
58-93-5 | Hydrochlorothiazide | 1 | 50, 108, 137 | 2026 online | 2024
581-89-5 | 2-Nitronaphthalene | 3 | 46 | 1989 | 1988
59-05-2 | Methotrexate | 3 | 26, Sup 7 | 1987 | 1987
59-87-0 | Nitrofural (Nitrofurazone) | 3 | 50 | 1990 | 1989
59-89-2 | N-Nitrosomorpholine | 2B | 17, Sup 7 | 1987 | 1987
592-62-1 | Methylazoxymethanol acetate | 2B | 10, Sup 7 | 1987 | 1987
59277-89-3 | Aciclovir | 3 | 76 | 2000 | 1999
593-60-2 | Vinyl bromide | 2A | 39, Sup 7, 71, 97 | 2008 | 2007 | NB (1) Overall evaluation upgraded to Group 2A based on mechanistic and other relevant data; (2) For practical purposes, vinyl bromide should be considered to act similarly to the human carcinogen vinyl chloride
593-70-4 | Chlorofluoromethane | 3 | 41, Sup 7, 71 | 1999 | 1998
59536-65-1 | Polybrominated biphenyls | 2A | 41, Sup 7, 107 | 2016 | 2013 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data, namely mechanistic similarity with polychlorinated biphenyls classified in Group 1
598-55-0 | Methyl carbamate | 3 | 12, Sup 7 | 1987 | 1987
59820-43-8 | HC Yellow No. 4 | 3 | 57 | 1993 | 1992
59865-13-3, 79217-60-0 | Ciclosporin | 1 | 50, 100A | 2012 | 2008
5989-27-5 | d-Limonene | 3 | 56, 73 | 1999 | 1998 | NB Overall evaluation downgraded to Group 3 with supporting evidence from other relevant data
599-79-1 | Sulfasalazine | 2B | 108 | 2016 | 2013
60-09-3 | para-Aminoazobenzene | 2B | 8, Sup 7 | 1987 | 1987
60-11-7 | para-Dimethylaminoazobenzene | 2B | 8, Sup 7 | 1987 | 1987
60-35-5 | Acetamide | 2B | 7, Sup 7, 71 | 1999 | 1998
60-56-0 | Methimazole | 3 | 79 | 2001 | 2000
60-57-1 | Dieldrin (see Dieldrin, and aldrin metabolized to dieldrin)
60-57-1, 309-00-2 | Dieldrin, and aldrin metabolized to dieldrin | 2A | 5, Sup 7, 117 | 2019 | 2016
60102-37-6 | Petasitenine | 3 | 31, Sup 7 | 1987 | 1987
60153-49-3 | 3-(N-Nitrosomethylamino)propionitrile | 2B | Sup 7, 85 | 2004 | 2003
602-60-8 | 9-Nitroanthracene | 3 | 33, Sup 7 | 1987 | 1987
602-87-9 | 5-Nitroacenaphthene | 2B | 16, Sup 7 | 1987 | 1987
603-48-5 | Leucogentian violet (see also Gentian violet) | 3 | 129 | 2022 online | 2021
604-75-1 | Oxazepam | 2B | Sup 7, 66 | 1996 | 1996
606-20-2 | 2,6-Dinitrotoluene | 2B | 65 | 1996 | 1995
607-57-8 | 2-Nitrofluorene | 2B | 46, 105 | 2014 | 2012
609-20-1 | 2,6-Dichloro-para-phenylenediamine | 3 | 39, Sup 7 | 1987 | 1987
61-57-4 | Niridazole | 2B | 13, Sup 7 | 1987 | 1987
61-82-5 | Amitrole | 3 | 79, Sup 7 | 2001 | 2000 | NB Overall evaluation downgraded to Group 3 with supporting evidence from other relevant data
611-06-3 | 2,4-Dichloro-1-nitrobenzene | 2B | 123 | 2020 online | 2018
613-35-4 | N,N'-Diacetylbenzidine | 2B | 16, Sup 7 | 1987 | 1987
615-05-4 | 2,4-Diaminoanisole | 2B | Sup 7, 79 | 2001 | 2000
615-28-1 | ortho-Phenylenediamine dihydrochloride | 2B | 123 | 2020 online | 2018
615-53-2 | N-Methyl-N-nitrosourethane | 2B | 4, Sup 7 | 1987 | 1987
6164-98-3 | Chlordimeform | 3 | 30, Sup 7 | 1987 | 1987
618-85-9 | 3,5-Dinitrotoluene | 3 | 65 | 1996 | 1995
62-44-2 | Phenacetin | 1 | 24, Sup 7, 100A | 2012 | 2008 | NB Overall evaluation upgraded to Group 1 with supporting evidence from other relevant data
62-50-0 | Ethyl methanesulfonate | 2B | 7, Sup 7 | 1987 | 1987
62-53-3 | Aniline (see also Aniline hydrochloride) | 2A | 27, Sup 7, 127 | 2021 online | 2020
62-55-5 | Thioacetamide | 2B | 7, Sup 7 | 1987 | 1987
62-56-6 | Thiourea | 3 | Sup 7, 79 | 2001 | 2000
62-73-7 | Dichlorvos | 2B | Sup 7, 53 | 1991 | 1990
62-75-9 | N-Nitrosodimethylamine | 2A | 17, Sup 7 | 1987 | 1987 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
621-64-7 | N-Nitrosodi-n-propylamine | 2B | 17, Sup 7 | 1987 | 1987
62450-06-0 | Trp-P-1 (3-Amino-1,4-dimethyl-5H-pyrido[4,3-b]indole) | 2B | 31, Sup 7 | 1987 | 1987
62450-07-1 | Trp-P-2 (3-Amino-1-methyl-5H-pyrido[4,3-b]indole) | 2B | 31, Sup 7 | 1987 | 1987
627-12-3 | n-Propyl carbamate | 3 | 12, Sup 7 | 1987 | 1987
63-25-2 | Carbaryl | 3 | 12, Sup 7 | 1987 | 1987
63-92-3 | Phenoxybenzamine hydrochloride | 2B | 24, Sup 7 | 1987 | 1987
630-20-6 | 1,1,1,2-Tetrachloroethane | 2B | 41, Sup 7, 71, 106 | 2014 | 2012
63041-90-7 | 6-Nitrobenzo[a]pyrene | 3 | Sup 7, 46 | 1989 | 1988
631-64-1 | Dibromoacetic acid | 2B | 101 | 2013 | 2011
632-99-5 | Magenta | 2B | Sup 7, 57, 99, 100F | 2012 | 2009
6358-53-8 | Citrus Red No. 2 | 2B | 8, Sup 7 | 1987 | 1987
6368-72-5 | Sudan Red 7B | 3 | 8, Sup 7 | 1987 | 1987
637-07-0 | Clofibrate | 3 | Sup 7, 66 | 1996 | 1996
637-92-3 | Ethyl tert-butyl ether | 2B | 138 | In prep. | 2025
6373-74-6 | CI Acid Orange 3 | 3 | 57 | 1993 | 1992
64-17-5 | Ethanol in alcoholic beverages | 1 | 96, 100E | 2012 | 2009
64-67-5 | Diethyl sulfate | 2A | 54, 71 | 1999 | 1998 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
641-48-5 | 1,2-Dihydroaceanthrylene | 3 | 92 | 2010 | 2005
6416-57-5 | Sudan Brown RR | 3 | 8, Sup 7 | 1987 | 1987
64436-13-1 | Arsenobetaine and other organic arsenic compounds that are not metabolized in humans | 3 | 100C | 2012 | 2009
6459-94-5 | CI Acid Red 114 | 2B | 57 | 1993 | 1992
64742-93-4 | Bitumens, occupational exposure to oxidized bitumens and their emissions during roofing | 2A | 103 | 2013 | 2011
65271-80-9 | Mitoxantrone | 2B | 76 | 2000 | 1999
65996-93-2 | Coal-tar pitch | 1 | 35, Sup 7, 100F | 2012 | 2009
66-27-3 | Methyl methanesulfonate | 2A | 7, Sup 7, 71 | 1999 | 1998 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
66-75-1 | Uracil mustard | 2B | 9, Sup 7 | 1987 | 1987
66733-21-9 | Erionite | 1 | 42, Sup 7, 100C | 2012 | 2009
67-20-9 | Nitrofurantoin | 3 | 50 | 1990 | 1989
67-45-8 | Furazolidone | 3 | 31, Sup 7 | 1987 | 1987
67-63-0 | Isopropyl alcohol | 3 | 15, Sup 7, 71 | 1999 | 1998
67-66-3 | Chloroform | 2B | Sup 7, 73 | 1999 | 1998
67-72-1 | Hexachloroethane | 2B | 73 | 1999 | 1998
67730-10-3 | Glu-P-2 (2-Aminodipyrido[1,2-a:3',2'-d]imidazole) | 2B | 40, Sup 7 | 1987 | 1987
67730-11-4 | Glu-P-1 (2-Amino-6-methyldipyrido[1,2-a:3',2'-d]imidazole) | 2B | 40, Sup 7 | 1987 | 1987
68-12-2 | N,N-Dimethylformamide | 2A | 47, 71, 115 | 2018 | 2016
68-76-8 | Tris(aziridinyl)-para-benzoquinone (Triaziquone) | 3 | 9, Sup 7 | 1987 | 1987
680-31-9 | Hexamethylphosphoramide | 2B | 15, Sup 7, 71 | 1999 | 1998
68006-83-7 | MeA-alpha-C (2-Amino-3-methyl-9H-pyrido[2,3-b]indole) | 2B | 40, Sup 7 | 1987 | 1987
68308-34-9 | Shale oils | 1 | 35, Sup 7, 100F | 2012 | 2009
684-93-5 | N-Methyl-N-nitrosourea | 2A | 17, Sup 7 | 1987 | 1987 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
68603-42-9 | Coconut oil diethanolamine condensate | 2B | 101 | 2013 | 2011
6870-67-3 | Jacobine | 3 | 10, Sup 7 | 1987 | 1987
69-53-4 | Ampicillin | 3 | 50 | 1990 | 1989
693-98-1 | 2-Methylimidazole | 2B | 101 | 2013 | 2011
69655-05-6 | Didanosine | 3 | 76 | 2000 | 1999
70-25-7 | N-Methyl-N´-nitro-N-nitrosoguanidine (MNNG) | 2A | 4, Sup 7 | 1987 | 1987 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
70-30-4 | Hexachlorophene | 3 | 20, Sup 7 | 1987 | 1987
7099-43-6 | 5,6-Cyclopenteno-1,2-benzanthracene | 3 | 92 | 2010 | 2005
71-43-2 | Benzene | 1 | 29, Sup 7, 100F, 120 | 2018 | 2017
71-55-6 | 1,1,1-Trichloroethane | 2A | 20, Sup 7, 71, 130 | 2022 online | 2021
71-58-9 | Medroxyprogesterone acetate | 2B | 21, Sup 7 | 1987 | 1998
712-68-5 | 2-Amino-5-(5-nitro-2-furyl)-1,3,4-thiadiazole | 2B | 7, Sup 7 | 1987 | 1987
71267-22-6 | N'-Nitrosoanatabine (NAT) | 3 | 37, Sup 7, 89 | 2007 | 2004
72-20-8 | Endrin | 3 | 5, Sup 7 | 1987 | 1987
72-43-5 | Methoxychlor | 3 | 20, Sup 7 | 1987 | 1987
72-57-1 | Trypan blue | 2B | 8, Sup 7 | 1987 | 1987
7220-79-3 | Methylene blue | 3 | 108 | 2016 | 2013
723-46-6 | Sulfamethoxazole | 3 | Sup 7, 79 | 2001 | 2000
73459-03-7 | 5-Methylangelicin plus ultraviolet A radiation | 3 | Sup 7 | 1987 | 1987
74-83-9 | Methyl bromide | 3 | 41, Sup 7, 71 | 1999 | 1998
74-85-1 | Ethylene | 3 | Sup 7, 60 | 1994 | 1998
74-87-3 | Methyl chloride | 3 | 41, Sup 7, 71 | 1999 | 1998
74-88-4 | Methyl iodide | 3 | 41, Sup 7, 71 | 1999 | 1998
74-96-4 | Bromoethane | 3 | 52, 71 | 1999 | 1998
7439-92-1 | Lead | 2B | 23, Sup 7 | 1987 | 1987
7439-97-6 | Mercury and inorganic mercury compounds | 3 | 58 | 1993 | 1993
7440-02-0 | Nickel, metallic | 2B | Sup 7, 49 | 1990 | 1989
7440-07-5 | Plutonium | 1 | 78, 100D | 2012 | 2009
7440-29-1 | Thorium-232 and its decay products | 1 | 78, 100D | 2012 | 2009
7440-38-2 | Arsenic and inorganic arsenic compounds | 1 | 23, Sup 7, 100C | 2012 | 2009
7440-41-7 | Beryllium and beryllium compounds | 1 | Sup 7, 58, 100C | 2012 | 2009
7440-43-9 | Cadmium and cadmium compounds | 1 | 58, 100C | 2012 | 2009
7440-47-3 | Chromium, metallic | 3 | Sup 7, 49 | 1990 | 1989
7440-48-4 | Cobalt metal (without tungsten carbide or other metal alloys) | 2A | 52, 86, 131 | 2023 online | 2022
7440-48-4, 12070-12-1 | Cobalt metal with tungsten carbide | 2A | 86 | 2006 | 2003
7440-61-1 | Uranium, mixture of isotopes (see Radionuclides, alpha-particle-emitting, internally deposited) | 1 | 78, 100D | 2012 | 2009 | NB Uranium was listed as a Group 1 agent as a member of the family, Radionuclides, alpha-particle-emitting, internally deposited. There was limited evidence for the carcinogenicity of mixtures of uranium isotopes, with positive associations for lung cancer
7446-09-5 | Sulfur dioxide | 3 | 54 | 1992 | 1991
7460-84-6 | Glycidyl stearate | 3 | 11, Sup 7 | 1987 | 1987
7481-89-2 | Zalcitabine | 2B | 76 | 2000 | 1999
7496-02-8 | 6-Nitrochrysene | 2A | Sup 7, 46, 105 | 2014 | 2012 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
75-00-3 | Chloroethane | 3 | 52, 71 | 1999 | 1998
75-01-4 | Vinyl chloride | 1 | Sup 7, 97, 100F | 2012 | 2009
75-02-5 | Vinyl fluoride | 2A | Sup 7, 63, 97 | 2008 | 2007 | NB (1) Overall evaluation upgraded to Group 2A based on mechanistic and other relevant data; (2) For practical purposes, vinyl fluoride should be considered to act similarly to the human carcinogen vinyl chloride
75-07-0 | Acetaldehyde | 2B | 36, Sup 7, 71 | 1999 | 1998
75-07-0 | Acetaldehyde associated with consumption of alcoholic beverages | 1 | 100E | 2012 | 2009
75-09-2 | Dichloromethane (Methylene chloride) | 2A | Sup 7, 71, 110 | 2017 | 2014
75-21-8 | Ethylene oxide | 1 | Sup 7, 60, 97, 100F | 2012 | 2009 | NB Overall evaluation upgraded to Group 1 based on mechanistic and other relevant data
75-25-2 | Bromoform | 3 | 52, 71 | 1999 | 1998
75-26-3 | 2-Bromopropane | 2A | 133 | 2024 online | 2023
75-27-4 | Bromodichloromethane | 2B | 52, 71 | 1999 | 1998
75-35-4 | Vinylidene chloride | 2B | 39, Sup 7, 71, 119 | 2019 | 2017
75-38-7 | Vinylidene fluoride | 3 | 39, Sup 7, 71 | 1999 | 1998
75-45-6 | Chlorodifluoromethane | 3 | 41, Sup 7, 71 | 1999 | 1998
75-52-5 | Nitromethane | 2B | 77 | 2000 | 2000
75-55-8 | 2-Methylaziridine (Propyleneimine) | 2B | 9, Sup 7, 71 | 1999 | 1998
75-56-9 | Propylene oxide | 2B | Sup 7, 60 | 1994 | 1994
75-60-5 | Dimethylarsinic acid | 2B | 100C | 2012 | 2009
75-65-0 | tert-Butyl alcohol | 3 | 138 | In prep. | 2025
75-87-6 | Chloral | 2A | 63, 84, 106 | 2014 | 2012
75-88-7 | 2-Chloro-1,1,1-trifluoroethane | 3 | 41, Sup 7, 71 | 1999 | 1998
7519-36-0 | N-Nitrosoproline | 3 | 17, Sup 7 | 1987 | 1987
75321-20-9 | 1,3-Dinitropyrene | 2B | 46, 105 | 2014 | 2012
7572-29-4 | Dichloroacetylene | 3 | 39, Sup 7, 71 | 1999 | 1998
759-73-9 | N-Ethyl-N-nitrosourea | 2A | 17, Sup 7 | 1987 | 1987 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
76-01-7 | Pentachloroethane | 3 | 41, Sup 7, 71 | 1999 | 1998
76-03-9 | Trichloroacetic acid | 2B | 63, 84, 106 | 2014 | 2012
76-44-8 | Heptachlor | 2B | Sup 7, 53, 79 | 2001 | 2000
76180-96-6 | IQ (2-Amino-3-methylimidazo[4,5-f]quinoline) | 2A | Sup 7, 56 | 1993 | 1992 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
7631-86-9 | Silica, amorphous | 3 | Sup 7, 68 | 1997 | 1996
7647-01-0 | Hydrochloric acid | 3 | 54 | 1992 | 1991
765-34-4 | Glycidaldehyde | 2B | 11, Sup 7, 71 | 1999 | 1998
7664-93-9 | Strong-inorganic-acid mists containing sulfuric acid (see Acid mists)
77-09-8 | Phenolphthalein | 2B | 76 | 2000 | 1999
77-78-1 | Dimethyl sulfate | 2A | 4, Sup 7, 71 | 1999 | 1998 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
77094-11-2 | MeIQ (2-Amino-3,4-dimethylimidazo[4,5-f]quinoline) | 2B | Sup 7, 56 | 1993 | 1992
7722-84-1 | Hydrogen peroxide | 3 | 36, Sup 7, 71 | 1999 | 1998
77439-76-0 | 3-Chloro-4-(dichloromethyl)-5-hydroxy-2(5H)-furanone | 2B | 84 | 2004 | 2002
77500-04-0 | MeIQx (2-Amino-3,8-dimethylimidazo[4,5-f]quinoxaline) | 2B | Sup 7, 56 | 1993 | 1992
7758-01-2 | Potassium bromate | 2B | Sup 7, 73 | 1999 | 1998
7758-19-2 | Sodium chlorite | 3 | 52 | 1991 | 1990
7782-49-2 | Selenium and selenium compounds | 3 | 9, Sup 7 | 1987 | 1987
78-59-1 | Isophorone | 2B | 130 | 2022 online | 2021
78-79-5 | Isoprene | 2B | 60, 71 | 1999 | 1998
78-87-5 | 1,2-Dichloropropane | 1 | 41, Sup 7, 71, 110 | 2017 | 2014
78-98-8 | Methylglyoxal | 3 | 51 | 1991 | 1990
789-07-1 | 2-Nitropyrene | 3 | 46 | 1989 | 1988
79-00-5 | 1,1,2-Trichloroethane | 3 | 52, 71 | 1999 | 1998
79-01-6 | Trichloroethylene | 1 | Sup 7, 63, 106 | 2014 | 2012
79-06-1 | Acrylamide | 2A | 60, Sup 7 | 1994 | 1994 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
79-10-7 | Acrylic acid | 3 | 19, Sup 7, 71 | 1999 | 1998
79-34-5 | 1,1,2,2-Tetrachloroethane | 2B | 20, Sup 7, 71, 106 | 2014 | 2012
79-43-6 | Dichloroacetic acid | 2B | 63, 84, 106 | 2014 | 2012
79-44-7 | Dimethylcarbamoyl chloride | 2A | 12, Sup 7, 71 | 1999 | 1998 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
79-46-9 | 2-Nitropropane | 2B | 29, Sup 7, 71 | 1999 | 1998
79-94-7 | Tetrabromobisphenol A | 2A | 115 | 2018 | 2016 | NB Overall evaluation upgraded to Group 2A with supporting evidence from mechanistic data
794-93-4 | Dihydroxymethylfuratrizine (see also Panfuran S) | 3 | 24, Sup 7 | 1987 | 1987
794-93-4 | Panfuran S (containing dihydroxymethylfuratrizine) | 2B | 24, Sup 7 | 1987 | 1987
80-08-0 | Dapsone | 3 | 24, Sup 7 | 1987 | 1987
80-15-9 | Cumyl hydroperoxide | 2B | 141 | In prep. | 2026
80-62-6 | Methyl methacrylate | 3 | Sup 7, 60 | 1994 | 1994
800-24-8 | Aziridyl benzoquinone | 3 | 9, Sup 7 | 1987 | 1987
8001-35-2 | Toxaphene (Polychlorinated camphenes) | 2B | Sup 7, 79 | 2001 | 2000
8001-50-1 | Terpene polychlorinates (Strobane®) | 3 | 5, Sup 7 | 1987 | 1987
8001-58-9 | Creosotes | 2A | Sup 7, 92 | 2010 | 2005
8002-05-9 | Crude oil | 3 | 45 | 1989 | 1988
8007-45-2 | Coal tars (see Coal-tar distillation) |  | 35, Sup 7 | 1987 | 1987
8007-45-2 | Coal-tar distillation | 1 | 92, 100F | 2012 | 2009
8018-07-3 | Acriflavinium chloride | 3 | 13, Sup 7 | 1987 | 1987
804-36-4 | Nitrovin | 3 | 31, Sup 7 | 1987 | 1987
8047-67-4 | Saccharated iron oxide | 3 | 2, Sup 7 | 1987 | 1987
8052-42-4 | Bitumens, extracts of steam-refined and air-refined; steam-refined, cracking-residue and air-refined bitumens (see Bitumens, occupational exposures)
8052-42-4, 64741-56-6 | Bitumens, occupational exposure to straight-run bitumens and their emissions during road paving | 2B | 103 | 2013 | 2011
81-07-2 | Saccharin and its salts | 3 | Sup 7, 73 | 1999 | 1998 | NB Overall evaluation downgraded to Group 3 with supporting evidence from other relevant data
81-15-2 | Musk xylene | 3 | 65 | 1996 | 1995
81-49-2 | 1-Amino-2,4-dibromoanthraquinone | 2B | 101 | 2013 | 2011
81-88-9 | Rhodamine B | 3 | 16, Sup 7 | 1987 | 1987
817-09-4 | Trichlormethine (Trimustine hydrochloride) | 2B | Sup 7, 50 | 1990 | 1989
82-28-0 | 1-Amino-2-methylanthraquinone | 3 | 27, Sup 7 | 1987 | 1987
82-68-8 | Quintozene (Pentachloronitrobenzene) | 3 | 5, Sup 7 | 1987 | 1987
822-36-6 | 4-Methylimidazole | 2B | 101 | 2013 | 2011
82413-20-5 | Droloxifene | 3 | 66 | 1996 | 1996
828-00-2 | Dimethoxane | 3 | 15, Sup 7 | 1987 | 1987
83-32-9 | Acenaphthene | 3 | 92 | 2010 | 2005
83-63-6 | Diacetylaminoazotoluene | 3 | 8, Sup 7 | 1987 | 1987
83-66-9 | Musk ambrette | 3 | 65 | 1996 | 1995
83-67-0 | Theobromine | 3 | 51 | 1991 | 1990
832-69-9 | 1-Methylphenanthrene | 3 | Sup 7, 92 | 2010 | 2005
83463-62-1 | Bromochloroacetonitrile | 3 | 52, 71 | 1999 | 1998
838-88-0 | 4,4'-Methylene bis(2-methylaniline) | 2B | 4, Sup 7 | 1987 | 1987
84-65-1 | Anthraquinone | 2B | 101 | 2013 | 2011
842-07-9 | Sudan I | 3 | 8, Sup 7 | 1987 | 1987
846-50-4 | Temazepam | 3 | 66 | 1996 | 1996
85-01-8 | Phenanthrene | 3 | Sup 7, 92 | 2010 | 2005
85-68-7 | Butyl benzyl phthalate | 3 | Sup 7, 73 | 1999 | 1998
85-83-6 | Scarlet Red | 3 | 8, Sup 7 | 1987 | 1987
85-84-7 | Yellow AB | 3 | 8, Sup 7 | 1987 | 1987
85-86-9 | Sudan III | 3 | 8, Sup 7 | 1987 | 1987
85502-23-4 | 3-(N-Nitrosomethylamino)propionaldehyde | 3 | Sup 7, 85 | 2004 | 2003
85878-62-2 | Pyrido[3,4-c]psoralen | 3 | 40, Sup 7 | 1987 | 1987
85878-63-3 | 7-Methylpyrido[3,4-c]psoralen | 3 | 40, Sup 7 | 1987 | 1987
86-30-6 | N-Nitrosodiphenylamine | 3 | 27, Sup 7 | 1987 | 1987
86-54-4 | Hydralazine | 3 | 24, Sup 7 | 1987 | 1987
86-57-7 | 1-Nitronaphthalene | 3 | 46 | 1989 | 1988
86-73-7 | Fluorene | 3 | Sup 7, 92 | 2010 | 2005
86-74-8 | Carbazole | 2B | 32, Sup 7, 71, 103 | 2013 | 2011
86-88-4 | 1-Naphthylthiourea (ANTU) | 3 | 30, Sup 7 | 1987 | 1987
868-85-9 | Dimethyl hydrogen phosphite | 2B | 48, 71, 133 | 2024 online | 2023
87-29-6 | Cinnamyl anthranilate | 3 | Sup 7, 77 | 2000 | 2000
87-62-7 | 2,6-Dimethylaniline (2,6-Xylidine) | 2B | 57 | 1993 | 1992
87-68-3 | Hexachlorobutadiene | 3 | 73 | 1999 | 1998
87-86-5 | Pentachlorophenol (see also Polychlorophenols) | 1 | 53, 71, 117 | 2019 | 2016
87625-62-5 | Ptaquiloside | 3 | 40, Sup 7 | 1987 | 1987
88-05-1 | 2,4,6-Trimethylaniline | 3 | 27, Sup 7 | 1987 | 1987
88-06-2 | 2,4,6-Trichlorophenol (see also Polychlorophenols) | 2B | 117 | 2019 | 2016
88-12-0 | N-Vinyl-2-pyrrolidone | 3 | 19, Sup 7, 71 | 1999 | 1998
88-72-2 | 2-Nitrotoluene | 2A | 101 | 2013 | 2011 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
88-73-3 | 2-Chloronitrobenzene | 2B | 65, 123 | 2020 online | 2018
88-73-3, 121-73-3, 100-00-5 | Chloronitrobenzenes (see 2-Chloronitrobenzene, 3-Chloronitrobenzene, 4-Chloronitrobenzene)
89-61-2 | 1,4-Dichloro-2-nitrobenzene | 2B | 65, 123 | 2020 online | 2018
89-82-7 | Pulegone | 2B | 108 | 2016 | 2013
892-21-7 | 3-Nitrofluoranthene | 3 | 33, Sup 7 | 1987 | 1987
89778-26-7 | Toremifene | 3 | 66 | 1996 | 1996
90-04-0 | ortho-Anisidine (see also ortho-Anisidine hydrochloride) | 2A | Sup 7, 73, 127 | 2021 online | 2020
90-43-7 | ortho-Phenylphenol | 3 | 73 | 1999 | 1998
90-65-3 | Penicillic acid | 3 | 10, Sup 7 | 1987 | 1987
90-94-8 | Michler’s ketone [4,4´-Bis(dimethylamino)­benzophenone] | 2B | 99 | 2010 | 2008
9000-07-1 | Carrageenan, native | 3 | 31, Sup 7 | 1987 | 1987
9000-38-8 | Kava extract | 2B | 108 | 2016 | 2013
9002-84-0 | Polytetrafluoroethylene | 3 | 19, Sup 7 | 1987 | 1987
9002-86-2 | Polyvinyl chloride | 3 | 19, Sup 7 | 1987 | 1987
9002-88-4 | Polyethylene | 3 | 19, Sup 7 | 1987 | 1987
9002-89-5 | Polyvinyl alcohol | 3 | 19, Sup 7 | 1987 | 1987
9003-01-4 | Polyacrylic acid | 3 | 19, Sup 7 | 1987 | 1987
9003-07-0 | Polypropylene | 3 | 19, Sup 7 | 1987 | 1987
9003-20-7 | Polyvinyl acetate | 3 | 19, Sup 7 | 1987 | 1987
9003-22-9 | Vinyl chloride-vinyl acetate copolymers | 3 | 19, Sup 7 | 1987 | 1987
9003-39-8 | Polyvinyl pyrrolidone | 3 | 19, Sup 7, 71 | 1999 | 1998
9003-53-6 | Polystyrene | 3 | 19, Sup 7 | 1987 | 1987
9003-54-7 | Styrene-acrylonitrile copolymers | 3 | 19, Sup 7 | 1987 | 1987
9003-55-8 | Styrene-butadiene copolymers | 3 | 19, Sup 7 | 1987 | 1987
9004-51-7 | Iron-dextrin complex | 3 | 2, Sup 7 | 1987 | 1987
9004-66-4 | Iron-dextran complex | 2B | 2, Sup 7 | 1987 | 1987
90045-36-6 | Ginkgo biloba extract | 2B | 108 | 2016 | 2013
9009-54-5 | Polyurethane foams | 3 | 19, Sup 7 | 1987 | 1987 | NB This CAS No. has been deleted.
9010-98-4 | Polychloroprene | 3 | 19, Sup 7 | 1987 | 1987
9011-06-7 | Vinylidene chloride-vinyl chloride copolymers | 3 | 19, Sup 7 | 1987 | 1987
9011-14-7 | Polymethyl methacrylate | 3 | 19, Sup 7 | 1987 | 1987
9016-87-9 | Polymethylene polyphenyl isocyanate | 3 | 19, Sup 7 | 1987 | 1987
90370-29-9 | 4,4',6-Trimethylangelicin plus ultraviolet A radiation | 3 | Sup 7 | 1987 | 1987
91-20-3 | Naphthalene | 2B | 82 | 2002 | 2002
91-22-5 | Quinoline | 2B | 121 | 2019 | 2018
91-23-6 | ortho-Nitroanisole | 2A | 65, 127 | 2021 online | 2020 | NB Originally evaluated as 2-Nitroanisole
91-59-8 | 2-Naphthylamine | 1 | 4, Sup 7, 99, 100F | 2012 | 2009
91-64-5 | Coumarin | 3 | Sup 7, 77 | 2000 | 2000
91-93-0 | 3,3'-Dimethoxybenzidine-4,4'-diisocyanate | 3 | 39, Sup 7 | 1987 | 1987
91-94-1 | 3,3'-Dichlorobenzidine | 2B | 29, Sup 7 | 1987 | 1987
915-67-3 | Amaranth | 3 | 8, Sup 7 | 1987 | 1987
92-67-1 | 4-Aminobiphenyl | 1 | 1, Sup 7, 99, 100F | 2012 | 2009
92-87-5 | Benzidine | 1 | 29, Sup 7, 99, 100F | 2012 | 2009
92-93-3 | 4-Nitrobiphenyl | 3 | 4, Sup 7 | 1987 | 1987
924-16-3 | N-Nitrosodi-n-butylamine | 2B | 17, Sup 7 | 1987 | 1987
924-42-5 | N-Methylolacrylamide | 2B | 60, 130 | 2022 online | 2021
93-15-2 | Methyleugenol | 2A | 101, 134 | 2024 online | 2023
930-55-2 | N-Nitrosopyrrolidine | 2B | 17, Sup 7 | 1987 | 1987
94-36-0 | Benzoyl peroxide | 3 | 36, Sup 7, 71 | 1999 | 1998
94-58-6 | Dihydrosafrole | 2B | 10, Sup 7 | 1987 | 1987
94-59-7 | Safrole | 2B | 10, Sup 7 | 1987 | 1987
94-75-7 | 2,4-D (2,4-dichlorophenoxyacetic acid) (See also Chlorophenoxy herbicides) | 2B | 113 | 2018 | 2015
95-06-7 | Sulfallate | 2B | 30, Sup 7 | 1987 | 1987
95-50-1 | ortho-Dichlorobenzene | 3 | Sup 7, 73 | 1999 | 1998
95-53-4 | ortho-Toluidine | 1 | Sup 7, 77, 99, 100F | 2012 | 2009
95-54-5 | ortho-Phenylenediamine | 2B | 123 | 2020 online | 2018
95-68-1 | 2,4-Xylidine | 3 | 16, Sup 7 | 1987 | 1987
95-69-2 | 4-Chloro-ortho-toluidine | 2A | 77, 99 | 2010 | 2008
95-70-5 | 2,5-Diaminotoluene | 3 | 16, Sup 7 | 1987 | 1987
95-78-3 | 2,5-Xylidine | 3 | 16, Sup 7 | 1987 | 1987
95-79-4 | 5-Chloro-ortho-toluidine | 3 | 77 | 2000 | 2000
95-80-7 | 2,4-Diaminotoluene | 2B | 16, Sup 7 | 1987 | 1987
95-83-0 | 4-Chloro-ortho-phenylenediamine | 2B | 27, Sup 7 | 1987 | 1987
95-85-2 | 2-Amino-4-chlorophenol | 2B | 123 | 2020 online | 2018
96-09-3 | Styrene-7,8-oxide | 2A | Sup 7, 60, 121 | 2019 | 2018 | NB Overall evaluation upgraded to Group 2A with supporting evidence from other relevant data
96-12-8 | 1,2-Dibromo-3-chloropropane | 2B | 20, Sup 7, 71 | 1999 | 1998
96-13-9 | 2,3-Dibromopropan-1-ol | 2B | 77 | 2000 | 2000
96-18-4 | 1,2,3-Trichloropropane | 2A | 63 | 1995 | 1995
96-23-1 | 1,3-Dichloro-2-propanol | 2B | 101 | 2013 | 2011
96-24-2 | 3-Monochloro-1,2-propanediol | 2B | 101 | 2013 | 2011
96-33-3 | Methyl acrylate | 2B | 39, Sup 7, 71, 122 | 2019 | 2018
96-45-7 | Ethylenethiourea | 3 | Sup 7, 79 | 2001 | 2000 | NB Overall evaluation downgraded to Group 3 with supporting evidence from other relevant data
96-48-0 | gamma-Butyrolactone | 3 | 11, Sup 7, 71 | 1999 | 1998
97-53-0 | Eugenol | 3 | 36, Sup 7 | 1987 | 1987
97-54-1 | Isoeugenol | 2B | 134 | 2024 online | 2023
97-56-3 | ortho-Aminoazotoluene | 2B | 8, Sup 7 | 1987 | 1987
97-77-8 | Disulfiram | 3 | 12, Sup 7 | 1987 | 1987
97-88-1 | Butyl methacrylate | 2B | 133 | 2024 online | 2023
98-00-0 | Furfuryl alcohol | 2B | 119 | 2019 | 2017
98-01-1 | Furfural | 3 | 63 | 1995 | 1995
98-56-6 | 4-Chlorobenzotrifluoride | 2B | 125 | 2020 online | 2019
98-82-8 | Cumene | 2B | 101 | 2013 | 2011
98-83-9 | α-Methylstyrene | 2B | 101 | 2013 | 2011
98-87-3, 98-07-7, 100-44-7, 98-88-4 | alpha-Chlorinated toluenes (benzal chloride, benzotrichloride, benzyl chloride) and benzoyl chloride (combined exposures) | 2A | 29, Sup 7, 71 | 1999 | 1998
98-95-3 | Nitrobenzene | 2B | 65 | 1996 | 1995
989-38-8 | Rhodamine 6G | 3 | 16, Sup 7 | 1987 | 1987
99-08-1, 99-99-0 | Nitrotoluenes | 3 | 65 | 1996 | 1995
99-55-8 | 5-Nitro-ortho-toluidine | 3 | 48 | 1990 | 1989
99-56-9 | 1,2-Diamino-4-nitrobenzene | 3 | 16, Sup 7 | 1987 | 1987
99-57-0 | 2-Amino-4-nitrophenol | 3 | 57 | 1993 | 1992
99-59-2 | 5-Nitro-ortho-anisidine | 3 | 27, Sup 7 | 1987 | 1987
99-80-9 | N-Methyl-N,4-dinitrosoaniline | 3 | 1, Sup 7 | 1987 | 1987
99-97-8 | N,N-Dimethyl-p-toluidine | 2B | 115 | 2018 | 2016
994-05-8 | tert-Amyl methyl ether | 3 | 138 | In prep. | 2025

## ED Assessment (ECHA)
Endocrine disruptor assessment
Note: Group entries are split in different rows.
https://echa.europa.eu/ed-assessment
Last table update: 2026-05-04
Substance name | DISLIST_ED_description | EC / List no | CAS no | Authority | Status | Outcome | Latest update | First published
Terephthalic acid |  | 202-830-0 | 100-21-0 | France | Concluded | not ED | 11-Dec-2015 | 03-Oct-2018
Tributyl O-acetylcitrate |  | 201-067-0 | 77-90-7 | France | Concluded | not ED | 19-Dec-2016 | 12-Oct-2018
p-cresol |  | 203-398-6 | 106-44-5 | United Kingdom | Concluded | not ED | 07-Mar-2017 | 12-Oct-2018
Thiram |  | 205-286-2 | 137-26-8 | Sweden | Concluded | not ED | 24-Oct-2018 | 12-Oct-2018
1,3,4,6,7,8-hexahydro-4,6,6,7,8,8-hexamethylindeno[5,6-c]pyran |  | 214-946-9 | 1222-05-5 | France | Under development | Under development (other) | 03-Feb-2021 | 24-Oct-2018
2,2',6,6'-tetra-tert-butyl-4,4'-methylenediphenol |  | 204-279-1 | 118-82-1 | Austria | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 12-Oct-2018
2,2',6,6'-tetrabromo-4,4'-isopropylidenediphenol |  | 201-236-9 | 79-94-7 | Denmark | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 12-Oct-2018
2,4-di-tert-butylphenol |  | 202-532-0 | 96-76-4 | Belgium | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 12-Oct-2018
2,6-di-tert-butyl-p-cresol |  | 204-881-4 | 128-37-0 | France | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 24-Oct-2018
3-methylpyrazole |  | 215-925-7 | 1453-58-3 | Belgium | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 24-Oct-2018
4,4'-isopropylidenediphenol |  | 201-245-8 | 80-05-7 | France | Concluded | ED HH | 03-Feb-2021 | 12-Oct-2018
4,4'-methylenedi-2,6-xylenol |  | 226-378-9 | 5384-21-4 | France | Under development | Under development (other) | 03-Feb-2021 | 12-Oct-2018
6,6'-di-tert-butyl-4,4'-butylidenedi-m-cresol |  | 201-618-5 | 85-60-9 | France | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 06-Nov-2019
6,6'-di-tert-butyl-4,4'-thiodi-m-cresol |  | 202-525-2 | 96-69-5 | Austria | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 12-Oct-2018
alphachloralose |  | 240-016-7 | 15879-93-3 | Poland | Under development | Under development (BPR) | 03-Feb-2021 | 06-Nov-2019
Copper |  | 231-159-6 | 7440-50-8 | France | Under development | Under development (BPR) | 03-Feb-2021 | 20-Oct-2020
Cyclohexylhydroxydiazene 1-oxide, potassium salt (This entry also covers the hydrated forms of K-HDO) |  | 613-953-8 | 66603-10-9 | Austria | Under development | Under development (BPR) | 03-Feb-2021 | 14-May-2019
Dicyclohexyl phthalate |  | 201-545-9 | 84-61-7 | Sweden | Concluded | ED HH | 03-Feb-2021 | 04-Oct-2018
Diethyl phthalate |  | 201-550-6 | 84-66-2 | ECHA | Under development | Under development (other) | 03-Feb-2021 | 07-Jan-2021
Methyl 4-hydroxybenzoate |  | 202-785-7 | 99-76-3 | France | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 12-Oct-2018
Peracetic acid |  | 201-186-8 | 79-21-0 | Austria | Under development | Under development (BPR) | 03-Feb-2021 | 24-Sep-2019
resorcinol; 1,3-benzenediol |  | 203-585-2 | 108-46-3 | France | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 25-Oct-2018
resorcinol; 1,3-benzenediol |  | 203-585-2 | 108-46-3 | Finland | Concluded | inconclusive | 03-Feb-2021 | 12-Oct-2018
resorcinol; 1,3-benzenediol |  | 203-585-2 | 108-46-3 | France | Under development | Under development (other) | 03-Feb-2021 | 06-Nov-2019
Silver copper zeolite |  | 868-573-7 | 130328-19-7 | Sweden | Under development | Under development (BPR) | 03-Feb-2021 | 08-Mar-2019
Silver sodium zirconium hydrogenphosphate | ALPHASAN ANTIMICROBIAL RC 2000; ALPHASAN ANTIMICROBIAL RC 5000; NOVARON; SILVER SODIUM HYDROGEN ZIRCONIUM PHOSPHATE | 422-570-3 | 265647-11-8 | Sweden | Under development | Under development (BPR) | 03-Feb-2021 | 12-Oct-2018
Triclosan |  | 222-182-2 | 3380-34-5 | Denmark | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 06-Nov-2019
Tris(2-ethylhexyl) benzene-1,2,4-tricarboxylate |  | 222-020-0 | 3319-31-1 | Austria | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 25-Jul-2019
Ziram |  | 205-288-3 | 137-30-4 | Denmark | Under development under SEV | Under development (SEV) | 03-Feb-2021 | 12-Oct-2018
2,2-dibromo-2-cyanoacetamide |  | 233-539-7 | 10222-01-2 | Denmark | Concluded | ED ENV#ED HH | 08-Mar-2021 | 08-Mar-2019
Cyanamide |  | 206-992-3 | 420-04-2 | Germany | Concluded | ED ENV#ED HH | 08-Mar-2021 | 14-May-2019
Silver |  | 231-131-3 | 7440-22-4 | Sweden | Under development | Under development (BPR) | 02-Jul-2021 | 02-Jul-2021
dichloromethane; methylene chloride |  | 200-838-9 | 75-09-2 | Italy | Concluded | postponed | 03-Sep-2021 | 12-Oct-2018
4,4'-(1-methylpropylidene)bisphenol; bisphenol B |  | 201-025-1 | 77-40-7 | France | Concluded | ED ENV#ED HH | 27-Sep-2021 | 24-Oct-2018
butyl 4-hydroxybenzoate |  | 202-318-7 | 94-26-8 | Denmark | Concluded | ED HH | 27-Sep-2021 | 19-Sep-2019
Phenol, alkylation products (mainly in para position) with C12-rich branched alkyl chains from oligomerisation, covering any individual isomers and/ or combinations thereof (PDDP) |  | 799-972-3 | - | Germany | Concluded | ED ENV#ED HH | 27-Sep-2021 | 19-Oct-2020
4-isododecylphenol |  | - | 27459-10-5 | Germany | Concluded | ED ENV#ED HH | 27-Sep-2021 | 19-Oct-2020
Phenol, tetrapropylene- |  | - | 57427-55-1 | Germany | Concluded | ED ENV#ED HH | 27-Sep-2021 | 19-Oct-2020
Phenol, 4-dodecyl, branched |  | 640-104-9 | 210555-94-5 | Germany | Concluded | ED ENV#ED HH | 27-Sep-2021 | 19-Oct-2020
Phenol, (tetrapropenyl) derivatives |  | 616-100-8 | 74499-35-7 | Germany | Concluded | ED ENV#ED HH | 27-Sep-2021 | 19-Oct-2020
Phenol, 4-isododecyl- |  | 608-055-8 | 27147-75-7 | Germany | Concluded | ED ENV#ED HH | 27-Sep-2021 | 19-Oct-2020
Phenol, dodecyl-, branched |  | 310-154-3 | 121158-58-5 | Germany | Concluded | ED ENV#ED HH | 27-Sep-2021 | 19-Oct-2020
Chloramide |  | 234-217-9 | 10599-90-3 | Sweden | Under development | Under development (BPR) | 22-Oct-2021 | 22-Oct-2021
2-methyl-1-(4-methylthiophenyl)-2-morpholinopropan-1-one | ACETOCURE 97; GENOCURE*PMP; IGM 4817; IRGACURE 907; SPEEDCURE 97 | 400-600-6 | 71868-10-5 | Austria | Concluded | postponed | 05-Nov-2021 | 14-May-2019
Nonylphenol, branched, ethoxylated | 1 - 2.5 moles ethoxylated | 500-209-1 | 68412-54-4 | United Kingdom | Concluded | postponed | 05-Nov-2021 | 05-Nov-2021
Isopentyl p-methoxycinnamate |  | 275-702-5 | 71617-10-2 | Germany | Under development under SEV | Under development (SEV) | 08-Dec-2021 | 12-Oct-2018
Melamine |  | 203-615-4 | 108-78-1 | France | Under development | Under development (other) | 10-Feb-2022 | 10-Feb-2022
Sodium fluoride |  | 231-667-8 | 7681-49-4 | France | Under development | Under development (other) | 10-Feb-2022 | 10-Feb-2022
2-(4-tert-butylbenzyl)propionaldehyde |  | 201-289-8 | 80-54-6 | Sweden | Concluded | inconclusive | 10-Mar-2022 | 12-Oct-2018
tert-butyl methyl ether |  | 216-653-1 | 1634-04-4 | France | Concluded | inconclusive | 10-Mar-2022 | 12-Oct-2018
2,2'-[(1-methylethylidene)bis(4,1-phenyleneoxymethylene)]bisoxirane |  | 216-823-5 | 1675-54-3 | Denmark | Concluded | inconclusive | 15-Mar-2022 | 12-Oct-2018
Carbon disulphide |  | 200-843-6 | 75-15-0 | France | Concluded | not ED | 18-May-2022 | 21-Sep-2021
(±)-1,7,7-trimethyl-3-[(4-methylphenyl)methylene]bicyclo[2.2.1]heptan-2-one covering any of the individual isomers and/or combinations thereof (4-MBC) |  | - | - | Denmark | Concluded | ED HH | 04-Aug-2022 | 20-Oct-2020
(±)-1,7,7-trimethyl-3-[(4-methylphenyl)methylene]bicyclo[2.2.1]heptan-2-one |  | 253-242-6 | 36861-47-9 | Denmark | Concluded | ED HH | 04-Aug-2022 | 20-Oct-2020
(3E)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | 701-394-3 | 1782069-81-1 | Denmark | Concluded | ED HH | 04-Aug-2022 | 20-Oct-2020
(1R,3E,4S)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | - | 95342-41-9 | Denmark | Concluded | ED HH | 04-Aug-2022 | 20-Oct-2020
(1S,3E,4R)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | - | 852541-30-1 | Denmark | Concluded | ED HH | 04-Aug-2022 | 20-Oct-2020
(1R,3Z,4S)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | - | 852541-21-0 | Denmark | Concluded | ED HH | 04-Aug-2022 | 20-Oct-2020
(1R,4S)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | - | 741687-98-9 | Denmark | Concluded | ED HH | 04-Aug-2022 | 20-Oct-2020
(1S,3Z,4R)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | - | 852541-25-4 | Denmark | Concluded | ED HH | 04-Aug-2022 | 20-Oct-2020
diuron (ISO); 3-(3,4-dichlorophenyl)-1,1-dimethylurea |  | 206-354-4 | 330-54-1 | Denmark | Under development | Under development (BPR) | 20-Oct-2022 | 20-Oct-2022
1,2,4-triazole |  | 206-022-9 | 288-88-0 | Belgium | Concluded | postponed | 23-Jan-2023 | 12-Oct-2018
4,4'-(1,3-phenylene-bis(1-methylethylidene))bisphenol |  | 428-970-4 | 13595-25-0 | Belgium | Concluded | postponed | 23-Jan-2023 | 12-Oct-2018
4,4'-Isopropylidenediphenol, propoxylated | 1 - 4.5 moles propoxylated | 500-097-4 | 37353-75-6 | Denmark | Concluded | postponed | 23-Jan-2023 | 19-Sep-2019
Propargite |  | 219-006-1 | 2312-35-8 | Netherlands | Concluded | postponed | 23-Jan-2023 | 19-Sep-2019
Bronopol |  | 200-143-0 | 52-51-7 | Spain | Under development | Under development (BPR) | 15-Mar-2023 | 27-Feb-2023
Phenol, styrenated |  | 262-975-0 | 61788-44-1 | Germany | Concluded | ED ENV | 20-Mar-2023 | 12-Oct-2018
Reaction mass of 2,6-bis(1-phenylethyl)phenol and 2,4,6-tris(1-phenylethyl)phenol |  | 701-171-0 | - | Germany | Concluded | ED ENV | 20-Mar-2023 | 23-Mar-2022
Reaction products of phosphoryl trichloride and 2-methyloxirane |  | 807-935-0 | 1244733-77-4 | Denmark | Under development under SEV | Under development (SEV) | 03-Apr-2023 | 03-Apr-2023
4,4'-sulphonyldiphenol |  | 201-250-5 | 80-09-1 | Belgium | Concluded | ED ENV#ED HH | 04-May-2023 | 12-Oct-2018
Triphenyl phosphate |  | 204-112-2 | 115-86-6 | France | Concluded | ED ENV | 12-Jun-2023 | 12-Oct-2018
Butanoic acid, 4-amino-4-oxosulfo-, N-coco alkyl derivs., monosodium salts, compds. with triethanolamine |  | 308-662-5 | 98171-53-0 | France | Under development under SEV | Under development (SEV) | 19-Jul-2023 | 19-Jul-2023
Dioctyltin oxide |  | 212-791-1 | 870-08-6 | Austria | Under development under SEV | Under development (SEV) | 19-Jul-2023 | 19-Jul-2023
Propyl 4-hydroxybenzoate |  | 202-307-7 | 94-13-3 | Belgium | Concluded | ED ENV | 01-Sep-2023 | 12-Oct-2018
Salicylic acid |  | 200-712-3 | 69-72-7 | Netherlands | Under development | Under development (BPR) | 15-Sep-2023 | 15-Sep-2023
2-Propenoic acid, methyl ester, reaction products with mixed O,O-bis(branched and linear pentyl and iso-Bu) phosphorodithioates |  | 300-340-2 | 93925-38-3 | Sweden | Under development under SEV | Under development (SEV) | 25-Sep-2023 | 25-Sep-2023
Dicyclohexyl phthalate |  | 201-545-9 | 84-61-7 | Sweden | Concluded | ED ENV | 04-Oct-2023 | 03-Nov-2020
4,4'-methylenedi-2,6-xylenol |  | 226-378-9 | 5384-21-4 | Germany | Under development | Under development (other) | 16-Oct-2023 | 16-Oct-2023
2-ethylhexyl (2E)-3-(4-methoxyphenyl)acrylate |  | 629-661-9 | 83834-59-7 | Germany | Concluded | inconclusive | 23-Nov-2023 | 12-Oct-2018
3-(3-biphenyl-4-yl-1,2,3,4-tetrahydro-1-naphthyl)-4-hydroxycoumarin |  | 259-978-4 | 56073-07-5 | Finland | Under development | Under development (BPR) | 16-Jan-2024 | 16-Jan-2024
Diethylmethylbenzenediamine |  | 270-877-4 | 68479-98-1 | Denmark | Concluded | inconclusive | 22-Feb-2024 | 12-Oct-2018
p-(1,1-dimethylpropyl)phenol |  | 201-280-9 | 80-46-6 | Germany | Concluded | ED ENV#inconclusive | 19-Mar-2024 | 04-Oct-2018
chlorfenapyr (ISO); 4-bromo-2-(4-chlorophenyl)-1- ethoxymethyl-5-trifluoromethylpyrrole-3-carbonitrile |  | 602-782-4 | 122453-73-0 | Portugal | Under development | Under development (BPR) | 04-Apr-2024 | 19-Mar-2024
Mono- and di-ortho-phthalate esters with at least one of the following:-a linear or branched alkyl moiety with a C4-C6 backbone* with possible substitution by methyl and/or ethyl groups;-a benzyl moiety with methylene possibly substituted by methyl and/or ethyl groups;-a cyclopentyl or cyclohexyl moiety with possible substitution by methyl and/or ethyl groups.[*The backbone is defined as the longest linear carbon chain from the ester function] |  | - | - | France | Under development | Under development (other) | 18-Apr-2024 | 18-Apr-2024
Oligomerisation and alkylation reaction products of 2-phenylpropene and phenol |  | 700-960-7 | - | Denmark | Concluded | ED ENV | 04-Jun-2024 | 25-Oct-2018
Phenol, methylstyrenated |  | 270-966-8 | 68512-30-1 | Denmark | Concluded | ED ENV | 04-Jun-2024 | 25-Oct-2018
Formic acid |  | 200-579-1 | 64-18-6 | Belgium | Concluded | not ED | 06-Jun-2024 | 14-May-2019
diuron (ISO); 3-(3,4-dichlorophenyl)-1,1-dimethylurea |  | 206-354-4 | 330-54-1 | Finland | Concluded | ED ENV#inconclusive | 05-Jul-2024 | 12-Oct-2018
2,2'-[(1-methylethylidene)bis(4,1-phenyleneoxy)]bisethyl diacetate |  | 242-895-2 | 19224-29-4 | Denmark | Under development under SEV | Under development (SEV) | 16-Aug-2024 | 16-Aug-2024
2,4,6-tribromophenol |  | 204-278-6 | 118-79-6 | France | Under development under SEV | Under development (SEV) | 16-Aug-2024 | 16-Aug-2024
3-aminopropyldiethylamine |  | 203-236-4 | 104-78-9 | Austria | Under development | Under development (other) | 16-Aug-2024 | 19-Mar-2024
4,4'-isopropylidenedi-o-cresol |  | 201-240-0 | 79-97-0 | Germany | Under development under SEV | Under development (SEV) | 16-Aug-2024 | 21-Sep-2021
(1-methylethylidene)bis(4,1-phenyleneoxy-3,1-propanediyl) bismethacrylate |  | 248-607-1 | 27689-12-9 | Denmark | Under development under SEV | Under development (SEV) | 05-Sep-2024 | 05-Sep-2024
4,4'-isopropylidenebis[2-allylphenol] |  | 217-121-1 | 1745-89-7 | Denmark | Under development under SEV | Under development (SEV) | 05-Sep-2024 | 05-Sep-2024
[1,3(or 1,4)-phenylenebis(1-methylethylidene)]bis[tert-butyl] peroxide |  | 246-678-3 | 25155-25-3 | Netherlands | Under development under SEV | Under development (SEV) | 05-Sep-2024 | 14-Aug-2023
Ethyl N-acetyl-N-butyl-β-alaninate |  | 257-835-0 | 52304-36-6 | Belgium | Under development | Under development (BPR) | 14-Oct-2024 | 14-Oct-2024
2-tert-butyl-4-methoxyphenol |  | 204-442-7 | 121-00-6 | France | Concluded | ED ENV#ED HH | 08-Nov-2024 | 12-Oct-2018
Chrysanthemum cinerariaefolium, ext. | Extractives and their physically modified derivatives such as tinctures, concretes, absolutes, essential oils, oleoresins, terpenes, terpene-free fractions, distillates, residues, etc., obtained from Chrysanthemum cinerariaefolium, Compositae. | 289-699-3 | 89997-63-7 | Spain | Concluded | not ED#inconclusive | 13-Jan-2025 | 14-May-2019
Chrysanthemum cinerariaefolium, ext. | Extractives and their physically modified derivatives such as tinctures, concretes, absolutes, essential oils, oleoresins, terpenes, terpene-free fractions, distillates, residues, etc., obtained from Chrysanthemum cinerariaefolium, Compositae. | 289-699-3 | 89997-63-7 | Spain | Concluded | not ED#inconclusive | 13-Jan-2025 | 31-Jul-2019
Climbazole |  | 253-775-4 | 38083-17-9 | Belgium | Concluded | ED ENV 1 | 13-Jan-2025 | 12-Oct-2018
dinotefuran (ISO); 1-methyl-2-nitro-3-(tetrahydro-3-furylmethyl)guanidine |  | 605-399-0 | 165252-70-0 | Belgium | Concluded | not ED | 13-Jan-2025 | 10-Feb-2022
Ethylene oxide |  | 200-849-9 | 75-21-8 | Norway | Concluded | not ED#inconclusive | 13-Jan-2025 | 25-Oct-2018
Ozone |  | 233-069-2 | 10028-15-6 | Germany | Concluded | not ED | 13-Jan-2025 | 14-May-2019
Silver zinc zeolite |  | 603-404-0 | 130328-20-0 | Sweden | Concluded | inconclusive | 13-Jan-2025 | 12-Oct-2018
Sulphur dioxide |  | 231-195-2 | 7446-09-5 | Germany | Concluded | not ED | 13-Jan-2025 | 14-May-2019
Sodium perfluoroheptanoate |  | 243-518-4 | 20109-59-5 | Belgium | Concluded | inconclusive | 21-Jan-2025 | 20-Jan-2021
α-cyano-3-phenoxybenzyl 3-(2,2-dichlorovinyl)-2,2-dimethylcyclopropanecarboxylate |  | 257-842-9 | 52315-07-8 |  | Under development | Under development (BPR) | 21-Jan-2025 | 19-Mar-2024
Sodium chlorate |  | 231-887-4 | 7775-09-9 | Austria | Under development | Under development (other) | 20-Feb-2025 | 20-Feb-2025
imidacloprid (ISO); 1-(6-chloropyridin-3-ylmethyl)-N-nitroimidazolidin-2-ylidenamine |  | 428-040-8 | 138261-41-3 | Germany | Under development | Under development (BPR) | 25-Feb-2025 | 25-Feb-2025
Tributyl citrate |  | 201-071-2 | 77-94-1 | France | Concluded | inconclusive | 02-Jun-2025 | 12-Oct-2018
3,5,5-trimethylcyclohex-2-enone |  | 201-126-0 | 78-59-1 | France | Concluded | inconclusive | 26-Jun-2025 | 12-Oct-2018
Methyl salicylate |  | 204-317-7 | 119-36-8 | France | Concluded | inconclusive | 26-Jun-2025 | 12-Oct-2018
Reaction mass of 3-[(diphenoxyphosphoryl)oxy]phenyl triphenyl 1,3-phenylene bis(phosphate) and tetraphenyl 1,3-phenylene bis(phosphate) |  | 701-337-2 | - | France | Concluded | inconclusive | 26-Jun-2025 | 12-Oct-2018
1,2-benzisothiazol-3(2H)-one |  | 220-120-9 | 2634-33-5 | Spain | Under development | Under development (BPR) | 14-Aug-2025 | 14-Aug-2025
Didecyldimethylammonium chloride |  | 230-525-2 | 7173-51-5 | Italy | Under development | Under development (BPR) | 14-Aug-2025 | 14-Aug-2025
Quaternary ammonium compounds, benzyl-C12-16-alkyldimethyl, chlorides |  | 270-325-2 | 68424-85-1 | Italy | Under development | Under development (BPR) | 14-Aug-2025 | 14-Aug-2025
2-phenoxyethanol |  | 204-589-7 | 122-99-6 | Italy | Under development | Under development (BPR) | 21-Aug-2025 | 21-Aug-2025
3-phenoxybenzyl-2-(4-ethoxyphenyl)-2-methylpropyl ether | ETHOFENPROX; TREBON | 407-980-2 | 80844-07-1 | Austria | Concluded | inconclusive | 23-Oct-2025 | 14-May-2019
6,6'-di-tert-butyl-4,4'-thiodi-m-cresol |  | 202-525-2 | 96-69-5 | Austria | Under development | Under development (SEV) | 25-Nov-2025 | 25-Nov-2025
Dioctyltin oxide |  | 212-791-1 | 870-08-6 | Austria | Under development | Under development (other) | 25-Nov-2025 | 25-Nov-2025
perchloric acid and its salts |  | - | - | Germany | Under development | Under development (other) | 25-Nov-2025 | 25-Nov-2025
Potassium dicyanoargentate |  | 208-047-0 | 506-61-6 | Netherlands | Under development | Under development (SEV) | 25-Nov-2025 | 25-Nov-2025
Triethyl phosphate |  | 201-114-5 | 78-40-0 | France | Under development | Under development (SEV) | 25-Nov-2025 | 25-Nov-2025
Tris(2-ethylhexyl) phosphate |  | 201-116-6 | 78-42-2 | France | Under development | Under development (SEV) | 25-Nov-2025 | 25-Nov-2025
3,5,5-trimethylcyclohex-2-enone |  | 201-126-0 | 78-59-1 | France | Under development | Under development (other) | 18-Dec-2025 | 18-Dec-2025
4,4'-isopropylidenedi-2,6-xylol |  | 227-033-5 | 5613-46-7 | Denmark | Concluded | ED ENV#ED HH | 18-Dec-2025 | 05-Sep-2024
Geraniol |  | 203-377-1 | 106-24-1 | France | Concluded | inconclusive | 18-Dec-2025 | 16-Sep-2022
Oxybenzone |  | 205-031-5 | 131-57-7 | Denmark | Concluded | ED ENV#ED HH | 18-Dec-2025 | 12-Oct-2018
Isobutyl 4-hydroxybenzoate |  | 224-208-8 | 4247-02-3 | Denmark | Concluded | ED HH | 01-Mar-2026 | 29-Apr-2022
Medetomidine |  | 811-718-6 | 86347-14-0 | Norway | Concluded | ED ENV#ED HH | 01-Mar-2026 | 10-Mar-2023
(1,3,4,5,6,7-hexahydro-1,3-dioxo-2H-isoindol-2-yl)methyl (1R-trans)-2,2-dimethyl-3-(2-methylprop-1-enyl)cyclopropanecarboxylate |  | 214-619-0 | 1166-46-7 | Germany | Under development | Under development (BPR) | 27-Apr-2026 | 09-Mar-2020
(3E)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one |  | 701-394-3 | 1782069-81-1 | Germany | Concluded | inconclusive | 27-Apr-2026 | 12-Oct-2018
1,1'-(isopropylidene)bis[3,5-dibromo-4-(2,3-dibromo-2-methylpropoxy)benzene] |  | 306-832-3 | 97416-84-7 | Germany | Concluded | postponed | 27-Apr-2026 | 12-Oct-2018
1,1'-(isopropylidene)bis[3,5-dibromo-4-(2,3-dibromopropoxy)benzene] |  | 244-617-5 | 21850-44-2 | Germany | Concluded | postponed | 27-Apr-2026 | 12-Oct-2018
1,3-dihydro-4(or 5)-methyl-2H-benzimidazole-2-thione |  | 258-904-8 | 53988-10-6 | Germany | Concluded | ED ENV | 27-Apr-2026 | 12-Oct-2018
1,3-dihydro-4(or 5)-methyl-2H-benzimidazole-2-thione, zinc salt |  | 262-872-0 | 61617-00-3 | Germany | Concluded | ED ENV | 27-Apr-2026 | 12-Oct-2018
1,7,7-trimethyl-3-(phenylmethylene)bicyclo[2.2.1]heptan-2-one |  | 239-139-9 | 15087-24-8 | Germany | Concluded | ED ENV | 27-Apr-2026 | 12-Oct-2018
1-(5,6,7,8-tetrahydro-3,5,5,6,8,8-hexamethyl-2-naphthyl)ethan-1-one | EC No. 216-133-4 and EC No. 244-240-6 | - | - | Germany | Under development under SEV | Under development (SEV) | 27-Apr-2026 | 06-Nov-2019
1-(5,6,7,8-tetrahydro-3,5,5,6,8,8-hexamethyl-2-naphthyl)ethan-1-one |  | 244-240-6 | 21145-77-7 | Germany | Under development under SEV | Under development (SEV) | 27-Apr-2026 | 06-Nov-2019
1-(5,6,7,8-tetrahydro-3,5,5,6,8,8-hexamethyl-2-naphthyl)ethan-1-one |  | 216-133-4 | 1506-02-1 | Germany | Under development under SEV | Under development (SEV) | 27-Apr-2026 | 06-Nov-2019
1-[2-(allyloxy)-2-(2,4-dichlorophenyl)ethyl]-1H-imidazole |  | 252-615-0 | 35554-44-0 | Germany | Concluded | postponed | 27-Apr-2026 | 12-Oct-2018
1-[4-(1,1-dimethylethyl)phenyl]-3-(4-methoxyphenyl)propane-1,3-dione |  | 274-581-6 | 70356-09-1 | Germany | Under development under SEV | Under development (SEV) | 27-Apr-2026 | 02-Aug-2022
1-[[2-(2,4-dichlorophenyl)-4-propyl-1,3-dioxolan-2-yl]methyl]-1H-1,2,4-triazole |  | 262-104-4 | 60207-90-1 | Finland | Concluded | ED ENV#ED HH | 27-Apr-2026 | 06-Nov-2019
2,2'-dimethyl-4,4'-methylenebis(cyclohexylamine) |  | 229-962-1 | 6864-37-5 | Germany | Concluded | not ED | 27-Apr-2026 | 12-Oct-2018
2-bromo-3,3,3-trifluoroprop-1-ene |  | 627-872-0 | 1514-82-5 | Spain | Concluded | postponed | 27-Apr-2026 | 06-Nov-2019
3,3,4,4,5,5,6,6,7,7,8,8,8-tridecafluorooctyl acrylate |  | 241-527-8 | 17527-29-6 | Germany | Concluded | ED ENV | 27-Apr-2026 | 12-Oct-2018
3,3,4,4,5,5,6,6,7,7,8,8,8-tridecafluorooctyl methacrylate |  | 218-407-9 | 2144-53-8 | Germany | Concluded | ED ENV | 27-Apr-2026 | 12-Oct-2018
3-iodo-2-propynyl butylcarbamate |  | 259-627-5 | 55406-53-6 | Denmark | Under development | Under development (BPR) | 27-Apr-2026 | 27-Aug-2020
4,4'-isopropylidenediphenol |  | 201-245-8 | 80-05-7 | Germany | Concluded | ED ENV | 27-Apr-2026 | 12-Oct-2018
4-heptylphenol, branched and linear | substances with a linear and/or branched alkyl chain with a carbon number of 7 covalently bound predominantly in position 4 to phenol, covering also UVCB- and well-defined substances which include any of the individual isomers or a combination thereof | - | - | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
Phenol, heptyl derivs. |  | 276-743-1 | 72624-02-3 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-heptylphenol |  | 217-862-0 | 1987-50-4 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(3-ethylpentan-3-yl)phenol |  | - | 37872-24-5 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(2-methylhexan-2-yl)phenol |  | - | 30784-31-7 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(3,3-dimethylpentan-2-yl)phenol |  | - | 911371-06-7 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(3-methylhexan-2-yl)phenol |  | - | 854904-93-1 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(4,4-dimethylpentan-2-yl)phenol |  | - | 911371-07-8 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(4-methylhexan-2-yl)phenol |  | - | 71945-81-8 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(5-methylhexan-2-yl)phenol |  | - | 857629-71-1 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(2,2-dimethylpentan-3-yl)phenol |  | - | 861010-65-3 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(5-methylhexan-3-yl)phenol |  | - | 854904-92-0 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(heptan-3-yl)phenol |  | - | 6465-74-3 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(heptan-2-yl)phenol |  | - | 6863-24-7 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(heptan-4-yl)phenol |  | - | 6465-71-0 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(3-ethylpentyl)phenol |  | - | 911370-98-4 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(3-methylhexyl)phenol |  | - | 102570-52-5 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(4-methylhexyl)phenol |  | - | 1139800-98-8 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(5-methylhexyl)phenol |  | - | 100532-36-3 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(2,4-dimethylpentan-3-yl)phenol |  | - | 1824346-00-0 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
Phenol, 4-(1-ethyl-1,2-dimethylpropyl)- |  | - | 30784-27-1 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
Phenol, 4-tert-heptyl- |  | - | 288864-02-8 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(2,3-dimethylpentan-2-yl)phenol |  | 854-135-2 | 861011-60-1 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(3-methylhexan-3-yl)phenol |  | 854-958-7 | 30784-32-8 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(2,3,3-trimethylbutan-2-yl)phenol |  | - | 72861-06-4 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-(2,4-dimethylpentan-2-yl)phenol |  | - | 33104-11-9 | Austria | Concluded | ED ENV | 27-Apr-2026 | 04-Oct-2018
4-tert-butylphenol |  | 202-679-0 | 98-54-4 | Germany | Concluded | ED ENV#inconclusive | 27-Apr-2026 | 04-Oct-2018
Ammonium perchlorate |  | 232-235-1 | 7790-98-9 | Germany | Concluded | ED ENV | 27-Apr-2026 | 12-Oct-2018
Benzimidazole-2-thiol |  | 209-502-6 | 583-39-1 | Germany | Under development | Under development (other) | 27-Apr-2026 | 27-Apr-2026
Benzotriazole |  | 202-394-1 | 95-14-7 | Germany | Concluded | ED ENV | 27-Apr-2026 | 12-Oct-2018
Bis(2-propylheptyl) phthalate |  | 258-469-4 | 53306-54-0 | Germany | Under development under SEV | Under development (SEV) | 27-Apr-2026 | 23-Jul-2020
Dapsone |  | 201-248-4 | 80-08-0 | Germany | Concluded | postponed | 27-Apr-2026 | 12-Oct-2018
Nitrobenzene |  | 202-716-0 | 98-95-3 | Austria | Concluded | inconclusive | 27-Apr-2026 | 12-Oct-2018
Phenol, dodecyl-, branched |  | 310-154-3 | 121158-58-5 | Germany | Concluded | ED ENV#ED HH | 27-Apr-2026 | 12-Oct-2018
Reaction mass of p-t-butylphenyldiphenyl phosphate and bis(p-t-butylphenyl) phenyl phosphate |  | 939-505-4 | - | France | Under development under SEV | Under development (SEV) | 27-Apr-2026 | 15-May-2024
Reaction products of paraformaldehyde and 2-hydroxypropylamine (ratio 3:2) |  | - | - | Austria | Concluded | inconclusive | 27-Apr-2026 | 02-Jun-2021
reaction products of paraformaldehyde with 2-hydroxypropylamine (ratio 1:1);[formaldehyde released from α,α,α-trimethyl-1,3,5-triazine-1,3,5(2H,4H,6H)-triethanol];[HPT] |  | - | - | Austria | Concluded | inconclusive | 27-Apr-2026 | 02-Jun-2021
Sodium perchlorate |  | 231-511-9 | 7601-89-0 | Germany | Concluded | ED ENV | 27-Apr-2026 | 12-Oct-2018
Terbutryn |  | 212-950-5 | 886-50-0 | Slovakia | Under development | Under development (BPR) | 27-Apr-2026 | 26-Nov-2020
Tetramethrin |  | 231-711-6 | 7696-12-0 | Germany | Under development | Under development (BPR) | 27-Apr-2026 | 09-Mar-2020

## TEDX
TEDX — The Endocrine Disruption Exchange — Potential Endocrine Disruptors
https://endocrinedisruption.org/interactive-tools/tedx-list-of-potential-endocrine-disruptors/search-the-tedx-list#sname=&searchfor=any&sortby=chemname&action=search&searchcats=all&sortby=chemname
⚠️ REGULATORY STATUS: TEDX is NOT an EU Official list. Findings → 🟠 MAJOR only (suspected — not confirmed ED).
Last updated September 2018
Chemical Name | Alternative Names | CAS Nos. | Endocrine Disruptor References | Categories of Use | Use References | Date Added | Date Updated
(+)-mono-1-tert-butyl-3-methylbutyl-phthalate |  | 109591-02-8 | Lampen A, Zimnik S, Nau H. 2003. Teratogenic phthalate esters and metabolites activate the nuclear receptors PPARs and induce differentiation of F9 cells. Toxicol Appl Pharmacol 188(1):14-23. DOI: 10.1016/S0041-008X(03)00014-0. URL: https://www.sciencedirect.com/science/article/pii/S0041008X03000140 |  | Corton JC, Lapinskas PJ. 2005. Table 3 In: Peroxisome proliferator-activated receptors: Mediators of phthalate ester-induced effects in the male reproductive tract? Toxicol Sci 83(1):4-17. DOI: 10.1093/toxsci/kfi011. URL: https://academic.oup.com/toxsci/article/83/1/4/1658906 Lampen A, Zimnik S, Nau H. 2003. Teratogenic phthalate esters and metabolites activate the nuclear receptors PPARs and induce differentiation of F9 cells. Toxicol Appl Pharmacol 188(1):14-23. DOI: 10.1016/S0041-008X(03)00014-0. URL: https://www.sciencedirect.com/science/article/pii/S0041008X03000140 | 07/15/2011 | 09/18/2017
(4-nonylphenoxy)acetic acid | NP1EC; 4-nonylphenoxyacetic acid; 2-(4-nonylphenoxy)acetic acid; 4-nonylphenoxycarboxylic acid; 4-nonylphenol carboxylic acid; (p-nonylphenoxy)acetic acid | 3115-49-9 | Jobling S, Sheahan D, Osborne JA, Matthiessen P, Sumpter JP. 1996. Inhibition of testicular growth in rainbow trout (Oncorhynchus mykiss) exposed to estrogenic alkylphenolic chemicals. Environ Toxicol Chem 15(2):194-202. DOI: 10.1002/etc.5620150218. URL: http://onlinelibrary.wiley.com/doi/10.1002/etc.5620150218/abstract Jobling S, Sumpter JP. 1993. Detergent components in sewage effluent are weakly oestrogenic to fish: An in vitro study using rainbow trout (Oncorhynchus mykiss) hepatocytes. Aquatic Toxicology 27(3-4):361-372. DOI: 10.1016/0166-445X(93)90064-8. URL: https://www.sciencedirect.com/science/article/pii/0166445X93900648 White R, Jobling S, Hoare SA, Sumpter JP, Parker MG. 1994. Environmentally persistent alkylphenolic compounds are estrogenic. Endocrinology 135(1):175-182. DOI: 10.1210/endo.135.1.8013351. URL: https://academic.oup.com/endo/article-abstract/135/1/175/3035834?redirectedFrom=fulltext | Household product ingredient, Pesticide ingredient | Dow Chemical Co. 2010 Oct 11. Nonylphenol Ethoxylate Surfactants. URL: http://msdssearch.dow.com/PublishedLiteratureDOWCOM/dh_07f3/0901b803807f336c.pdf?filepath=productsafety/pdfs/noreg/233-00257.pdf&fromPage=GetDoc White R, Jobling S, Hoare SA, Sumpter JP, Parker MG. 1994. Environmentally persistent alkylphenolic compounds are estrogenic. Endocrinology 135(1):175-182. DOI: 10.1210/endo.135.1.8013351. URL: https://academic.oup.com/endo/article-abstract/135/1/175/3035834?redirectedFrom=fulltext | 05/05/2011 | 07/11/2016
(E)-crotonaldehyde | E-2-butenal; crotonaldehyde | 123-73-9 | Grundfest CC, Chang J, Newcombe D. 1982. Acrolein: a potent modulator of lung macrophage arachidonic acid metabolism. Biochim Biophys Acta 713(1):149-159. DOI: 10.1016/0005-2760(82)90177-1. URL: https://www.sciencedirect.com/science/article/pii/0005276082901771 Jha AM, Kumar M. 2006. In vivo evaluation of induction of abnormal sperm morphology in mice by an unsaturated aldehyde crotonaldehyde. Mutat Res 603(2):159-163. DOI: 10.1016/j.mrgentox.2005.11.010. URL: https://www.sciencedirect.com/science/article/pii/S1383571805002962?via%3Dihub | Byproduct/Intermediate/Reactant, Food additives and contact material, Household product ingredient, Pesticide ingredient, Plastic/Rubber, Solvent, Unconventional Oil and Gas | US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB Waxman HA, Markey EJ, and DeGette D. 2011. Chemicals used in hydraulic fracturing. United States House of Representatives Committee on Energy and Commerce Minority Staff. URL: https://hero.epa.gov/hero/index.cfm/reference/details/reference_id/2079174 | 03/22/2017 | 03/22/2017
1-hydroxy-2,4,7,8-tetrachlorodibenzofuran |  |  | Kester MH, Bulduk S, van Toor H, Tibboel D, Meinl W, Glatt H, Falany CN, Coughtrie MW, Schuur AG, Brouwer A, Visser TJ. 2002. Potent inhibition of estrogen sulfotransferase by hydroxylated metabolites of polyhalogenated aromatic hydrocarbons reveals alternative mechanism for estrogenic activity of endocrine disrupters. J Clin Endocrinol Metab 87(3):1142-1150. DOI: 10.1210/jcem.87.3.8311. URL: https://academic.oup.com/jcem/article/87/3/1142/2846988 | Metabolite/Degradate | Kester MH, Bulduk S, van Toor H, Tibboel D, Meinl W, Glatt H, Falany CN, Coughtrie MW, Schuur AG, Brouwer A and Visser TJ. 2002. Potent inhibition of estrogen sulfotransferase by hydroxylated metabolites of polyhalogenated aromatic hydrocarbons reveals alternative mechanism for estrogenic activity of endocrine disrupters. Journal of Clinical Endocrinology and Metabolism 87(3): 1142-1150. DOI: 10.1210/jcem.87.3.8311. URL: https://academic.oup.com/jcem/article/87/3/1142/2846988 | 01/20/2017 | 01/20/2017
1-hydroxyanthraquinone | 1-hATQ | 129-43-1 | Kurihara R, Shiraishi F, Tanaka N, Hashimoto S. 2005. Presence and estrogenicity of anthracene derivatives in coastal Japanese waters. Environ Toxicol Chem 24(8):1984-1993. DOI: 10.1897/04-310R.1. URL: http://onlinelibrary.wiley.com/doi/10.1897/04-310R.1/abstract | Byproduct/Intermediate/Reactant | US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 01/20/2017 | 01/20/2017
1-hydroxybenzo[a]pyrene | 1-OH-B[a]P | 13345-23-8 | Fertuck KC, Matthews JB, Zacharewski TR. 2001. Hydroxylated benzo[a]pyrene metabolites are responsible for in vitro estrogen receptor-mediated gene expression induced by benzo[a]pyrene, but do not elicit uterotrophic effects in vivo. Toxicol Sci 59(2):231-240. DOI: 10.1093/toxsci/59.2.231. URL: https://academic.oup.com/toxsci/article/59/2/231/1667180 van Lipzig MM, Vermeulen NP, Gusinu R, Legler J, Frank H, Seidel A, Meerman JH. 2005. Formation of estrogenic metabolites of benzo[a]pyrene and chrysene by cytochrome P450 activity and their combined and supra-maximal estrogenic activity. Environ Toxicol Pharmacol 19(1):41-55. DOI: 10.1016/j.etap.2004.03.010. URL: https://www.sciencedirect.com/science/article/pii/S1382668904001279 | Metabolite/Degradate | Fertuck KC, Matthews JB and Zacharewski TR. 2001. Hydroxylated benzo[a]pyrene metabolites are responsible for in vitro estrogen receptor-mediated gene expression induced by benzo[a]pyrene, but do not elicit uterotrophic effects in vivo. Toxicol Sci 59(2): 231-240. DOI: 10.1093/toxsci/59.2.231. URL: https://academic.oup.com/toxsci/article/59/2/231/1667180 van Lipzig MM, Vermeulen NP, Gusinu R, Legler J, Frank H, Seidel A and Meerman JH. 2005. Formation of estrogenic metabolites of benzo[a]pyrene and chrysene by cytochrome P450 activity and their combined and supra-maximal estrogenic activity. Environ Toxicol Pharmacol 19(1): 41-55. DOI: 10.1016/j.etap.2004.03.010. URL: https://www.sciencedirect.com/science/article/pii/S1382668904001279?via%3Dihub | 01/20/2017 | 01/20/2017
1-hydroxychlordene |  | 2597-11-7 | Soto AM, Sonnenschein C, Chung KL, Fernandez MF, Olea N, Serrano FO. 1995. The E-SCREEN assay as a tool to identify estrogens: An update on estrogenic environmental pollutants. Environ Health Perspect 103 (Suppl. 7):113-122. DOI: 10.2307/3432519 . URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1518887/ | Metabolite/Degradate | ATSDR. 2007. Toxicological Profile for Heptachlor and Heptachlor Epoxide. URL: http://www.atsdr.cdc.gov/toxprofiles/tp12.pdf | 05/05/2011 | 05/05/2011
1-hydroxychrysene | 1-OHCHN | 63019-38-5 | van Lipzig MM, Vermeulen NP, Gusinu R, Legler J, Frank H, Seidel A, Meerman JH. 2005. Formation of estrogenic metabolites of benzo[a]pyrene and chrysene by cytochrome P450 activity and their combined and supra-maximal estrogenic activity. Environ Toxicol Pharmacol 19(1):41-55. DOI: 10.1016/j.etap.2004.03.010. URL: https://www.sciencedirect.com/science/article/pii/S1382668904001279 | Metabolite/Degradate | van Lipzig MM, Vermeulen NP, Gusinu R, Legler J, Frank H, Seidel A and Meerman JH. 2005. Formation of estrogenic metabolites of benzo[a]pyrene and chrysene by cytochrome P450 activity and their combined and supra-maximal estrogenic activity. Environ Toxicol Pharmacol 19(1): 41-55. DOI: 10.1016/j.etap.2004.03.010. URL: https://www.sciencedirect.com/science/article/pii/S1382668904001279?via%3Dihub | 01/20/2017 | 01/20/2017
1-hydroxypyrene |  | 5315-79-7 | Jin L, Tran DQ, Ide CF, McLachlan JA, Arnold SF. 1997. Several synthetic chemicals inhibit progesterone receptor-mediated transactivation in yeast. Biochemical & Biophysical Research Communications 233(1):139-146. DOI: 10.1006/bbrc.1997.6417. URL: https://www.sciencedirect.com/science/article/pii/S0006291X97964172 | Metabolite/Degradate | Jongeneelen FJ. 2001. Benchmark guideline for urinary 1-hydroxypyrene as biomarker of occupational exposure to polycyclic aromatic hydrocarbons. Ann Occup Hyg 45(1):3-13. DOI: 10.1016/S0003-4878(00)00009-0. URL: http://www.industox.nl/AnnOccupHyg452001FJ.PDF | 05/05/2011 | 05/05/2011
1-iodoperfluorohexane | PFHxI | 355-43-1 | Wang C, Wang T, Liu W, Ruan T, Zhou Q, Liu J, Zhang A, Zhao B, Jiang G. 2012. The in vitro estrogenic activities of polyfluorinated iodine alkanes. Environ Health Perspect 120(1):119-125. DOI: 10.1289/ehp.1103773. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC3261944/ | Byproduct/Intermediate/Reactant | Wang C, Wang T, Liu W, Ruan T, Zhou Q, Liu J, Zhang A, Zhao B and Jiang G. 2012. The in vitro estrogenic activities of polyfluorinated iodine alkanes. Environ Health Perspect 120(1):119-125. DOI: 10.1289/ehp.1103773. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC3261944/ | 01/20/2017 | 01/20/2017
1-iodoperfluorooctane | PFOI | 507-63-1 | Wang C, Wang T, Liu W, Ruan T, Zhou Q, Liu J, Zhang A, Zhao B, Jiang G. 2012. The in vitro estrogenic activities of polyfluorinated iodine alkanes. Environ Health Perspect 120(1):119-125. DOI: 10.1289/ehp.1103773. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC3261944/ | Byproduct/Intermediate/Reactant | Wang C, Wang T, Liu W, Ruan T, Zhou Q, Liu J, Zhang A, Zhao B and Jiang G. 2012. The in vitro estrogenic activities of polyfluorinated iodine alkanes. Environ Health Perspect 120(1):119-125. DOI: 10.1289/ehp.1103773. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC3261944/ | 01/20/2017 | 01/20/2017
1-methoxy-2-propanol | propylene glycol methyl ether; propylene glycol monomethyl ether; 1-methoxy-2-hydroxypropane; 2-methoxy-1-methylethanol; methyl proxitol | 107-98-2 | Carney EW, Crissman JW, Liberacki AB, Clements CM, Breslin WJ. 1999. Assessment of adult and neonatal reproductive parameters in Sprague-Dawley rats exposed to propylene glycol monomethyl ether vapors for two generations. Toxicol Sci 50(2):249-258. DOI: 10.1093/toxsci/50.2.249. URL: https://academic.oup.com/toxsci/article/50/2/249/2256963 | Byproduct/Intermediate/Reactant, Household product ingredient, Industrial additive, Pesticide ingredient, Solvent, Unconventional Oil and Gas | US Dept of Health & Human Services. Household Products Database. URL: http://hpd.nlm.nih.gov/ US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB Waxman HA, Markey EJ, and DeGette D. 2011. Chemicals used in hydraulic fracturing. United States House of Representatives Committee on Energy and Commerce Minority Staff. URL: https://hero.epa.gov/hero/index.cfm/reference/details/reference_id/2079174 | 03/22/2017 | 09/18/2017
1-methylpyrene |  | 2381-21-7 | Machala M, Vondracek J, Blaha L, Ciganek M, Neca JV. 2001. Aryl hydrocarbon receptor-mediated activity of mutagenic polycyclic aromatic hydrocarbons determined using in vitro reporter gene assay. Mutat Res 497(1-2):49-62. DOI: 10.1016/S1383-5718(01)00240-6. URL: https://www.sciencedirect.com/science/article/pii/S1383571801002406 | Byproduct/Intermediate/Reactant, Unconventional Oil and Gas | US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB Zielinska B, Campbell D, Samburova V. 2014. Impact of emissions from natural gas production facilities on ambient air quality in the Barnett Shale area: a pilot study. J Air Waste Manag Assoc 64(12):1369-1383. DOI: 10.1080/10962247.2014.954735. URL: http://www.tandfonline.com/doi/abs/10.1080/10962247.2014.954735 | 08/07/2014 | 03/20/2017
1-monochloronaphthalene | PCN-1; 1-MoCN | 90-13-1 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1-naphthol | naphthol-1 | 90-15-3 | Casale GP, Vennerstrom JL, Bavari S, Wang T. 1993. Inhibition of interleukin 2 driven proliferation of mouse CTLL2 cells, by selected carbamate and organophosphate insecticides and congeners of carbaryl. Immunopharmacology & Immunotoxicology 15(2 & 3):199-215. DOI: 10.3109/08923979309025994. URL: http://www.tandfonline.com/doi/abs/10.3109/08923979309025994?journalCode=iipi20 Mueller GC, Kim UH. 1978. Displacement of estradiol from estrogen receptors by simple alkyl phenols. Endocrinology 102(5):1429-1435. DOI: 10.1210/endo-102-5-1429. URL: https://academic.oup.com/endo/article-abstract/102/5/1429/2618306?redirectedFrom=fulltext | Byproduct/Intermediate/Reactant, Household product ingredient, Industrial additive, Medical/Veterinary/Research, Metabolite/Degradate, Personal care product/Cosmetic ingredient | Meeker JD, Barr DB, Serdar B, Rappaport SM, Hauser R. 2007. Utility of urinary 1-naphthol and 2-naphthol levels to assess environmental carbaryl and naphthalene exposure in an epidemiology study. J Expo Sci Environ Epidemiol 17(4):314-320. DOI: 10.1038/sj.jes.7500502. URL: https://www.nature.com/articles/7500502 US Dept of Health & Human Services. Household Products Database. URL: http://hpd.nlm.nih.gov/ US National Library of Medicine. Haz-Map. URL: http://hazmap.nlm.nih.gov/index.php | 05/05/2011 | 09/23/2013
1-naphthyl-carbamate |  | n/a | Casale GP, Vennerstrom JL, Bavari S, Wang T. 1993. Inhibition of interleukin 2 driven proliferation of mouse CTLL2 cells, by selected carbamate and organophosphate insecticides and congeners of carbaryl. Immunopharmacology & Immunotoxicology 15(2 & 3):199-215. DOI: 10.3109/08923979309025994. URL: http://www.tandfonline.com/doi/abs/10.3109/08923979309025994?journalCode=iipi20 | Medical/Veterinary/Research | Casale GP, Vennerstrom JL, Bavari S, Wang T. 1993. Inhibition of interleukin 2 driven proliferation of mouse CTLL2 cells, by selected carbamate and organophosphate insecticides and congeners of carbaryl. Immunopharmacol Immunotoxicol 15(2 & 3):199-215. DOI: 10.3109/08923979309025994. URL: http://www.tandfonline.com/doi/abs/10.3109/08923979309025994?journalCode=iipi20 | 05/05/2011 | 05/05/2011
1-naphthyl-N-methylacetamide |  | n/a | Casale GP, Vennerstrom JL, Bavari S, Wang T. 1993. Inhibition of interleukin 2 driven proliferation of mouse CTLL2 cells, by selected carbamate and organophosphate insecticides and congeners of carbaryl. Immunopharmacology & Immunotoxicology 15(2 & 3):199-215. DOI: 10.3109/08923979309025994. URL: http://www.tandfonline.com/doi/abs/10.3109/08923979309025994?journalCode=iipi20 | Medical/Veterinary/Research | Casale GP, Vennerstrom JL, Bavari S, Wang T. 1993. Inhibition of interleukin 2 driven proliferation of mouse CTLL2 cells, by selected carbamate and organophosphate insecticides and congeners of carbaryl. Immunopharmacol Immunotoxicol 15(2 & 3):199-215. DOI: 10.3109/08923979309025994. URL: http://www.tandfonline.com/doi/abs/10.3109/08923979309025994?journalCode=iipi20 | 05/05/2011 | 05/05/2011
1-naphthyl-N,N-dimethylcarbamate | N,N-dimethylcarbamic acid 1-naphthalenyl ester | 2619-00-3; 91384-85-9 | Casale GP, Vennerstrom JL, Bavari S, Wang T. 1993. Inhibition of interleukin 2 driven proliferation of mouse CTLL2 cells, by selected carbamate and organophosphate insecticides and congeners of carbaryl. Immunopharmacology & Immunotoxicology 15(2 & 3):199-215. DOI: 10.3109/08923979309025994. URL: http://www.tandfonline.com/doi/abs/10.3109/08923979309025994?journalCode=iipi20 | Medical/Veterinary/Research | Casale GP, Vennerstrom JL, Bavari S, Wang T. 1993. Inhibition of interleukin 2 driven proliferation of mouse CTLL2 cells, by selected carbamate and organophosphate insecticides and congeners of carbaryl. Immunopharmacol Immunotoxicol 15(2 & 3):199-215. DOI: 10.3109/08923979309025994. URL: http://www.tandfonline.com/doi/abs/10.3109/08923979309025994?journalCode=iipi20 | 05/05/2011 | 05/05/2011
1,1-bis(4-hydroxyphenyl)propane |  | 1576-13-2 | Perez P, Pulgar R, Olea-Serrano F, Villalobos M, Rivas A, Metzler M, Pedraza V, Olea N. 1998. The estrogenicity of bisphenol A-related diphenylalkanes with various substituents at the central carbon and the hydroxy groups. Environ Health Perspect 106(3):167-174. DOI: 10.2307/3434318 . URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1533034/ | Plastic/Rubber | Perez P, Pulgar R, Olea-Serrano F, Villalobos M, Rivas A, Metzler M, Pedraza V, Olea N. 1998. The estrogenicity of bisphenol A-related diphenylalkanes with various substituents at the central carbon and the hydroxy groups. Environ Health Perspect 106(3):167-174. DOI: 10.1289/ehp.98106167. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1533034/ | 05/05/2011 | 09/19/2013
1,2 butylene oxide | 1,2-epoxybutane | 106-88-7 | National Toxicology Program. NTP Toxicology and Carcinogenesis Studies of 1, 2-Epoxybutane (CAS No. 106-88-7) in F344/N Rats and B6C3F1 Mice (Inhalation Studies). National Toxicology Program technical report series. 1988 Mar;329:1. URL: https://ntp.niehs.nih.gov/results/pubs/longterm/reports/longterm/tr300399/abstracts/tr329/index.html | Byproduct/Intermediate/Reactant, Household product ingredient, Industrial additive, Pesticide ingredient | US Dept of Health & Human Services. Household Products Database. URL: http://hpd.nlm.nih.gov/ US EPA. 2017. Inert Ingredients Overview and Guidance. URL: https://www.epa.gov/pesticide-registration/inert-ingredients-overview-and-guidance US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 07/13/2017 | 09/18/2017
1,2-dichlorobenzene | o-dichlorobenzene | 95-50-1 | den Besten C, Vet JJRM, Besselink HT, Kiel GS, Beems R, van Bladeren PJV. 1991. The liver, kidney, and thyroid toxicity of chlorinated benzenes. Toxicol Appl Pharmacol 111(1):69-81. DOI: 10.1016/0041-008X(91)90135-2. URL: https://www.sciencedirect.com/science/article/pii/0041008X91901352 | Antimicrobial, Byproduct/Intermediate/Reactant, Household product ingredient, Industrial additive, Pesticide ingredient, Solvent | US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 05/05/2011 | 05/05/2011
1,2-dihydroxychrysene | 1,2-OHCHN | 64920-31-6 | van Lipzig MM, Vermeulen NP, Gusinu R, Legler J, Frank H, Seidel A, Meerman JH. 2005. Formation of estrogenic metabolites of benzo[a]pyrene and chrysene by cytochrome P450 activity and their combined and supra-maximal estrogenic activity. Environ Toxicol Pharmacol 19(1):41-55. DOI: 10.1016/j.etap.2004.03.010. URL: https://www.sciencedirect.com/science/article/pii/S1382668904001279 | Metabolite/Degradate | van Lipzig MM, Vermeulen NP, Gusinu R, Legler J, Frank H, Seidel A and Meerman JH. 2005. Formation of estrogenic metabolites of benzo[a]pyrene and chrysene by cytochrome P450 activity and their combined and supra-maximal estrogenic activity. Environ Toxicol Pharmacol 19(1): 41-55. DOI: 10.1016/j.etap.2004.03.010. URL: https://www.sciencedirect.com/science/article/pii/S1382668904001279?via%3Dihub | 01/20/2017 | 01/20/2017
1,2,3-trichloropropane | allyl trichloride; glycerol trichlorohydrin; glyceryl trichlorohydrin; trichlorohydrin | 96-18-4 | Johannsen, F.R., Levinskas, G.J., Rusch, G.M., Terrill, J.B. and Schroeder, R.E., 1988. Evaluation of the subchronic and reproductive effects of a series of chlorinated propanes in the rat. I. Toxicity of 1, 2, 3‐trichloropropane. Journal of Toxicology and Environmental Health, Part A Current Issues, 25(3), pp.299-315. DOI: 10.1080/15287398809531211. URL: https://www.tandfonline.com/doi/abs/10.1080/15287398809531211 National Toxicology Program, 1993. NTP Toxicology and Carcinogenesis of 1, 2, 3-Trichloropropane (CAS No. 96-18-4) in F344/N Rats and B6C3F1 Mice (Gavage Studies). National Toxicology Program technical report series, 384, p.1. URL: https://www.ncbi.nlm.nih.gov/pubmed/12692652 | Byproduct/Intermediate/Reactant, Food additives and contact material, Household product ingredient, Pesticide ingredient, Plastic/Rubber, Solvent | Groh, K.J., Backhaus, T., Carney-Almroth, B., Geueke, B., Inostroza, P.A., Lennquist, A., Maffini, M., Leslie, H.A., Slunge, D., Trasande, L. and Warhurst, M., 2018. Chemicals associated with plastic packaging: Inventory and hazards. PeerJ Preprints, 6, p.e27036v1. DOI: 10.7287/peerj.preprints.27036v1. URL: https://peerj.com/preprints/27036/ US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 09/26/2018
1,2,3-trihydroxybenzene | pyrogallol | 87-66-1 | Lindsay RH, Hill JB, Gaitan E, Cooksey RC, Jolley RL. 1992. Antithyroid effects of coal-derived pollutants. J Toxicol Environ Health 37(4):467-481. DOI: 10.1080/15287399209531686. URL: http://www.tandfonline.com/doi/abs/10.1080/15287399209531686?journalCode=uteh19 Van den Berg KJ, van Raaij JAGM, Bragt PC, Notten WRF. 1991. Interactions of halogenated industrial chemicals with transthyretin and effects on thyroid hormone levels in vivo. Arch Toxicol 65(1):15-19. DOI: 10.1007/BF01973497. URL: https://link.springer.com/article/10.1007/BF01973497 | Household product ingredient, Industrial additive, Medical/Veterinary/Research | US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 05/05/2011 | 05/05/2011
1,2,3,4-tetraclorodibenzo-p-dioxin | 1,2,3,4-TCDD | 30746-58-8 | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L, Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1):21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | Byproduct/Intermediate/Reactant | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L and Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1): 21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | 01/20/2017 | 01/20/2017
1,2,3,4,5,6,7-heptachloronaphthalene | PCN-73; 1,2,3,4,5,6,7-HpCN | 58863-14-2 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,2,3,4,5,6,7,8-octachloronaphthalene | PCN-75; OCN | 60983-69-9 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,2,3,4,6-pentachloronaphthalene | PCN-50; 1,2,3,4,6-PeCN | 67922-26-3 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,2,3,4,6,7-hexachloronaphthalene |  | 103426-96-6 | Omura M, Masuda Y, Hirata M, Tanaka A, Makita Y, Ogata R, Inoue N. 2000. Onset of spermatogenesis is accelerated by gestational administration of 1,2,3,4,6,7-hexachlorinated naphthalene in male rat offspring. Environ Health Perspect 108(6):539-544. DOI: 10.1289/ehp.00108539. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1638139/ | Byproduct/Intermediate/Reactant | Omura M, Masuda Y, Hirata M, Tanaka A, Makita Y, Ogata R, Inoue N. 2000. Onset of spermatogenesis is accelerated by gestational administration of 1,2,3,4,6,7-hexachlorinated naphthalene in male rat offspring. Environ Health Perspect 108(6):539-544. DOI: 10.1289/ehp.00108539. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1638139/ | 05/05/2011 | 05/05/2011
1,2,3,4,7-pentaclorodibenzo-p-dioxin | 1,2,3,4,7-pentaCDD | 39227-61-7 | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L, Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1):21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | Byproduct/Intermediate/Reactant | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L and Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1): 21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | 01/20/2017 | 01/20/2017
1,2,3,4,7,8-hexabromodibenzo-p-dioxin | 1,2,3,4,7,8-HxBDD | 110999-44-5 | Hornung MW, Zabel EW, Peterson RE. 1996. Toxic equivalency factors of polybrominated dibenzo-p-dioxin, dibenzofuran, biphenyl, and polyhalogenated diphenyl ether congeners based on rainbow trout early life stage mortality. Toxicol Appl Pharmacol 140(2):227-234. DOI: 10.1006/taap.1996.0217. URL: https://www.sciencedirect.com/science/article/pii/S0041008X96902173 | Byproduct/Intermediate/Reactant | Hornung MW, Zabel EW, Peterson RE. 1996. Toxic equivalency factors of polybrominated dibenzo-p-dioxin, dibenzofuran, biphenyl, and polyhalogenated diphenyl ether congeners based on rainbow trout early life stage mortality. Toxicol Appl Pharmacol 140(2):227-234. DOI: 10.1006/taap.1996.0217. URL: https://www.sciencedirect.com/science/article/pii/S0041008X96902173?via%3Dihub WHO. 1994. Brominated diphenyl ethers. Environmental Health Criteria 162. URL: http://www.inchem.org/documents/ehc/ehc/ehc162.htm | 05/05/2011 | 05/05/2011
1,2,3,4,7,8-hexabromodibenzofuran | 1,2,3,4,7,8-HxBDF | 129880-08-6 | Hornung MW, Zabel EW, Peterson RE. 1996. Toxic equivalency factors of polybrominated dibenzo-p-dioxin, dibenzofuran, biphenyl, and polyhalogenated diphenyl ether congeners based on rainbow trout early life stage mortality. Toxicol Appl Pharmacol 140(2):227-234. DOI: 10.1006/taap.1996.0217. URL: https://www.sciencedirect.com/science/article/pii/S0041008X96902173 | Byproduct/Intermediate/Reactant | Hornung MW, Zabel EW, Peterson RE. 1996. Toxic equivalency factors of polybrominated dibenzo-p-dioxin, dibenzofuran, biphenyl, and polyhalogenated diphenyl ether congeners based on rainbow trout early life stage mortality. Toxicol Appl Pharmacol 140(2):227-234. DOI: 10.1006/taap.1996.0217. URL: https://www.sciencedirect.com/science/article/pii/S0041008X96902173?via%3Dihub WHO. 1994. Brominated diphenyl ethers. Environmental Health Criteria 162. URL: http://www.inchem.org/documents/ehc/ehc/ehc162.htm | 05/05/2011 | 05/05/2011
1,2,3,4,7,8-hexachlorodibenzofuran | 1,2,3,4,7,8-HxCDF | 70648-26-9 | Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: http://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | Byproduct/Intermediate/Reactant | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,3,4,7,8-hexaclorodibenzo-p-dioxin | 1,2,3,4,7,8-hexaCDD | 39227-28-6 | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L, Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1):21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | Byproduct/Intermediate/Reactant | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L and Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1): 21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | 01/20/2017 | 01/20/2017
1,2,3,4,8-pentachlorodibenzofuran | 1,2,3,4,8-PeCDF | 67517-48-0 | Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: http://www.sciencedirect.com/science/article/pii/0300483X8490132X | Byproduct/Intermediate/Reactant | ATSDR. 1994. Toxicological profile for Chlorodibenzofurans. URL: http://www.atsdr.cdc.gov/ToxProfiles/tp32.pdf Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: https://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011
1,2,3,5,6,7-hexachloronaphthalene | PCN-67; 1,2,3,5,6,7-HxCN | 103426-97-7 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,2,3,5,6,8-hexachloronaphthalene | PCN-68; 1,2,3,4,6,8-HxCN | 103426-95-5 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,2,3,5,7,8-hexachloronaphthalene | PCN-69; 1,2,3,5,7,8-HxCN | 103426-94-4 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,2,3,6-tetrachlorodibenzofuran | 1,2,3,6-TCDF | 83704-21-6 | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | Byproduct/Intermediate/Reactant | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,3,6,7-pentachloronaphthalene | PCN-54; 1,2,3,6,7-PeCN | 150224-16-1 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,2,3,6,7,8-hexachlorodibenzofuran | 1,2,3,6,7,8-HxCDF | 57117-44-9 | Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: http://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | Byproduct/Intermediate/Reactant | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,3,6,7,8-hexachloronaphthalene | PCN-70; 1,2,3,6,7,8-HxCN | 17062-87-2 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,2,3,7-tetrachlorodibenzofuran | 1,2,3,7-TCDF | 83704-22-7 | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | Byproduct/Intermediate/Reactant | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,3,7,8-pentabromodibenzo-p-dioxin | 1,2,3,7,8-PeBDD | 109333-34-8 | Hornung MW, Zabel EW, Peterson RE. 1996. Toxic equivalency factors of polybrominated dibenzo-p-dioxin, dibenzofuran, biphenyl, and polyhalogenated diphenyl ether congeners based on rainbow trout early life stage mortality. Toxicol Appl Pharmacol 140(2):227-234. DOI: 10.1006/taap.1996.0217. URL: https://www.sciencedirect.com/science/article/pii/S0041008X96902173 | Byproduct/Intermediate/Reactant | Hornung MW, Zabel EW, Peterson RE. 1996. Toxic equivalency factors of polybrominated dibenzo-p-dioxin, dibenzofuran, biphenyl, and polyhalogenated diphenyl ether congeners based on rainbow trout early life stage mortality. Toxicol Appl Pharmacol 140(2):227-234. DOI: 10.1006/taap.1996.0217. URL: https://www.sciencedirect.com/science/article/pii/S0041008X96902173?via%3Dihub WHO. 1994. Brominated diphenyl ethers. Environmental Health Criteria 162. URL: http://www.inchem.org/documents/ehc/ehc/ehc162.htm | 05/05/2011 | 05/05/2011
1,2,3,7,8-pentabromodibenzofuran | 1,2,3,7,8-PeBDF | 107555-93-1 | Hornung MW, Zabel EW, Peterson RE. 1996. Toxic equivalency factors of polybrominated dibenzo-p-dioxin, dibenzofuran, biphenyl, and polyhalogenated diphenyl ether congeners based on rainbow trout early life stage mortality. Toxicol Appl Pharmacol 140(2):227-234. DOI: 10.1006/taap.1996.0217. URL: https://www.sciencedirect.com/science/article/pii/S0041008X96902173 | Byproduct/Intermediate/Reactant | Hornung MW, Zabel EW, Peterson RE. 1996. Toxic equivalency factors of polybrominated dibenzo-p-dioxin, dibenzofuran, biphenyl, and polyhalogenated diphenyl ether congeners based on rainbow trout early life stage mortality. Toxicol Appl Pharmacol 140(2):227-234. DOI: 10.1006/taap.1996.0217. URL: https://www.sciencedirect.com/science/article/pii/S0041008X96902173?via%3Dihub WHO. 1994. Brominated diphenyl ethers. Environmental Health Criteria 162. URL: http://www.inchem.org/documents/ehc/ehc/ehc162.htm | 05/05/2011 | 05/05/2011
1,2,3,7,8-pentachlorodibenzofuran | 1,2,3,7,8-PeCDF | 57117-41-6 | Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: http://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | Byproduct/Intermediate/Reactant | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,3,7,8-pentaclorodibenzo-p-dioxin | 1,2,3,7,8-pentaCDD | 40321-76-4 | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L, Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1):21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | Byproduct/Intermediate/Reactant | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L and Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1): 21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | 01/20/2017 | 01/20/2017
1,2,3,7,9-pentachlorodibenzofuran | 1,2,3,7,9-PeCDF | 83704-53-4 | Harris M, Zacharewski T, Safe S. 1990. Effects of 2,3,7,8-tetrachlorodibenzo-p-dioxin and related compounds on the occupied nuclear estrogen receptor in MCF-7 human breast cancer cells. Cancer Res 50(12):3579-3584. URL: http://cancerres.aacrjournals.org/content/50/12/3579.long Krishnan V, Safe S. 1993. Polychlorinated biphenyls (PCBs), dibenzo-p-dioxins (PCDDs), and dibenzofurans (PCDFs) as antiestrogens in MCF-7 human breast cancer cells: quantitative structure-activity relationships. Toxicol Appl Pharmacol 120(1):55-61. DOI: 10.1006/taap.1993.1086. URL: https://www.sciencedirect.com/science/article/pii/S0041008X83710860?via%3Dihub Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | Byproduct/Intermediate/Reactant | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,4-triclorodibenzo-p-dioxin | 1,2,4-triCDD | 39227-58-2 | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L, Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1):21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | Byproduct/Intermediate/Reactant | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L and Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1): 21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | 01/20/2017 | 01/20/2017
1,2,4,5-tetrachlorobenzene |  | 95-94-3; 63697-22-3 | den Besten C, Vet JJRM, Besselink HT, Kiel GS, Beems R, van Bladeren PJV. 1991. The liver, kidney, and thyroid toxicity of chlorinated benzenes. Toxicol Appl Pharmacol 111(1):69-81. DOI: 10.1016/0041-008X(91)90135-2. URL: https://www.sciencedirect.com/science/article/pii/0041008X91901352 | Byproduct/Intermediate/Reactant | US National Library of Medicine. Haz-Map. URL: http://hazmap.nlm.nih.gov/index.php | 05/05/2011 | 05/05/2011
1,2,4,5,7,8-hexaachloronaphthalene | PCN-72; 1,2,4,5,7,8-HxCN | 103426-92-2 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,2,4,6,7-pentachlorodibenzofuran | 1,2,4,6,7-PeCDF | 83704-50-1 | Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: http://www.sciencedirect.com/science/article/pii/0300483X8490132X | Byproduct/Intermediate/Reactant | ATSDR. 1994. Toxicological profile for Chlorodibenzofurans. URL: http://www.atsdr.cdc.gov/ToxProfiles/tp32.pdf Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: https://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,4,6,7,8-hexachlorodibenzofuran | 1,2,4,6,7,8-HxCDF | 67562-40-7 | Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: http://www.sciencedirect.com/science/article/pii/0300483X8490132X | Byproduct/Intermediate/Reactant | ATSDR. 1994. Toxicological profile for Chlorodibenzofurans. URL: http://www.atsdr.cdc.gov/ToxProfiles/tp32.pdf Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: https://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,4,6,8-pentachlorodibenzofuran | 1,2,4,6,8-PeCDF | 69698-57-3 | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | Byproduct/Intermediate/Reactant | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,4,7,8-pentabromodibenzo-p-dioxin | 1,2,4,7,8-PeBDD | 109333-35-9 | Mason G, Zacharewski T, Denomme MA, Safe L, Safe S. 1987. Polybrominated dibenzo-p-dioxins and related compounds: Quantitative in vivo and in vitro structure-activity relationships. Toxicology 44(3):245-255. DOI: 10.1016/0300-483X(87)90027-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X87900278 | Byproduct/Intermediate/Reactant | Mason G, Zacharewski T, Denomme MA, Safe L, Safe S. 1987. Polybrominated dibenzo-p-dioxins and related compounds: Quantitative in vivo and in vitro structure-activity relationships. Toxicology 44(3):245-255. DOI: 10.1016/0300-483X(87)90027-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X87900278 | 05/05/2011 | 05/05/2011
1,2,4,7,8-pentachlorodibenzofuran | 1,2,4,7,8-PeCDF | 58802-15-6 | Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: http://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | Byproduct/Intermediate/Reactant | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,4,7,8-pentaclorodibenzo-p-dioxin | 1,2,4,7,8-pentaCDD | 58802-08-7 | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L, Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1):21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | Byproduct/Intermediate/Reactant | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L and Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1): 21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 | 01/20/2017 | 01/20/2017
1,2,4,7,9-pentachlorodibenzofuran | 1,2,4,7,9-PeCDF | 71998-74-8 | Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: http://www.sciencedirect.com/science/article/pii/0300483X8490132X | Byproduct/Intermediate/Reactant | ATSDR. 1994. Toxicological profile for Chlorodibenzofurans. URL: http://www.atsdr.cdc.gov/ToxProfiles/tp32.pdf Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: https://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,4,8-tetrachlorodibenzofuran | 1,2,4,8-TCDF | 64126-87-0 | Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: http://www.sciencedirect.com/science/article/pii/0300483X8490132X | Byproduct/Intermediate/Reactant | ATSDR. 1994. Toxicological profile for Chlorodibenzofurans. URL: http://www.atsdr.cdc.gov/ToxProfiles/tp32.pdf Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: https://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,2,7,8-tetrachlorodibenzofuran | 1,2,7,8-TCDF | 58802-20-3 | Harper N, Wang X, Liu H, Safe S. 1994. Inhibition of estrogen-induced progesterone receptor in MCF-7 human breast cancer cells by aryl hydrocarbon (Ah) receptor agonists. Molecular & Cellular Endocrinology 104(1):47-55. DOI: 10.1016/0303-7207(94)90050-7. URL: http://www.sciencedirect.com/science/article/pii/0303720794900507?via%3Dihub Liu H, Biegel L, Narasimhan TR, Rowlands C, Safe S. 1992. Inhibition of insulin-like growth factor-I responses in MCF-7 cells by 2,3,7,8-tetrachlorodibenzo-p-dioxin and related compounds. Mol Cell Endocrinol 87(1-3):19-28. DOI: 10.1016/0303-7207(92)90229-Y. URL: https://www.sciencedirect.com/science/article/pii/030372079290229Y?via%3Dihub | Byproduct/Intermediate/Reactant | ATSDR. 1994. Toxicological profile for Chlorodibenzofurans. URL: http://www.atsdr.cdc.gov/ToxProfiles/tp32.pdf Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: https://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,3-butadiene | buta-1,3-diene; biethylene; bivinyl, butadiene; divinyl, erythrene; vinylethylene | 106-99-0 | Miller, R.A. and Boorman, G.A., 1990. Morphology of neoplastic lesions induced by 1, 3-butadiene in B6C3F1 mice. Environmental health perspectives, 86, p.37. DOI: 10.1289/ehp.908637. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1567728/ Owen, P.E., Glaister, J.R., Gaunt, I.F. and Pullinger, D.H., 1987. Inhalation toxicity studies with 1, 3-butadiene 3 two year toxicity/carcinogenicity study in rats. American Industrial Hygiene Association Journal, 48(5), pp.407-413. DOI: 10.1080/15298668791384959. URL: https://www.tandfonline.com/doi/abs/10.1080/15298668791384959 | Byproduct/Intermediate/Reactant, Food additives and contact material, Industrial additive, Plastic/Rubber | Groh, K.J., Backhaus, T., Carney-Almroth, B., Geueke, B., Inostroza, P.A., Lennquist, A., Maffini, M., Leslie, H.A., Slunge, D., Trasande, L. and Warhurst, M., 2018. Chemicals associated with plastic packaging: Inventory and hazards. PeerJ Preprints, 6, p.e27036v1. DOI: 10.7287/peerj.preprints.27036v1. URL: https://peerj.com/preprints/27036/ US Dept of Health & Human Services. Household Products Database. URL: http://hpd.nlm.nih.gov/ US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 09/26/2018
1,3-dichlorobenzene | m-dichlorobenzene | 541-73-1 | McCauley PT, Robinson M, Daniel FB, Olson GR. 1995. Toxicity studies of 1,3-dichlorobenzene in Sprague-Dawley rats. Drug & Chemical Toxicology 18(2-3):201-221. DOI: 10.3109/01480549509014321. URL: http://www.tandfonline.com/doi/abs/10.3109/01480549509014321?journalCode=idct20 | Byproduct/Intermediate/Reactant, Pesticide ingredient | US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 05/05/2011 | 05/05/2011
1,3-dicyclohexylurea |  | 2387-23-7 | Ahn KC, Zhao B, Chen J, Cherednichenko G, Sanmarti E, Denison MS, Lasley B, Pessah IN, Kultz D, Chang DPY, Gee SJ, Hammock BD. 2008. In vitro biologic activities of the antimicrobials triclocarban, its analogs, and triclosan in bioassay screens: Receptor-based bioassay screens. Environ Health Perspect 116(9):1203-1210. DOI: 10.1289/ehp.11200. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC2535623/ | Byproduct/Intermediate/Reactant, Medical/Veterinary/Research | Ahn KC, Zhao B, Chen J, Cherednichenko G, Sanmarti E, Denison MS, Lasley B, Pessah IN, Kultz D, Chang DPY, Gee SJ, Hammock BD. 2008. In vitro biologic activities of the antimicrobials triclocarban, its analogs, and triclosan in bioassay screens: Receptor-based bioassay screens. Environ Health Perspect 116(9):1203-1210. DOI: 10.1289/ehp.11200. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC2535623/ | 05/05/2011 | 05/05/2011
1,3-diphenyltetramethyldisiloxane | 1,3-diphenyl-1,1,3,3-tetramethyldisiloxane; PhMe2SiOSiMe2Ph | 56-33-7 | Fang H, Tong WD, Branham WS, Moland CL, Dial SL, Hong HX, Xie Q, Perkins R, Owens W, Sheehan DM. 2003. Study of 202 natural, synthetic, and environmental chemicals for binding to the androgen receptor. Chem Res Toxicol 16(10):1338-1358. DOI: 10.1021/tx030011g. URL: http://pubs.acs.org/doi/abs/10.1021/tx030011g Hayden JF, Barlow SA. 1972. Structure-activity relationships of organosiloxanes and the female reproductive system. Toxicol Appl Pharmacol 21(1):68-79. DOI: 10.1016/0041-008X(72)90028-2. URL: https://www.sciencedirect.com/science/article/pii/0041008X72900282 | Industrial additive, Medical/Veterinary/Research | LookChem URL: http://www.lookchem.com/1-3-diphenyl-1-1-3-3-tetramethyldisiloxane/ Santa Cruz BioTechnology. 1,3-Diphenyl-1,1,3,3-tetramethyldisiloxane (CAS 56-33-7) URL: http://www.scbt.com/datasheet-481116.html | 09/07/2015 | 09/18/2017
1,3,4,7,8-pentachlorodibenzofuran | 1,3,4,7,8-PeCDF | 58802-16-7 | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | Byproduct/Intermediate/Reactant | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,3,5,7-tetrachloronaphthalene | PCN-42; 1,3,5,7-TCN | 53555-64-9 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,3,6-trichlorodibenzofuran | 1,3,6-TriCDF | 83704-39-6 | Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: http://www.sciencedirect.com/science/article/pii/0300483X8490132X | Byproduct/Intermediate/Reactant | ATSDR. 1994. Toxicological profile for Chlorodibenzofurans. URL: http://www.atsdr.cdc.gov/ToxProfiles/tp32.pdf Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: https://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,3,6,7,8-pentachlorodibenzofuran | 1,3,6,7,8-PeCDF; 2,3,4,7,9-pentachlorodibenzofuran; 2,3,4,7,9-PeCDF | 70648-21-4 | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | Byproduct/Intermediate/Reactant | Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 04/24/2015
1,3,6,8-tetrachlorodibenzofuran | 1,3,6,8-TCDF | 71998-72-6 | Harris M, Zacharewski T, Safe S. 1990. Effects of 2,3,7,8-tetrachlorodibenzo-p-dioxin and related compounds on the occupied nuclear estrogen receptor in MCF-7 human breast cancer cells. Cancer Res 50(12):3579-3584. URL: http://cancerres.aacrjournals.org/content/50/12/3579.long Krishnan V, Safe S. 1993. Polychlorinated biphenyls (PCBs), dibenzo-p-dioxins (PCDDs), and dibenzofurans (PCDFs) as antiestrogens in MCF-7 human breast cancer cells: quantitative structure-activity relationships. Toxicol Appl Pharmacol 120(1):55-61. DOI: 10.1006/taap.1993.1086. URL: https://www.sciencedirect.com/science/article/pii/S0041008X83710860?via%3Dihub | Byproduct/Intermediate/Reactant | ATSDR. 1994. Toxicological profile for Chlorodibenzofurans. URL: http://www.atsdr.cdc.gov/ToxProfiles/tp32.pdf Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: https://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,3,7,8-tetrabromodibenzo-p-dioxin | 1,3,7,8-TBDD | 109333-31-5 | Hornung MW, Zabel EW, Peterson RE. 1996. Toxic equivalency factors of polybrominated dibenzo-p-dioxin, dibenzofuran, biphenyl, and polyhalogenated diphenyl ether congeners based on rainbow trout early life stage mortality. Toxicol Appl Pharmacol 140(2):227-234. DOI: 10.1006/taap.1996.0217. URL: https://www.sciencedirect.com/science/article/pii/S0041008X96902173 Mason G, Zacharewski T, Denomme MA, Safe L, Safe S. 1987. Polybrominated dibenzo-p-dioxins and related compounds: Quantitative in vivo and in vitro structure-activity relationships. Toxicology 44(3):245-255. DOI: 10.1016/0300-483X(87)90027-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X87900278 | Byproduct/Intermediate/Reactant | Mason G, Zacharewski T, Denomme MA, Safe L, Safe S. 1987. Polybrominated dibenzo-p-dioxins and related compounds: Quantitative in vivo and in vitro structure-activity relationships. Toxicology 44(3):245-255. DOI: 10.1016/0300-483X(87)90027-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X87900278 | 05/05/2011 | 05/05/2011
1,3,7,8-tetrachlorodibenzo-p-dioxin | 1,3,7,8-TCDD | 50585-46-1 | Harper N, Wang X, Liu H, Safe S. 1994. Inhibition of estrogen-induced progesterone receptor in MCF-7 human breast cancer cells by aryl hydrocarbon (Ah) receptor agonists. Molecular & Cellular Endocrinology 104(1):47-55. DOI: 10.1016/0303-7207(94)90050-7. URL: http://www.sciencedirect.com/science/article/pii/0303720794900507?via%3Dihub | Byproduct/Intermediate/Reactant, Medical/Veterinary/Research | Mason G, Farrell K, Keys B, Piskorska-Pliszczynska J, Safe L, Safe S. 1986. Polychlorinated dibenzo-p-dioxins: Quantitative in vitro and in vivo structure-activity relationships. Toxicology 41(1):21-31. DOI: 10.1016/0300-483X(86)90101-0. URL: https://www.sciencedirect.com/science/article/pii/0300483X86901010 US EPA. 2011 Apr 6. Landfill Methane Outreach Program FAQ: Public Health, Safety, and the Environment. URL: http://www.epa.gov/lmop/faq/public.html US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 05/05/2011 | 09/12/2013
1,3,7,9-tetrabromodibenzo-p-dioxin | 1,3,7,9-TBDD; 2,4,6,8-tetrabromodibenzo-p-dioxin | 109333-30-4 | Mason G, Zacharewski T, Denomme MA, Safe L, Safe S. 1987. Polybrominated dibenzo-p-dioxins and related compounds: Quantitative in vivo and in vitro structure-activity relationships. Toxicology 44(3):245-255. DOI: 10.1016/0300-483X(87)90027-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X87900278 | Byproduct/Intermediate/Reactant | Mason G, Zacharewski T, Denomme MA, Safe L, Safe S. 1987. Polybrominated dibenzo-p-dioxins and related compounds: Quantitative in vivo and in vitro structure-activity relationships. Toxicology 44(3):245-255. DOI: 10.1016/0300-483X(87)90027-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X87900278 | 05/05/2011 | 05/05/2011
1,3,8-trichlorodibenzofuran | 1,3,8-TriCDF | 76621-12-0 | Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: http://www.sciencedirect.com/science/article/pii/0300483X8490132X | Byproduct/Intermediate/Reactant | ATSDR. 1994. Toxicological profile for Chlorodibenzofurans. URL: http://www.atsdr.cdc.gov/ToxProfiles/tp32.pdf Bandiera S, Sawyer T, Romkes M, Zmudzka B, Safe L, Mason G, Keys B, Safe S. 1984. Polychlorinated dibenzofurans (PCDFs): effects of structure on binding to the 2,3,7,8-TCDD cytosolic receptor protein, AHH induction and toxicity. Toxicology 32(2):131-144. DOI: 10.1016/0300-483X(84)90132-X. URL: https://www.sciencedirect.com/science/article/pii/0300483X8490132X Mason G, Sawyer T, Keys B, Bandiera S, Romkes M, Piskorska-Pliszczynska J, Zmudzka B, Safe S. 1985. Polychlorinated dibenzofurans (PCDFs): correlation between in vivo and in vitro structure-activity relationships. Toxicology 37(1-2):1-12. DOI: 10.1016/0300-483X(85)90108-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X85901088 | 05/05/2011 | 05/05/2011
1,4-dichloronaphthalene | PCN-5; 1,4-DiCN | 1825-31-6 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,6-hexanediamine | hexamethylenediamine; 1,6-diamino-n-hexane | 124-09-4 | Kohno H, Sakurada T, Kisara K. 1981. Effects of intracerebroventricular administration of aliphatic diamines on ingestive behavior in the rat. Jpn J Pharmacol 31(2):211-220. DOI: 10.1254/jjp.31.211. URL: https://www.jstage.jst.go.jp/article/jphs1951/31/2/31_2_211/_article | Byproduct/Intermediate/Reactant, Unconventional Oil and Gas | US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB Waxman HA, Markey EJ, and DeGette D. 2011. Chemicals used in hydraulic fracturing. United States House of Representatives Committee on Energy and Commerce Minority Staff. URL: https://hero.epa.gov/hero/index.cfm/reference/details/reference_id/2079174 | 03/22/2017 | 03/22/2017
1,7-phenanthroline |  | 230-46-6 | Benisek M, Kubincova P, Blaha L, Hilscherova K. 2011. The effects of PAHs and N-PAHs on retinoid signaling and Oct-4 expression in vitro. Toxicol Lett 200(3):169-175. DOI: 10.1016/j.toxlet.2010.11.011. URL: http://www.sciencedirect.com/science/article/pii/S0378427410017716?via%3Dihub | Use/Source not available |  | 02/18/2015 | 02/18/2015
1,8-dichloronaphthalene | PCN-9; 1,8-DiCN | 2050-74-0 | Behnisch PA, Hosoe K, Sakai S. 2003. Brominated dioxin-like compounds: In vitro assessment in comparison to classical dioxin-like compounds and other polyaromatic compounds. Environment International 29(6):861-877. DOI: 10.1016/S0160-4120(03)00105-3. URL: https://www.sciencedirect.com/science/article/pii/S0160412003001053 | Byproduct/Intermediate/Reactant | Bidleman T F, Helm PA, Braune BM and Gabrielsen GW. 2010. Polychlorinated naphthalenes in polar environments--a review. Sci Total Environ 408(15):2919-2935. DOI: 10.1016/j.scitotenv.2009.09.013. URL: https://www.sciencedirect.com/science/article/pii/S0048969709008572?via%3Dihub | 01/20/2017 | 01/20/2017
1,8-dihydroxyanthraquinone |  | 117-10-2 | Matsuda H, Shimoda H, Morikawa T, Yoshikawa M. 2001. Phytoestrogens from the roots of Polygonum cuspidatum (Polygonaceae): structure-requirement of hydroxyanthraquinones for estrogenic activity. Bioorganic & Medicinal Chemistry Letters 11(14):1839-1842. DOI: 10.1016/S0960-894X(01)00318-3. URL: https://www.sciencedirect.com/science/article/pii/S0960894X01003183 | Byproduct/Intermediate/Reactant, Industrial additive, Medical/Veterinary/Research, Pesticide ingredient | National Pesticide Information Retrieval System. URL: http://ppis.ceris.purdue.edu/Default.aspx US National Library of Medicine. Haz-Map. URL: http://hazmap.nlm.nih.gov/index.php US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 05/05/2011 | 05/05/2011
1,9-dimethylphenanthrene |  | 20291-73-0 | Cook JW, Dodds EC, Hewett CL, Lawson W. 1934. The oestrogenic activity of some condensed-ring compounds in relation to their other biological activities. Proc Roy Soc Ser B 114(788):272-286. DOI: 10.1098/rspb.1934.0006 . URL: http://rspb.royalsocietypublishing.org/content/114/788/272 | Medical/Veterinary/Research | Cook JW, Dodds EC, Hewett CL, Lawson W. 1934. The oestrogenic activity of some condensed-ring compounds in relation to their other biological activities. Proc R Soc Lond B 114(788):272-286. DOI: 10.1098/rspb.1934.0006. URL: http://rspb.royalsocietypublishing.org/content/114/788/272 | 05/05/2011 | 05/05/2011
10-hydroxybenzo[a]pyrene | 10-OHBaP | 56892-31-0 | van Lipzig MM, Vermeulen NP, Gusinu R, Legler J, Frank H, Seidel A, Meerman JH. 2005. Formation of estrogenic metabolites of benzo[a]pyrene and chrysene by cytochrome P450 activity and their combined and supra-maximal estrogenic activity. Environ Toxicol Pharmacol 19(1):41-55. DOI: 10.1016/j.etap.2004.03.010. URL: https://www.sciencedirect.com/science/article/pii/S1382668904001279 | Metabolite/Degradate | van Lipzig MM, Vermeulen NP, Gusinu R, Legler J, Frank H, Seidel A and Meerman JH. 2005. Formation of estrogenic metabolites of benzo[a]pyrene and chrysene by cytochrome P450 activity and their combined and supra-maximal estrogenic activity. Environ Toxicol Pharmacol 19(1): 41-55. DOI: 10.1016/j.etap.2004.03.010. URL: https://www.sciencedirect.com/science/article/pii/S1382668904001279?via%3Dihub | 01/20/2017 | 01/20/2017
10:2 dialkylated phosphate surfactants | 10:2 diPAP | 1895-26-7 | Rosenmai AK, Nielsen FK, Pedersen M, Hadrup N, Trier X, Christensen JH, Vinggaard AM. 2013 Jan 1. Fluorochemicals used in food packaging inhibit male sex hormone synthesis. Toxicol Appl Pharmacol 266(1):132-142. DOI: 10.1016/j.taap.2012.10.022. URL: https://www.sciencedirect.com/science/article/pii/S0041008X12004644?via%3Dihub | Food additives and contact material, Industrial additive | Rosenmai AK, Nielsen FK, Pedersen M, Hadrup N, Trier X, Christensen JH and Vinggaard AM. 2013. Fluorochemicals used in food packaging inhibit male sex hormone synthesis. Toxicol Appl Pharmacol 266(1):132-142. DOI: 10.1016/j.taap.2012.10.022. URL: https://www.sciencedirect.com/science/article/pii/S0041008X12004644?via%3Dihub Rosenmai AK, Taxvig C, Svingen T, Trier X, Vugt‐Lussenburg BMA, Pedersen M, Lesné L, Jégou B and Vinggaard AM. 2016. Fluorinated alkyl substances and technical mixtures used in food paper‐packaging exhibit endocrine‐related activity in vitro. Andrology 4(4):662-672. DOI: 10.1111/andr.12190. URL: https://www.ncbi.nlm.nih.gov/pubmed/27152447 | 01/20/2017 | 12/06/2017
1a-phenyl-4a-(1'-phenylethyl)tetralin | styrene trimer ST-2 | 28213-80-1 | Satoh K, Nagai F, Aoki N. 2001. Several environmental pollutants have binding affinities for both androgen receptor and estrogen receptor α. Journal of Health Science 47(5):495-501. DOI: 10.1248/jhs.47.495. URL: https://www.jstage.jst.go.jp/article/jhs/47/5/47_5_495/_article | Byproduct/Intermediate/Reactant, Plastic/Rubber | Ohyama K-I, Nagai F, Tsuchiya Y. 2001. Certain styrene oligomers have proliferative activity on MCF-7 human breast tumor cells and binding affinity for human estrogen receptor. Environ Health Perspect 109(7):699-703. DOI: 10.2307/3454786. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1240373/ | 05/05/2011 | 09/24/2018
1a-phenyl-4e-(1'-phenylethyl)tetralin | styrene trimer ST-3 | 28213-80-1 | Ohyama K-I, Nagai F, Tsuchiya Y. 2001. Certain styrene oligomers have proliferative activity on MCF-7 human breast tumor cells and binding affinity for human estrogen receptor. Environ Health Perspect 109(7):699-703. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1240373/ Satoh K, Nagai F, Aoki N. 2001. Several environmental pollutants have binding affinities for both androgen receptor and estrogen receptor α. Journal of Health Science 47(5):495-501. DOI: 10.1248/jhs.47.495. URL: https://www.jstage.jst.go.jp/article/jhs/47/5/47_5_495/_article | Byproduct/Intermediate/Reactant, Plastic/Rubber | Ohyama K-I, Nagai F, Tsuchiya Y. 2001. Certain styrene oligomers have proliferative activity on MCF-7 human breast tumor cells and binding affinity for human estrogen receptor. Environ Health Perspect 109(7):699-703. DOI: 10.2307/3454786. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1240373/ | 05/05/2011 | 09/24/2018
1e-phenyl-4a-(1'-phenylethyl)tetralin | styrene trimer ST-4 | 28213-80-1 | Ohyama K-I, Nagai F, Tsuchiya Y. 2001. Certain styrene oligomers have proliferative activity on MCF-7 human breast tumor cells and binding affinity for human estrogen receptor. Environ Health Perspect 109(7):699-703. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1240373/ Satoh K, Nagai F, Aoki N. 2001. Several environmental pollutants have binding affinities for both androgen receptor and estrogen receptor α. Journal of Health Science 47(5):495-501. DOI: 10.1248/jhs.47.495. URL: https://www.jstage.jst.go.jp/article/jhs/47/5/47_5_495/_article | Byproduct/Intermediate/Reactant, Plastic/Rubber | Ohyama K-I, Nagai F, Tsuchiya Y. 2001. Certain styrene oligomers have proliferative activity on MCF-7 human breast tumor cells and binding affinity for human estrogen receptor. Environ Health Perspect 109(7):699-703. DOI: 10.2307/3454786. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1240373/ | 05/05/2011 | 09/24/2018
1e-phenyl-4e-(1'-phenylethyl)tetralin | styrene trimer ST-5 | 28213-80-1 | Ohyama K-I, Nagai F, Tsuchiya Y. 2001. Certain styrene oligomers have proliferative activity on MCF-7 human breast tumor cells and binding affinity for human estrogen receptor. Environ Health Perspect 109(7):699-703. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1240373/ Satoh K, Nagai F, Aoki N. 2001. Several environmental pollutants have binding affinities for both androgen receptor and estrogen receptor α. Journal of Health Science 47(5):495-501. DOI: 10.1248/jhs.47.495. URL: https://www.jstage.jst.go.jp/article/jhs/47/5/47_5_495/_article | Byproduct/Intermediate/Reactant, Plastic/Rubber | Ohyama K-I, Nagai F, Tsuchiya Y. 2001. Certain styrene oligomers have proliferative activity on MCF-7 human breast tumor cells and binding affinity for human estrogen receptor. Environ Health Perspect 109(7):699-703. DOI: 10.2307/3454786. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1240373/ | 05/05/2011 | 09/24/2018
1H,1H,2H,2H-perfluorooctan-1-ol | 6:2 FTOH | 647-42-7 | Ishibashi H, Ishida H, Matsuoka M, Tominaga N and Arizono K. 2007. Estrogenic effects of fluorotelomer alcohols for human estrogen receptor isoforms α and β in vitro. Biological and Pharmaceutical Bulletin 30(7):1358-1359. DOI: 10.1248/bpb.30.1358. URL: https://www.ncbi.nlm.nih.gov/pubmed/17603182 Ishibashi H, Yamauchi R, Matsuoka M, Kim JW, Hirano M, Yamaguchi A, Tominaga N and Arizono K. 2008. Fluorotelomer alcohols induce hepatic vitellogenin through activation of the estrogen receptor in male medaka (Oryzias latipes). Chemosphere 71(10):1853-1859. DOI: 10.1016/j.chemosphere.2008.01.065. URL: https://www.ncbi.nlm.nih.gov/pubmed/18334264 Liu C, Du Y and Zhou B. 2007. Evaluation of estrogenic activities and mechanism of action of perfluorinated chemicals determined by vitellogenin induction in primary cultured tilapia hepatocytes. Aquatic toxicology 85(4):267-277. DOI: 10.1016/j.aquatox.2007.09.009. URL: https://www.ncbi.nlm.nih.gov/pubmed/17980923 Maras M, Vanparys C, Muylle F, Robbens J, Berger U, Barber JL, Blust R and De Coen W. 2006. Estrogen-like properties of fluorotelomer alcohols as revealed by MCF-7 breast cancer cell proliferation. Environ Health Perspect 114(1):100-5. DOI: 10.1289/ehp.8149. URL: https://www.ncbi.nlm.nih.gov/pubmed/16393665 | Food additives and contact material | Rosenmai AK, Taxvig C, Svingen T, Trier X, Vugt‐Lussenburg BMA, Pedersen M, Lesné L, Jégou B and Vinggaard AM. 2016. Fluorinated alkyl substances and technical mixtures used in food paper‐packaging exhibit endocrine‐related activity in vitro. Andrology 4(4):662-672. DOI: 10.1111/andr.12190. URL: https://www.ncbi.nlm.nih.gov/pubmed/27152447 | 12/06/2017 | 12/06/2017
1R-cis-bifenthrin |  | 82657-04-3 | Wang L, Liu W, Yang C, Pan Z, Gan J, Xu C, Zhao M, Schlenk D. 2007. Enantioselectivity in estrogenic potential and uptake of bifenthrin. Environ Sci Technol 41(17):6124-6128. DOI: 10.1021/es070220d. URL: http://pubs.acs.org/doi/abs/10.1021/es070220d | Pesticide ingredient | Wang L, Liu W, Yang C, Pan Z, Gan J, Xu C, Zhao M, Schlenk D. 2007. Enantioselectivity in estrogenic potential and uptake of bifenthrin. Environ Sci Technol 41(17):6124-6128. DOI: 10.1021/es070220d. URL: https://pubs.acs.org/doi/abs/10.1021/es070220d | 05/16/2011 | 01/27/2017
2-(4'-hydroxyphenyl)-2,4,4-trimethyl chroman |  | 5026-12-0 | Terasaki M, Shiraishi F, Nishikawa T, Edmonds JS, Morita M, Makino M. 2005. Estrogenic activity of impurities in industrial grade bisphenol A. Environ Sci Technol 39(10):3703-3707. DOI: 10.1021/es048932g. URL: http://pubs.acs.org/doi/abs/10.1021/es048932g?journalCode=esthag | Byproduct/Intermediate/Reactant | Chapin RE, Adams J, Boekelheide K, Gray Jr. LE , Hayward SW, Lees PSJ, Mcintyre BS, Portier KM, Schnorr TM, Selevan SG, Vandenbergh JG and Woskie SR. 2008. NTP-CERHR expert panel report on the reproductive and developmental toxicity of bisphenol A. Birth Defects Research Part B-Developmental and Reproductive Toxicology 83(3): 157-395. DOI: 10.1002/bdrb.20147. URL: http://onlinelibrary.wiley.com/doi/10.1002/bdrb.20147/abstract Terasaki M, Shiraishi F, Nishikawa T, Edmonds JS, Morita M and Makino M. 2005. Estrogenic activity of impurities in industrial grade bisphenol A. Environ Sci Technol 39(10):3703-3707. DOI: 10.1021/es048932g. URL: https://pubs.acs.org/doi/abs/10.1021/es048932g?journalCode=esthag | 01/20/2017 | 01/20/2017
2-(thiocyanomethylthio)benzothiazole | TCMTB; TCMTBT | 21564-17-0 | Nakama A, Funasaka K, Shimizu M. 2007. Evaluation of estrogenic activity of organic biocides using ER-binding and YES assay. Food Chem Toxicol 45(9):1558-1564. DOI: 10.1016/j.fct.2007.02.014. URL: https://www.sciencedirect.com/science/article/pii/S0278691507000725?via%3Dihub | Pesticide ingredient | National Pesticide Information Retrieval System . URL: http://ppis.ceris.purdue.edu/ | 01/20/2017 | 01/20/2017
2-aminoethanol | monoethanolamine; beta-aminoethanol; 2-ethanolamine; ethanolamine | 141-43-5 | Barseghian G, Zak I, Hwang DL, Roitman A, Lev-Ran A. 1986. In vitro effects of ethanolamine on insulin secretion. Life Sci 38(7):645-651. DOI: 10.1016/0024-3205(86)90058-5. URL: http://www.sciencedirect.com/science/article/pii/0024320586900585 | Antimicrobial, Byproduct/Intermediate/Reactant, Household product ingredient, Industrial additive, Medical/Veterinary/Research, Personal care product/Cosmetic ingredient, Pesticide ingredient, Plastic/Rubber, Unconventional Oil and Gas | US Dept of Health & Human Services. Household Products Database. URL: http://hpd.nlm.nih.gov/ US National Library of Medicine. Haz-Map. URL: http://hazmap.nlm.nih.gov/index.php US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB Waxman HA, Markey EJ, and DeGette D. 2011. Chemicals used in hydraulic fracturing. United States House of Representatives Committee on Energy and Commerce Minority Staff. URL: https://hero.epa.gov/hero/index.cfm/reference/details/reference_id/2079174 | 03/22/2017 | 03/22/2017
2-benzotriazol-2-yl-4,6-di-tert-butylphenol | 2-benzotriazol-2-yl-4,6-di-tert-butylphenol; 2-benzotriazol-2-yl-4,6-di-tert-butylphenol; 2-(Benzotriazol-2-yl)-4,6-ditert-butylphenol; HDBB | 3846-71-7 | Hirata-Koizumi, M., Ise, R., Kato, H., Matsuyama, T., Nishimaki-Mogami, T., Takahashi, M., Ono, A., Ema, M. and Hirose, A., 2016. Transcriptome analyses demonstrate that Peroxisome Proliferator-Activated Receptor α (PPARα) activity of an ultraviolet absorber, 2-(2’-hydroxy-3’, 5’-di-tert-butylphenyl) benzotriazole, as possible mechanism of their toxicity and the gender differences. The Journal of toxicological sciences, 41(5), pp.693-700. DOI: 10.2131/jts.41.693. URL: https://www.jstage.jst.go.jp/article/jts/41/5/41_693/_article/-char/ja/ Hirata-Koizumi, M., Matsuyama, T., Imai, T., Hirose, A., Kamata, E. and Ema, M., 2008. Gonadal influence on the toxicity of 2-(2′-hydroxy-3′, 5′-di-tert-butylphenyl) benzotriazole in rats. Drug and chemical toxicology, 31(1), pp.115-126. DOI: 10.1080/01480540701688808. URL: https://www.tandfonline.com/doi/abs/10.1080/01480540701688808?casa_token=iAXhbXUYWlQAAAAA:VIIEUJAMS7oDKRUYpvGbj9FjyCsVt1EJhfHmbX68U7DHE3n5U8HE_zJzCBIiSonjYYmgJNORxR3l Hirata-Koizumi, M., Watari, N., Mukai, D., Imai, T., Hirose, A., Kamata, E. and Ema, M., 2007. A 28-day repeated dose toxicity study of ultraviolet absorber 2-(2′-hydroxy-3′, 5′-di-tert-butylphenyl) benzotriazole in rats. Drug and chemical toxicology, 30(4), pp.327-341. DOI: 10.1080/01480540701522254. URL: https://www.tandfonline.com/doi/abs/10.1080/01480540701522254?casa_token=kQ1FrESb0_4AAAAA:joQP1MS9ZojKULZ9ScPGCrJDVR2tGQEstjBHhVkx0zhSqaEYqSHGvgMbKPxYJTORBqwvl7ashuCu | Industrial additive, Plastic/Rubber | Groh, K.J., Backhaus, T., Carney-Almroth, B., Geueke, B., Inostroza, P.A., Lennquist, A., Maffini, M., Leslie, H.A., Slunge, D., Trasande, L. and Warhurst, M., 2018. Chemicals associated with plastic packaging: Inventory and hazards. PeerJ Preprints, 6, p.e27036v1. DOI: 10.7287/peerj.preprints.27036v1. URL: https://peerj.com/preprints/27036/ | 09/26/2018
2-bromo-3,7,8-trichlorodibenzo-p-dioxin | 2,3,7-trichloro-8-bromodibenzodioxin | 109333-33-7 | Hornung MW, Zabel EW, Peterson RE. 1996. Toxic equivalency factors of polybrominated dibenzo-p-dioxin, dibenzofuran, biphenyl, and polyhalogenated diphenyl ether congeners based on rainbow trout early life stage mortality. Toxicol Appl Pharmacol 140(2):227-234. DOI: 10.1006/taap.1996.0217. URL: https://www.sciencedirect.com/science/article/pii/S0041008X96902173 Mason G, Zacharewski T, Denomme MA, Safe L, Safe S. 1987. Polybrominated dibenzo-p-dioxins and related compounds: Quantitative in vivo and in vitro structure-activity relationships. Toxicology 44(3):245-255. DOI: 10.1016/0300-483X(87)90027-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X87900278 Schulz-Schalge T, Koch E, Schwind KH, Hutzinger O, Neubert D. 1991. Inductive potency of TCDD, TBDD and three 2,3,7,8-mixed-halogenated dioxins in liver microsomes of male rats. Enzyme kinetic considerations. Chemosphere 23(11-12):1925-1931. DOI: 10.1016/0045-6535(91)90040-K. URL: https://www.sciencedirect.com/science/article/pii/004565359190040K | Byproduct/Intermediate/Reactant, Medical/Veterinary/Research | Mason G, Zacharewski T, Denomme MA, Safe L, Safe S. 1987. Polybrominated dibenzo-p-dioxins and related compounds: Quantitative in vivo and in vitro structure-activity relationships. Toxicology 44(3):245-255. DOI: 10.1016/0300-483X(87)90027-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X87900278 Schulz-Schalge T, Koch E, Schwind KH, Hutzinger O, Neubert D. 1991. Inductive potency of TCDD, TBDD and three 2,3,7,8-mixed-halogenated dioxins in liver microsomes of male rats. Enzyme kinetic considerations. Chemosphere 23(11-12):1925-1931. DOI: 10.1016/0045-6535(91)90040-K. URL: https://www.sciencedirect.com/science/article/pii/004565359190040K | 05/05/2011 | 09/10/2013
2-bromo-4-(2,4,6-tribromo-phenoxy)phenol | "T3-HO-BDE" | n/a | Meerts IATM, Letcher RJ, Hoving S, Marsh G, Bergman Å, Lemmen JG, van der Burg B, Brouwer A. 2001. In vitro estrogenicity of polybrominated diphenyl ethers, hydroxylated PBDEs, and polybrominated bisphenol A compounds. Environ Health Perspect 109(4):399-407. DOI: 10.2307/3454900 . URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1240281/ Meerts IATM, van Zanden JJ, Luijks EAC, van Leeuwen-Bol I, Marsh G, Jakobsson E, Bergman Å, Brouwer A. 2000. Potent competitive interactions of some brominated flame retardants and related compounds with human transthyretin in vitro. Toxicol Sci 56(1):95-104. DOI: 10.1093/toxsci/56.1.95. URL: https://academic.oup.com/toxsci/article/56/1/95/1646061 | Medical/Veterinary/Research | Meerts IATM, Letcher RJ, Hoving S, Marsh G, Bergman Å, Lemmen JG, van der Burg B, Brouwer A. 2001. In vitro estrogenicity of polybrominated diphenyl ethers, hydroxylated PBDEs, and polybrominated bisphenol A compounds. Environ Health Perspect 109(4):399-407. DOI: 10.1289/ehp.01109399. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1240281/ Meerts IATM, van Zanden JJ, Luijks EAC, van Leeuwen-Bol I, Marsh G, Jakobsson E, Bergman Å, Brouwer A. 2000. Potent competitive interactions of some brominated flame retardants and related compounds with human transthyretin in vitro. Toxicol Sci 56(1):95-104. DOI: 10.1093/toxsci/56.1.95. URL: https://academic.oup.com/toxsci/article/56/1/95/1646061 | 05/05/2011 | 09/10/2013
2-bromodibenzo-p-dioxin | 2-MBDD | 105906-36-3 | Mason G, Zacharewski T, Denomme MA, Safe L, Safe S. 1987. Polybrominated dibenzo-p-dioxins and related compounds: Quantitative in vivo and in vitro structure-activity relationships. Toxicology 44(3):245-255. DOI: 10.1016/0300-483X(87)90027-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X87900278 | Byproduct/Intermediate/Reactant | Mason G, Zacharewski T, Denomme MA, Safe L, Safe S. 1987. Polybrominated dibenzo-p-dioxins and related compounds: Quantitative in vivo and in vitro structure-activity relationships. Toxicology 44(3):245-255. DOI: 10.1016/0300-483X(87)90027-8. URL: https://www.sciencedirect.com/science/article/pii/0300483X87900278 | 05/05/2011 | 05/05/2011
2-chloro-4-methylphenol | 2-chloro-4-cresol | 6640-27-3 | Blair RM, Fang H, Branham WS, Hass BS, Dial SL, Moland CL, Tong W, Shi L, Perkins R, Sheehan DM. 2000. The estrogen receptor relative binding affinities of 188 natural and xenochemicals: structural diversity of ligands. Toxicol Sci 54(1):138-153. DOI: 10.1093/toxsci/54.1.138. URL: https://academic.oup.com/toxsci/article/54/1/138/1670624 | Byproduct/Intermediate/Reactant | Alibaba.com. 2-Chloro-4-methylphenol. URL: http://www.alibaba.com/product-gs/481554487/2_Chloro_4_methylphenol.html | 05/05/2011 | 05/05/2011
2-chlorobiphenyl | PCB-1 | 2051-60-7 | Gierthy JF, Arcaro KF, Floyd M. 1997. Assessment of PCB estrogenicity in a human breast cancer cell line. Chemosphere 34(5-7):1495-1505. DOI: 10.1016/S0045-6535(97)00446-3. URL: http://www.sciencedirect.com/science/article/pii/S0045653597004463?via%3Dihub | Industrial additive | CDPH. 1988. Polychlorinated Biphenyls (PCBs). Fact Sheet. URL: http://www.cdph.ca.gov/programs/hesis/Documents/pcbs.pdf UNEP. 1999. Guidlines for the Identification of PCBs and Materials Containing PCBs. URL: http://wedocs.unep.org/handle/20.500.11822/8221 US EPA. 2012. Aroclor and other PCB mixtures. URL: http://www.epa.gov/osw//hazard/tsd/pcbs/pubs/aroclor.htm | 05/05/2011 | 05/05/2011
2-chlorophenol |  | 95-57-8 | Van den Berg KJ, van Raaij JAGM, Bragt PC, Notten WRF. 1991. Interactions of halogenated industrial chemicals with transthyretin and effects on thyroid hormone levels in vivo. Arch Toxicol 65(1):15-19. DOI: 10.1007/BF01973497. URL: https://link.springer.com/article/10.1007/BF01973497 van den Berg KJ. 1990. Interaction of chlorinated phenols with thyroxine binding sites of human transthyretin, albumin and thyroid binding globulin. Chem Biol Interact 76(1):63-75. DOI: 10.1016/0009-2797(90)90034-K. URL: https://www.sciencedirect.com/science/article/pii/000927979090034K | Antimicrobial, Byproduct/Intermediate/Reactant, Household product ingredient, Industrial additive, Pesticide ingredient | US National Library of Medicine. Haz-Map. URL: http://hazmap.nlm.nih.gov/index.php US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 05/05/2011 | 05/05/2011
2-ethoxyethyl acetate | ethylene glycol monoethyl ether acetate; cellosolve acetate; oxitol acetate; ethoxyethyl acetate; ethylglycol acetate; ethyl cellosolve acetate; oxytol acetate | 111-15-9 | Nagano K, Nakayama E, Oobayashi H, Nishizawa T, Okuda H, Yamazaki K. 1984. Experimental studies on toxicity of ethylene glycol alkyl ethers in Japan. Environ Health Perspect 57:75-84. DOI: 10.2307/3429901. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1568303/ Reproductive toxicology. Ethylene glycol monoethyl ether acetate. 1997. Environmental Health Perspectives, 105(Suppl 1), 221–222. URL: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC1470290/ | Byproduct/Intermediate/Reactant, Household product ingredient, Personal care product/Cosmetic ingredient, Solvent | European Commisson. 2016. Commission Staff Working Document Impact Assessment. Defining criteria for identifying endocrine disruptors in the context of the implementation of the plant protection products regulation and biocidal products regulation. URL: https://ec.europa.eu/health/sites/health/files/endocrine_disruptors/docs/2016_impact_assessment_en.pdf US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 04/04/2018
2-ethylhexanoic acid | Hexanoic acid, 2-ethyl- | 149-57-5 | Lampen A, Zimnik S, Nau H. 2003. Teratogenic phthalate esters and metabolites activate the nuclear receptors PPARs and induce differentiation of F9 cells. Toxicol Appl Pharmacol 188(1):14-23. DOI: 10.1016/S0041-008X(03)00014-0. URL: https://www.sciencedirect.com/science/article/pii/S0041008X03000140 Maloney EK, Waxman DJ. 1999. trans-Activation of PPARalpha and PPARgamma by structurally diverse environmental chemicals. Toxicol Appl Pharmacol 161(2):209-218. DOI: 10.1006/taap.1999.8809. URL: https://www.sciencedirect.com/science/article/pii/S0041008X99988099?via%3Dihub | Byproduct/Intermediate/Reactant, Industrial additive, Pesticide ingredient, Plastic/Rubber, Solvent | US EPA. InertFinder. URL: https://iaspub.epa.gov/apex/pesticides/f?p=INERTFINDER:1:0::NO:1:: US National Library of Medicine. Haz-Map. URL: http://hazmap.nlm.nih.gov/index.php | 07/15/2011 | 04/24/2015
2-ethylhexanol | 2-ethyl-1-hexanol | 104-76-7 | Kassotis CD, Klemp KC, Vu DC, Lin CH, Meng CX, Besch-Williford CL, Pinatti L, Zoeller RT, Drobnis EZ, Balise VD, et al. 2015 in press. Endocrine-disrupting activity of hydraulic fracturing chemicals and adverse health outcomes after prenatal exposure in male mice. Endocrinology:en20151375. DOI: 10.1210/en.2015-1375. URL: https://academic.oup.com/endo/article/156/12/4458/2422671 Kassotis CD, Tillitt DE, Davis JW, Hormann AM, Nagel SC. 2014. Estrogen and androgen receptor activities of hydraulic fracturing chemicals and surface and ground water in a drilling-dense region. Endocrinology 155(3):897-907. DOI: 10.1210/en.2013-1697. URL: https://academic.oup.com/endo/article/155/3/897/2843268 | Antimicrobial, Byproduct/Intermediate/Reactant, Household product ingredient, Industrial additive, Metabolite/Degradate, Pesticide ingredient, Plastic/Rubber, Solvent, Unconventional Oil and Gas | Macey GP, Breech R, Chernaik M, Cox C, Larson D, Thomas D, Carpenter DO. 2014. Air concentrations of volatile compounds near oil and gas production: a community-based exploratory study. Environ Health 13:82. DOI: 10.1186/1476-069X-13-82. URL: https://ehjournal.biomedcentral.com/articles/10.1186/1476-069X-13-82 US Dept of Health & Human Services. Household Products Database. URL: http://hpd.nlm.nih.gov/ US EPA. InertFinder. URL: https://iaspub.epa.gov/apex/pesticides/f?p=INERTFINDER:1:0::NO:1:: US EPA. Pesticide Chemical Search. URL: http://www.epa.gov/pesticides/chemicalsearch US National Library of Medicine. Haz-Map. URL: http://hazmap.nlm.nih.gov/index.php US National Library of Medicine. Hazardous Substances Data Bank. URL: http://toxnet.nlm.nih.gov/cgi-bin/sis/htmlgen?HSDB | 08/27/2015 | 03/21/2017